Supplementary Material

Table 1 – Activity concentration, bias and z-score of 210Pb obtained when using direct TDCR instead of efficiency for different time intervals of water sample from intercomparison program.

| Sample | Days after precipitation | Activity concentration (Bq/L) | Reference value (Bq/L) | Bias (%) | z-score |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 2.02 ± 0.10 | 1.866 ± 0.373 | 8.3 | 0.4 |
|  | 2 | 2.17 ± 0.10 |  | 16.0 | 0.8 |
|  | 3 | 2.44 ± 0.11 |  | 30.9 | 1.5 |
|  | 5 | 2.94 ± 0.11 |  | 57.4 | 2.9 |
|  | 10 | 3.45 ± 0.11 |  | 84.7 | 4.2 |
| 2 | 1 | 2.12 ± 0.10 | 1.866 ± 0.373 | 13.8 | 0.7 |
|  | 2 | 2.25 ± 0.10 |  | 20.5 | 1.0 |
|  | 3 | 2.41 ± 0.11 |  | 29.1 | 1.5 |
|  | 5 | 2.89 ± 0.11 |  | 54.8 | 2.7 |
|  | 10 | 3.39 ± 0.11 |  | 81.7 | 4.1 |
| 3 | 1 | 2.20 ± 0.10 | 1.866 ± 0.373 | 18.0 | 0.9 |
|  | 2 | 2.38 ± 0.10 |  | 27.5 | 1.4 |
|  | 3 | 2.44 ± 0.11 |  | 30.6 | 1.5 |
|  | 5 | 3.16 ± 0.11 |  | 69.3 | 3.5 |
|  | 10 | 3.45 ± 0.11 |  | 84.7 | 4.2 |

Table 2 – Activity concentration, bias and z-score of 90Sr obtained when using direct TDCR for different time intervals of water sample from intercomparison program.

| Sample | Days after precipitation | Activity concentration (Bq/L) | Reference value (Bq/L) | Bias (%) | z-score |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 1.70 ± 0.10 | 1.331 ± 0.200 | 28.1 | 1.9 |
|  | 2 | 1.93 ± 0.10 |  | 44.9 | 3.0 |
|  | 3 | 2.11 ± 0.11 |  | 30.9 | 3.9 |
|  | 5 | 2.45 ± 0.11 |  | 58.5 | 5.6 |
|  | 10 | 2.57 ± 0.11 |  | 84.2 | 6.2 |
| 2 | 1 | 1.65 ± 0.10 | 1.331 ± 0.200 | 23.7 | 1.6 |
|  | 2 | 1.80 ± 0.10 |  | 35.6 | 2.4 |
|  | 3 | 1.94 ± 0.10 |  | 45.4 | 3.0 |
|  | 5 | 2.38 ± 0.11 |  | 79.0 | 5.3 |
|  | 10 | 2.51 ± 0.11 |  | 88.3 | 5.9 |
| 3 | 1 | 1.56 ± 0.09 | 1.331 ± 0.200 | 17.5 | 1.2 |
|  | 2 | 1.84 ± 0.10 |  | 38.5 | 2.6 |
|  | 3 | 1.94 ± 0.10 |  | 45.9 | 3.1 |
|  | 5 | 2.31 ± 0.11 |  | 73.5 | 4.9 |
|  | 10 | 2.37 ± 0.11 |  | 77.8 | 5.2 |

Table 3 – Activity concentration, bias and z-score of 210Pb obtained when using Bateman-corrected TDCR instead of efficiency for different time intervals of water sample from intercomparison program.

| Sample | Days after precipitation | Activity concentration (Bq/L) | Reference value (Bq/L) | Bias (%) | z-score |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 1.76 ± 0.08 | 1.866 ± 0.373 | - 5.4 | - 0.3 |
|  | 2 | 1.75 ± 0.07 |  | - 6.0 | - 0.3 |
|  | 3 | 1.86 ± 0.07 |  | - 0.3 | 0.0 |
|  | 5 | 1.95 ± 0.07 |  | 4.3 | 0.2 |
|  | 10 | 1.97 ± 0.06 |  | 5.8 | 0.3 |
| 2 | 1 | 1.85 ± 0.08 | 1.866 ± 0.373 | - 0.6 | 0.0 |
|  | 2 | 1.82 ± 0.07 |  | - 2.3 | - 0.1 |
|  | 3 | 1.83 ± 0.07 |  | - 1.7 | - 0.1 |
|  | 5 | 1.91 ± 0.07 |  | 2.5 | 0.1 |
|  | 10 | 1.94 ± 0.06 |  | 4.1 | 0.2 |
| 3 | 1 | 1.91 ± 0.08 | 1.866 ± 0.373 | 2.5 | 0.1 |
|  | 2 | 1.92 ± 0.07 |  | 3.0 | 0.1 |
|  | 3 | 1.85 ± 0.07 |  | - 0.9 | 0.0 |
|  | 5 | 2.09 ± 0.07 |  | 11.9 | 0.6 |
|  | 10 | 1.97 ± 0.06 |  | 5.8 | 0.3 |

Table 4 – Activity concentration, bias and z-score of 90Sr obtained when using TDCR corrected from Bateman equation instead of efficiency for different time intervals of water sample from intercomparison program.

| Sample | Days after precipitation | Activity concentration (Bq/L) | Reference value (Bq/L) | Bias (%) | z-score |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | 1.40 ± 0.08 | 1.331 ± 0.200 | 5.4 | 0.4 |
|  | 2 | 1.36 ± 0.07 |  | 1.8 | 0.1 |
|  | 3 | 1.35 ± 0.07 |  | 1.3 | 0.1 |
|  | 5 | 1.41 ± 0.07 |  | 6.0 | 0.4 |
|  | 10 | 1.34 ± 0.06 |  | 0.5 | 0.0 |
| 2 | 1 | 1.35 ± 0.08 | 1.331 ± 0.200 | 1.5 | 0.1 |
|  | 2 | 1.27 ± 0.07 |  | - 4.9 | - 0.3 |
|  | 3 | 1.24 ± 0.07 |  | - 7.2 | - 0.5 |
|  | 5 | 1.37 ± 0.07 |  | 2.9 | 0.2 |
|  | 10 | 1.31 ± 0.06 |  | - 1.8 | - 0.1 |
| 3 | 1 | 1.38 ± 0.08 | 1.331 ± 0.200 | 3.8 | 0.3 |
|  | 2 | 1.38 ± 0.07 |  | 3.5 | 0.2 |
|  | 3 | 1.31 ± 0.07 |  | -1.5 | - 0.1 |
|  | 5 | 1.40 ± 0.07 |  | 5.3 | 0.4 |
|  | 10 | 1.32 ± 0.06 |  | -0.8 | - 0.1 |