Supplementary File 1. Optical modelling inputs and design disclosure (SV, BF, and AR)

Purpose: All three lens (SV, BF and AR) designs were modelled in OpticStudio (Ansys, Inc, USA) with identical base parameters in the flexed state. This folder contains the optical modelling inputs (ZOS files) used to generate the wide-field retinal images in the manuscript, and a concise disclosure of the SV, BF, and AR lens implementations that are deemed sufficient for reproducibility. The intent is to allow an independent reader to (1) open the provided optical models, (2) understand what each design represents, and (3) reproduce the modelling conditions across residual-error states.

Summary of the folder contents

OpticStudio (Ansys Inc, USA) model files for SV and BF contact lens designs coupled to schematic myopic eyes can be found in the folder. The schematic myopic eyes are provided in three residual-error states, to evaluate whether residual accommodative error influences optical outcomes. The states are:

- RE0: baseline model-eye configuration

- RE+: positive residual error state

- RE−: negative residual error state

The optical model files are named to indicate:

- Lens type (SV or BF)

- Model eye type (3D myopic eye)

- Residual-error state (RE0, RE+, or RE-)

Representative file patterns in the zipped folder include:

- SV on 3D myopic eye RE0; SV on 3D myopic eye RE+; SV on 3D myopic eye RE-

- BF on 3D myopic eye RE0; BF on 3D myopic eye RE+; BF on 3D myopic eye RE-

How to use this folder

- Choose a design and select the desired residual-error state (RE0, RE+, or RE-).

- Open the corresponding OpticStudio model file.

- Use the pupil sizes and other optical simulation settings described in the manuscript methods to replicate the modelling conditions.

- If reproducing the full dynamic workflow, apply the similar contact lens movement emulation protocol described in the manuscript and generate the time-resolved montage of wide-field retinal images consistently across designs.

Design summary table

| Design label | Model representation | Design specification | Folder/File details |
| --- | --- | --- | --- |
| SV (single vision) | Baseline refractive correction, used as control condition for comparison. | -3.00D SV Rx (minimal spherical aberration over 8 mm), on same base as BF and AR | SV files on the schematic myopic eye for RE0, RE+, and RE- |
| BF (concentric-ring bifocal) | Concentric-ring bifocal design established for myopia control. | In-built OpticStudio super-conic surface type optimised to match a target power profile via damped least-squares optimisation. | BF model files on schematic myopic eyes for RE0, RE+, and RE-. Nominal distance Rx -3.00D |
| AR (opaque non-refractive features) | Base Rx matching SV design combined with non-refractive opaque features to introduce temporal modulation of spatial contrast. | User-defined surface with a spatial transmittance function that encodes the opaque feature pattern. Central 5 mm optic zone contains 28 opaque features (~350 µm × 100 µm), covering ~5% of optic zone; transmittance (0.0 opaque, 1.0 transparent); features absorb >95% of light | The AR model specification details are provided here that facilitate reproducibility. Nominal distance Rx -3.00D. |

Table S1. Design summary table (SV, BF, and AR). Summary of how each design is represented, what is disclosed for reproducibility, and what is included in this folder.

| Item | SV (and AR) | BF (concentric-ring bifocal) |
| --- | --- | --- |
| Surface label (comment) | Surf 2, Anterior contact lens surface | Surf 2, Anterior contact lens surface |
| Surface type | STANDARD (SV), USER-DEFINED (AR) | SUPERCON (BF) |
| Front surface radius | 8.262 mm | 8.262 mm |
| Front surface conic | -0.136 (Q) | -0.136 (Q) |
| Centre thickness | 0.135 mm | 0.135 mm |
| Front surface semi-diam | 4 mm | 4 mm |
| Lens material Abbe # | n = 1.42, Abbe = 50.2 | n = 1.42, Abbe = 50.2 |
| Super Conic terms | Not applicable | 48 terms |
| Coefficients | Not applicable | Coefficients XDAT 2 to 49. |
| Surface label (comment) | Surf 3, Posterior Surface | Surf 3, Posterior Surface |
| Surface type | STANDARD | STANDARD |
| Back surface radius | 7.75 mm | 7.75 mm |
| Back surface conic | -0.25 (Q) | -0.25 (Q) |
| Back surface half-chord, (mm) | 4 | 4 |

Table S2. Contact lens refractive surface specification (front and back surfaces). Surface-level parameters extracted from the OpticStudio files provided for SV and BF. The table highlights that the posterior contact lens surface is shared across SV, AR and BF designs, and that the BF anterior surface is defined by a SUPERCON surface with 48 coefficients.

The main manuscript summarises the BF design in brief to manage space. This section aims to remove ambiguity by specifying the BF target profile, the OpticStudio implementation (super-conic anterior surface optimised by damped least-squares method), and the verification method (target-versus-achieved power-profile comparison).

The BF target represents an established concentric-ring, dual-focus myopia-control class used as an active comparator and is intended to be reproducible by an independent reader.

We derived the BF target profile using metrology: the optic-zone power profile of -3.00 D MiSight contact lenses (CooperVision, USA) was measured with a NIMOevo instrument (Lambda-X, Nivelles), imported into OpticStudio, and used directly as the merit-function operand for optimisation of the BF anterior surface.

The resulting BF front-surface parameterisation (SUPERCON surface type with its coefficient set) reproduces the measured target profile, with Figure SF1 showing the target-versus-achieved agreement.

Figure S1.BF target-versus-achieved radial power profile. Overlay plot showing the BF target power profile used as the optimisation objective and the achieved BF design power profile generated by the utilising supersonic anterior surface in the OpticStudio.