Supplemental Information

Supplemental Mehods: The methods of ddPCR, qPCR, NGS assay

Supplemental Tables

Supplemental Table S1: LOB and LOD of CRISPR-CPPC for T790M mutation

Supplemental Table S2: Application of CRISPR-CPPC on patient samples

Supplemental Table S3: Analytical performance of assays for detecting T790M mutation in cfDNA from patients with NSCLC presenting disease progression on 1st or 2nd generation EGFR-TKI

Supplemental Table S4: Comparison of the test results of qPCR to CRISPR-CPPC and ddPCR for EGFR T790M in cell-free plasma DNA

Supplemental Table S5: Comparison of the test results of ddPCR to those of CRISPR-CPPC for EGFR T790M in cell-free plasma DNA

Supplemental Table S6: Cases with different ddPCR and CRISPR-CPPC results

Supplemental Table S7: Application of CRISPR-CPPC on follow-up patient samples

Supplemental Table S8: ddPCR results after pre-amplification step

Supplemental Figure 1: Study flow chart

Supplemental Figure 2: Schematic illustration of the CRISPR target site around the human EGFR T790 locus

Supplemental Figure 3: Distribution of T790M copies from sixty samples in this study

Supplemental Methods: The methods of ddPCR, qPCR, NGS assay

ddPCR assay

The number of T790M mutant copies in cfDNA samples before and after CRISPR-CPPC was quantified using ddPCR with the PrimePCRTM ddPCRTM Mutation Detection Assay kit (Bio-Rad, Hercules, CA, USA) according to the manufacturer’s instructions. Amplification was performed in a reaction volume of 20 µL using a QX100 Droplet Digital PCR System (Bio-Rad, Hercules, CA, USA). The PCR mix was composed of 10 µL of Bio-Rad Super mix TaqMan, 2 µL of T790M primer/probe mix, and 8 µL of post-CRISPR-CPPC cfDNA, which was ahead diluted 100-times for the optimal separation of false-positive and true-positive. Thermal cycling conditions were as follows: 10 min at 95 °C, followed by 40 cycles of 95 °C for 30 s and 55 °C for 60 s. Results were analyzed with Quantasoft v.1.7.2 software (Bio-Rad, Hercules, CA, USA).

Roche cobas® EGFR Mutation Test v2 (qPCR assay)

The qPCR assay was executed using cobas® EGFR Mutation Test v2 according to manufacturer’s instructions. 75 uL of DNA from each plasma sample (2 mL) was loaded into three reaction wells (25 uL DNA per well). Amplification and detection were performed using the Cobas z 480 analyzer (Roche Molecular Systems, Inc., Branchburg, NJ, USA). The data were interpreted using the Cobas z 480 software.

NGS assay

The NGS results are referenced from the NGS report using either Grardant360 CDx (Guardant Health, Redwood City, CA, USA) or Oncomine Pan-Cancer Cell Free Assay (Thermo Fisher Scientific). Patient samples from Severance Hospital were sent to a laboratory in the United States where Guardant360 CDx (Guardant Health, Redwood City, CA, USA) was used, following the manufacturer’s instructions. Patient samples from Gangnam Severance Hospital was prepared using the Oncomine Pan-Cancer Cell-Free Assay (Thermo Fisher Scientific). The NGS preparation using Oncomine Pan-Cancer Cell-Free Assay followed the Kim et al.’s method  (1). The libraries were prepared using >5 ng input of nucleci acid according to manufacturer’s instructions. Templating and sequencing were performed using the Ion 540TM Kit on the Ion ChefTM and on the Ion S5 XL system (Thermo Fisher Scientific). The alignment to the hg19 human reference genome and variant calling were executed using Torrent SuiteTM and Ion ReporterTM software version 5.10, respectively. The molecular coverage depth and read coverage depth at the target base were provided by the Torrent SuiteTM software. The average of median reac coverage and median molecular coverage were 36,095 x and 1934 x, respectively. Variants with an allele frequency of >0.1% were reported, and the measured allele frequency (%) was calculated from the division of mutant coverage depth by the total coverage depth.

1.	Kim Y, Shin S, Lee KA. Exosome-based detection of egfr t790m in plasma and pleural fluid of prospectively enrolled non-small cell lung cancer patients after first-line tyrosine kinase inhibitor therapy. Cancer Cell Int 2021;21:50.

Supplemental Table 1. LOB and LOD of CRISPR-CPPC for T790M mutation

| Number of samples | Blank* | T790M (-) Healthy control cfDNA† |
| --- | --- | --- |
|  | Positive events | Positive events |
| 1 | 0 | 1 |
| 2 | 0 | 0 |
| 3 | 0 | 2 |
| 4 | 0 | 2 |
| 5 | 0 | 0 |
| 6 | 0 | 1 |
| 7 | 0 | 1 |
| 8 | 0 | 5 |
| 9 | 0 | 2 |
| 10 | 0 | 2 |
| 11 | 0 | 5 |
| Event | 0 | 21 |
| Mean | 0 | 1.9 |
| SD | 0 | 1.6 |
| LOB | 0 |  |
| LOD (CLSI EP17-A2) |  | 3 |
| 95% CI upper bound (one-tail Poisson distribution) |  | 2.5 |
| *Analysis of post CRISPR-CPPC blank samples without DNA (n = 11) |  |  |  |
| †11 healthy subjects were anonymized and studied as control samples |  |  |  |
| Abbreviations: cfDNA, cell-free DNA; CI, confidence interval; LOB, limit of blank; LOD, limit of detection; SD, standard deviation |  |  |  |

Supplemental Table 2. Application of CRISPR-CPPC on patient samples

| Sample No. | T790M qPCR‡ | T790M ddPCR (cfDNA) Detection positive (≥2 events/assay) |  |  |  | T790M CRISPR-CPPC (cfDNA) Detection positive (≥6 events/assay) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Result | Events | Wild | Allele frequency (%) | Result | Events | Wild | Allele frequency (%) | Result |
| 1 | Positive | 17 | 293 | 5.5 | Positive | 224 | 1257 | 15.1 | Positive |
| 2 | ND | 0 | 423 | 0.0 | ND | 5 | 2336 | 0.2 | ND |
| 3 | ND | 3 | 560 | 0.5 | Positive | 353 | 4570 | 7.2 | Positive |
| 4 | Positive | 1 | 238 | 0.4 | ND | 102 | 1811 | 5.3 | Positive |
| 5 | Positive | 22 | 423 | 4.9 | Positive | 47 | 2727 | 1.7 | Positive |
| 6 | Positive | 758 | 1864 | 28.9 | Positive | 6932 | 12176 | 36.3 | Positive |
| 7 | Positive | 0 | 533 | 0.0 | ND | 6 | 1228 | 0.5 | Positive |
| 8 | Positive | 3 | 430 | 0.7 | Positive | 7 | 1556 | 0.4 | Positive |
| 9 | Positive | 5 | 438 | 1.1 | Positive | 8 | 1215 | 0.7 | Positive |
| 10 | Positive | 16 | 362 | 4.2 | Positive | 88 | 2333 | 3.6 | Positive |
| 11 | ND | 1 | 191 | 0.5 | ND | 3 | 3138 | 0.1 | ND |
| 12* | ND | 3 | 194 | 1.5 | Positive | 3 | 2579 | 0.1 | ND |
| 13 | Positive | 18 | 2068 | 0.9 | Positive | 62 | 2402 | 2.5 | Positive |
| 14 | ND | 0 | 648 | 0.0 | ND | 4 | 2536 | 0.2 | ND |
| 15 | Positive | 42 | 1221 | 3.3 | Positive | 87 | 807 | 9.7 | Positive |
| 16 | ND | 0 | 32 | 0.0 | ND | 4 | 683 | 0.6 | ND |
| 17 | ND | 0 | 401 | 0.0 | ND | 2 | 1870 | 0.1 | ND |
| 18 | ND | 0 | 590 | 0.0 | ND | 5 | 2071 | 0.2 | ND |
| 19 | ND | 0 | 602 | 0.0 | ND | 4 | 3796 | 0.1 | ND |
| 20 | ND | 0 | 146 | 0.0 | ND | 4 | 3036 | 0.1 | ND |
| 21 | ND | 0 | 412 | 0.0 | ND | 8 | 3589 | 0.2 | Positive |
| 22 | ND | 0 | 1886 | 0.0 | ND | 4 | 872 | 0.5 | ND |
| 23 | ND | 0 | 5416 | 0.0 | ND | 1 | 1181 | 0.1 | ND |
| 24 | ND | 0 | 292 | 0.0 | ND | 2 | 3114 | 0.1 | ND |
| 25 | ND | 0 | 481 | 0.0 | ND | 14 | 2776 | 0.5 | Positive |
| 26 | ND | 0 | 513 | 0.0 | ND | 0 | 1180 | 0.0 | ND |
| 27 | ND | 0 | 39 | 0.0 | ND | 1 | 736 | 0.1 | ND |
| 28 | ND | 0 | 29 | 0.0 | ND | 4 | 1012 | 0.4 | ND |
| 29 | ND | 3 | 1170 | 0.3 | Positive | 9 | 2568 | 0.3 | Positive |
| 30 | ND | 2 | 386 | 0.5 | Positive | 38 | 2766 | 1.4 | Positive |
| 31 | ND | 0 | 162 | 0.0 | ND | 6 | 3995 | 0.1 | Positive |
| 32 | ND | 0 | 302 | 0.0 | ND | 2 | 1978 | 0.1 | ND |
| 33 | ND | 0 | 542 | 0.0 | ND | 4 | 4108 | 0.1 | ND |
| 34 | ND | 1 | 310 | 0.3 | ND | 24 | 3816 | 0.6 | Positive |
| 35 | ND | 1 | 2432 | 0.0 | ND | 4 | 2013 | 0.2 | ND |
| 36 | ND | 0 | 1111 | 0.0 | ND | 5 | 4803 | 0.1 | ND |
| 37 | ND | 0 | 134 | 0.0 | ND | 0 | 1326 | 0.0 | ND |
| 38 | ND | 0 | 155 | 0.0 | ND | 5 | 2460 | 0.2 | ND |
| 39 | ND | 1 | 207 | 0.5 | ND | 2 | 1440 | 0.1 | ND |
| 40 | ND | 0 | 172 | 0.0 | ND | 12 | 1394 | 0.9 | Positive |
| 41 | Positive | 51 | 721 | 6.6 | Positive | 91 | 912 | 9.1 | Positive |
| 42 | ND | 0 | 601 | 0.0 | ND | 4 | 8345 | 0.0 | ND |
| 43 | ND | 0 | 110 | 0.0 | ND | 6 | 5712 | 0.1 | Positive |
| 44 | ND | 0 | 1541 | 0.0 | ND | 8 | 9580 | 0.1 | Positive |
| 45 | Positive | 6 | 205 | 2.8 | Positive | 28 | 1220 | 2.2 | Positive |
| 46 | Positive | 2 | 114 | 1.7 | Positive | 158 | 7316 | 2.1 | Positive |
| 47† | ND | 12 | 410 | 2.8 | Positive | 5 | 4709 | 0.1 | ND |
| 48 | ND | 0 | 71 | 0.0 | ND | 7 | 2119 | 0.3 | Positive |
| 49 | ND | 0 | 98 | 0.0 | ND | 2 | 3618 | 0.1 | ND |
| 50 | Positive | 28 | 762 | 3.5 | Positive | 319 | 3713 | 7.9 | Positive |
| 51 | ND | 0 | 112 | 0.0 | ND | 2 | 726 | 0.3 | ND |
| 52 | ND | 0 | 44 | 0.0 | ND | 2 | 2870 | 0.1 | ND |
| 53 | ND | 0 | 44 | 0.0 | ND | 3 | 4128 | 0.1 | ND |
| 54 | Positive | 0 | 31 | 0.0 | ND | 274 | 3874 | 6.6 | Positive |
| 55 | ND | 0 | 63 | 0.0 | ND | 13 | 7324 | 0.2 | Positive |
| 56 | ND | 0 | 201 | 0.0 | ND | 16 | 7243 | 0.2 | Positive |
| 57 | ND | 0 | 19 | 0.0 | ND | 5 | 3641 | 0.1 | ND |
| 58 | ND | 0 | 139 | 0.0 | ND | 12 | 6052 | 0.2 | Positive |
| 59 | ND | 0 | 11 | 0.0 | ND | 10 | 5294 | 0.2 | Positive |
| 60 | Positive | 1 | 137 | 0.7 | ND | 562 | 6343 | 8.1 | Positive |

* T790M was negative by next-generation sequencing (NGS)

† T790M was positive with an allele frequency of 0.2% by NGS

‡T790M was detected using Roche cobas® EGFR Mutation Test v2

Abbreviations: cfDNA, Cell-free DNA; CRISPR-CPPC, CRISPR system combined post-PCR cfDNA; ddPCR, droplet digital PCR; ND, not detected; qPCR, real-time PCR; SQI, semiquantitative index

Supplemental Table 3. Analytical performance of assays for detecting T790M mutation in cfDNA from patients with NSCLC presenting disease progression on 1st or 2nd generation EGFR-TKI

| Method | T790M mutation was confirmed with multiple assays* |  |  | Sensitivity (95%CI) | Specificity (95%CI) | Accuracy (95%CI) |
| --- | --- | --- | --- | --- | --- | --- |
|  | Results | Pos (n=25) | Neg (n=35) |  |  |  |
| ddPCR | Pos | 16 | 1 | 64.0% ( 42.5% to 82.0%) | 97.1% ( 85.1% to 99.9%) | 83.3% (71.5% to 91.7%) |
|  | Neg | 9 | 34 |  |  |  |
| CRISPR-CPPC | Pos | 23 | 8 | 92.0% (74.0% to 99.0%) | 77.1% (59.9% to 89.6%) | 83.3% (71.5% to 91.7%) |
|  | Neg | 2 | 27 |  |  |  |

*T790M detected by more than two methods (qPCR from cfDNA or tissue, NGS, ddPCR, CRISPR-CPPC) simultaneously is considered “true positive”.

Abbreviations: cfDNA, cell-free DNA; NSCLC, non-small cell lung cancer; EGFR-TKI, epidermal growth factor receptor-tyrosine kinase inhibitor; Pos, positive; Neg, negative; CI, confidence interval; CRISPR-CPPC, CRISPR system combined post-PCR cfDNA; ddPCR, droplet digital PCR; qPCR, real-time PCR; NGS, next-generation sequencing

Supplemental Table 4. Comparison of test results of qPCR to those of CRISPR-CPPC and ddPCR for EGFR T790M in cell-free plasma DNA

|  |  | qPCR |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Positive | Negative |  |  |  |  |
| CRISPR-CPPC | Positive | 16 | 15 | PPA (%) (95% CI) |  | 100.0 (80.6 – 100) |  |
|  | Negative | 0 | 29 | NPA (%) (95% CI) |  | 65.9 (51.1 – 78.1) |  |
|  |  |  |  | OPA (%) (95% CI) |  | 75.0 (62.8 – 84.2) |  |
| ddPCR | Positive | 12 | 5 | PPA (%) (95% CI) |  | 75.0 (50.5 – 89.8) |  |
|  | Negative | 4 | 39 | NPA (%) (95% CI) |  | 88.6 (76.0 – 95.0) |  |
|  |  |  |  | OPA (%) (95% CI) |  | 85.0 (73.9 – 91.9) |  |

Abbreviations: CI, confidence interval; CRISPR-CPPC, CRISPR system combined with post-PCR cfDNA; ddPCR, droplet digital polymerase chain reaction; NPA, negative percent agreement; OPA, overall percent agreement; PPA, positive percent agreement; qPCR, real-time PCR

Supplemental Table 5. Comparison of the test results of ddPCR to those of CRISPR-CPPC for EGFR T790M in cell-free plasma DNA

|  |  | ddPCR |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Positive | Negative |  |  |  |  |  |
| CRISPR-CPPC | Positive | 15 | 16 | PPA (%) (95% CI) |  | 88.2 (65.7 – 96.7) |  |  |
|  | Negative | 2* | 27 | NPA (%) (95% CI) |  | 62.8 (47.9 – 75.6) |  |  |
|  |  |  |  | OPA (%) (95% CI) |  | 70.0 (57.5 – 80.1) |  |  |

* NGS was performed, and it was confirmed that one sample was T790M-positive with an allele frequency of 0.2% and the other sample was T790M-negative.

Abbreviations: CI, confidence Interval; CRISPR-CPPC, CRISPR system combined with post-PCR cfDNA; ddPCR, droplet digital PCR; NPA, negative percent agreement; OPA, overall percent agreement; PPA, positive percent agreement

Supplemental Table 6. Cases with different ddPCR and CRISPR-CPPC results

| Sample No. | Sex | Age | First- or second-generation EGFR-TKIs | Tissue Genotyping at diagnosis | DNA input (ng) | Plasma collection date (Months since TKI) | CRISPR-CPPC |  |  |  | qPCR (SQI) | ddPCR |  | NGS | Tissue genotyping at disease progression | Image (CT, MRI, PET-CT) interpretation at disease progression* |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | Positive event (≥6 events/assay) |  | Wild event | AF (%) |  | Positive event (≥2 events/assay) |  |  |  |  |
| CRISPR-CPPC-Positive |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 | M | 55 | Afatinib | Exon19 deletion | 0.96 | 14 | Pos | 102 | 1811 | 5.3 | Pos (8.45) | Neg | 1 | T790M 2.4% | N/T |  |
| 7 | M | 77 | Gefitinib | L858R | 0.58 | 7 | Pos | 6 | 1228 | 0.5 | Pos (4.00) | Neg | 0 | Not tested | N/T |  |
| 21* | F | 69 | Afatinib | Exon19 deletion | 1.41 | 8 | Pos | 8 | 3589 | 0.2 | Neg | Neg | 0 | Not tested | N/T | *Hepatic metz R/O bone metz |
| 25* | F | 59 | Afatinib | L858R | 1.07 | 12 | Pos | 14 | 2776 | 0.5 | Neg | Neg | 0 | Not tested | N/T | *New metz at right adrenal gland R/O malignant pleural effusion |
| 31* | F | 75 | Gefitinib | L858R | 0.58 | 18 | Pos | 6 | 3995 | 0.1 | Neg | Neg | 0 | Not tested | N/T | *Aggravation of lung cancer Pleural seeding metz with malignant effusion Metz in both lungs |
| 34 | F | 85 | Gefitinib | Exon19 deletion | 1.28 | 16 | Pos | 24 | 3816 | 0.6 | Neg | Neg | 1 | Not tested | T790M (+) |  |
| 40* | F | 68 | Gefitinib | Exon19 deletion | 0.71 | 17 | Pos | 12 | 1394 | 0.9 | Neg | Neg | 0 | Not tested | N/T | *Pleural seeding at metastatic lesions |
| 43*† | M | 43 | Gefitinib | Exon19 deletion | 2.26 | 46 | Pos | 6 | 5712 | 0.1 | Neg | Neg | 0 | Not tested | N/T | *PD Brain metz |
| 44*† | M | 44 | Gefitinib | Exon19 deletion | 14.5 | 48 | Pos | 8 | 9580 | 0.1 | Neg | Neg | 0 | Not tested | N/T | *PD |
| 48 | F | 59 | Gefitinib | Exon19 deletion | 1.66 | 14 | Pos | 7 | 2119 | 0.3 | Neg | Neg | 0 | T790M 0.1% | N/T |  |
| 54 | F | 72 | Gefitinib | L858R | <0.1 | 31 | Pos | 274 | 3874 | 6.6 | Pos (4.99) | Neg | 0 | Not tested | N/T |  |
| 55* | M | 59 | Gefitinib | L858R | <0.1 | 9 | Pos | 13 | 7324 | 0.2 | Neg | Neg | 0 | Not tested | N/T | *R/O metz at brain Diffused metz in entire both lungs |
| 56* | F | 70 | Gefitinib | G719S | 1.38 | 13 | Pos | 16 | 7243 | 0.2 | Neg | Neg | 0 | Not tested | N/T | *PD Endobronchial invasion and pleural seeding metz Left pleural metz with pleural effusion. Increased metastatic lesion in the left temporal lobe of brain. |
| 58 | F | 46 | Afatinib | Exon19 deletion | <0.1 | 8 | Pos | 12 | 6052 | 0.2 | Neg | Neg | 0 | T790M 0.2% | N/T |  |
| 59 | F | 49 | Afatinib | L858R | <0.1 | 54 | Pos | 10 | 5294 | 0.2 | Neg | Neg | 0 | Neg | T790M (+) |  |
| 60 | F | 60 | Afatinib | Exon19 deletion | 0.68 | 11 | Pos | 562 | 6343 | 8.1 | Pos (8.36) | Neg | 1 | T790M 2.1% | N/T |  |
| CRISPR-CPPC-Negative |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 | F | 61 | Gefitinib | L858R | 1.09 | 24 | Neg | 3 | 2579 | 0.1 | Neg | Pos | 3 | Neg | N/T |  |
| 47 | M | 57 | Gefitinib | Exon19 deletion | <0.1 | 6 | Neg | 5 | 4709 | 0.1 | Neg | Pos | 12 | T790M 0.2% | N/T |  |

* Clinical history and image interpretation support that a positive CRISPR-CPPC T790M result would be close to a true positive.

† Same patient. The second test was performed three months after the first test.

Abbreviations: AF, allele frequency; CRISPR-CPPC, CRISPR system combined post-PCR cfDNA; CT, computed tomography; ddPCR, droplet digital PCR; metz, metastasis; MRI, magnetic resonance imaging; NGS, next-generation sequencing; N/T, not tested; PD, progressive disease; PET-CT, positron emission tomography-computed tomography; Pos, positive; qPCR, real-time PCR; R/O, rule out; SQI, semiquantitative index; TKI, tyrosine kinase inhibitor.

Supplemental Table 7. Application of CRISPR-CPPC on follow-up patient samples

|  | T790M qPCR | T790M ddPCR (cfDNA) |  |  | T790M CRISPR-CPPC (cfDNA) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | SQI* | Events | Wild | Allele frequency (%) | Events | Wild | Allele frequency (%) |
| Patient A | ND | 1 | 191 | 0.5 | 3 | 3138 | 0.1 |
| Patient A | ND | 3 | 194 | 1.5 | 3 | 2579 | 0.1 |
| Patient A | 9.29 | 18 | 2068 | 0.9 | 62 | 2402 | 2.5 |
| Patient B | ND | 0 | 648 | 0.0 | 4 | 2536 | 0.2 |
| Patient B | 10.74 | 42 | 1221 | 3.3 | 87 | 807 | 9.7 |
| Patient C | ND | 0 | 32 | 0.0 | 4 | 683 | 0.6 |
| Patient C | ND | 0 | 401 | 0.0 | 2 | 1870 | 0.1 |
| Patient D | ND | 0 | 590 | 0.0 | 5 | 2071 | 0.2 |
| Patient D | ND | 0 | 602 | 0.0 | 4 | 3796 | 0.1 |
| Patient E | ND | 0 | 146 | 0.0 | 4 | 3036 | 0.1 |
| Patient E | ND | 0 | 412 | 0.0 | 8 | 3589 | 0.2 |
| Patient F | ND | 0 | 1886 | 0.0 | 4 | 872 | 0.5 |
| Patient F | ND | 0 | 5416 | 0.0 | 1 | 1181 | 0.1 |
| Patient G | ND | 0 | 29 | 0.0 | 4 | 1012 | 0.4 |
| Patient G | ND | 3 | 1170 | 0.3 | 9 | 2568 | 0.3 |
| Patient H | ND | 0 | 110 | 0.0 | 6 | 5712 | 0.1 |
| Patient H | ND | 0 | 1541 | 0.0 | 8 | 9580 | 0.1 |
| * SQI was measured with Roche cobas® EGFR Mutation Test v2 |  |  |  |  |  |  |  |
| Abbreviations: cfDNA, Cell-free DNA; CRISPR-CPPC, CRISPR system combined post-PCR cfDNA; ddPCR, droplet digital PCR; ND, not detected; qPCR, real-time PCR; SQI, semiquantitative index Supplementary Table 8. ddPCR results after pre-amplification step Sample No.; T790M qPCR*; T790M ddPCR (cfDNA) Detection positive (≥2 events/assay); T790M ddPCR (pre-amplified cfDNA) Detection positive (≥2 events/assay); T790M ddPCR (pre-amplified cfDNA with 1000x dilution) Detection positive (≥2 events/assay); T790M CRISPR-CPPC (cfDNA) Detection positive (≥6 events/assay). Result; Events; Wild; Allele frequency (%); Result; Events; Wild; Allele frequency (%); Result; Events; Wild; Allele frequency (%); Result; Events; Wild; Allele frequency (%); Result. 7; Pos; 0; 533; 0.0; ND; 0; 19140; 0.0; ND; 1; 16483; 0.0; ND; 6; 1228; 0.5; Pos. 34; ND; 1; 310; 0.3; ND; 0; 17735; 0.0; ND; 0; 16426; 0.0; ND; 24; 3816; 0.6; Pos. 54; Pos; 0; 31; 0.0; ND; 0; 14852; 0.0; ND; 0; 17624; 0.0; ND; 274; 3874; 6.6; Pos. 60; Pos; 1; 137; 0.7; ND; 0; 15181; 0.0; ND; 0; 18986; 0.0; ND; 562; 6343; 8.1; Pos. |

*T790M was detected using Roche cobas® EGFR Mutation Test v2

Abbreviations: cfDNA, Cell-free DNA; CRISPR-CPPC, CRISPR system combined post-PCR cfDNA; ddPCR, droplet digital PCR; ND, not detected; Pos, positive; qPCR, real-time PCR; SQI, semiquantitative index

Supplemental Figure 1. Study Flow chart

Supplemental Figure 2. Schematic illustration of the CRISPR target site around the human EGFR T790 locus

Supplemental Figure 3. Distribution of T790M copies from sixty samples in this study