# Supplementary Information

## Analytical Validation

Methods

To ensure the technical accuracy of StrokeVista before clinical validation, analytical validation was performed using the digital phantom dataset established by Kudo et al. [51]. This digital phantom has been widely used to evaluate the performance of CTP analysis software [52, 53]. The dataset is based on widely accepted tracer kinetic theory and includes ground-truth values for cerebral blood flow (CBF), cerebral blood volume (CBV), mean transit time (MTT), and time to maximum (Tmax).

Validation Range and Performance Metrics

Validation was performed using CBV, CBF, MTT, and Tmax as evaluation metrics to assess whether StrokeVista achieved ≥90% agreement with the RAPID results. The validation range and acceptance criteria for each parameter are presented in Table S1. Validation ranges were established on the basis of clinically relevant perfusion parameter ranges, ensuring coverage of the full spectrum from normal to severely ischemic brain tissue [18].

Analytical Validation Results

StrokeVista achieved 100% agreement with the RAPID success criteria for MTT, CBF, and CBV and 99.07% agreement for Tmax, meeting all validation acceptance criteria (Table S2). Therefore, StrokeVista demonstrates high accuracy and reliability in processing CTP images and generating precise perfusion parameters [54].

## Bland–Altman Analysis of Agreement Between RAPID and StrokeVista CTP Parameters

Mean differences were minimal and nonsignificant for all parameters (Tmax: 2.03 mL, P = 0.28; core volume: 0.88 mL, P = 0.07; mismatch volume: 1.15 mL, P = 0.53), indicating no systematic bias (Table S3). More than 90% of all measurements fell within the limits of agreement for all parameters (Tmax: 91.8%, core volume: 93.8%, and mismatch volume: 91.4%). Weak to moderate positive correlations between differences and mean values (r = 0.32–0.57) revealed proportional bias, with slightly larger differences at higher measurement ranges; this is a common phenomenon in perfusion imaging, and it did not affect clinical decision concordance (overall agreement: 97.4%).

## Detailed Analysis of Discordant Decision Cases

Overall Characteristics and Clinical Presentation

The analysis revealed eight discordant cases (Table S4), representing 2.6% of the total sample. Three of these cases comprised positive RAPID and negative StrokeVista decisions, and the other five cases comprised the opposite decision pattern. All discordant cases involved large vessel occlusion , with middle cerebral artery first segment occlusion being the most common concern (50%), followed by internal carotid artery (37.5%) and basilar or vertebral artery (12.5%) occlusions. Stroke severity in the cases comprising positive RAPID and negative StrokeVista decisions was relatively homogeneous (National Institutes of Health Stroke Scale score: 15.0 ± 1.0, age: 73.3 ± 2.1 years, and onset-to-imaging time: 6.2 ± 2.6 hours) and less variable than that in cases comprising negative RAPID and positive StrokeVista decisions (National Institutes of Health Stroke Scale score: 18.6 ± 6.3, age: 75.2 ± 12.0 years, and onset-to-imaging time: 3.2 ± 2.5 hours), reflecting a broader severity spectrum and earlier presentation.

Differences in Imaging Parameters

Systematic differences in core volume measurement drove decision discordance. In RAPID+/StrokeVista− cases, StrokeVista measured larger cores than did RAPID (26.2 ± 45.3 vs. 17.7 ± 30.6 mL), yielding smaller mismatch volumes and nonrecommendation for thrombectomy. Conversely, in RAPID−/StrokeVista+ cases, RAPID measured larger cores than did StrokeVista (78.2 ± 47.3 vs. 47.6 ± 28.5 mL), leading to treatment nonrecommendation. Notably, despite these core differences, final mismatch volumes were similar in the RAPID−/StrokeVista+ cases (RAPID vs. StrokeVista: 91.4 ± 79.7 vs. 91.5 ± 83.0 mL), suggesting that systematic biases in Tmax and core measurements cancel each other out. High standard deviations in both groups reflected that parameters were near decision thresholds. Case 004, a hyperacute basilar occlusion (0.6 hours), exhibited extreme Tmax differences (RAPID vs. StrokeVista: 9 vs. 77.29 mL), possibly representing incomplete perfusion deficit establishment in the ultra-early phase.

Clinical Outcome Analysis

Clinicians did not strictly follow software recommendations: 33.3% (1/3) of all patients with RAPID+/StrokeVista− recommendations and 60.0% (3/5) of those with RAPID−/StrokeVista+ recommendations underwent thrombectomy. Overall prognosis was poor, with 87.5% patients (7/8) achieving an mRS of >4 at discharge and only 25% (2/8) achieving functional independence (mRS ≤ 2). No significant differences were noted in outcomes between the RAPID+/StrokeVista− group and the RAPID−/StrokeVista+ group (favorable outcome: 33.3% vs. 20.0%, P = 1.000; median mRS: 5 vs. 5, P = 0.634). Notably, the two cases with favorable outcomes (Cases 003 and 004; mRS ≤ 2 for both) underwent thrombectomy, with Case 004 benefiting from ultra-early treatment (0.6 hours from symptom onset).

Table S1. Performance metrics for analytical validation

| Test parameter | RAPID success criteria | Validation range | Sample size (N) | Acceptance criteria |
| --- | --- | --- | --- | --- |
| MTT (SNR = 5) | Absolute error < 5.0 seconds | 3.4–12 seconds | 630 | Agreement > 90% (567) |
| CBF ≤ 60 (SNR = 5) | Normalized CBF absolute error < 0.35 | 2.5–60 mL/100g/minute | 651 | Agreement > 90% (586) |
| CBF 61–87.5 (SNR = 5) | Normalized CBF absolute error < 0.45 | 61–87.5 mL/100g/minute | 84 | Agreement > 90% (76) |
| CBV (SNR = 5) | Absolute error < 1.5% | 1.0–5.0 mL/100g | 735 | Agreement > 90% (662) |
| Tmax (SNR = 5) | Absolute error < 4.5 seconds | 1.7–12.0 seconds | 645 | Agreement > 90% (581) |

Table S2. Results of analytical validation

| Test parameter | Acceptance criteria | Validation results | Actual error range | Result |
| --- | --- | --- | --- | --- |
| MTT | Agreement > 90% | 100% (630) | -4.198 to 4.069 sec | PASS |
| CBF ≤ 60 | Agreement > 90% | 100% (651) | -0.273 to 0.294 | PASS |
| CBF 61–87.5 | Agreement > 90% | 100% (84) | -0.098 to 0.227 | PASS |
| CBV | Agreement > 90% | 100% (735) | -1.241 to 0.632% | PASS |
| Tmax | Agreement > 90% | 99.07% (639) | -6.954 to 1.331 sec | PASS |

Table S3. Bland–Altman analysis of agreement between RAPID and StrokeVista perfusion measurements

| Parameter | Mean Diff (mL) | 95% CI (mL) | Upper LoA (mL) | Lower LoA (mL) | Within LoA (%) | Systematic bias P | Proportional bias (r) | P value |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Tmax >6 seconds | 2.03 | -1.64 to 5.70 | 66.00 | -61.94 | 91.8 | 0.279 | 0.330 | <0.001 |
| Core | 0.88 | -0.07 to 1.83 | 17.41 | -15.65 | 93.8 | 0.071 | 0.573 | <0.001 |
| Mismatch | 1.15 | -2.41 to 4.71 | 63.21 | -60.91 | 91.4 | 0.527 | 0.323 | <0.001 |

Diff = difference. Mean differences were assessed using paired t tests (or one-sample t tests against 0).

Systematic bias P: P value from testing whether the mean difference significantly differs from 0.

Proportional bias: Pearson correlation between differences and mean values, with significant correlation indicating increasing measurement discordance with parameter magnitude.

All mean differences were nonsignificant (P > 0.05), indicating no systematic bias. Proportional bias was detected for all parameters but did not affect clinical decision concordance (overall agreement: 97.4%, kappa: 0.93).

Table S4. Detailed characteristics of individual discordant cases

| Case ID | Age/ sex | NIHSS score | Onset-to-CT time (h) | Vessel occlusion location | RAPID parameters |  |  | StrokeVista parameters |  |  | Actual treatment | mRS at discharge |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Tmax (mL) | Cor volume (mL) | MV (mL) | Tmax (mL) | Core volume (mL) | MV (mL) |  |  |
| RAPID recommended EVT, StrokeVista did not (n = 3) |  |  |  |  |  |  |  |  |  |  |  |  |
| 001 | 75/F | 14 | 8.1 | M1 | 14 | 0 | 14 | 2.6 | 0 | 2.6 | EVT- | 5 |
| 002 | 74/M | 15 | 3.2 | M1 | 196 | 53 | 143 | 231.44 | 78.51 | 152.93 | EVT- | 6 |
| 003 | 71/M | 16 | 7.4 | M1 | 19 | 0 | 19 | 8.75 | 0 | 8.75 | EVT+ | 2 |
| Mean |  | 15.0 | 6.2 |  | 76.3 | 17.7 | 58.7 | 80.9 | 26.2 | 54.8 |  | 4.3 |
| ± |  | ± | ± |  | ± | ± | ± | ± | ± | ± |  | ± |
| SD |  | 1.0 | 2.6 |  | 103.7 | 30.6 | 73.1 | 130.4 | 45.3 | 85.1 |  | 2.1 |
| StrokeVista recommended EVT, RAPID did not (n = 5) |  |  |  |  |  |  |  |  |  |  |  |  |
| 004 | 55/M | 9 | 0.6 | BA/VA | 9 | 0 | 9 | 77.29 | 0 | 77.29 | EVT+ | 1 |
| 005 | 85/F | 26 | 4.7 | ICA | 192 | 94 | 98 | 134 | 66.67 | 67.33 | EVT- | 5 |
| 006 | 83/F | 21 | 6.4 | ICA | 284 | 126 | 158 | 305.36 | 69.85 | 235.51 | EVT+ | 5 |
| 007 | 79/F | 17 | 3.5 | M1 | 88 | 76 | 12 | 67.88 | 43.9 | 23.98 | EVT+ | 4 |
| 008 | 74/M | 20 | 1.0 | ICA | 275 | 95 | 180 | 110.77 | 57.6 | 53.17 | EVT- | 5 |
| Mean |  | 18.6 | 3.2 |  | 169.6 | 78.2 | 91.4 | 139.1 | 47.6 | 91.5 |  | 4.0 |
| ± |  | ± | ± |  | ± | ± | ± | ± | ± | ± |  | ± |
| SD |  | 6.3 | 2.5 |  | 119.5 | 47.3 | 79.7 | 96.6 | 28.5 | 83.0 |  | 1.7 |

MV, mismatch volume; Tmax, Tmax > 6.0 seconds.

Cases are grouped by direction of discordance.

Mean values in each group are provided for comparison.