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Fig. S1. Comparison of the field Biosensor with reference laboratory spectrophotometry A) Scatter plot and B) Bland-Altman plot. Inconsistent results: results with field not similar to laboratory; A) rs = 0.73; p<0.001, n = 120, red lines indicate 4U/gHb and 6U/gHb cut off. B) Mean difference: 0.13 U/gHb, 95% LoA: -4.31 to 4.57 U/gHb (grey shaded area), green dashed line indicates mean difference.
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Fig. S2. Comparison of the field Biosensor with reference laboratory spectrophotometry A) Scatter plot and B) Bland-Altman plot. Inconsistent results: results with field not similar to laboratory; A) rs = 0.80; p<0.001, n = 119, red lines indicate 4U/gHb and 6U/gHb cut off. B) Mean difference: -0.04 U/gHb, 95% LoA: -3.08 to 2.99 U/gHb (grey shaded area), green dashed line indicates mean difference. 
Table S1 Discordant categorical G6PD results by assay

	Sex
	Field Biosensor in U/gHb (category)
	Laboratory Biosensor in U/gHb (category)
	Spectrophotometry in U/gHb (category)

	Male
	5.1 (Normal)
	1.2 (Deficient)
	0.33 (Deficient)

	Female
	5.3 (Intermediate)
	6.6 (Normal)
	6.1 (Normal)

	Female
	5.0 (Intermediate)
	4.9 (Intermediate)
	5.8 (Normal)

	Female
	5.2 (Intermediate)
	7.2 (Normal)
	7.2 (Normal)

	Female
	5.8 (Intermediate)
	10.8 (Normal)
	10.7 (Normal)

	Female
	5.8 (Intermediate)
	8.1 (Normal)
	7.4 (Normal)

	Female
	6.1 (Normal)
	5.7 (Intermediate)
	7.8 (Normal)

	Female
	7.6 (Normal)
	6.2 (Normal)
	4.8 (Intermediate)


Biosensor categories: deficient: <4.1 U/gHb, intermediate: 4.1 to 6.0 U/gHb, normal: >6.0U/gHb
Spectrophotometry categories: deficient: <2.1 U/gHb, intermediate: 2.1 to 4.9 U/gHb, normal: >4.9 U/gHb.
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