Supporting Information for Scientific Reports

A Part Per Trillion Isotope Ratio Analysis of 90Sr/88Sr using Energy-Filtered Thermal ionization mass spectrometry

Shigeyuki Wakaki1†, Jo Aoki2†, Ryoya Shimode2, Katsuhiko Suzuki3†, Takashi Miyazaki4†, Jenny Roberts5, Hauke Vollstaedt5, Satoshi Sasaki6 and Yoshitaka Takagai2,7†*

1 Kochi Institute for Core Sample Research, Japan Agency for Marine-Earth Science and Technology (JAMSTEC), 200 Monobe Otsu, Nankoku, Kochi 783-8502, JAPAN

2 Faculty of Symbiotic Systems Science, Cluster of Science and Technology, Fukushima University, 1 Kanayagawa, Fukushima 960-1296, Japan

3 Submarine Resources Research Center, JAMSTEC, 2-15 Natushima, Yokosuka, Kanagawa 237-0061, Japan

4 Volcanoes and Earth’s Interior Research Center, JAMSTEC, 2-15 Natushima, Yokosuka, Kana-gawa 237-0061, Japan

5 Thermo Fisher Scientific Bremen GmbH, Hanna-Kunath-Str. 11, Bremen 28199, Germany

6 Thermo Fisher Scientific K.K, 4-2-8 Shibaura, Minato-ku, Tokyo 208-0023, Japan

7 Institute of Environmental Radioactivity, Fukushima University, 1 Kanayagawa, Fukushima, 960-1296 Japan.

*To whom correspondence should be addressed: E-mail: s015@ipc.fukushima-u.ac.jp

† These authors contributed equally to this work

Table of Contents

Section											 Page

- Mass spectrum of 88Sr and the effect of RPQ lenses (Figure S1)		  S2

- Mass spectrum in the vicinity of m/z = 90 (Figure S2)				  S3

- Transition of Noise signal during Sr isotope ratio measurement (Figure S3)	  S4

- Measured 90Sr/88Sr ratios of NIST SRM987 (Figure S4)			  S5

- 90Sr activity parameters of the environmental samples (Table S1)	 	  S6

- 90Sr/88Sr ratios of the environmental samples (Table S2)		 	  S7

- Measured 90Sr/88Sr ratios of the environmental samples (Figure S5)		  S8

- Detector configuration of TIMS measurements (Table S3)			  S9

Figure S1. Mass spectrum in the vicinity of m/z 88 showing the flat-top peak of 88Sr measured by the center Faraday cup (black lines). The 88Sr peak tails on the low- and high-mass sides is also shown. Peak tails measured by SEM without the RPQ are shown in blue lines and peak tails measured by SEM with the RPQ are shown in red lines.

Figure S2. Mass spectrum of 90Sr noise observed in 90Sr-free NIST SRM987. The black line represents the 88Sr spectrum as measured by the L1 Faraday cup, and the red line represents the noise spectrum as measured by RPQ-SEM. The signal at m/z 89.7–89.78 corresponds to the recoil of the 88Sr ions in the collector housing. The intensity level of the 88Sr peak tail is represented by signals on m/z 89.78–89.82 and 90.1–90.2. With intensity levels lower than 1cps, the mass spectrum of 88SrH2+ with m/z 89.908 and those of the organics with m/z 90.0 and 90.05 were observed.

Figure S3. The transition of the noise signal intensities on 90Sr during measurement of 90Sr-free NIST SRM987 is plotted against the elapsed time. Small circles are raw data of each measurement cycle and pink diamonds are the average of 10 cycles. Changes in filament temperature during the measurement were plotted together in a red line.

Figure S4. The noise corrected 90Sr/88Sr ratio of NIST SRM987 was measured in (a) Bremen, (b) Fukushima-1, and (c) Fukushima-2 sessions. White diamonds represent a single measurement. The blue line and blue band represent the average and 2 SD range of the multiple measurements. The red dotted line represents the detection limit of 90Sr/88Sr, defined as the 3-standard deviation of the multiple NIST SRM987 measurements.

Table S1. Summary of literature 90Sr activity parameters of the environmental samples1).

| sample | sample type | 90Sr | Sr | status of | reference date |
| --- | --- | --- | --- | --- | --- |
|  |  | [Bq/kg] | [mg/kg] | 90Sr activity | for 90Sr activity |
| NASS-6 | seawater | ― | 7.482) | ― | ― |
| JCp-1 | coral | ― | 7271 (218) | ― | ― |
| IAEA 156 | clover | 14.8 (3.0) | 33.5 (2.0) | recommended value | August 1 1986 |
| IAEA 330 | spinach | 20.1 (4.2) | 44.0 (2.6) | certified value | October 15 2007 |
| Ash 1 | Crayfish | 110 (8) | 1344 (66) | measured value | July 12 2013 |
| Ash2 | Smallmouth bass | 25 (2) | 414 (28) | measured value | July 12 2013 |

- All the errors in the parenthesis are 2SD.

- Wakaki, S.; Obata, H.; Tazoe, H.; Ishikawa, T. Precise and Accurate Analysis of Deep and Surface Seawater Sr Stable Isotopic Composition by Double-Spike Thermal Ionization Mass Spectrometry. Geochem. J. 2017, 51 (3), 227–239. https://doi.org/10.2343/geochemj.2.0461.

Table S2. Detailed results of 90Sr/88Sr measurements of the environmental samples.

| sample | session | reference date | Reference 90Sr/88Sr 1) | 88Sr [V] | 90Sr [cps] 2) | 90Sr/88Sr 2) | number of analysis |
| --- | --- | --- | --- | --- | --- | --- | --- |
| NASS-6 | Fukushima-2 | June 5th, 2021 | ― | 25.3 | 0.0000 (37) | -0.1 (2.5) × 10-12 | 5 |
| JCp-1 | Fukushima-1 | April 3rd, 2021 | ― | 26.9 | 0.0003 (48) | 0.3 (2.9) × 10-12 | 8 |
|  | Fukushima-2 | June 5th, 2021 | ― | 25.3 | 0.0008 (54) | 0.5 (3.4) × 10-12 | 5 |
| IAEA 156 | Fukushima-2 | June 5th, 2021 | 4.4 (0.9) × 10-11 | 26.8 | 0.0546 (51) | 3.26 (0.30) × 10-12 | 6 |
| IAEA 330 | Bremen | January 22nd, 2020 | 7.9 (1.7) × 10-11 | 20.6 | 0.0886 (192) | 6.7 (1.3) × 10-11 | 4 |
|  | Fukushima-2 | June 5th, 2021 | 7.6 (1.7) × 10-11 | 25.9 | 0.0928 (43) | 5.76 (0.43) × 10-11 | 5 |
| Crayfish | Fukushima-1 | April 3rd, 2021 | 1.57 (0.15) × 10-11 | 28.0 | 0.0241 (90) | 1.38 (0.53) × 10-11 | 6 |
|  | Fukushima-2 | June 5th, 2021 | 1.56 (0.15) × 10-11 | 25.7 | 0.0210 (44) | 1.32 (0.42) × 10-11 | 6 |
| Smallmouth bass | Fukushima-1 | April 3rd, 2021 | 1.16 (0.11) × 10-11 | 27.9 | 0.0244 (71) | 1.40 (0.39) × 10-11 | 6 |
|  | Fukushima-2 | June 5th, 2021 | 1.15 (0.11) × 10-11 | 26.3 | 0.0217 (44) | 1.33 (0.27) × 10-11 | 6 |

All the errors in the parenthesis are 2SD.

1) Decay corrected to the reference date.

2) Noise corrected values using the average 90Sr count rate of NIST SRM 987 measurements.

Figure S5. The noise corrected 90Sr/88Sr ratio of the environmental samples. White circles represent a single measurement. Blue lines and blue bands represent the average and 2 SD range of the multiple measurements. Grey square represents the decay corrected reference 90Sr/88Sr ratios.

Table S3 Detector configuration for 90Sr/88Sr measurement

| Collector | L4 | L3 | L2 | L1 | C |
| --- | --- | --- | --- | --- | --- |
| Isotope | 84Sr* | 86Sr | 87Sr | 88Sr | 90Sr |
| Detector type | FC | FC | FC | FC | RPQ-SEM |

* 84Sr is not measured with TIMS TRITON XT at Bremen.