| Supplementary Table 1. Further intra-device agreement results of the devices for each of their illumination levels. |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | AXL WO | AXL WG | PupX 0 | PupX 0.06 | PupX 0.12 | PupX 4 | PupX 64 | MS 0.04 | MS 4 | MS 50 |
| COV | 0.041 | 0.043 | 0.0248 | 0.037 | 0.0295 | 0.0373 | 0.0377 | 0.0261 | 0.0313 | 0.0303 |
| ITT | 0.6156 | 0.36099 | 0.44793 | 0.53298 | 0.41526 | 0.37611 | 0.30942 | 0.44982 | 0.3969 | 0.32886 |
| Intra-session test-retest variability was calculated by multiplying intra-device standard deviation (IDSD) by 2.77 and coefficient of viaration (COV) was calculated by dividing IDSD by the mean of the measurements. (AXL = Pentacam AXL Wave; COV = coefficient of variation; ITT = intra-session test-retest variability; MS = MS-39; PupX = PupilX; WG = with glare; WO = without glare) |  |  |  |  |  |  |  |  |  |  |

| Supplementary Table 2. RM-ANOVA results of three consecutive measurements of the devices for each of their illumination levels. |  |  |
| --- | --- | --- |
|  | P value | Sphericity |
| AXL WO | 0.194 | Assumed |
| AXL WG | 0.143 | Assumed |
| PupX 0 | 0.125 | Huynh-Feldt adjustment |
| PupX 0.06 | 0.002 | Assumed |
| PupX 0.12 | <0.001 | Assumed |
| PupX 4 | <0.001 | Assumed |
| PupX 64 | 0.098 | Assumed |
| MS 0.04 | 0.206 | Assumed |
| MS 4 | 0.293 | Huynh-Feldt adjustment |
| MS 50 | 0.014 | Assumed |
| To additionally compare intra-device measurements, we used 2-factor repeated measurements-ANOVA (RM-ANOVA). If necessary, Huynh-Feldt adjustment was used to correct for violations of sphericity. (AXL = Pentacam AXL Wave; MS = MS-39; PupX = PupilX; RM-ANOVA = repeated measures-analysis of variance; WG = with glare; WO = without glare) |  |  |