Supplementary material

Additional File 1: Description of analytical procedure for each endocrine disruptor for the assay in urine and colostrum. PREVED Project.

We use Ultra High-Performance Liquid Chromatography coupled with tandem Mass Spectrometry to assay BPA and ClxBPA in urine samples according to the method developed and validated by our team [1]. Calibration standards ranged from 0.25 to 16 ng/mL and 0.025 to 1.6 ng/mL for BPA and ClxBPA respectively. The limits of quantification (LoQ) in urine were 0.25 ng/mL and 0.025 ng/mL for BPA and ClxBPA, respectively. During analysis, quality controls were performed at different concentration levels: 0.25, 2 and 8 ng/mL for BPA and 0.025, 0.2 and 0.8 ng/mL for ClxBPA. BPA and ClxBPA in colostrum samples were assayed according to the method developed by our team [2,3]. Calibration standards ranged from 0.4 to 12.8 ng/mL and 0.04 to 1.28 ng/mL for BPA and ClxBPA respectively. The LoQ in colostrum were 0.4 ng/mL and 0.04 ng/mL for BPA and ClxBPA, respectively. Each analytical run included three quality controls prepared in matrices at concentration levels: 0.4, 1.6 and 3.2 ng/mL for BPA and 0.04, 0.16 and 0.64 ng/ml for ClxBPA. For PB determination, an original LC-MS/MS was developed and fully validated to determine MePB, EtPB, PrPB and BuPB presence or concentrations in urine and colostrum. In urine, calibration standards ranged from 0.025 to 3.2 ng/mL. Whatever the paraben, the LoQ in urine was validated at 0.025 ng/mL. During analysis, quality control prepared in matrices were analyzed at different concentration levels: 0.025, 0.05, 0.2, and 1.6 ng/mL. In colostrum, calibration standards ranged from 0.025 to 3.2 ng/mL. The LoQ in colostrum were validated at 0.025 ng/mL for EtPB and PrPB and 0.05 ng/mL for MePB and BuPB. During analysis, quality control prepared in matrices were analyzed at different concentration levels: 0.025, 0.05, 0.4 and 3.2 ng/mL. For both urine and colostrum samples analysis, the limits of detection (LoD) were defined as three times the standard deviation of peak area obtained in the analysis of five blanks for each day of analysis. A compound was considered as detected when the corresponding peak area obtained in a sample was above this value [1]. References [1] Grignon C, Venisse N, Rouillon S, et al. Ultrasensitive determination of bisphenol A and its chlorinated derivatives in urine using a high-throughput UPLC-MS/MS method. Anal Bioanal Chem 2016; 408: 2255–2263. [2] Cariot A, Dupuis A, Albouy-Llaty M, et al. Reliable quantification of bisphenol A and its chlorinated derivatives in human breast milk using UPLC-MS/MS method. Talanta 2012; 100: 175–182. [3] Migeot V, Dupuis A, Cariot A, et al. Bisphenol a and its chlorinated derivatives in human colostrum. Environ Sci Technol 2013; 47: 13791–13797.