Supplementary Table 1: Robustness evaluation of the in-house gE ELISA under different incubation conditions.

Supplementary Table 1A. Experimental conditions used to evaluate robustness of the in-house gE ELISA:

Incubation times of different assay steps were varied around the standard protocol to evaluate the effect of minor deviations in assay conditions on test performance.

| Condition | Sample | mab | Conjugate | Substrate |
| --- | --- | --- | --- | --- |
| Minimum | 18 hour | 55 min | 55 min | 14 min |
| Normal | 22 hour | 60 min | 60 min | 15 min |
| Maximum | 24 hour | 65 min | 65 min | 16 min |

Supplementary Table 1B. Effect of incubation time variations on assay performance: Average blocking percentage values obtained for four reference sera when assay incubation times were varied as described in Supplementary Table S1A.

| Condition | High positive | Moderate positive | Low positive | Negative |
| --- | --- | --- | --- | --- |
| Minimum | 83.72 | 65.12 | 38.31 | 4.93 |
| Normal | 81.62 | 64.34 | 34.61 | 1.12 |
| Maximum | 81.67 | 66.20 | 36.74 | 1.47 |

Supplementary table 2: Diagnostic sensitivity and specificity of the in-house gE ELISA at different candidate cut-off values derived from ROC analysis.

| Target Se | Cut-point | Sensitivity | Se Lower 95% CL | Se Upper 95% CL | Specificity | Sp Lower 95% CL | Sp Upper 95% CL |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0.999 | 2.79 | 1 | 0.954 | 1 | 0.15 | 0.113 | 0.195 |
| 0.995 | 2.79 | 1 | 0.954 | 1 | 0.15 | 0.113 | 0.195 |
| 0.99 | 2.79 | 1 | 0.954 | 1 | 0.15 | 0.113 | 0.195 |
| 0.98 | 6.16 | 0.987 | 0.932 | 0.998 | 0.381 | 0.327 | 0.438 |
| 0.95 | 23.23 | 0.962 | 0.894 | 0.987 | 1 | 0.987 | 1 |
| 0.9 | 53.67 | 0.911 | 0.828 | 0.956 | 1 | 0.987 | 1 |
| 0.8 | 74.85 | 0.81 | 0.71 | 0.881 | 1 | 0.987 | 1 |

| Target Sp | Cut-point | Specificity | Sp Lower 95% CL | Sp Upper 95% CL | Sensitivity | Se Lower 95% CL | Se Upper 95% CL |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0.999 | 22.8 | 1 | 0.987 | 1 | 0.975 | 0.912 | 0.993 |
| 0.995 | 19.7 | 0.997 | 0.981 | 0.999 | 0.975 | 0.912 | 0.993 |
| 0.99 | 19.43 | 0.993 | 0.976 | 0.998 | 0.975 | 0.912 | 0.993 |
| 0.98 | 18.88 | 0.983 | 0.961 | 0.993 | 0.975 | 0.912 | 0.993 |
| 0.95 | 17.67 | 0.952 | 0.922 | 0.971 | 0.975 | 0.912 | 0.993 |
| 0.9 | 16.11 | 0.901 | 0.862 | 0.93 | 0.975 | 0.912 | 0.993 |
| 0.8 | 13.5 | 0.803 | 0.753 | 0.844 | 0.975 | 0.912 | 0.993 |

Supplementary table 3: Results of external proficiency testing for detection of antibodies against BoHV-1.

|  | Sample ID | Sample dilution details | Intended result | PI value | Test result |
| --- | --- | --- | --- | --- | --- |
| Round 1 (14644/SE PT0053 | 21/7541 | G 1/90 | Positive | 36.0 | Positive |
|  | 21/7542 | G 1/30 | Positive | 68.5 | Positive |
|  | 21/7543 | G 1/90 | Positive | 36.8 | Positive |
|  | 21/7544 | Negative | Negative | 7.27 | Negative |
|  | 21/7545 | G 1/60 | Positive | 49.4 | Positive |
|  | 21/7546 | G 1/120 | Positive | 32.5 | Positive |
| Round 2 (14801/SE PT 0053 | 21/7929 | G 1/120 | Positive | 37.8 | Positive |
|  | 21/7930 | G 1/60 | Positive | 54.4 | Positive |
|  | 21/7931 | G 1/30 | Positive | 72.9 | Positive |
|  | 21/7932 | G 1/60 | Positive | 53.8 | Positive |
|  | 21/7933 | Negative | Negative | 1.77 | Negative |
|  | 21/7932 | G 1/90 | Positive | 40.8 | Positive |
| Round 3 (14935/SE PT0053 | 22/7342 | G 1/120 | Positive | 34.3 | Positive |
|  | 22/7343 | G 1/30 | Positive | 67.1 | Positive |
|  | 22/7344 | Negative | Negative | 14.90 | Negative |
|  | 22/7345 | G 1/90 | Positive | 37.4 | Positive |
|  | 22/7346 | G 1/60 | Positive | 43.9 | Positive |
|  | 22/7347 | G 1/120 | Positive | 34.2 | Positive |

Supplementary table 4: Batch-to-batch consistency of recombinant gE antigen and monoclonal antibody reagents used in the in-house ELISA.

Three batches of gE antigen and Two batches of gE mAb were prepared and compared with the current batch for consistency study. 10 sera samples (04 IBR positive and 04 IBR negative, 02 vaccinated negative sera) were evaluated

| Sera | Old Ag-Old mAb | Old Ag-New mAb B1 | Old Ag-New mAb B2 | New Ag B1- Old mAb | New Ag B1-new mAb B1 | New Ag B1-new mAb B2 | New Ag B2-Old mAb | New Ag B2-New mAb B1 | New Ag B2-New mAb B2 | New Ag B3- Old mAb | New Ag B3-new mAb B1 | New Ag B3-new mAb B2 | average | SD | CV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| IBR +ve 1 | 78.7 | 75.3 | 68.9 | 75 | 68.6 | 80.1 | 77.4 | 71.8 | 72.3 | 69.8 | 68.8 | 69.1 | 72.98333 | 4.007666 | 5.491207 |
| IBR +ve 2 | 59.3 | 54.6 | 54.1 | 56.4 | 58.8 | 59.6 | 66.8 | 61.2 | 62.8 | 55.8 | 58.9 | 53.9 | 58.51667 | 3.699061 | 6.321381 |
| IBR +ve 3 | 88.2 | 85.6 | 83.4 | 72.5 | 81.2 | 81.4 | 89.6 | 85.1 | 84.4 | 78 | 82 | 85.7 | 83.09167 | 4.407656 | 5.30457 |
| IBR +ve 4 | 82.7 | 78 | 78.7 | 77.2 | 76.3 | 78.3 | 86.7 | 82.1 | 81.4 | 73.3 | 77.4 | 81 | 79.425 | 3.383077 | 4.259461 |
| IBR -ve 1 | 13.9 | 10.8 | 17.5 | 15.8 | 17 | 17.8 | 18.5 | 3.7 | 18.9 | 14.1 | 17.1 | 13.4 | 14.875 | 4.085773 | 27.46738 |
| IBR -ve 2 | 2.3 | 4.4 | 6.6 | 10.9 | 7.6 | 14.6 | 12.8 | 6.5 | 14.5 | 10.8 | 10.9 | 8.9 | 9.233333 | 3.718049 | 40.26768 |
| IBR -ve 3 | 11 | 11.4 | 6.3 | 10.6 | 13.1 | 21.6 | 11.7 | 5.8 | 8.6 | 10.9 | 13.9 | 11.1 | 11.33333 | 3.868319 | 34.13222 |
| IBR -ve 4 | 3.8 | 9.3 | 0.7 | 5 | 8.2 | 6.8 | 9.7 | 4.7 | 0.1 | 7.8 | 4.3 | 2 | 5.2 | 3.094888 | 59.51708 |
| IBR -ve 5 | 10.2 | 13.6 | 13.5 | 12.5 | 15.6 | 31.4 | 21.5 | 15.4 | 11.7 | 16.4 | 21.5 | 5.9 | 15.76667 | 6.298721 | 39.94961 |
| IBR -ve 6 | 6.6 | 2.2 | 2.8 | 5.8 | 9 | 12.8 | 19.3 | 4.9 | 11.3 | 1.4 | 0.8 | 2.9 | 6.65 | 5.317032 | 79.95537 |