Supplementary figures

Supplementary Figure 1. Fluorescence LAMP amplification plots for a dilution series of papGII-positive E. coli with primers for (A) uidA and (B) papGII. Measurements were taken from three biological replicates. Fluorescence values were normalized to the first time point (fold-change) to reduce baseline differences. Curves are stratified by bacterial load. Solid lines indicate mean, shaded bands indicate  standard deviation.

Supplementary Figure 2. Fluorescence kinetics of LAMP assays for the detection of (A) uidA (E. coli) and (B) papGII in clinical urine samples (n = 191). Fluorescence values were normalized to the first time point (fold-change). Curves are stratified based on the reference method (red = positive, blue = negative; culture for E. coli, WGS for papGII). Within each plot, solid lines indicate mean and shaded bands indicate variability ( standard deviation).

Supplementary tables

| Time [min] | Fluorescence fold-change | Youden’s index score | Sensitivity | Specificity | PPV | NPV |
| --- | --- | --- | --- | --- | --- | --- |
| 39 | 11.40 | 0.82 | 0.95 | 0.87 | 0.64 | 0.99 |
| 30 | 2.88 – 3.25 | 0.82 | 0.87 | 0.95 | 0.80 | 0.97 |

Supplementary Table 2. Diagnostics metrics for the best 2 thresholds ranked by Youden’s index score for uidA (E. coli).

| Time [min] | Fluorescence fold-change | Youden’s index score | Sensitivity | Specificity | PPV | NPV |
| --- | --- | --- | --- | --- | --- | --- |
| 30 (fixed) | 3.00 (fixed) | 0.95 | 1.00 | 0.95 | 0.38 | 1.00 |
| 30 (fixed) | 11.69 – 19.58 | 1.00 | 1.00 | 1.00 | 1.00 | 1.00 |
| 20 - 25 | 3.00 (fixed) | 1.00 | 1.00 | 1.00 | 1.00 | 1.00 |

Supplementary Table 3. Diagnostics metrics for papGII with different thresholds fixed.

| Metric | Equation |
| --- | --- |
| Sensitivity | TPTP+FN |
| Specificity | TNTN+FP |
| Youden’s index score | Sensitivity+Specificity-1 |
| Positive Predictive Value (PPV) | TPTP+FP |
| Negative Predictive Value (NPV) | TNTN+FN |

Supplementary Table 4. Diagnostics metrics used in this study.