Supplementary Information Methods

Kinematic apparatus

Each kick was recorded by two synchronised cameras Casio® Ex FH-20 (sampling frequency of 210 Hz each one), positioned perpendicularly (Figure 1) 0.3 m above the floor and 3.0 m away from the Force Platform (FP). The synchronisation of the two cameras was carried out through a light emitting system.

-------------------------------------------

Figure 1 near here

--------------------------------------------

A spatial three-dimensional (x, y, z) calibrator composed by six plumb lines (three at the front and three at the back, 120 and 70 cm away from each other in the x and y axes respectively) that were hanging from the ceiling, was used to determine an adequate amount of data for capturing the image of the camera during the kick. Cartesian coordinates were used to calibrate the camera images. The two extreme Cartesian coordinate points of each plumb line were selected [53]. Furthermore, a commercial  triaxial acceleration sensor was used to measure various frequencies and acceleration (g) levels.

Kinetic apparatus

A non- commercial FP, designed according to the International Standards of Measurement, was used to capture the CoP data related to the postural stability of the players during the kick. The FP was used to measure the horizontal and vertical components of the GRF, Fx, Fy and Fz (anteroposterior, mediolateral and vertical directions respectively), the coordinates of the CoP, and the moment of force components Mx, My and Mz acting directly on the platform.

Kinematic calibration procedure

The “Av_Rec_2cam.m” routine (developed in the Biomechanics Laboratory, UNESP) was employed to carry out the spatial calibration in Matlab and the analysis of the CoM displacement during the kick. To perform the direct linear transformation (DLT), i.e., transformation of pixels into mm, 57 reference points were used that consisted of spheres attached to the plumb lines (Figure 2). The calibrator images were captured by the cameras, divided into forty frames by Quintic Player® and saved as Matlab files.

-------------------------------------------

Figure 2 near here

--------------------------------------------

The acceleration sensor (Silicon Designs, INC., Model: 2210-025); Capacitive Type; Scale: ± 25 g; Frequency: 0-1000 Hz; Sensitivity: Differential Mode = 160 mV/g, Single-Ended Mode = 80 mV/g; Nominal output voltage = 2.5 V, was attached to the proximal tibial tuberosity beneath all participants’ right lateral malleolus to ensure consistency and repeatability between participants; where the x, y and z axes correspond respectively to the lateral (from right to left), horizontal (from back to front) and vertical (from up to down) directions of the participants’ ankle. A double- sided adhesive tape and elastic straps were used to fix the accelerometer on the participants’ limb to avoid undesired movement due to skin vibration [54].

To enable the measurement of frequencies and acceleration levels, the test accelerometer was coupled directly with the vibration exciter, a portable battery-operated built-in servo-stabilisation installed in a handheld shaker (Sound and Vibration Measuring Instruments).

Supplementary Equations

From the Cartesian coordinates the following equation was used to reconstruct the images:

%-(max.diffdb2+h²)                                                                                                                           (1)

where max. diff is the difference between the farthest and nearest point in the calibrator system; db is the workload volume, and h is the total height. From the Equation 1 all data were expressed in cm; thus, the following values were found:

%-2.7284257.68-1.06%

-------------------------------------------

Table 1 near here

--------------------------------------------

Kinetic calibration procedure

The non- commercial force platform, made of a 500 mm-side length square piece of 10 mm-thick 5052 F aluminium plate, weighing 6.30 kg, was positioned on a solid structural foundation. The material properties of the plate are: Young’s Modulus of Elasticity (E) 7.31 x 1010 N/m²; Poisson’s coefficient (v) 0.3333; Density (p) 2.70 x 10³ kg/m³; Yield Strength (Re) 8.07 x 107 N/m².

The static calibration procedure of FP consisted of positioning standard weights with known mass in the central area of the force platform (Figure 3 a) and in each of the four corners for 10 s per trial.

-------------------------------------------

Figure 3 a near here

--------------------------------------------

A load was applied at four different points to simulate the load of the human feet during stabilometric evaluations. This method is considered the closest to the actual stabilometric measurements [55]. With the four points (P1, P2, P3, and P4) positioned between the platform channels 0,1,2,3 respectively and 100 mm away from the centre of the platform (Figure 3 b), it was possible to measure the output of each load cell and compare it to the weight implemented at each point (P1, P2, P3, P4) separately.

The four load cell signals were processed by an HBM Spider-8 conditioner with a CATMAN software at an acquisition frequency of 100 Hz. The hysteresis of the transducers was obtained through the load overlapping procedure. For both directions, x and y, previously known loads (from zero to 900 N) were applied without disconnecting the system at any time.

Through the vertical and horizontal calibration, the linear adjustment was obtained. This was defined as the division of the upper point of the calibration curve and its respective point of the curve deformation by the full scale of each transducer, equivalent to 900 N. Regarding the horizontal calibration, we used a steel cable attached to a rigid structure fixed to the ground (Figure 3 c). Another cable end was connected to a load cell (Alfa Instruments® with load capacity of 100 Kg and sensitivity of 2 mV/V) fixed at the bottom of the force platform. The cable was tensioned in the horizontal direction (known values from zero to 900 N), allowing the calibration.

--------------------------------------------------

Figure 3 b, c near here

--------------------------------------------------

Upon placing the four load cells on the platform and performing their calibration a dynamic impact analysis was performed, during which an acceleration sensor (Silicon Designs, INC., Model: 2210-025) was fixed on the top of the platform. An impact load was applied directly at the centre of the platform using a rubber mallet to produce the desired thrust. The Fast Fourier Transform (FFT) was calculated, through the Matlab software, from the signals of the force platform to obtain the resonance of the material and the frequency response of the platform for the applied impulse.

Supplementary Equations for the calculation of CoP location

The CoP location was calculated from the signals measured through the force platform in the anteroposterior (AP) and mediolateral (ML) directions. Fz is equal to the sum of vertical reaction forces of each load cell. Considering that Σ Moments = 0 and Σ Forces = 0, the CoP must be located exactly at the centre of the system. The CoP location was calculated by the following equations:

(2)

(3)

where Mx and My (N.m) are x and y torque components, Fx and Fy (N) are x and y force components, Fz is the sum of F1, F2, F3 and F4, h (m) is the height between the top of the plat-form and the strain-gauge, CoPap (m) and CoPml (m) are CoP coordinates on the top of the platform (centred at mid-distance from the four sensors) with the x and y axes as previously defined. These equations are applied to the analysis of systems with four uniaxial load cells arranged in a rectangular shape [55]. The usable area of the force platform corresponds to the rectangular area that has the centres of the four load cells as corners.

Equations of the force elements

VMx =A+B-C-D

VMy =A-B-C+D                                                                                                                     (4)

VFz =A+B+C+D

where A =F1, B =F2, C =F3, D =F4; Fz is the sum of F1, F2, F3 and F4.

The values of loading and unloading of each force and moment element for each axis were treated with linear regression and the average value of each trial was calculated and adopted as a reference. The force elements (Mx, My and Fz) were calculated, according to Equation 4 [56,57].

This procedure was adopted to analyse the force and moment elements for each axis and to calculate the parameter or angular coefficient [56,57], enabling the construction of both the calibration and the sensitivity matrices. Furthermore, the sensitivity and hysteresis of the z axis and the moments Mx and My were calculated. The matrix of calibration was based on Equation 5 and on the calculation of the slopes of the linear regression curves  [56,57].

(5)

References

[53] Lafontaine, D., & Lamontagne, M. (2003). 3-D Kinematics using moving cameras. Part 1: Development and validation of the mobile data acquisition system. Journal of Applied Biomechanics, 19, 372-377. https://doi.org/10.1123/jab.19.4.372.

[54] Avsar-Arpinar., P., &  Soylu., A. R. (2010). Consistency in acceleration patterns of football players with different skill levels. Journal of Sports Science & Medicine,  9(3), 382-387.

[55] Dias, J. A., Borges, L., Mattos, D. J. S., Wentz, M. D., Domenech, S. C., Kaufmann, P., & Borges Gomes Junior, N.  (2011). Validity of a new stabilometric force platform for postural balance evaluation. Revista Brasileira de Cineantropometria e Desempenho Humano, 13(5), 367-372. https://doi:10.5007/1980-0037.2011v13n5p367.

[56] Gurgel, J. L, Porto, F., Castro, L., Russomano, T., Beck, J., & Schroeder, I. (2005). Desenvolvimento e avaliação de um procedimento de calibração para pedal sensor de forças e momentos. In: Anais do XI Congresso Brasileiro de Biomecânica, João Pessoa, Brazil (pp.19-22).

[57] Gurgel, J. L., Russomano, T., Porto, F., Cambraia, R., Azevedo, D. F. G., Glock, F. S., Pinheiro, B. J. C., & Heledga, S. (2006). Development and calibration of a pedal with force and moment sensors. In IEEE: Proceedings of the 28th International Conference of the Engineering in Medicine and Biology Society (pp. 4144-4146). New York City, USA: IEEE-EMBS. https://doi:10.1109/IEMBS.2006.4398363.