ADDITIONAL FILE 1

Retained methodological specification and aggregate exploratory machine-learning results

Associated manuscript: A wearable multi-task prefrontal functional near-infrared spectroscopy methodology for studying mild traumatic brain injury: a human-subject pilot implementation

Purpose. Part I consolidates technical details preserved in the retained manuscript and thesis documentation. Part II preserves the aggregate machine-learning results for transparency. This combined file does not reconstruct missing participant-level data, executable code, quality-control decisions, fold assignments, or the complete 864-column feature dictionary.

# Part I. Retained methodological specification

# Table M1. Wearable fNIRS system specification

| Parameter | Retained specification |
| --- | --- |
| Modality | Continuous-wave fNIRS |
| Source module | Marubeni America L4*730/4*805/4*850-40Q96; 730 and 850 nm used |
| Reported optical output | Approximately 2-4 mW average output; retained document reports 7-V drive |
| Detector | Texas Instruments OPT101 monolithic photodetector with integrated operational amplifier; photosensitive area 2.29 × 2.29 mm² |
| Source-detector separation | 30 mm, fixed |
| Channels and locations | Four channels, approximately F7, Fp1, Fp2, and F8 (international 10-20 system) |
| Sampling rate | 17 Hz |
| Controller and data link | ESP32 microcontroller; Bluetooth transmission to laptop |
| Host software functions | Connection control, source-intensity adjustment, real-time signal display, task selection/control, and CSV storage with task timing |
| Headform | Padded modified commercial VR headset with adjustable rear dial and elastic straps; exact base model, dimensions, weight, and setup time not retained |
| Power | Mobile power supply through micro-USB, 5 V; retained documentation also identifies HW-183 boost and TLC5916 driver modules |
| Not retained | ADC resolution, absolute calibration protocol, quantitative coupling comparison, channel-yield criteria, artifact burden, and setup time |

# Figure M1. Retained wearable system layout

Retained engineering image showing the modified VR-headset structure, four 30-mm source-detector channels, and approximate prefrontal locations. The structure was designed to support fit and contact; no comparative artifact-reduction or signal-quality claim is made.

# Table M2. Programmed four-task cognitive protocol

| Task | Programmed timing | Stimulus and response details |
| --- | --- | --- |
| VFT | 30-s eyes-closed rest; 60-s task; 70-s eyes-closed recovery (160 s) | Three 20-s overt phonetic-cue periods. One cue was selected from each of three retained Bopomofo phonetic pools. Participants generated as many words as possible; a 3-s auditory warning preceded the task. |
| CT | Five rounds × 60 s: 5-s eyes-open rest; 30-s task; 5-s reinforcement; 20-s eyes-closed recovery (300 s) | An airplane moved across the screen; its height was mapped to frontal ΔHbO2. A rise threshold of 0.25 μM above the resting mean caused ascent. Participants attempted to maintain the airplane in the upper third. A green circle or red cross provided reinforcement; a 3-s piano cue marked the end of recovery. |
| MAT | 30-s eyes-closed rest; 180-s task; 60-s eyes-closed recovery (270 s) | Three-digit ± two-digit multiple-choice arithmetic. Each item comprised 1-s display, 3-s thinking/countdown, 1-s left/right response, and immediate green-circle, red-cross, or question-mark feedback. |
| SSST | Six rounds × 50 s: 5-s fixation; 30-s task; 15-s eyes-closed recovery (300 s) | Participants clicked randomly positioned numbered targets in order. The first two rounds used 4 and 6 targets; the remaining four used 8. Correct selections produced an auditory cue; repeated or incorrect selections were logged. A 3-s piano cue marked the end of recovery. |
| Session total | 1030 s (approximately 17.2 min), excluding setup and transitions | Retained documentation describes the sequence VFT, CT, MAT, and SSST. Setup time, inter-task transition time, tolerability measures, and complete behavioral-performance data were not retained. |

# Table M3. Stage-aware processing and representation workflow

| Step | Retained method | Reproducibility boundary |
| --- | --- | --- |
| 1 | Raw optical intensity → optical density | Exact software implementation not retained |
| 2 | TDDR motion correction [1] applied to optical density | Quality-control thresholds and channel-rejection decisions not retained |
| 3 | Modified Beer-Lambert law → ΔHbO2 and ΔHHb; 30-mm distance; DPF 6.0172 at 730 nm and 4.9399 at 850 nm | Extinction-coefficient table and exact code not retained |
| 4 | Fourth-order Butterworth low-pass at 0.1 Hz using forward-backward zero-phase filtering | No separate high-pass filter is documented |
| 5 | Baseline correction by subtracting the mean of the corresponding pre-task rest interval | Defined in retained documentation |
| 6 | Min-max normalization: (x − xmin)/(xmax − xmin), with extrema described across the analysed time series | Exact participant/channel/task grouping used by the original software not retained |
| 7 | Reconstructable vector-derived signals [2,3]: ΔCBV=(ΔHHb+ΔHbO2)/√2; ΔCOE=(ΔHHb−ΔHbO2)/√2; ¦L¦=√[(ΔCOE)²+(ΔCBV)²]. The retained thesis also records Angle K=arctan(ΔCOE/ΔCBV). | Angle K cannot be fully reconstructed: the positive-ΔO reference-axis conversion, quadrant handling, phase assignment, and behavior at ΔCBV=0 were not preserved. |
| 8 | Stage segmentation: rest, task, reinforcement where applicable, and recovery | Programmed timing retained in Table M2 |
| 9 | Documented feature families were applied to ΔHbO2, ΔHHb, ΔCBV, ΔCOE, ¦L¦, and the legacy Angle K feature across four channels. | The original manuscript reports 864 candidates, but the Angle K implementation and a one-to-one 864-column map cannot be reconstructed. |
| 10 | Secondary exploratory univariate screen and classifier evaluation | Participant-level data, executable pipeline, seeds, folds, grids, versions, and fully nested validation unavailable |

# Reproducibility statement

The retained records permit partial methodological reconstruction of the acquisition hardware, programmed task timing, principal signal-processing sequence, and the ΔCBV, ΔCOE, and |L| equations. They do not permit reconstruction of Angle K phase handling or computational reproduction of the participant-level machine-learning analysis. The feature count of 864 is preserved as an aggregate result and is not presented as a reconciled column-level dictionary.

# References supporting the retained processing description

1. Fishburn FA, Ludlum RS, Vaidya CJ, Medvedev AV. Temporal derivative distribution repair (TDDR): a motion correction method for fNIRS. NeuroImage. 2019;184:171-179.

2. Oka N, Yoshino K, Yamamoto K, et al. Greater activity in the frontal cortex on left curves: a vector-based fNIRS study of left and right curve driving. PLoS One. 2015;10(5):e0127594.

3. Nazeer H, Naseer N, Khan RA, et al. Enhancing classification accuracy of fNIRS-BCI using features acquired from vector-based phase analysis. J Neural Eng. 2020;17(5):056025.

# Part II. Retained aggregate exploratory machine-learning results

The following results are preserved from the original exploratory analysis. Feature and model selection was not fully nested, the holdout set contained only four participants, and participant-level scores were unavailable for reanalysis. These results do not establish model superiority, diagnostic validation, or incremental value beyond demographics.

# Table S1. Task-wise retained descriptive performance

| Task | Retained model | Train (%) | Validation (%) | Holdout (%) |
| --- | --- | --- | --- | --- |
| VFT | SVM | 100 +/- 0 | 67 +/- 17 | 75 |
| CT | SVM | 100 +/- 0 | 78 +/- 5 | 50 |
| MAT | XGB | 90 +/- 5 | 67 +/- 8 | 75 |
| SSST | MLP | 85 +/- 12 | 81 +/- 13 | 50 |
| VFT, CT, MAT | KNN | 99 +/- 2 | 72 +/- 5 | 50 |
| VFT, CT, MAT, SSST | MLP | 93 +/- 9 | 75 +/- 0 | 75 |

All holdout percentages are based on four participants. Model and task-combination selection was not nested within an independent evaluation framework. The retained record did not specify whether the “+/-” values represented standard deviation or another dispersion measure; they are reproduced descriptively as originally reported.

# Table S2. Retained combined-task MLP internal-validation results

| Fold | Accuracy (%) | Precision (%) | Recall (%) | F1-score (%) | AUC |
| --- | --- | --- | --- | --- | --- |
| 1 | 75 | 83 | 75 | 73 | 0.75 |
| 2 | 75 | 76 | 75 | 75 | 0.69 |
| 3 | 75 | 83 | 75 | 73 | 0.81 |

Each fold contained 12 participants (6 control and 6 mTBI); 9/12 were correctly classified in each fold (exact 95% CI for accuracy, 42.8%-94.5%). Pooled out-of-fold accuracy was 27/36 (75%; exact 95% CI, 57.8%-87.9%). Precision, recall, and F1-score in Table S2 are macro-averaged across the two classes, as reconstructed from the retained confusion matrices.

# Table S3. Retained aggregate combined-task MLP performance

| Dataset | Accuracy (%) | Precision (%) | Recall (%) | F1-score (%) |
| --- | --- | --- | --- | --- |
| Train | 93 +/- 9 | 93 +/- 8 | 93 +/- 9 | 93 +/- 9 |
| Validation | 75 +/- 0 | 75 +/- 0 | 75 +/- 0 | 74 +/- 1 |
| 4-participant holdout | 75 | Not reconstructable | Not reconstructable | Not reconstructable |

The holdout accuracy was 3/4 (75%; exact 95% CI, 19.4%-99.4%). Holdout precision, recall, and F1-score could not be reconstructed and are therefore not reported. For the train and validation rows, the retained record did not specify whether the “+/-” values represented standard deviation or another dispersion measure; they are reproduced descriptively as originally reported.

# Figure S1. Retained fold-level confusion matrices

Aggregate confusion matrices for the three original internal-validation folds, relabelled using “Control” rather than “Normal.” Each fold contained 12 participants and 9/12 correct classifications. Retained fold-specific AUC values reported in the original analysis were 0.75, 0.69, and 0.81; these values could not be independently reconstructed because participant-level scores were unavailable. ROC curves were therefore removed.