4035 | 1220648-2025-47624 | 23364699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23364699&pc=OZD | 09/22/2025 | 11/20/2025 |  | 10/23/2025 |  |  | 09/22/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723020 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was prepped for impella support due to asystole. during preparation, the impella was started outside of the patient¿s body. an impella defective alarm was displayed. a second pump was opened and the same message was displayed. a new aic was turned on, the second pump was connected with no alarms, and the case proceeded. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was inadvertently captured under the wrong product complaint. the event associated with this device is captured under mwr-19112025-0001976792.
4036 | 1220648-2025-47720 | 23364715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23364715&pc=OZD | 09/22/2025 | 12/05/2025 |  | 10/23/2025 |  |  | 09/22/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729400 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Inflammation | Pain | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had pain, inflammation and fever during impella 5.5 patient support. while on support, it was noted that there were still concerns over right upper extremity (rue) swelling and pain. there was a plan for an ultrasound as well as an arterial and venous duplex of rue and chest area. the patient spiked a small temperature and started on antibiotics. the rue picc was removed as well. | [5 empty] | The investigation for the reported pain/ fever/ inflammation has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of pain/ fever/ inflammation could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4037 | 1220648-2025-47721 | 23366363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366363&pc=OZD | 09/26/2025 |  |  | 10/23/2025 |  |  | 09/26/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724354 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced sudden nausea and vomiting at home and then became unconscious. emergency medical services (ems) was called, and it was noted that the patient had a down time of five minutes before ems arrived. the patient was found to be in ventricular fibrillation arrest and cardiopulmonary resuscitation was initiated and the patient was intubated. the patient was transferred to the hospital in cardiac arrest. upon arrival to the hospital, the patient was found to be in st elevation myocardial infarction. the patient was emergently taken to the cardiac catheterization laboratory where the impella cp was implanted without complications. the patient had a successful percutaneous coronary intervention performed with two stents placed in the left anterior descending coronary artery. a swan was attempted to be placed, however due to recurring ventricular tachycardia, it was aborted. the patient was sent to the unit to rest and recover. later the same day, the patient¿s family decided to withdraw care. care was withdrawn and the patient expired. | [5 empty] | Medical safety reviewed the event. there is no clear evidence to dissociate the ventricular tachycardia (vt). this event occurred during impella use and although it is believed the vt to be from a difficulty swan placement versus the impella, this cannot be conclusively determined. this is a death type of reportable event as we cannot dissociate (not that the impella contributed to the demise outcome given the details of the event). the patient's underlying condition contributed to the patient's outcome of death. the patient never recovered from the cardiac event that occurred 13 hours prior. family members determined to stop all care, which was withdrawn and this led to the patient expiring. investigation summary: the investigation has been completed. no product was returned for investigation. arrhythmia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. h6 type of investigation, investigation findings, conclusion codes and component code were updated accordingly based on the completed investigation. | The reported ventricular tachycardia reported while the patient was on device support. at this time, there is no evidence of a direct causal relationship between the device and the reported arrhythmia, or the patient outcome. however, based on the available information, a device contribution cannot be ruled out. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4038 | 1220648-2025-47722 | 23366511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366511&pc=OZD | 09/04/2025 |  |  | 10/23/2025 |  |  | 09/04/2025 | 10/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683596 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the 5.5 was advanced through the graft and into axillary artery but the team was unable to advance the pump past innominate artery. after multiple attempts by the doctors to advance the pump, the decision was made to place impella cp via axillary cutdown. the placement of 5.5 was aborted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. the patient was a 73-year-old with a history of peripheral artery disease, coronary artery disease, hypertension, hyperlipidemia, and active tobacco use who presented with exertional chest pain. a nuclear stress test was positive, and cardiac catheterization revealed an 80% occlusion of the left anterior descending (lad) artery. during percutaneous coronary intervention (pci), a lad dissection occurred, and the patient went asystole requiring approximately two minutes of cpr, after which return of spontaneous circulation was achieved. the patient was intubated, and three stents were placed in the lad; however, the dissection could not be fully covered. the patient remained hemodynamically unstable with ongoing cardiogenic shock requiring high-dose vasopressors and was taken emergently for coronary artery bypass grafting (cabg) with planned impella support (indication: pccs/lcos). an unsuccessful attempt was made to implant an impella 5.5 via right axillary artery. after multiple attempts, the impella 5.5 could not be advanced past the innominate artery, and the procedure was aborted. an impella cp was subsequently implanted via axillary cutdown and was successfully placed. impost impella cp implant same day, the patient experienced a major bleeding event with significant chest tube output, requiring transfusion of approximately 12 units of blood products intraoperatively. over the subsequent days, the patient developed atrial fibrillation with rapid ventricular response, acute kidney injury requiring nephrology consultation and planned continuous renal replacement therapy (crrt), white blood cells elevated with concern for sepsis, and thrombocytopenia necessitating interruption of heparin therapy. the patient remained critically ill, intubated, and on combined mechanical circulatory support including impella (p4) and veno-venous extracorporeal membrane oxygenation (vv ecmo). it was noted despite ongoing support; there was no meaningful clinical improvement. a decision was made to withdraw life-sustaining support leading to the patient's subsequent demise. h6. investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | H6: updated codes for health effect impact code, type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy at innominate artery preventing successful delivery of impella. device history review device passed all post sterile inspection checks.
4039 | 1220648-2025-47723 | 23366522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366522&pc=OZD | 09/25/2025 |  |  | 10/23/2025 |  |  | 09/25/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682303 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the patient was in septic shock. no other information was provided. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the sepsis, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the sepsis was unable to be determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4040 | 1220648-2025-47724 | 23366633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366633&pc=OZD | 09/17/2025 |  |  | 10/23/2025 |  |  | 09/05/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662801 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 141 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Material Split, Cut or Torn | [5 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Necrosis | Thrombosis/Thrombus | Hemodynamic instability | [7 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported impella 5.5 was successfully implanted across the aortic valve (5.2 cm depth), pointed toward apex. initial flow setting started at p-2 and ramped to p-4. upon cardiovascular intensive care unit (cvicu) arrival, flow was only 0.8 l/min at p-3. while on support, the patient experienced ongoing chest tube bleeding, requiring overnight surgical washout and leaving the chest temporarily open. the patient experienced intermittent ectopy. volume was given to the patient, electrocardiogram machine confirmed patient in atrial fibrillation (afib) with left axis deviation (l-axis deviation). additionally, a gastrointestinal bleed was treated with 1 unit of packed red blood cells intraoperatively. imaging raised concern for hemorrhagic pancreatitis, and later revealed abdominal hematoma and necrotizing hepatitis. product sleeve damage was also reported. plan to withdraw care was discussed. extracorporeal membrane oxygenation circuit clotted off and patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product returned for investigation. the cause of the low pump flow was patient condition related as the patient was also on va ecmo and blood administration was improving flows and hemodynamics. the cause of the anticontamination sleeve damage was not determined since insufficient clinical details were provided and no product was returned. the cause of the arrythmia and thrombosis was unable to be determined due to insufficient clinical details. the cause of the hemodynamic instability, major bleeding, hematoma, and necrosis cannot be determined as insufficient clinical details were provided. the device passed all post-sterile inspection checks.
4041 | 1220648-2025-47725 | 23366726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366726&pc=PYX | 09/08/2025 |  |  | 10/23/2025 |  |  | 09/08/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025533985 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced lower than expected flows. the pump was explanted and fibrin was found in the catheter outflow of the catheter. the patient survived. | [5 empty] | The root cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the low pump flow was most likely biomaterial ingestion, as clinical details reported finding thrombus in the pump, and data logs showed signs of ingestion, but the cause of the biomaterial could not be determined due to insufficient clinical details. the device passed all post sterile inspection criteria. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4042 | 1220648-2025-47728 | 23366960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366960&pc=OZD | 09/18/2025 |  |  | 10/23/2025 |  |  | 09/18/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645708 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | False Alarm | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp encountered intermittent placement signal low alarms and pump in aorta alarms. an ultrasound scan and an echocardiogram confirmed the impella pump position was normal. in addition to the alarms, there was frequent suction above p2. the pump was exchanged for an impella 5.5 to continue patient support. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Because the product was not returned and insufficient clinical details were provided, the cause of the optical signal issue and the suction/false alarms could not be determined. the device passed all post-sterile inspection checks.
4043 | 1220648-2025-47729 | 23366997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23366997&pc=OZD | 09/18/2025 |  |  | 10/23/2025 |  |  | 09/18/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811668 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 71 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 and supported with an automated impella controller experienced placement and optical signal discrepancies. no patient harm was reported. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 and h6 updated.
4044 | 1220648-2025-47727 | 23367099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367099&pc=OZD | 09/10/2025 |  |  | 10/23/2025 |  |  | 09/10/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026707216 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 87 KG | Non Hispanic | White | Injury | Yes | No | False Alarm | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  | Anemia | Atrial Fibrillation | Fever | Hemolysis | Unspecified Infection | Hematuria | Renal Impairment | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had severe right ventricle dysfunction, acute tubular necrosis with dialysis, and unknown negative gram bacteremia from an unknown source prior to and during pump support. the patient experienced position unknown alarms and suction alarms with higher p-level support. the patient presented with fever requiring antibiotics. hemolysis was suspected, and the patient had atrial fibrillation, dark stools, and was anuric. multiple units of packed red blood cells were transfused in response to low hemoglobin. the patient was successfully weaned from impella support and survived. | [5 empty] | No product or logs were returned for evaluation. the cause of the fever, infection, an atrial fibrillation was unable to be determined due to insufficient clinical details. the cause of the kidney failure, hematuria, and anemia was not determined due to insufficient clinical details. the cause of the false alarms could not be determined due to no product or logs returned and limited clinical details. the cause of the pump suction was most likely right ventricle dysfunction and inotropes and p-levels were balanced to resolve suction alarms. the cause of the hemolysis could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. the device passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4045 | 1220648-2025-47731 | 23367524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367524&pc=OZD | 09/10/2025 | 10/23/2025 |  | 10/23/2025 |  |  | 09/12/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026707226 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Atrial Fibrillation | Fever | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [7 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. there was a small hematoma which developed at the axillary cutdown site, this was managed with manual pressure. as such anticoagulation was held over night. patient¿s urine turned pink/red overnight. staff stating that the catheter was ¿adjusted¿ overnight and then urine turned red after. the patient received diuretics and during the check-in, the urine cleared up. physicians are not convinced that it is true hemolysis as there were no alarms on automated impella controller and all labs are resulting. patient shocked last night due to patient going into atrial fibrillation, resolved with cardioversion. white blood cell continues to rise. staff stating ¿clots¿ in urine, urinalysis showed red blood cells. the patient talked with the doctors and nurses. he made the decision to withdraw care and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. product was not returned for review. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal ps or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. atrial fibrillation/thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4046 | 1220648-2025-47733 | 23367789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367789&pc=OZD | 09/23/2025 |  |  | 10/23/2025 |  |  | 09/23/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584647 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had hemolyzed labs and urine after foley placement. in response, the pump was repositioned under echocardiogram. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The product was not returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the device passed all post sterile inspection checks.
4047 | 1220648-2025-47732 | 23367804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367804&pc=OZD | 09/10/2025 |  |  | 10/23/2025 |  |  | 09/10/2025 | 10/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701876 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient placed on continuous renal replacement therapy due to renal failure and treated with multiple medications while on impella cp support. it was also reported that the patient had paroxysmal junction rhythm. the patient was escalated to impella 5.5 for continued support. this report is for the fist pump (serial number (b)(6)). there will be additional reports for this patient. pump 2: impella 5.5 smartassist, (serial number (b)(6)). pump 3: impella rp flex smartassist, (serial number (b)(6)). automated impella controller, (serial number (b)(6)). | [5 empty] | E1 revised initial reporter facility name as it was entered incorrectly on the initial report that was submitted. h6 added b13 code as it was omitted from the initial report that was submitted. h10 added related report number. investigation summary: the investigation has been completed. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | Medical safety/clinical reviewed the event. a 70-year-old female with a medical history of persistent atrial flutter, paroxysmal atrial fibrillation, mixed hyperlipidemia, hypertension, and morbid obesity presented to the emergency department with several weeks of fatigue, dizziness, palpitations, and associated shortness of breath. the patient was newly diagnosed with heart failure with reduced ejection fraction (hfref) with a severely reduced ejection fraction of 10¿15% and severe global hypokinesis. the patient underwent left heart catheterization for ischemic evaluation, which demonstrated two-vessel coronary artery disease involving the mid left anterior descending artery and right coronary artery. right heart catheterization demonstrated worsening lactic acidosis, with lactate levels trending up to 12.9 mmol/l. the advanced heart failure team was consulted, and the patient was initiated on left-sided impella support, which was subsequently escalated to biventricular (bipella) support with impella devices: impella 5.5 and rp flex. during impella cp support, the patient developed acute renal failure requiring continuous renal replacement therapy and was treated with multiple medications. paroxysmal junctional rhythm was also reported. the patient was escalated to an impella 5.5, and bipella support was initiated with the addition of an impella rp flex. while on bipella support, the patient experienced multiple adverse clinical events. additionally, the automated impella controller (aic) connected to the impella rp flex displayed a frozen screen with a persistent ¿complete the procedure¿ alarm. replacement of the purge cassette did not resolve the issue. the aic was exchanged with no noted consequence to the patient, and the patient continued to receive support without further interruption. after approximately eight days of mechanical circulatory support, care was withdrawn, and the patient expired. h6. investigation type/findings/conclusion remain unchanged related medical device reports: this is one of four devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4048 | 1220648-2025-47403 | 23367821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367821&pc=OZD | 09/10/2025 | 10/10/2025 |  | 10/23/2025 |  |  | 09/10/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675449 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during impella cp support the patient's movement triggered an alarm within the aortic pump. the catheter fell out of the aorta. upon checking the pump position, an ultrasound scan revealed a left ventricular thrombus. impella reinsertion was abandoned and the pump was replaced with an intra-aortic balloon pump. patient is stable post explant. | [5 empty] | Correction b1 the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position during patient movement. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4049 | 1220648-2025-47734 | 23367933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367933&pc=OZD | 09/10/2025 | 01/07/2026 |  | 10/23/2025 |  |  | 09/10/2025 | 10/23/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738858 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | Cardiogenic Shock | Hematuria | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | A notification received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry noted additional patient conditions. | The us complainant reported that a patient implanted with an impella cp experienced hematuria. the patient was escalated to an impella 5.5 and survived. |  |  |  |  | B5 updated. h6 health effect - clinical codes added - anemia, hemolysis, cardiogenic shock. the investigation is still ongoing. | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. investigation summary: hematuria : the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The device or logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hematuria : the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the hematuria was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
4050 | 1220648-2025-47735 | 23367934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23367934&pc=OZD | 10/01/2025 |  |  | 10/23/2025 |  |  | 10/01/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544167 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 89 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Cardiac Arrest | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant placed this second impella cp after the first had persistent suction. this pump too developed suction that could not be resolved and the patient coded and expired after just under 13 hours of support. this report represents the second impella cp. another report was submitted to represent the first pump. refer to section h.10 for the related report number. | [5 empty] | Impella cp was returned for evaluation. pump suction/low pump flows: clinical details report that the pump experienced suction while at p-6. data logs were reviewed and the suction was confirmed. returned product was visually inspected and there was no damage to the impeller or biomaterial. there was a minor kink noted in the cannula near the radiopaque marker. the pump was run in a bucket of water and low flows did not reproduce. the cannula kink is a potential cause for the suction and low flows, however, clinical details report that there was no cannula kink on imaging, so it could not be confirmed. since data log review was inconclusive, returned product did not present any definitive cause, and limited clinical details were provided, the cause of the low pump flows and suction could not be determined. health effect - clinical code was updated from e0601 to e0602. | Medical safety/clinical review: medical safety/clinical reviewed the event. the patient was a 77-year-old female with a past medical history of hypertension, hyperlipidemia, and congestive heart failure. she presented to the emergency department with a non-st-elevation myocardial infarction (nstemi) and was noted to have mildly elevated troponin levels and a lactate of 3.2 mmol/l. on the following day, the patient was taken to the cardiac catheterization laboratory for left heart catheterization. after administration of sedation, the patient became profoundly hypotensive. the decision was made to initiate mechanical circulatory support. left ventriculography demonstrated an estimated ejection fraction of approximately 25%, and severe multivessel coronary artery disease was identified. the patient was referred and accepted for surgical intervention at a later date. an impella cp was placed for mechanical circulatory support. following placement, persistent suction alarms were noted that could not be resolved with repositioning or fluid administration. the treating physician suspected possible clot ingestion. the impella cp (pump 1) was subsequently removed and replaced with a second impella cp (pump 2). approximately 12 hours later, csc was contacted to assist with recurrent suction alarms. abrupt suction events were observed, with motor current flattening and flows decreasing to approximately 0.1 l/min. the performance level was reduced from p-6 to p-2, with some pulsatility noted on motor current; however, no improvement in flows was observed. the patient was anticoagulated but not at therapeutic levels at that time. based on the clinical presentation, clot ingestion was again suspected. an echocardiogram was ordered to confirm device position. subsequently, motor current flattened again and flows abruptly decreased to approximately 0.1¿0.3 l/min. suction could not be resolved despite additional troubleshooting measures. following these events, the patient experienced cardiac arrest. advanced cardiac life support was initiated; however, the family elected to discontinue resuscitative efforts and transition the patient to comfort care. the patient continued to deteriorate and subsequently expired. the impella cp pump 1 is a serious injury, and the impella cp pump 2 is a death type of reportable event. h6 investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4051 | 1220648-2025-47736 | 23368078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23368078&pc=OZD | 08/13/2025 |  |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025581115 |  | 000000000001000459 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant implanted impella cp for a patient experiencing cardiogenic shock, subsequently escalating to the impella 5.5 pump for a patient with biventricular failure. the 5.5 provided support for a duration of 56 days until a signal loss was detected, which led the team to explant the device and replace it with another 5.5 pump. no harm or injury was observed. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D6b: corrected explant date.
4052 | 1220648-2025-47737 | 23368334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23368334&pc=OZD | 08/11/2025 |  |  | 10/23/2025 |  |  | 10/09/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1317752 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 99 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during an impella cp implant to a patient with acute myocardial infarction/cardiogenic shock for mechanical circulatory support (mcs) the impella cp device was unable to connect to the automated impella controller (aic) despite multiple attempts. during investigation of the impella pump, the aic was identified as the issue. per the reporter, the impella cp device was connected according to standard practice. once the pump was connected yellow to yellow and was primed, the white cable was connected to the aic. however, the aic did not detect the pump. several disconnects with reconnects were executed without success. care was taken to assess the connection cable for moisture as well as the pump terminal, and both were dry, appeared to be in standard configuration. the case was cancelled and restarted. the purge tubing was reinserted; however, again the pump was not recognized. consequently, the impella cp device was replaced, connected to the aic and was recognized in standard fashion without an issue. there was no patient harm as a result of this event. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. related medical device reports: this automated impella controller report is one of two devices associated with the event. refer to section h.10 for the related report number that represents the pump.
4053 | 1220648-2025-47738 | 23368555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23368555&pc=OZD | 09/11/2025 |  |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103522 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The reporter stated that the automated impella controller (aic) was unable to detect a purge disc during prep, which was resolved by replacing the aic. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4054 | 1220648-2025-47739 | 23368607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23368607&pc=OZD | 09/25/2025 |  |  | 10/23/2025 |  |  | 09/25/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732851 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | Multiple Organ Dysfunction Syndrome | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient was placed onto impella cp support (left ventricle) due to acute myocardial infarction / cardiogenic shock. while on support, patient experienced right ventricular dysfunction and was implanted with impella rp in the right ventricle. the patient eventually developed flash pulmonary edema with rapid cardiopulmonary decompensation. family withdrew care and patient expired. this report is for the impella cp (serial number: (b)(6)) there is an additional report for the impella rp (serial number: (b)(6)). | [5 empty] | B.2 added selection that was omitted from the initial report that was submitted. d.4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. h.6 added code b13 as it was omitted from the initial report that was submitted. medical safety reviewed the event. there is not enough information to dissociate the impella devices used for bipella support as an associated factor to the patient's outcome. however, in this case, patient underlying condition likely contributed most to an outcome of death. care was withdrawn which led to the patient expiring. investigation summary: the investigation has been completed. no product was returned for investigation. pulmonary edema/ cardiac failure/ organ failure: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passes all post sterile inspection checks. h6 type of investigation, investigation findings, conclusion codes and component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella rp flex. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4055 | 1220648-2025-47741 | 23369000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369000&pc=OZD | 09/25/2025 |  |  | 10/23/2025 |  |  | 09/25/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026742660 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during insertion of impella cp, that there was no anticoagulation, due to suspected heparin-induced thrombocytopenia. the day after observed increased lactate dehydrogenase and urine turning red/pink on rounds. point-of-care ultrasound with device deep (~6cm). retracted and slack removed to 4cm from aortic valve with poor windows. team to evaluate hemolytic profile. not a candidate for surgical revasc or percutaneous coronary intervention. plan to wean to explant over weekend and proceed with home hospice. urine "looks like it is clearing". the pump was explanted the next day. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were placement signal (ps) alarms triggered in early of the case but resolved quickly. there were no motor current (mc) spikes or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in proper position since it was confirmed pump was deep and hemolysis was improved after positioning of the pump. the cause of positioning issue was unknown. thrombocytopenia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4056 | 1220648-2025-47743 | 23369636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369636&pc=OZD | 09/27/2025 |  |  | 10/23/2025 |  |  | 09/27/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 69 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on day three of impella cp support experienced atrial fibrillation with rapid ventricular response. amiodarone bolus was given. the patient continued impella support until impella was explanted as planned and the patient survived. | [5 empty] | The investigation was completed. the impella device was discarded by the customer. paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4057 | 1220648-2025-31881 | 23369654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369654&pc=OZD | 07/23/2024 | 10/23/2025 |  | 10/23/2025 |  |  | 07/23/2024 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522432 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during an impella cp implant, the first automated impella controller (aic) got an optical sensor error. the aic was exchanged, and the procedure was completed successfully. there was no patient harm, and the patient remained stable. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. the field logs were reviewed and all optical metrics were stable throughout support. there was no pump issue found during the case. the pump functioned/operated to specification. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4058 | 1220648-2025-47742 | 23369680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369680&pc=OZD | 09/11/2025 | 10/15/2025 |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627686 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 58 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | Heart Block | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp placed to support the patient admitted in with a non-st-elevation myocardial infarction. the pump was placed and was supporting when there was heart block while on support and the team had to add the pacer for the patient support. after 2 days the pump was explanted and there was thrombus observed at the time of explant. the patient survived the issues and weaned successfully. | Protamine was given to assist with achieving hemostasis at the site of access. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.3 revised date of event as it was entered incorrectly on the initial report that was submitted. b.5 added information that was omitted from the initial report that was submitted. b.7 added information that was omitted from the initial report that was submitted. e.1 added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.6 added health effect - impact code that was omitted from the initial report that was submitted. investigation summary: no product was returned for investigation. arrhythmia/ thrombosis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. heart block: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4059 | 1220648-2025-47745 | 23369700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369700&pc=OZD | 09/26/2025 | 10/31/2025 |  | 10/23/2025 |  |  | 09/26/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665149 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 70 KG |  |  | Death | Yes | No | False Alarm | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp. three days after start of the impella cp support, the patient developed limb ischemia (impella leg). there were no perfusion and elevation of lactate. additionally, during the impella cp support there was bleeding at the extracorporeal membrane oxygenation (ecmo) site, but no noted intervention. the decision was made to remove the impella cp and escalate to an impella 5.5 with aortic access. when escalated to the 5.5 the patient experienced a shunt closing failure. approximately 16 days after start of the impella 5.5 support, it was reported the placement signal was found to be unreliable. position signal was recalibrated without repositioning of the device. the physician reported a sudden drop in the ao position signal on the aic without any changes in patient status, medicines given or ecmo flow/rpm. aortic position signal was recalibrated without repositioning of the device. instantly parameters improved and became similar as before the sudden change. there was no harm to the patient as a result of this event. additionally, it was noted the patient had oeasophagic bleeding and developed sepsis. there were no advance therapy options (left ventricular assist device, heart transplant) to treat the patient. a multidisciplinary decision for withdrawal of care was made, which was completed three days later. the patient expired. | It was reported that a patient implanted with an impella 5.5 experienced bleeding at the access site requiring surgical intervention. the pump placement signal was also found to be unreliable. the signal was recalibrated. later, the patient developed sepsis and the decision was made to withdraw care. the patient expired. |  |  |  |  | An update was made to section b5 event description to include complete details regarding the events. additional information was subsequently received and noted the following: the patient had very sick arteries, already before the implantation, the toes were blue before the puncture, and it was of course worse afterwards. the axillary artery was also way too narrow for an impella 5.5, and this is why the physician opted for direct access. the physician's decided to withdraw the therapy because they had no other therapeutic option since the patient had a very low left ventricular ejection fraction (5%) and an atrial septal defect that they could not treat. a definitive left ventricular assist device (lvad) was not possible because of the asd, and the patient was no candidate for a heart transplant. medical safety review: medical safety reviewed the event. the pump was placed but the impella cp limb became ischemic, and this prompted pump explant and escalation to a 5.5 pump after 3 days of support. additionally, during the impella cp support there was bleeding at the extracorporeal membrane oxygenation (ecmo) site, but no noted intervention. the patient was successfully escalated to an impella 5.5. this is a serious injury type of reportable event. while on impella 5.5 the patient had oeasophagic bleeding. it cannot be determined if the use of anticoagulant for use of the impella devices impacted/exacerbated the out of the bleeds for the patient. following, the patient developed sepsis. there were no advance therapy options (left ventricular assist device, heart transplant) to treat the patient. ultimately the patient expired. a multidisciplinary decision was made to withdraw care leading to the patient's demise. it is unknown how or if the impella 5.5 pump played a role involving the sepsis and given this information at hand there is not enough evidence to dissociate the impella 5.5 device from the demise outcome. note: placement signal issue or noted drop in the aortic position signal is not association to the demise outcome. there were no changes in patient status, medicines given or ecmo flow/rpm. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4060 | 1220648-2025-47746 | 23369715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369715&pc=OZD | 09/27/2025 |  |  | 10/23/2025 |  |  | 09/27/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026745623 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Inaccurate Flow Rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella cp placement the impella catheter appeared to have a left ventricle waveform above the placement signal. the impella was exchanged for a new cp and the patient survived. | [5 empty] | D.9 the device was returned and date was added. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 codes were updated due to the product bein returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A2, a3, a4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was returned for investigation. the root cause of the optical sensor issue and the high flow is due to ingested biomaterial. the device passed all post sterile inspection checks.
4061 | 1220648-2025-47749 | 23369739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369739&pc=OZD | 09/26/2025 |  |  | 10/23/2025 |  |  | 09/26/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675443 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that there was bleeding and a hematoma during impella cp patient support. while on support, there was bleeding from the access site after the peel-away was exchanged for the repositioning sheath. the doctor ordered manual pressure, a forward tension suture, and a fem-stop. a large hematoma was noted. care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleeding and hematoma was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks.
4062 | 1220648-2025-47750 | 23369757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369757&pc=OZD | 09/28/2025 |  |  | 10/23/2025 |  |  | 09/28/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685939 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Decrease in Pressure | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn |  |  |  |  | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella cp purge line on the pump side was accidentally cut, so it was replaced. the patient developed a build up of cerebrospinal fluid accumulation and ischemia. the decision was made to withdraw care and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the pump was not returned for investigation. data log was not provided or could not be retrieved from the remote server. product damage (tubing split/cut): according to the clinical details, the pump catheter was accidentally cut by the nurse, and the pump was later replaced. the cause of the tubing split/cut was most likely user error, specifically, accidental cutting of the purge line by clinical staff during routine care in the intensive care unit (icu). purge pressure low: data log was not provided to verify the purge parameter trend during the field run. the cause of the purge pressure low was most likely user error, specifically, accidental cutting of the purge line by clinical staff during routine care in the icu. clinical symptom (ischemia): in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical details: (b)(6) 2025: impella cp c8 implanted via femoral artery. the purge line on the pump catheter side was accidentally cut by the nurse (on the kk pump catheter side). after the line was cut, purge pressure dropped and an alarm was issued indicating a purge system failure, but there was no abnormality in the equipment itself. according to the doctor's comments, the pump continued to operate normally, so there was no causal relationship between the equipment and the event. the purge line on the pump side was replaced in the icu. additional information (29-sep-2025): the cut location was confirmed to be on the purge line in the purge side arm (see attached screenshot). additional information (10-oct-2025): no plans to return the product. the patient's health continued to decline, and they developed acute cerebrospinal fluid accumulation/ischemia. (b)(6) 2025: pump explant day. device history review the pump passed all post sterile inspection checks.
4063 | 1220648-2025-47751 | 23369767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369767&pc=OZD | 09/27/2025 |  |  | 10/23/2025 |  |  | 09/27/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025540280 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 YR |  | Male | 33 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Arrhythmia | Ventricular Fibrillation | Asystole | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, patient ventricular fibrillation arrested and was placed onto venoarterial extracorporeal membrane oxygenation. the next day, patient had a few suction events related to torsades which required defibrillation. a few days later, purge pressure was high with a blockage of the purge system. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4064 | 1220648-2025-47747 | 23369772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369772&pc=PYX | 09/11/2025 |  |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - |  |  | US |  | 2026679959 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 127 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Patient was implanted with impella rp for heart support post cardiac surgery.during the support the patient remains intubated, on multiple vasopressors inotropes, and continuous renal replacement therapy (crrt). purge changed to tissue plasminogen activator (tpa) with purge pressure and flow improvement. patient was successfully weaned as planned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4065 | 1220648-2025-47752 | 23369845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369845&pc=OZD | 09/27/2025 |  |  | 10/23/2025 |  |  | 09/27/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026626909 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to heart failure / cardiogenic shock. while on support, patient was bleeding from the groin and was taken to operating room for vascular repair and hematoma evacuation. patient has received several units of blood products and albumin. the patient was weaned successfully and survived. | [5 empty] | Investigation was completed and no product was returned. hematoma/ major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review was completed and passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4066 | 1220648-2025-47753 | 23369850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369850&pc=PYX | 09/27/2025 |  |  | 10/23/2025 |  |  | 09/27/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026647350 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient, previously implanted with impella cp in the left ventricle, was placed onto impella rp flex support due to right ventricular dysfunction. the patient eventually developed flash pulmonary edema with rapid cardiopulmonary decompensation. family withdrew care and patient expired. this report is for the imeplla rp flex (serial number: (b)(6)). an additional report will be sent for the impella cp (serial number: (b)(6)). | [5 empty] | D4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. h6 added b13 code as it was omitted from the initial report that was submitted. medical safety reviewed the event. there is not enough information to dissociate the impella devices used for bipella support as an associated factor to the patient's outcome. however, in this case, patient underlying condition likely contributed most to an outcome of death. care was withdrawn which led to the patient expiring. nvestigation summary: the investigation has been completed. pulmonary edema/ organ failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. h6 type of investigation, investigation findings, conclusion codes and component code were updated accordingly based on the completed investigation. related medical device reports: this impella rp flex device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4067 | 1220648-2025-47754 | 23369863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369863&pc=OZD | 09/27/2025 |  |  | 10/23/2025 |  |  | 09/27/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648633 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that there was hemolysis during impella cp patient support. although the patient's urine was clear, the ldh level was elevated. no patient harm was reported. | [5 empty] | Neither the product nor data logs were returned for review. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the device passed all post sterile inspection checks. | A4 is unknown.
4068 | 1220648-2025-47755 | 23369874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369874&pc=OZD | 09/28/2025 |  |  | 10/23/2025 |  |  | 09/28/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722852 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 70 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella cp support due to cardiomyopathy. after several days the patient was escalated to an impella 5.5 for continued support. while on impella 5.5 support the patient experienced renal failure as he was placed onto continuous renal replacement therapy. four days later, patient was taken off of heparin due to bleeding in the mouth. the following day, patient had ventricular tachycardia (vt) for which he was given one shock and placed onto amio and lido. three days later, patient experienced additional runs of vt with no resolution specified. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and device was not returned. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. the necessary materials were not returned for analysis so the serial number could not be identified to complete a device history review inspection.
4069 | 1220648-2025-47757 | 23369886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23369886&pc=OZD | 09/28/2025 |  |  | 10/23/2025 |  |  | 09/28/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716917 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had bradycardic episode appearing to be complete heart block. amiodarone stopped and epinephrine was increased. the patient eventually converted back to sinus rhythm. an echocardiogram at bedside confirmed that the impella position was appropriate, and no adjustments were made. the patient was successfully explanted as planned and survived. | [5 empty] | The investigation was completed and no product was returned. bradycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. heart block: the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4070 | 1220648-2025-47756 | 23370502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23370502&pc=OZD | 09/11/2025 |  |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693051 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp, the patient developed gastrointestinal bleed (gib) overnight, blood count stable. continuous renal replacement therapy started with continued renal insufficiency. don't think that patient has a gib but maybe just some rectal bleeding. heparin drip was previously stopped due to patient having bloody stools and urine, and platelets steadily dropping. argatriban drip is to be started. later it was reported that the patient went into pulseless electrical activity and the staff was able to get return of spontaneous circulation on patient after high quality cardiopulmonary resuscitation and intubation. after the patient woke up decision was made to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added code b11. h11: added the results of the investigation. investigation summary: the device was not returned for investigation. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. renal failure/ cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1934648. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
4071 | 1220648-2025-47758 | 23371149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23371149&pc=OZD | 09/11/2025 |  |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723865 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 73 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient had multiple hypotensive episodes overnight. vasopressors increased. heart team is not planning for escalation/advanced therapies due to patient medical non-compliance. the patient was deemed not a candidate for escalation. the patient made a do not resuscitate and the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added type of investigation codes b13 and b24. h11: added the results of the investigation. investigation summary: the device was not returned for evaluation. hypotension: the cause of the injury was not determined due to insufficient clinical details. device history review: this pump sn (b)(6) passed all post sterile inspection checks.
4072 | 1220648-2025-47759 | 23371869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23371869&pc=OZD | 09/11/2025 |  |  | 10/23/2025 |  |  | 09/11/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654294 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | Multiple Organ Dysfunction Syndrome | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient started to decompensate after intra-aortic balloon pump was removed so decision for impella cp was made. the cp was successfully placed. during support systemic organ failure was reported. the patient was taken to the operating room for open abdomen exploration for concerns of bowel ischemia/necrosis. bowel appeared healthy but left open and wound vac placed. additionally, there was an onset of hematuria. the impella was deep so it was repositioned with being pulled back several centimeters. the patient was also given 1 liter of bolus for suction. in addition to 1 unit of red blood cells given for insertion site oozing. systemic heparin was held. the patient remained in critical condition, ultimately expiring on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction h6. health effect - impact code. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of bleed was most likely use issue since ptt was high and systemic heparin was paused. device in wrong position (positioning issue): the cause of the positioning issue was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4073 | 1220648-2025-47761 | 23372501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23372501&pc=OZD | 09/12/2025 |  |  | 10/23/2025 |  |  | 09/12/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738926 |  | 0048-0045 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was slight oozing at the impella access site. the nurse changed the dressing and tried to find a better angle match for the repositioning sheath. the team believed the patient may have been in septic shock based on elevated white blood cell count. the plan was to remove the impella. the patient was weaned off of impella support and explanted. patient survived the procedure. | [5 empty] | B.2 revised selection as it was previously entered incorrectly on the initial report that was submitted. h.6 code e0506 was reported incorrectly on the initial report that was submitted. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ investigation summary: the investigation of has been completed. no product was returned for investigation. sepsis: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the sepsis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4074 | 1220648-2025-47763 | 23373155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23373155&pc=OZD | 09/25/2025 |  |  | 10/23/2025 |  |  | 09/25/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709931 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Pericardial Effusion | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. patient was transferred from cape to holmes with wire left in left anterior descending artery (lad). he was brought straight to the or and coronary artery bypass graft x2 was performed. patient had issues with excessive bleeding with recent brillinta administration. in addition, patient struggled coming off pump and requiring multiple vasopressors. decision was made to place impella. the surgeon contemplated bringing patient back to or and leaving chest open due to significant bleeding resulting in tachycardia that occurred once chest was closed. 2 days after implant atrial fibrillation with rapid ventricular response started overnight amio. echo was done and new pericardial effusion on right side was noted. blood being given. patient successfully weaned from support and survived the procedure | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. tachycardia/ paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4075 | 1220648-2025-47764 | 23373186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23373186&pc=OZD | 09/23/2025 |  |  | 10/23/2025 |  |  | 09/23/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740376 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support patient developed hematuria. impella was weaned to p-5 from p-9, echocardiogram obtained with device inlet measuring 7cm from annulus. device was pulled back and repositioned successfully. however, no resolution of bloody urine/hemolysis. the cp was successfully weaned and explanted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | B.1 added selection as it was omitted from the initial report that was submitted. d.3 added manufacturer fax number as it was omitted from the initial report that was submitted. d.4 added model as it was omitted from the initial report that was submitted. h.6 added medical device problem code. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ investigation summary: the investigation has been completed. no product was returned. hematuria: the cause of the injury was not determined due to insufficient clinical details. hemolysis: no product was returned. the patient developed hematuria/dark red urine early in the am. after reviewing the logs, no suction alarms or motor current spike or any trend in placement signal was observed. the cause of hemolysis cannot be determined since insufficient clinical details were provided. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4076 | 1220648-2025-47765 | 23373193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23373193&pc=OZD | 09/23/2025 |  |  | 10/23/2025 |  |  | 09/23/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696505 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient presented with chest pains and coded in cath lab. decision was made to place impella cp. however, attempt to place the cp was unsuccessful due to anatomy of patient and previous grafting/stenting to bilateral illiac and aortic arteries physician unable to advance impella. decision was made to abort procedure and place the patient on an intra-aortic balloon pump. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy along with grafting/stenting to bilateral iliac and aortic arteries preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4077 | 1220648-2025-47766 | 23373199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23373199&pc=OZD | 09/23/2025 |  |  | 10/23/2025 |  |  | 09/23/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724280 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with no known medical history presented with chest pain times one month, however worsening. st-segment elevation myocardial infarction activation. angio revealed distal left main disease. impella placed pre-percutaneous coronary intervention via right femoral artery. it was noted that the patient coded post cp placement. max presser support and cardiopulmonary resuscitation initiated. return of spontaneous circulation achieved. echocardiogram performed with optimal placement noted for cp. tricuspid annular plane systolic excursion within normal limits. patient decompensated and code resumed. unfortunately, family decided against further intervention and patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4078 | 1220648-2025-31567 | 23373208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23373208&pc=OZD | 12/23/2024 | 10/23/2025 |  | 10/23/2025 |  |  | 12/23/2024 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574256 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, there were erratic, intermittent left ventricular and placement signals. pump position was verified; flows and motor current were within the normal range. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | D6b: the explant date is unknown. this report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. there were no data logs provided. the root cause of the optical signal issue was not determined as no product was returned and no data logs were provided for complaint occurrence date a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4079 | 1220648-2025-47767 | 23375734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23375734&pc=OZD | 09/26/2025 |  |  | 10/23/2025 |  |  | 09/26/2025 | 10/23/2025 |  |  |  |  | 10/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739082 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the impella was being transferred from automated impella controller (aic )to transport aic, impella was not recognized and could not be transferred back to original aic. the pump was pulled back by intensivist and replaced by interventional cardiologist. this mdr specifically addresses impella cp which is the subject of this report. a separate mdr was submitted for the associated aic with this complaint. | [5 empty] | Section a4 and a6 : the patient race and weight is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was received for investigation. d9 was updated with receipt date. one the investigation is complete a supplemental report will be submitted.
4080 | 1220648-2025-47661 | 23354615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354615&pc=OZD | 09/23/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738887 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed hematuria. the impella was weaned down from p9 to p7. additionally, the patient was transfused one unit of red blood cells in response to low hemoglobin values. the patient was successfully weaned and the pump was explanted as planned. no patient harm was reported. | [5 empty] | The cause of the hematuria and anemia was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4081 | 1220648-2025-47660 | 23354618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354618&pc=OZD | 09/19/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026678000 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during prep, a "defective impella catheter" alarm presented. the nurse attempted to remove and replace the plug, resulting in the same alert. the impella was removed and replaced with another impella. case proceeded with no further reported complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. h6 type of investigation, investigation findings, conclusion codes and component codes were updated accordingly based on the completed investigation. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, data logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the pump passed all the post sterile inspection checks. regarding, pump not detected, clinical details stated that when plugging in the impella cp, there was an "impella defective" alarm. unplugging and replugging the pump in did not resolve the issue. the log review confirmed an impella defective alarm and indicate that there was a lack of response from the epm because it did not receive any response from the pump eeprom. a new pump (sn (b)(6) was successfully connected to the console shortly after. the cause of the pump not being detected was not determined since product was not returned.
4082 | 1220648-2025-47662 | 23354622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354622&pc=OZD | 09/23/2025 | 10/21/2025 |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665408 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed hemolysis. the impella was repositioned. | [5 empty] | Correction d3, e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of the mechanical interaction with blood was unable to be determined since product was not returned and the pump was both repositioned and turned down to resolve the hemolysis. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4083 | 1220648-2025-47663 | 23354623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354623&pc=OZD | 09/24/2025 | 11/29/2025 |  | 10/22/2025 |  |  | 09/24/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1460542 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic), upon turning on, experienced a system self-check failed alarm. a new aic was used. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation for the reported sysyem self check error has been completed. the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
4084 | 1220648-2025-47664 | 23354625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354625&pc=OZD | 09/30/2025 |  |  | 10/22/2025 |  |  | 09/29/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced controller failure during a purge cassettes change. aic console was changed. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4085 | 1220648-2025-47665 | 23354628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354628&pc=OZD | 08/30/2025 | 10/21/2025 |  | 10/22/2025 |  |  | 10/10/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1615855 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a placement signal didn't have a reliable alarm. aberrant waveforms were noted. mc/purge related flows/pressures were monitored. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: data analysis: partial console logs obtained from the cloud are consistent with complaint and show large number of placement signal not reliable alarms (#1036, opm signal invalid). rtlog data shows some dependence of snr on motor current (as expected) but also significant correlation between values of snr and of measured pressure (placement) for constant motor current. device analysis: console ic9280 was not returned for the investigation despite being requested. it was flagged for further investigation for next time it is returned to field service.
4086 | 1220648-2025-47666 | 23354630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354630&pc=OZD | 09/20/2025 | 01/06/2026 |  | 10/22/2025 |  |  | 10/13/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1683279 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 104 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella aic had a controller error alarm while supporting the patient. the placement signal was lost but resolved within 20 seconds. the aic was not exchanged. the patient continued support and was bridge to left ventricular assist device. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4087 | 1220648-2025-47667 | 23354638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354638&pc=OZD | 08/01/2025 |  |  | 10/22/2025 |  |  | 10/13/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1636031 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support to bridge to heart transplant. the complainant reported that sensor failure occurred on the impella automated impella (aic) controller while supporting a patient. the placement monitoring was disable. support was not affected by the sensor issue. no patient harm was reported due to the issue. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data review: the logs revealed sudden fluctuations in placement signals, with pump flow oscillating between 0 and 5 l/min. further review of logs from the exchanged aic (b)(4) showed that retrograde flow (#426), placement signal not reliable (#1036), and placement signal not reliable (#1048) alarms continued for two days device analysis: the console was returned to field service for further investigation. the reported complaint could not be reproduced, even after the console was operated with the pump and purge cassette for a couple of hours. additionally, a visual inspection of the internal circuitry showed no abnormalities. further inspection of the console revealed slight distortion in the optical bench analog signal. however, this does not seem to be related to the reported issue. the optical bench 5s parameters were within specification when tested with the 2 different test cavities. root cause: the root cause of the ¿sensor failure¿ retrograde flow alarm could not be determined, as the issue could not be reproduced. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
4088 | 1220648-2025-47668 | 23356434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23356434&pc=OZD | 09/23/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1583419 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the white plug inside the automated impella controller (aic) console was fractured during transport. no patient involvement was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The returned console was tested and no abnormalities were observed. the console operated as intended while testing with a known good pump and purge cassette kit. the console was returned to the customer for use. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4089 | 1220648-2025-47669 | 23356647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23356647&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678047 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced runs of ventricular tachycardia. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the bleeding is determined to be use issue because the peel away introducer was left in place and not removed per the instructions for use. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4090 | 1220648-2025-47670 | 23356713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23356713&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741834 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 57 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hypertension and was treated with nipride. the patient was also treated with bivalirudin for a thrombus encasing the impella catheter. oozing from insertion site was noted so packed red blood cells were transfused. a thrombectomy was performed and the patient was escalated to an impella 5.5. no patient harm was reported. | [5 empty] | The cause of the bleeding was not determined since product was not returned and insufficient clinical details provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the hemodynamic instability was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
4091 | 1220648-2025-47671 | 23356858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23356858&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700177 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a delivery issue during impella 5.5 patient support. due to patient anatomy, the case was aborted, and the patient remained on veno-arterial extracorporeal membrane oxygenation. | [5 empty] | The device was discarded. the investigation was completed. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history reviewed was completed and passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
4092 | 1220648-2025-47673 | 23356951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23356951&pc=OZD | 09/16/2025 |  |  | 10/22/2025 |  |  | 09/16/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682786 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced septic shock. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4093 | 1220648-2025-47674 | 23357078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23357078&pc=OZD | 09/23/2025 | 02/23/2026 |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641071 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hemodynamic instability and bradycardia. the patient was treated with medication and a temporary pacemaker was implanted. later, the patient was successfully weaned from the pump and survived. | The complainant reported additional information upon request: "these were results of patients clinical condition and in no way related to impella use. pump is not available for return." |  |  |  |  | No product was returned for investigation. the cause of the hemodynamic instability was not determined due to insufficient clinical details. the root cause of the bradycardia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B5 updated.
4094 | 1220648-2025-47675 | 23357255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23357255&pc=OZD | 10/03/2025 | 10/29/2025 |  | 10/22/2025 |  |  | 10/03/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301536 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 had immediate priming/zero/controller failures from the supporting automated impella controller (aic). the aic was replaced to continue patient support. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report number.
4095 | 1220648-2025-47676 | 23357352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23357352&pc=OZD | 10/03/2025 |  |  | 10/22/2025 |  |  | 10/03/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716121 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Failure to Read Input Signal | Adverse Event Without Identified Device or Use Problem | Air/Gas in Device |  |  |  |  | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant had impella 5.5-automated impella controller (aic) support initiated for a patient admitted in with decompensation after a prior impella cp had been explanted the day prior. the impella cp was removed, an intra-aortic balloon pump was placed, and now an impella 5.5 was placed and the aic had immediate priming/zero/controller failure. the aic was replaced and support was initiated but after the pump was started there was an observation made of air. the team removed and replaced the impella 5.5 with a new pump. the second pump remains on for support to date despite the air observation. | [5 empty] | The returned product was visually inspected and there were no abnormalities. the cause of the pump not detected could not be determined. the cause of the embolism was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | E4 has been updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the pump. another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number. | Product complaint # (b)(4). the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4096 | 1220648-2025-47672 | 23357533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23357533&pc=OZD | 09/22/2025 | 10/22/2025 |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706540 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 138 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | Sepsis | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a fever and increased white blood cell count. the sepsis was unresolved and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The root cause of the infection and sepsis was unable to be determined due to insufficient clinical details. the cause of the fever was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
4097 | 1220648-2025-47677 | 23357568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23357568&pc=OZD | 09/15/2025 |  |  | 10/22/2025 |  |  | 09/15/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727181 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 124 KG |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient became unresponsive while driving their car. the patient was intubated and brought to the emergency room where return of spontaneous circulation was achieved. the patient was admitted to the intensive care unit and decompensated throughout the day. the patient was brought to the cardiac catheterization laboratory for a coronary angiogram and impella cp placement. on the third day of impella support, there was reported to be suction alarms occurring intermittently when the patient was turned, or when the medical team tried to pull more fluid with continuous renal replacement therapy. the medical team discussed giving the patient volume depending on results from laboratory bloodwork. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | A5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 revised health effect - impact code as it was entered incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned. pump suction - after three days on support, the patient was having intermittent suction alarms on p-6. fluid status was discussed with the bedside nurse. suction was noted to occur when the patient was moved or additional fluid was pulled off of continuous renal replacement therapy (crrt). fluid may be given at a later point. data log review showed no suction occurs below p-5. no motor current spikes outside of p-level changes. placement signal trends up toward case end but remains normal. pump position was noted to be ok throughout case. the cause of the pump suction was patient condition related due to the issue noting to occur during patient movement and crrt fluid being pulled off. data logs confirmed suction does not occur under p-5 and pump positioning ok. device history review: the pump passed all post sterile inspection checks.
4098 | 1220648-2025-47555 | 23357612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23357612&pc=OZD | 09/14/2025 | 10/22/2025 |  | 10/22/2025 |  |  | 09/14/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601424 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 60 KG |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. micro puncture and ultrasound used to achieve access to right common femoral artery. during complex percutaneous coronary intervention, the left anterior descending artery (lad) was dissected and patient went into ventricular tachycardia/ventricular fibrillation. multiple rounds of cardiopulmonary resuscitation and defibrillation (13x) were performed and delivered. the patient was intubated and multiple vasopressors started. at this point impella was inserted with flows in auto getting 2.8 l/min to 3.3 l/min of flow. patient continued to code multiple times requiring repeated defibrillations. lad fixed and flows adjusted to p6 with flows of 2.9 l/min. p9 attempted but due to size of ventricle and suction events impella weaned and continued at p6 with goals for hydration and blood products for patient. however, the patient continued to decline. abdomen continued to grow in size and becoming tauter. flows unable to be kept above p2. ultimately care was withdrawn and the patient expired | [5 empty] | Correction b1. the investigation of the reported failure mode has been completed. no product was received for evaluation. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. low pump flow: after reviewing logs, multiple suction alarms were observed along with couple of motor current spike and sudden drop in flows. the cause of low pump flow was most likely biomaterial ingestion per the log review. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4099 | 1220648-2025-47646 | 23358214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358214&pc=OZD | 09/28/2025 | 10/15/2025 |  | 10/22/2025 |  |  | 07/17/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695303 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the impella experienced ongoing yellow alarms for placement signal low. after reviewing the alarm history, the alarm had been an ongoing issue for days. after further inquiry, the registered nurse (rn) relayed that the patient was stable, and their monitor pressures did not match the automated impella console (aic). rn was guided through an ao adjustment and adjusted the aic 31 points. alarm resolved. i advised the rn that this adjustment should only be done this one time and should not be done ever again. i advised the rn to call us should the alarm return. additionally, the patient underwent thrombectomy secondary to a stroke while on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4100 | 1220648-2025-47681 | 23358381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358381&pc=OZD | 06/28/2025 |  |  | 10/22/2025 |  |  | 06/28/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624646 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Arrhythmia | Fever | Unspecified Infection | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. patient experienced delivery issues, optical signal issue, low pump flow, positioning issue, fever, infection and thrombus during impella 5.5 patient support. during the initial wireless attempt, the clinical staff were unable to cross the aortic valve. the doctor scrubbed in and proceeded with a wired placement. there were intermittent "placement signal unreliable¿ alarms. the placement signal returned but, but flows were lower than expected at p5. a formal bedside echocardiogram (echo) was done, and pump placement was deep. the pump was retracted under echo visualization 2.5cm, final left ventricle depth was 4.85cm. there was immediate improvement in flows. the patient had atrial fibrillation. the patient was febrile and was resolving on broad spectrum antibiotics and cultures were pending. there was a suspected clot. no immediate plans for removal of suspected pericardial clot. the impella was explanted. a few days following impella removal, the patient developed drainage from his graft site. it was cultured and positive, so he started on antibiotics. there was a slight elevation in white bleed cells, but the patient was not considered septic. he was eventually taken to the or for washout and inspection of his externalization site. it was cutdown on, washout out with antibiotic irrigation and graft trimmed further. no re-sternotomy was done and graft remained attached to ascending aorta. ultimately, the patient responded well to treatment. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical signal issue - data log review shows placement signal not reliable (psnr) alarms that occurred on complaint date but resolved. after cross functional review, no defect was found with the pump and the console is suspected of being the potential issue. low or blocked pump flow -the cause of the low pump flow was the pump being in too deep which was a user related issue during implantation of the device. device in wrong position - the cause of the positioning issue was user related due the initial placement of the pump being too deep. patient codes - arrhythmia, infection, thrombosis, fever: the root cause of these injuries was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4101 | 1220648-2025-47682 | 23358409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358409&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608444 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was escalated to another impella cp and they encountered false alarm and placement signal issue during impella cp support. it was noted that a red alarm with impella position in aorta appeared, despite good impella position in left ventricle (lv) which was confirmed via echocardiogram. the red alarm impella position in aorta was able to be reproduced once p level was reduced to p4. during this alarm, all metric curves were pulsatile. the first pump (serial number (b)(6)) will be reported in an additional report. this report is for the second pump (serial number (b)(6)). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4102 | 1220648-2025-47683 | 23358419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358419&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026707599 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered thrombus and pump suction during impella cp support. it was noted that during bypass the medical team got issues with suction and position of the pump. pump was then explanted and biomaterial was discovered in the inlet area. the pump was subsequently replaced with a new impella cp (sn: (b)(6)). this report is for the first impella cp pump and another report will be sent for the second pump (serial (b)(6)) (b)(4). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. clinical symptoms (thrombosis/thrombus): the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: suction there was no motor current spike noted to indicate biomaterial interaction. the cause of the suction issue was not determined without returned product investigation. device history review: the device passed all the post sterile inspection checks.
4103 | 1220648-2025-47685 | 23358421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358421&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731725 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 99 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient hemodynamically improved but during the support distal limb loss of pulse present, for these reasons, impella was explanted. impella successfully removed and explanted at bedside. bleed back allowed. manual pressure held. color and pulse improvement noted after removal. hemostasis achieved. patient survived the procedure and is stable. | [5 empty] | Correction d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4104 | 1220648-2025-47686 | 23358432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358432&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026692780 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 64 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the physician tried to place impella 5.5 but the 5.5 was unable to advance through graft anastomosis due to small vessel size and anatomy. the 5.5 placement was aborted and intra-aortic balloon pump was inserted through the vascular graft. | [5 empty] | B2 added selection that was omitted from the initial report that was submitted. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. h6 added b13 code as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had small vessel size and anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4105 | 1220648-2025-47688 | 23358434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358434&pc=OZD | 09/23/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729937 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support for elective placement for coronary artery bypass graft . during the support the patient developed facial droop and left side weakness, and stroke was called. stat ct of head showed subdural hematoma acute on chronic with midline shift. neurosurgery consulted. heparin was stopped. the patient was successfully weaned and explanted as planned . it is unknow if surgical intervention was done to address the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4106 | 1220648-2025-47684 | 23358435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358435&pc=OZD | 08/02/2025 |  |  | 10/22/2025 |  |  | 08/02/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694102 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 112 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp there was concern for bleeding issues so the physician reinserted the sheath. then, the patient developed ischemia in the right leg. the device was positioned too deep and was generating alarms. the device was pulled back via echocardiogram but it would go shallow and deep again. the p level was decreased and albumin was given. the pump was removed, pressure was held, and the patient started decompensating. the patient coded for an hour. it was suspected that the patient suffered a retroperitoneal bleed because their hemoglobin dropped from 8.1 to 4. the patient expired. | [5 empty] | No product was returned for investigation. data logs were returned for analysis. the root cause of the pump suction was most likely user related due to positioning issues based on the provided clinical details which mentioned that the pump was confirmed by echo to be in the wrong position and data log analysis which confirmed suction alarms occurring during the case around the time the pump was reported to be in the wrong position from the clinical details. the root cause of the positioning issues was most likely patient condition related based on the provided clinical details concerning the patient's anatomical abnormalities, which caused the pump to go shallow and then deep again after repositioning. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the root cause of the cardiac arrest was not determined due to insufficient clinical details. the cause of the major bleeding was not determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4107 | 1220648-2025-47687 | 23358438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358438&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741838 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 48 KG | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed to support the percutaneous coronary intervention (pci) procedure. during the pci, once shockwave was utilized, placement signal not reliable alarm came on, and remained on for the remainder of the case. flows remained about 3.7l throughout case. pci was completed successfully with impella support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. optical sensor issue: the pump was not returned for investigation. data logs show a psnr alarm with a spike to 3000 after 50 minutes of pump use. the snr and contrast dropped to 0, while the light, lamp and gain remained unchanged during the psnr event. the data log indicates a known pattern of a damaged sensor. according to the clinical details, the psnr issue occurred during shockwave initialization; however, the distance between the shockwave device and the pump optical sensor was unknown. the cause of the optical signal issue was most likely a damaged sensor due to shockwave device interaction, based on data log review and provided clinical details. device history review was completed and passed all post sterile inspection checks.
4108 | 1220648-2025-47680 | 23358609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358609&pc=- | 09/13/2025 |  |  | 10/22/2025 |  |  | 09/13/2025 | 10/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | - |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301112 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient for cardiac surgery was implanted with an impella 5.5 for mechanical circulatory support. while on support, there was a noted large discrepancy between the patient¿s arterial line and the placement signal of the automated impella controller support. the placement signal was thirty points lower than the actual arterial line pressure. due to complications, the patient expired and is reported under the related pump report. this report represents the automated impella controller. a separate report was submitted to represent the pump. refer to section h.10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety/clinical reviewed the event. the event describes a patient presenting with non¿st-elevation myocardial infarction and cardiogenic shock. the patient presented to the emergency department with two days of chest pain and wheezing. electrocardiogram demonstrated right bundle branch block, and there was concern for pulmonary edema and new-onset congestive heart failure. the patient was taken to the cardiac catheterization laboratory, where angiography revealed an occluded right coronary artery and diffuse disease of the left anterior descending artery (lad). during planned lad angioplasty, acute thrombosis occurred, left ventricular end-diastolic pressure increased to 44 mmhg, and ventriculography demonstrated severely reduced ejection fraction. the procedure was aborted, and a decision was made to place an impella cp and transfer the patient for surgical evaluation. percutaneous impella cp insertion was unsuccessful due to inability to advance the device across the iliac artery after multiple attempts. due to clinical deterioration, an intra-aortic balloon pump was placed. this represents a serious injury event involving the impella cp device. the patient was transferred to an outside hospital, where cardiothoracic surgery determined the vessels were suitable for impella 5.5 placement. an impella 5.5 was surgically implanted via the right axillary artery for mechanical circulatory support as a bridge to possible cardiac surgery. during impella support, the patient experienced multiple complications that can be associated with patients in cardiogenic shock, including arrhythmia, anemia, heart block, ischemia, inflammation, pain, hemodynamic instability, thrombosis, hypotension, and cardiac arrest with return of spontaneous circulation. after approximately 27 days of support, the impella 5.5 was surgically removed at the icu bedside. following device removal, the patient experienced immediate hemodynamic deterioration requiring maximal vasopressor support, including norepinephrine and epinephrine. loss of arterial line pulsatility was noted, and advanced cardiac life support, including defibrillation and cpr, was performed without successful resuscitation and the patient expired. related medical device reports: this is one of three devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
4109 | 1220648-2025-47689 | 23358633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358633&pc=OZD | 08/06/2025 | 10/22/2025 |  | 10/22/2025 |  |  | 08/06/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133147 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 63 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Decrease in Pressure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient with difficult anatomy experienced multiple events: ventricular tachycardia requiring resuscitation, and purge disc error alarm resolved by replacing the automated impella controller. the medical safety officer reviewed and determined the automated impella controller (aic) issue was initially managed with tissue plasminogen activator administration to the purge and is a serious injury. the aic was eventually exchanged due to an unresolved purge disc error alarm. the patient remained stable throughout the console change interruption. the demise outcome is associated with the pump as there is not enough evidence to exclude the impella 5.5 as an associated factor in the patient's outcome. | [5 empty] | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. there were no issues related to the complaint discovered when reviewing the product history of the console. console was tested after arriving in field service. the reported purge disc detection issue was able to be reproduced during testing. inspection of the console revealed that the purge disc flag was broken (missing at blue retainer). the root cause of the purge disc detection issues was a broken purge disc flag (missing at blue retainer). the root cause of the reported random decreased purge flows could not be conclusively determined but was likely caused by a temporary buildup in the purge line that resolved after tpa was added to the purge. | The automated impella controller (aic) was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the aic. another report was submitted for the pump. refer to section h.10 for the related report number.
4110 | 1220648-2025-47690 | 23358773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23358773&pc=OZD | 08/14/2025 |  |  | 10/22/2025 |  |  | 08/12/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706913 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp expired. no details related to patient death were provided. | [5 empty] | B2 revised date of death and b3 date of event as they were entered incorrectly on the initial report that was submitted. b7 added information as it was inadvertently omitted from the initial report that was submitted. d4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. g1 revised manufacturing site name as it was entered incorrectly on the initial report that was submitted. added manufacturer site information as it was inadvertently omitted from the initial report that was submitted. g3 date was reported incorrectly on the initial report that was submitted. the date should of reflected 14-aug-2025. medical safety reviewed the event. the event describes purge cassette leak resulting in an alarm of purge pressure low. the purge system was successfully replaced. the purge cassette had a malfunction and did not contribute to the demise outcome. there was low purge pressure but there was no interruption or compromised to support because it occurred before implantation of the impella cp. conservatively, the death will be reported as the impella cp was in use at time of expiration. withdrawing care led to the patient's demise. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the report on the purge cassette. | The investigation was completed. procedural adverse event: the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4111 | 1220648-2025-47692 | 23359028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359028&pc=OZD | 08/25/2025 |  |  | 10/22/2025 |  |  | 08/24/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612280 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 111 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Fever | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a fever. due to concerns for bleeding, systemic heparin was turned off. the patient also experienced stroke like symptoms, specifically arm and leg weakness. an mri identified approximately 10 microvascular events less than 1mm in size. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4112 | 1220648-2025-47693 | 23359101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359101&pc=OZD | 08/04/2025 |  |  | 10/22/2025 |  |  | 08/04/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678483 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 93 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Restricted Flow rate | Decrease in Pressure | Malposition of Device |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there were suction events, positioning issues, low purge pressure and product damage during impella 5.5 patient support. while on support, the impella required positioning due to noted ventricular waveform. the patient experienced multiple suction events. the p-level was decreased and impella pump was repositioned per the managing physician. there was a visible linear crack on the purge male port. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product returned. pump suction - after 37 days on support, the patient had multiple suction events, p-level was decreased and the pump was repositioned. data logs for this occurrence data were not recovered. the cause of the pump suction was not determined due to no product returned, insufficient data logs recovered, and insufficient clinical details. device in wrong position - after 37 days on support, the patient had multiple suction events, p-level was decreased and the pump was repositioned. the pump now was listed as left ventricle 4.9-5.2cm. two days prior, ventricular waveform was noted. data logs for this occurrence data were not recovered. the cause of the device in wrong position was not determined due to no product returned, insufficient data logs recovered, and insufficient clinical details. purge pressure low - after 72 days on support, visible linear crack on purge male port. the automated impella controller (aic) was alarming low purge pressure (pp). data logs for this occurrence data were not recovered. the cause of the purge pressure low was not determined due to no product returned, insufficient data logs recovered, and insufficient clinical details. product damage (purge cassette crack) - after 72 days on support, visible linear crack on purge male port was reported. the aic was alarming low pp. data logs for this occurrence data were not recovered. the cause of the product damage (purge cassette crack) was not determined due to no product returned, insufficient data logs recovered, and insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
4113 | 1220648-2025-47679 | 23359133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359133&pc=OZD | 08/26/2025 |  |  | 10/22/2025 |  |  | 08/26/2025 | 10/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707489 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Anemia | Arrhythmia | Cardiac Arrest | Low Blood Pressure/ Hypotension | Unspecified Infection | Inflammation | Ischemia | Pain | Sepsis | Thrombocytopenia | Thrombosis/Thrombus | Heart Block | Hemodynamic instability | [25 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient for cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. after insertion of the impella, laminar flow and low pulsatility were noted with the patient in second degree block in the 40's suction occurred only being able to achieve reduced flows at p-4, with pressors having to be increased to maintain mean arterial pressure >65. the decision was made to float a temporary pacer, after completion of the procedure wave forms appear to be more appropriate and able to start weaning pressors. while on support, the patient¿s hemoglobin dropped, getting as low as 6.5. one unit of blood product was administered. the patient had some rhythm changes while ambulating. the doctor advanced the temporary pacer and pressure to correct the rhythm. the impella flows were below normal range for associated p-level, the doctor at bedside believed it was a volume issue. the nurses were unable to get peripheral pulses as well, a systemic vascular resistance of 1550 was calculated. furthermore, the patient¿s white blood cell count had increased and the patient was on vancomycin and zosyn. the nurse stated that an ultrasound showed deep vein thrombosis. a clot in the bilateral upper extremities and internal jugular. there was a suspicion of the patient being septic so blood cultures were drawn. the patient was started on a low-dose of levophed to combat hypotension. the patient did receive a bolus of fluid, which improved the hypotension and allowed for levophed to be titrated down. the patient could not support herself while standing up due to pain and weakness. blood cultures were negative. the patient does have swelling in all extremities that was causing her pain and unable to rest. the patients extremities were tight and beginning to blister due to swelling and third spacing of fluid. lasix drop was started at a low rate to diurese patient. swelling was decreased. it was noted there was underlying suspected sepsis yet white blood cell count increased. there was a noted large discrepancy between the patient¿s arterial line and the placement signal of the automated impella controller. during the procedure, the doctor did have to use a clot retrieval device to remove the clot from inferior vena cava before able to advance the micra catheter. the patient was having suction events along with severe hypotension. the patient coded but return of spontaneous circulation was achieved. the resident reported pulseless electricity activity. an echocardiogram was at bedside right after code was done and showed the left ventricle was empty,. blood was being returned from the continuous renal replacement therapy machine, and after full return, the ventricle appeared to be full and the impella in the incorrect position. appropriate flows were reached with a higher p-level. labs showed hemoglobin, platelets and hematocrit were low. the patient was positive for heparin induced thrombocytopenia. before the p-level reduction, there were several self-limiting suction alarms. the impella did briefly alarm ¿placement signal low¿ during rounds. of note aorta signal has been reading lower than a-line consistently and the impella position appeared to be unchanged. care was withdrawn. as soon as the patient lost support of the impella 5.5, there was an immediate deterioration in her hemodynamics, which required the intensive care unit team to have to max out levophed as well as start epinephrine drops with rapid titration increases. multiple rounds of advanced cardiac life support including shocks and cardiopulmonary resuscitation were given. the patient expired. | [5 empty] | Medical safety/clinical reviewed the event. the event describes a patient presenting with non¿st-elevation myocardial infarction and cardiogenic shock. the patient presented to the emergency department with two days of chest pain and wheezing. electrocardiogram demonstrated right bundle branch block, and there was concern for pulmonary edema and new-onset congestive heart failure. the patient was taken to the cardiac catheterization laboratory, where angiography revealed an occluded right coronary artery and diffuse disease of the left anterior descending artery (lad). during planned lad angioplasty, acute thrombosis occurred, left ventricular end-diastolic pressure increased to 44 mmhg, and ventriculography demonstrated severely reduced ejection fraction. the procedure was aborted, and a decision was made to place an impella cp and transfer the patient for surgical evaluation. percutaneous impella cp insertion was unsuccessful due to inability to advance the device across the iliac artery after multiple attempts. due to clinical deterioration, an intra-aortic balloon pump was placed. this represents a serious injury event involving the impella cp device. the patient was transferred to an outside hospital, where cardiothoracic surgery determined the vessels were suitable for impella 5.5 placement. an impella 5.5 was surgically implanted via the right axillary artery for mechanical circulatory support as a bridge to possible cardiac surgery. during impella support, the patient experienced multiple complications that can be associated with patients in cardiogenic shock, including arrhythmia, anemia, heart block, ischemia, inflammation, pain, hemodynamic instability, thrombosis, hypotension, and cardiac arrest with return of spontaneous circulation. after approximately 27 days of support, the impella 5.5 was surgically removed at the icu bedside. following device removal, the patient experienced immediate hemodynamic deterioration requiring maximal vasopressor support, including norepinephrine and epinephrine. loss of arterial line pulsatility was noted, and advanced cardiac life support, including defibrillation and cpr, was performed without successful resuscitation and the patient expired. note: h6. component code and investigation type/findings/conclusion coding remain unchanged as previously reported. related medical device reports: this is one of three devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump suction - data log review showed suction events occur throughout case. the cause of the pump suction was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. low or blocked pump flow - data log review confirmed flows at p-7 dip below the expected 3.9-4.5 l/min. the cause of the low pump flow was patient condition related to low volume. device in wrong position - the cause of the device in wrong position was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. optical sensor issue - the cause of the optical signal issue was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. heart block, hemodynamic instability, anemia, hypotension, pain, inflammation, cardiac arrest, sepsis, thrombocytopenia: the causes of these injuries were not determined due to insufficient clinical details. arrhythmia, thrombosis, ischemia, infection: the root cause of the injuries was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report represents the pump. a separate report was submitted to represent the automated impella controller. refer to section h.10 for the related report number.
4114 | 1220648-2025-47694 | 23359306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359306&pc=OZD | 08/28/2025 |  |  | 10/22/2025 |  |  | 08/28/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1488898 |  | 0042-0010-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male | 106 KG |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced leaking from the purge sidearm. the supporting automated impella controller did not generated any alarms. the pump was explanted and replaced with a new impella cp. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D4 and h4 updated. the investigation is still ongoing.
4115 | 1220648-2025-47696 | 23359419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359419&pc=OZD | 08/28/2025 |  |  | 10/22/2025 |  |  | 08/30/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687526 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient presented from an outside after multiple shocks on the implantable cardioverter defibrillator. the patient was placed on an intra-aortic balloon pump (iabp) for support and was being worked up for heart transplant candidacy. while on unit the patient¿s right and left ventricles function continued to diminish. the physicians made the decision to remove the labp and place protek duo and impella 5.5 for more stability during heart transplant waiting. the patient was awaiting heart recovery. approximately seven days after start of the impella mechanical circulatory support, the impella connect was coming in and out and notifying local team that a new pump is turned on and then connecting. this created false alarms for the abiomed team, particularly overnight and when not on site. the is a non-patient effect issue. support continued uninterrupted and uncompromised. continuing support, the most recent ejection fraction showed 13%. the patient had been placed on extracorporeal membrane oxygenation and noted to have remained on high ecmo and impella 5.5 support (performance level p-9) for recovery. additional information updates were received and noted the patient experienced a gastrointestinal bleed. one unit of red blood cells and one unit of platelets were administered. support continued and a plan for recovery was still noted. however, the family discussed and eventually decided to go comfort care. they ultimately decided to remove the endotracheal tube for terminal withdrawal of care. mechanical circulatory support devices were left on. the patient expired within 20 minutes of extubation. | [5 empty] | The event describes gi bleeding and product bloods were administered. the patient would subsequently expire, care was withdrawn. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. there is not enough information to dissociate the impella 5.5 device. however, in this case, patient underlying condition likely contributed most to an outcome of death. plan for recovery was still noted. however, the family decided to end care for the patient leading to demise. medical safety review was completed considering the automated impella controller and demise association and noted: the event describes a false alarm/buzzer failure that had no (direct or indirect) impact on the patient's support. the automated impella controller remained in use supporting the patient. there was no interruption or compromised to the impella mcs. in this case the patient expired for reasons (underlying condition, termination of care) unrelated to the impella aic non-patient impacting event. the event will be reported as a malfunction type of reportable event. associate director of post market clinical review noted the following: agree this occurrence does not contribute to patient outcome. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller. patient-specific information a4: weight is unknown. investigation: the impella 5.5 device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the major bleed/retroperitoneal hemorrhage cause cannot be determined since insufficient clinical details were provided.
4116 | 1220648-2025-47698 | 23359623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359623&pc=OZD | 09/06/2025 |  |  | 10/22/2025 |  |  | 09/06/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675337 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 92 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on an impella cp there was poor neurology assessment per the bedside nurse. the patient was currently running on continuous renal replacement therapy. overnight, the patient had uncontrolled atrial fibrillation requiring cardioversion and amiodarone bolus x 2. the patient received one unit of packed red blood cells. there was significant internal and external bleeding around the impella site. initial insertion was a high stick. heparin was turned off and perclose was tightened but did not resolve issue. the impella was explanted and the patient survived. | [5 empty] | Investigation summary: the investigation was completed. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4117 | 1220648-2025-47697 | 23359806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359806&pc=OZD | 09/12/2025 |  |  | 10/22/2025 |  |  | 09/12/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689288 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding and positioning issues during explantation. a stitch was dislodged on the graft during explant which opened the pocket. the issue was quickly resolved and the patient survived. | [5 empty] | B.1 added selection as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name was entered incorrectly on the initial report that was submitted and should of been left blank. investigation summary: difficult to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had large resistance preventing successful delivery of impella. product damage (catheter kink):the cause of product damage catheter kink was most likely patient condition since large resistance was met during implant. patient anatomy or condition prior to therapy: the cause of patient anatomy or condition prior to therapy was most likely patient condition since large resistance was met during implant. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4118 | 1220648-2025-47700 | 23359918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23359918&pc=OZD | 09/06/2025 |  |  | 10/22/2025 |  |  | 09/06/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712247 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Patient-Device Incompatibility | Patient Device Interaction Problem |  |  |  | Hemolysis | Hemoptysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia, causing suction events. the pump was confirmed to be in a good position. decreased flow was observed. in response, tissue plasminogen activator was added to the purge. a suction event occurred while the patient was experiencing vomiting and coughing. a bronchoscopy revealed the left main stem was occluded with blood clots. the patient also experienced gastrointestinal bleeding and coughing up blood. the purge flow was inconsistent so the purge cassette was exchanged. the patient received a transplant and was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
4119 | 1220648-2025-47701 | 23360004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360004&pc=PYX | 09/11/2025 | 11/05/2025 |  | 10/22/2025 |  |  | 09/06/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025601200 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Anemia | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp experienced suction alarms. the alarms were resolved by transfusing support blood and albumin. the patient was placed onto continuous renal replacement therapy for anuria and was transfused packed red blood cells for low hemoglobin. while patient was awaiting removal, the placement signal waveform disappeared and the flows showed 0 although the impella was still running. the impella was subsequently removed as planned. no patient harm was reported. | [5 empty] | The pump was not returned for investigation. the cause of the anemia and renal impairment was not determined due to insufficient clinical details. the cause of the placement signal issue was not determined, as the product was not returned and sufficient clinical information was not provided. the cause of the suction issue was most likely related to the patient's condition, based on the given clinical information. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4120 | 1220648-2025-47702 | 23360395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360395&pc=OZD | 09/06/2025 |  |  | 10/22/2025 |  |  | 09/06/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680818 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Anemia | Arrhythmia | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella 5.5 support for intervention for abdominal aortic aneurysm. during support the patient became hypotensive and bradycardic on two occasions. cardene stopped, was on epinephrine drops at 4 briefly, increased to p9 and received 1l lr and 500mls albumin. overnight no issues, eventually landed on p6 due to mean arterial pressure goal. transthoracic echocardiography (tte) report reads pump is positioned towards mitral valve, measuring 5cm from aortic valve. having intermittent ectopy. remains on systemic amiodarone and holding diuretics. overnight/early morning hemoglobin dropped to 6.9. gave one unit and hemoglobin back up to 7.8. suction alarm over night. dropped to p-7, gave a bolus of albumin and went back up to p-8. patient experienced position unknown alarms, got an echocardiogram (echo) and echo showed a hemipericardium clot overlying the right ventricle. the patient was taken to operating room for clot evacuation and washout. while in operating room received 4 packed red blood cells, 2fresh frozen plasma, 1 platelet, 1l crystalloid. the nurse reports overnight suction alarms and patient received 250 albumin and 500lr. weaned dobutamine off overnight. the patient was successfully weaned from pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: data log review showed a slight decrease in placement signal (ps) and loss of pulsatility during the impella position unknown alarms. 5s parameters were stable at this time. the impella position unknown alarms were occurring consistently for 6 hours during the reported time. the ps pulsatility improved as the alarms resolved. the cause of the optical sensor issue was most likely patient condition related as data log review showed a loss of ps pulsatility and clinical details. pump suction: the cause of the suction was patient condition related since clinical details note low volume status and data logs show loss of ps pulsatility. thrombosis/tachycardia/arrythmia: the cause of the injuries was unable to be determined due to insufficient clinical details. anemia: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4121 | 1220648-2025-47703 | 23360637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360637&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026726876 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. during impella cp placement after the peel-away was removed and the repositioning sheath was advanced, the patient had an ongoing hematoma and bleeding above the level of the entry site requiring over an hour of manual pressure and a femstop for full resolution of bleed. the patient did have very diseased vasculature which likely contributed to this issue. in the night, the groin was oozing, and the physician placed another stitch and bleeding stopped. additionally, there was concern for hemolysis due to red tinged urine. an echocardiogram was done and the inlet appeared to be in contact with the mitral leaflet. the medical doctor tried repositioning the impella and were unable to torque towards the apex despite numerous attempts. they placed at 3.0 centimeters with good flows on the impella. the patient continued impella support and was successfully weaned off. stable hemodynamics were noted at the time of explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in ventricle. pump then was back in good position to the end of the case. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in proper position since echo confirmed pump inlet appears to be in contact with mitral leaflet and hemolysis was improved after positioning of the pump. the cause of positioning issue was unknown. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. hematoma/bleeding: the cause of hematoma/bleeding was most likely patient condition related since the patient have very diseased vasculature which likely contributed to this issue. device history review: the device passed all the post sterile inspection checks.
4122 | 1220648-2025-47704 | 23360848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360848&pc=OZD | 09/19/2025 | 11/11/2025 |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703785 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that there was a leak from the glucose side arm. no visible cracks were observed. the leakage amount was about one drop observed over several hours. the purge flow rate was approximately 14 milliliters and purge pressure was within the 400 mmhg range, the specified range. there was no patient harm. the hemodynamics was stable and the catheter was removed the next day after percutaneous coronary intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. purge leak - pump: returned pump and cassette were visually inspected and no abnormalities were observed. the pump and cassette were purged together and no leak was able to be reproduced. the purge values were within specification. data logs not returned for review. the cause of the purge leak - pump was not determined since the leak was unable to be reproduced. device history review: the device passed all the post sterile inspection checks.
4123 | 1220648-2025-47705 | 23360954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360954&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727853 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented to the emergency department with shortness of breath which had increased over the previous two days that was associated with chest pressure. the patient was brought to the cardiac catheterization laboratory for left heart catheterization which revealed multi-vessel coronary artery disease. the decision was made to proceed with a percutaneous coronary intervention (pci). during the pci on the left anterior descending artery, the patient started decompensating and coded twice and the decision was made to implant the impella cp. approximately seven hours into impella cp support, it was reported that the patient was bleeding from the impella access site and catheter site and there were ongoing suction alarms. the patient was transfused with three units of packed red blood cells and received fluids. it was noted that the patient was not oxygenating well. the following day, it was noted that the patient¿s pulmonary capillary wedge pressure was 4mmhg and the patient was receiving fluid and albumin. the medical team planned to increase impella support when the patient¿s fluid and pressures increased. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4124 | 1220648-2025-47695 | 23360981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360981&pc=OZD | 09/05/2025 |  |  | 10/22/2025 |  |  | 09/05/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692908 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered ventricular tachycardia (vt) and atrial fibrillation (af) during impella 5.5 support. the patient had non-sustained vt so an amiodarone bolus and lidocaine was given. a few days later, the patient had sustained vt and was shocked. later, the patient experienced af. no further intervention was noted and the patient remained stable through the explant. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4125 | 1220648-2025-47699 | 23360998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23360998&pc=OZD | 09/06/2025 |  |  | 10/22/2025 |  |  | 09/06/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706426 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding and a thrombus. the clot was removed and the pump was repositioned. the patient was transfused approximately (b)(4) units of blood products. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
4126 | 1220648-2025-47706 | 23361130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23361130&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710889 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a loss of placement signal during a percutaneous intervention. the procedure was successful. the patient was weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned. a review of the data logs confirmed the complaint. the cause of the optical sensor issue was a damaged sensor from shockwave interaction since data logs show evidence of sensor damage and clinical details note that the ps was lost during shockwave use. the device passed all post-sterile inspection checks.
4127 | 1220648-2025-47707 | 23361186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23361186&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026718383 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 support experienced a pump stop during patient movement. the pump was explanted and the patient was successfully bridged to a left ventricle assist device. the patient survived. | [5 empty] | A3, a4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4128 | 1220648-2025-47708 | 23361311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23361311&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/22/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710780 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 53 KG | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted for high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during the insertion of the impella cp, the physician was unable to cross the aortic valve. the decision was made to abort the implant and remove the device. the patient survived the device being explanted and the overall outcome of the patient is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | Investigation summary: the investigation has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
4129 | 1220648-2025-47709 | 23362848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23362848&pc=OZD | 09/23/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644185 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. once thhe impella cp was inserted and started and the patient arrested several times with cardiopulmonary resuscitation in between while on impella support. additionally, hematoma was noted on the impella site. manual pressure was held, and pressure bag was placed on site. one unit of blood was administered. the patient went to ct to look for bleeding on first attempt. on 2nd attempt, ct identified what looked like a splenic bleed liver embolization was done to stop bleeding. heparin was restarted the next day. the patient ejection fraction recovered and the impella was explanted. | [5 empty] | Correction: d4: the investigation of the reported failure mode has been completed. no product was received for evaluation. cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma: hematoma noted upon inspection of impella site. hold manual pressure/place pressure bag on site. the cause of hematoma is determined to be use issue due to patient movement because hematoma was noted after transfer to icu. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4130 | 1220648-2025-47710 | 23362882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23362882&pc=OZD | 09/22/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1250845 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) console experienced an optical sensor error during support. the aic was swapped for a new unit to continue patient support. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. root cause: the exact root cause of the optical sensor errors is undetermined, but the presence of an intermittent air gap between the patient pump cable and the optical pump connector could point to a potential use issue.
4131 | 1220648-2025-47711 | 23362957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23362957&pc=OZD | 09/23/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679129 |  | 000000000001000100 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced acute kidney injury and anemia. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the anemia and acute kidney failure was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4132 | 1220648-2025-47712 | 23362988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23362988&pc=OZD | 09/23/2025 |  |  | 10/22/2025 |  |  | 09/23/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709949 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to the patient¿s anatomy. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device passed all post sterile inspection checks.
4133 | 1220648-2025-47714 | 23363041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363041&pc=OZD | 07/22/2025 |  |  | 10/22/2025 |  |  | 09/24/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696190 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced thrombocytopenia. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4134 | 1220648-2025-47715 | 23363084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363084&pc=OZD | 09/26/2025 |  |  | 10/22/2025 |  |  | 09/26/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1587053 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) failed to boot up during routine maintenance. no patient is involved in this event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4135 | 1220648-2025-47716 | 23363113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363113&pc=OZD | 09/16/2025 |  |  | 10/22/2025 |  |  | 09/26/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673984 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fracture | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a purge system open alarm. the connection between the purge filter and the pressure reservoir broke into two pieces when the patient was being moved. the impella was replaced with another pump to continue patient support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4136 | 1220648-2025-47717 | 23363131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363131&pc=PYX | 09/27/2025 |  |  | 10/22/2025 |  |  | 09/27/2025 | 10/22/2025 |  |  | 02/28/2026 |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025518059 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 58 KG | Hispanic |  | Death | Yes | No | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pulmonary Edema | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced bleeding from the endotracheal tube and chest tube. as a result, heparin was withheld. additionally, a placement signal not reliable alarm occurred. the impella flow calculation was disabled and the position of the impella was confirmed. it was further reported that the patient went into flash pulmonary edema and was unable to oxygenate. later, the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleeding and edema was not determined due to insufficient clinical details. the cause of the placement signal issue was determined to be due to a sensor drift as evidenced by log pattern.
4137 | 1220648-2025-47718 | 23363146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363146&pc=OZD | 09/29/2025 |  |  | 10/22/2025 |  |  | 09/29/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629094 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while on impella cp support, the patient experienced hemolysis. continuous renal replacement therapy (crrt) was started and blood was given to the patient. hemolysis continued and there was a placement signal alarm. the following day care was withdrawn and the patient expired. an additional report will be sent for a automated impella controller (serial (b)(6)) for the same patient. | [5 empty] | B2 added required intervention as it was omitted from the initial report that was submitted. e1 added initial reporter phone number as it was omitted from the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product returned. mechanical interaction with blood - after three days of support, the patient experienced hemolysis. imaging showed accurate pump positioning. hemolyzed lab results occurred through explant. data log review showed no motor current spike noted outside of p-level changes. no suction alarms were noted. impella positioning unknown alarms triggered on (b)(6) 2025 and then resolved. the cause of the mechanical interaction with blood could not be determined due to no pump return, inconclusive data log review, and insufficient clinical details. optical sensor issue - after eight days on support, placement signal (ps) low alarms occurred. the patient noted to not be a surgical candidate and had poor prognosis at the time of ps alarms, and withdrawn from care one day later. data log review showed a period of low signal-to-noise ratio (snr) while the pump was run on an automated impella controller (aic). however, this had no impact to the ps nor ps low alarms. after snr recovery when exchanged on a different aic, case ends with frequent ps low alarms. ps low alarm ranging from as low as 6 mmhg diastolic to 7 0mmhg systolic. the cause of the optical signal issue was patient condition related due to ps deteriorate over the course of the case without impact on pump metrics. the patient noted to have poor prognosis during ps low alarms and withdrawn from care within one day of ps low compliant. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4138 | 1220648-2025-47719 | 23363264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363264&pc=OZD | 09/30/2025 |  |  | 10/22/2025 |  |  | 09/30/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739877 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant had impella cp pump placed to support the patient admitted in with a non-st-segment elevation myocardial infarction and history of underlying cardiac comorbidities. the cp was placed to allow for heart rest prior to a planned bypass in following days. the cp was explanted and replaced after just over 5 hours due to suction that could not be resolved. the 2nd cp is also captured by an additional report complaint: (b)(4). | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: data logs were reviewed and the suction was confirmed. the pump was flowing steadily at p-9 when there was a sudden drop in motor current (mc) (both max and min) and a systolic suction alarm triggered. there were no mc spikes or ms deviations. this event aligned with a drop in psmin values as well, though the waveform remained aortic. pump flows remained generally low for the rest of the time the patient was supported by the pump. device analysis: returned product was investigated and after soaking the pump, there was no evidence of biomaterial. there was a minor kink in the cannula noted near the motor housing. no other defects or damages were noted on the pump. the pump was then run in a bucket of water, and low flows did not reproduce. though there is potential that the cannula kink was the cause of low pump flows, clinical details report that the pump was repositioned with imaging a number of times and no cannula kink was noted at this time. since data logs were inconclusive, returned product did not reproduce low pump flows, and insufficient clinical details were provided, the cause of the reported pump suction could not be determined. device history review: the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | After review of the case by medical safety, it was determined the impella cp pump 1 is a serious injury type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | Medical safety/clinical reviewed the event. the patient was a 77-year-old female with a past medical history of hypertension, hyperlipidemia, and congestive heart failure. she presented to the emergency department with a non-st-elevation myocardial infarction (nstemi) and was noted to have mildly elevated troponin levels and a lactate of 3.2 mmol/l. on the following day, the patient was taken to the cardiac catheterization laboratory for left heart catheterization. after administration of sedation, the patient became profoundly hypotensive. the decision was made to initiate mechanical circulatory support. left ventriculography demonstrated an estimated ejection fraction of approximately 25%, and severe multivessel coronary artery disease was identified. the patient was referred and accepted for surgical intervention at a later date. an impella cp was placed for mechanical circulatory support. following placement, persistent suction alarms were noted that could not be resolved with repositioning or fluid administration. the treating physician suspected possible clot ingestion. the impella cp (pump 1) was subsequently removed and replaced with a second impella cp (pump 2). approximately 12 hours later, csc was contacted to assist with recurrent suction alarms. abrupt suction events were observed, with motor current flattening and flows decreasing to approximately 0.1 l/min. the performance level was reduced from p-6 to p-2, with some pulsatility noted on motor current; however, no improvement in flows was observed. the patient was anticoagulated but not at therapeutic levels at that time. based on the clinical presentation, clot ingestion was again suspected. an echocardiogram was ordered to confirm device position. subsequently, motor current flattened again and flows abruptly decreased to approximately 0.1¿0.3 l/min. suction could not be resolved despite additional troubleshooting measures. following these events, the patient experienced cardiac arrest. advanced cardiac life support was initiated; however, the family elected to discontinue resuscitative efforts and transition the patient to comfort care. the patient continued to deteriorate and subsequently expired. the impella cp pump 1 is a serious injury, and the impella cp pump 2 is a death type of reportable event. h6 investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
4139 | 1220648-2025-47713 | 23363600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23363600&pc=OZD | 08/05/2025 | 10/22/2025 |  | 10/22/2025 |  |  | 09/30/2025 | 10/22/2025 |  |  |  |  | 10/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026631783 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported on september 30, 2025, after a cross functional review with both the durable and disposable investigation teams, the following action has been identified. alleged ¿placement signal issue¿ (possibly false positioning alarms), ) could have been due to the pump. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | No product or data logs were returned for evaluation. the cause of the optical signal issue could not be determined as no product or logs were returned for evaluation and limited clinical details were provided.
4140 | 1220648-2025-47641 | 23346914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23346914&pc=OZD | 09/27/2025 |  |  | 10/21/2025 |  |  | 09/27/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727856 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced thrombotic thrombocytopenic purpura. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product returned. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report represents the pump. another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number.
4141 | 1220648-2025-47640 | 23346998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23346998&pc=OZD | 10/01/2025 |  |  | 10/21/2025 |  |  | 10/01/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299521 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Power Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when attempting to turn off the automated impella controller, the console would not turn off and needed a hard turn off. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report.
4142 | 1220648-2025-47642 | 23347625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23347625&pc=PYX | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661908 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 104 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella rp flex had a placement signal not reliable alarm and flow calculations were disabled. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was discarded h6 code updated. d6b was updated as additional information provided explant date.
4143 | 1220648-2025-47643 | 23347882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23347882&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661700 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 72 KG | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support during transfer from bed to chair when the patient¿s hand got caught on white impella catheter. downward force on catheter was noted. the catheter separated from the red impeller plug causing impella stoppage. the patient was hemodynamically stable and had no issues. the medical doctor explanted at bedside. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Investigation summary: the investigation has been completed. no product returned. pump stop/product damage (red plug detachment): clinical details noted that during transfer from the bed to a chair, the patient's hand got caught on the catheter and pulled down on it, which caused the catheter to separate from red pump plug and a pump stop. data log review showed an alarm pump stop with a motor current spike over 30000 ma. the cause of the pump stop was an electrical open as data logs showed an electrical pump stop and the clinical details noted that the patient inadvertently pulled on the catheter. the cause of the product damage (red plug detachment) was use related as clinical details noted that the patient inadvertently pulled on the catheter, separating it from the red pump plug. device history review: the pump passed all post-sterile inspection checks.
4144 | 1220648-2025-47644 | 23347966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23347966&pc=OZD | 09/23/2025 |  |  | 10/21/2025 |  |  | 09/23/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728960 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during placement of impella 5.5(the fourth pump for this patient), it was noted that the patient's chest was left open with retractor in due to irretractable ventricular tachycardia. the patient was later explanted and survived. this report represents the fourth pump additional reports will be submitted for the other pumps. | [5 empty] | The investigation was completed and no product was returned. arrhythmia: root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4145 | 1220648-2025-31319 | 23348064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23348064&pc=OZD | 07/04/2023 | 10/21/2025 |  | 10/21/2025 |  |  | 07/04/2023 | 10/21/2025 |  |  | 05/31/2025 |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351817 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump displayed a ¿placement signal not reliable¿ alarm. pump support continued without interruption, and there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned for evaluation. data logs were returned and evaluated. per the given clinical information and data log review, the root cause of the optical signal issue was most likely a damaged/bent optical fiber line. all pertinent information available to abiomed has been submitted at this time.
4146 | 1220648-2025-31879 | 23348108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23348108&pc=OZD | 01/05/2023 | 10/21/2025 |  | 10/21/2025 |  |  | 01/05/2023 | 10/21/2025 |  |  | 07/31/2025 |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023254772 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the pump experienced an optical sensor failure after the device was repositioned during support. both the placement signal and lv signal stopped working as a result of the failure. patient support continued with the installed device as the motor current, flows, and purge pressure remained within acceptable limits. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the returned pump was connected to an automated impella controller (aic) and placement signal not reliable reproduced. the pump was soaked and water but the alarm persisted. a loose patient cable to kink protector bond was observed with the patient cable able to rotate freely. the cause of the optical signal failure was a damaged optical fiber line during torqueing the device due to a loose patient cable to kink protector bond. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4147 | 1220648-2025-31880 | 23348132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23348132&pc=OZD | 12/14/2023 | 10/21/2025 |  | 10/21/2025 |  |  | 12/14/2023 | 10/21/2025 |  |  | 10/31/2025 |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411504 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that fifteen days into impella 5.5 support, the pump triggered high purge pressure alarms and stopped unexpectedly. the pump was removed as a result of the noted failure. the patient later succumbed to their condition and passed away. it is unknown if the pump stop contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4148 | 1220648-2025-47645 | 23348271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23348271&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738793 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella cp support due to acute myocardial infarction/ cardiogenic shock. while on support, an echocardiogram revealed a fixed mural mass consistent with a thrombus with ejection fraction of 15%. it was successfully managed as the patient was on systemic heparin. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation was completed and no product was returned. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and passed all the post sterile inspection checks.
4149 | 1220648-2025-47648 | 23348672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23348672&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629310 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support had arrhythmia and then started getting suction alarms. the left ventricle appeared to be decompressed, and volume helped. additionally, the patient developed hemolysis. the impella was repositioned to treat the hemolysis, but the patient had blood colored urine and worsening renal function. the team opted to explant the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The pump was not returned for review. the data logs confirmed the complaint. the cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of pump suction was most likely patient condition related since the suction was resolved after fluid compensation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the cause of the renal failure was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4150 | 1220648-2025-47647 | 23348777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23348777&pc=OZD | 08/23/2025 |  |  | 10/21/2025 |  |  | 08/23/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1587787 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella console (aic) was having intermittent alarms for purge pressure increase that quickly resolve. purge pressure was maintained. it was reported that the patient was rolling and moving a lot in bed and might be putting pressure on the purge line. registered nurse (rn) was instructed to try and keep the catheter clear and free from possible impingement and make sure the patient was not rolling over on the catheter. the patient continued on support until the patient was successfully bridged to orthotopic heart transplant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4151 | 1220648-2025-47649 | 23349186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23349186&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690026 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 65 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma in both the radial and groin sites. manual pressure was held to obtain hemostasis. a couple hours later a hematoma was noted post bowel movement. manual pressure was held again and then an angiogram was performed on the right common femoral access site. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the hematoma is determined to be patient condition because the patient developed hematoma at multiple locations. the cause of the anemia was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4152 | 1220648-2025-47650 | 23349434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23349434&pc=OZD | 08/26/2025 |  |  | 10/21/2025 |  |  | 08/26/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687163 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 68 KG | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Difficult to Advance | [5 empty] | Low Blood Pressure/ Hypotension | Heart Failure/Congestive Heart Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 encountered difficulty in delivery of the device and hypotension upon explantation. septic shock was suspected so inotropes were weaned and pressors were titrated. tissue plasminogen activator was added to the purge in response to high purge pressure alarms. a transesophageal echocardiogram identified a thrombus in the left atrium, which had collapsed. the pump's p level was adjusted to improve pulsatility. the patient eventually expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary no product was returned for evaluation. hypotension, cardiac failure: clinical details noted that transesophageal echocardiography (tee) showed left ventricle (lv) and left atrium (la) were collapsed. the cause of the injury was not determined due to insufficient clinical details. difficultly to deliver: clinical details noted that pump was unable to cross lv on a wire. pump was successfully able to be inserted wirelessly via direct aortic access. the cause of the difficulty to deliver pump could not be determined as limited clinical details were provided. high purge pressure: clinical details noted that overnight on day 12 of support, purge flows were below 3.5 and purge pressures were escalating. tissue-type plasminogen activator (tpa) was started in purge and in the morning, purge flows increased to 12.8ml/hr. data logs were reviewed, and lt logs showed a sudden increase in purge pressure and decrease in purge flow. rt logs show no "high purge pressure" alarms were triggered and a sudden step up in purge pressure was observed. purge flow ticks at that time became elongated. no motor current spike was observed. the cause of the high purge pressure was due to a kink in the purge tubing; however, limited clinical details were provided at time of purge pressure increase and the exact location and cause of the kink could not be confirmed as no product was returned. device history review pump passed all post-sterile inspection checks.
4153 | 1220648-2025-47653 | 23349730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23349730&pc=PYX | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026700176 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | Renal Impairment | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered acute kidney injury and embolism during impella rp flex support. during support, continuous renal replacement therapy was started as the patient had acute kidney injury and liver shock. additionally, a computed tomography scan showed the patient had multiple focal emboli in the brain and the patient had not been waking off of sedation. the decision was ultimately made to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation was completed and no product was returned for investigation. acute kidney failure/ embolism: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
4154 | 1220648-2025-47655 | 23350425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23350425&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727879 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that on the first day of impella cp support, following the patient being moved for a central line placement, the impella alarmed for placement signal low. a transthoracic echocardiogram (tte) was performed, and it confirmed that the impella remained in the correct position within the heart despite the alarm. the patient¿s care team contacted the abiomed customer support center to discuss options. based on the tte result and that the alarm was likely caused by the patient's movement, the team decided to leave the impella as it was and the team concluded that the alarm was aberrant. the following day, it was reported that the impella cp had intermittent suction alarms. it was noted that the patient had 2 liter positive fluid balance and the patient was transfused with a unit of packed red blood cells. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | B1 adverse event should not have been reported on the initial report as there was not adverse event. b2 required intervention should not have been reported on the initial report as there was not adverse event. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. code b17 should not have been reported on the initial report as the device was not kept for investigation. h11 the device was discarded and should of been reported on the initial report. investigation summary: the investigation has been completed. no product was returned. optical sensor issue: clinical details noted that there was a placement signal low alarm on the first day of support. transthoracic echocardiogram confirmed the impella was in good position. hemodynamics were stable. data log review confirmed placement signal low alarms intermittently throughout the case. there were no abnormalities found with the optical parameters. there were also no issues with g0 calibration. without returned product, the reading of values by the optical sensor cannot be tested. the cause of the optical sensor issue was not determined since product was not returned and insufficient clinical details were provided. pump suction: clinical details noted that there was intermittent suction alarms on (b)(6) 2025. data log review showed intermittent suction alarms throughout the case (alarm 14, diastolic). there were no motor current spikes and purge pressure was stable. there were no positioning alarms seen in the logs during the case. the cause of the suction was not determined since insufficient clinical details were provided and no product was returned. device history review: the pump passed all post-sterile inspection checks. | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4155 | 1220648-2025-47651 | 23350449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23350449&pc=PYX | 09/13/2025 |  |  | 10/21/2025 |  |  | 09/13/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025507450 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Decrease in Pressure | Device Sensing Problem | [5 empty] | Anemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with a pulmonary embolism was implanted with an impella rp flex for mechanical circulatory support. the impella rp flex was placed and ran for six days till right ventricle failure improved. the pump did have low purge pressure that prompted purge bag replacement, and there was optical signal loss. the patient also had one unit of red blood cells infused during the support for low anemia . | [5 empty] | E4 revised selection as it was previously entered incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. the pump was not returned for investigation. data log was not provided or could not be retrieved from the remote server. placement signal issue: the cause of the placement signal issue was not determined without returned product investigation and sufficient clinical information, including data logs. purge pressure low: the cause of the low purge pressure issue was not determined without returned product investigation and sufficient clinical information. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4156 | 1220648-2025-47652 | 23350453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23350453&pc=OZD | 09/10/2025 | 11/28/2025 |  | 10/21/2025 |  |  | 09/10/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722269 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Hematuria | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was tachycardia, vascular dissection, and hematuria during impella cp patient support. while on support, there was a noted dissection in the left main coronary artery. the patient was sent for cat scan surgery. additionally, the patient¿s urine was a light pink color. an echocardiogram was performed with no findings noted. the patient had non sustained run of supraventricular tachycardia and was put on amiodarone drops. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported vascular dissection, tachycardia, and hematuria has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular dissection, tachycardia, and hematuria could not be determined.
4157 | 1220648-2025-47656 | 23350718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23350718&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716279 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 86 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing at the access site which required frequent dressing changes. no further complications were reported. the patient's health declined and a decision to withdraw care was made. the patient expired. | [5 empty] | The cause of the bleeding and cardiovascular insufficiency was not determined since product was not returned and insufficient clinical details provided. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4158 | 1220648-2025-47658 | 23354297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354297&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729326 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was planned to be implanted with an impella cp device for mechanical circulatory support. the impella 5.5 was not able to be positioned properly due to limited vessel size. location of resistance was the aortic arch. a decision was made to replace 5.5 with impella cp. there was no patient harm. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had limited vessel size and small aortic arch preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4159 | 1220648-2025-47657 | 23354299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23354299&pc=OZD | 09/21/2025 |  |  | 10/21/2025 |  |  | 09/21/2025 | 10/21/2025 |  |  |  |  | 10/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731888 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 91 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was found unresponsive after no family contact for two days and taken to the hospital by emergency medical services. initial laboratory results indicated a critically low glucose and hematocrit, and a very high troponin result. it was noted that there was an absence of any cranial bleeding. the decision was made to transfer the patient to an outside hospital for further treatment. upon arrival, the patient was intubated and taken to the cardiac catheterization laboratory for impella cp placement. prior to impella placement, the patient was transfused with 3 units of red blood cells. the impella was inserted with no reported issues. a coronary angiogram was performed and revealed a moderate lesion in the left anterior descending coronary artery. the physician decided to not attempt a percutaneous intervention at that time, as there was concern for unknown origin of anemia. the decision was made to continue to assess the patient, and the patient was released to the intensive care unit (icu). upon arrival to the icu a bedside echocardiogram was done to verify impella placement. the impella position was found to be shallow with the pigtail likely under the papillary muscle. the physician opted to not reposition the device and it was noted the care plan was to transfuse the patient further and correct metabolic acidosis. on the third day of impella support, the patient was placed on continuous renal replacement therapy, with noted very high ammonia levels. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | D4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. e1 added initial reporter address line 1 as it was omitted from the initial report that was submitted. h6 added medical device problem code. investigation summary: the investigation has been completed. no product was returned. renal failure: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4160 | 1220648-2025-47620 | 23334772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23334772&pc=OZD | 09/22/2025 |  |  | 10/20/2025 |  |  | 09/21/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707218 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 68 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was escalated from an impella cp to an impella 5.5. during impella 5.5 support, the patient was off continuous renal replacement therapy due to line clotting and was unable to return blood. as a result, the patient received packed red blood cells. it was also reported that the patient had four to five runs of ventricular tachycardia and potassium was replaced. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. | B3 revised date of event as it was entered incorrectly on the initial report that was submitted. d6b explant date was added. g3 date was entered incorrectly on the initial report that was submitted and should of reflected 22-sep-2025. investigation summary: the investigation has been completed. thrombosis/ tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4161 | 1220648-2025-47621 | 23335077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23335077&pc=PYX | 09/23/2025 |  |  | 10/20/2025 |  |  | 09/23/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025567885 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 91 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp flex to support a patient admitted in with known right heart failure and transplant earlier in the month. the condition post transplant warranted the impella rp flex and it was placed but removed after just over two hours of support due to low flows that could not be successfully troubleshot. the team removed the impella rp flex and the decision was made to place a protek duo. | [5 empty] | The impella rp flex was returned for evaluation. the returned product revealed no kinks in the cannula or damages seen. there was biomaterial located underneath the impeller. the biomaterial was removed, and the pump ran with no issues. based on data log and returned product analysis, the root cause of the low pump flows is due to biological material ingestion. the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
4162 | 1220648-2025-47622 | 23335207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23335207&pc=OZD | 09/26/2025 |  |  | 10/20/2025 |  |  | 09/26/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687253 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 115 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient became hemodynamically unstable after going through impella cp support and subsequently impella 5.5 support. after a few hours of impella cp support, the patient was escalated to an impella 5.5 and transferred from south bay hospital to largo medical center. the patient encountered hematuria, ventricular fibrillation, ventricular tachycardia, and hypoxia during impella 5.5 support and will be noted in a report for the impella 5.5. ultimately, care was withdrawn and the patient expired. the death is currently associated with the first pump (impella cp). this report represents the first pump. another report will be submitted for the second pump (impella 5.5). | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Medical safety/clinical reviewed the event. a 71-year-old male presented to the emergency department with st-elevation myocardial infarction (stemi) and was emergently transferred to the cardiac catheterization laboratory. during the initial coronary angiogram, the patient decompensated and required one round of cardiopulmonary resuscitation (cpr), with return of spontaneous circulation (rosc). an impella cp was successfully implanted for mechanical circulatory support, resulting in hemodynamic stabilization. percutaneous coronary intervention of the left main into the left anterior descending artery was performed. the patient was subsequently escalated to an impella 5.5 for continued support. on (b)(6) 2025, the patient underwent a bedside left lower extremity fasciotomy for reported compartment syndrome related to an intraosseous line with wound vacuum placement. the patient later experienced ventricular arrhythmias, atrial flutter, fever (antibiotics regimen adjusted), hypoxia, and required mechanical ventilation and continuous renal replacement therapy. despite continued mechanical circulatory and supportive care, the patient¿s condition deteriorated. a decision was made to withdraw care, and the patient expired. the impella 5.5 will be reported as death type of reportable event. however, based on available information, the patient's underlying condition (including progression of cardiogenic shock [scai stage d], acute myocardial infarction) contributed most to the patient¿s outcome. correction: regarding this impella cp that was initially associated with the patient¿s death, this device was placed emergently with the intent to stabilize the patient prior to transfer to an outside facility for escalation to impella 5.5 support. there were no reports or indication of device malfunction, performance issues, or impella cp-related adverse events. in contrast, the impella 5.5 was in use at the time of the patient¿s death. as no complaint was identified for the impella cp and no device-related issues were reported, the event has been reassessed and determined to be not reportable for the impella cp. the medwatch follow-up correction is being submitted to redact this previously reported event in error. the death is associated with the impella 5.5 device. refer to the related manufacturer report number(s) as noted in section h10. no further information will be provided pertaining to this device. | The investigation was completed and no product was returned. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history review completed and passed all the post sterile inspection checks.
4163 | 1220648-2025-31873 | 23335209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23335209&pc=OZD | 05/13/2024 | 10/20/2025 |  | 10/20/2025 |  |  | 06/24/2024 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150574 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the pump lost optical signal. the team was discussing if the loss was from damage during insertion. the pump remained on for 31 hours and was then weaned/explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that intermittent ¿placement signal not reliable¿ condition started presenting shortly after pump start, and its occurrences intensified over time. in addition a ¿low signal-to-noise ratio (snr)¿ condition (and associated alarm # 130) kept recurring frequently. there were also two isolated instances of ¿loss of opm data¿ as evidenced by event 1045 and dropped samples of opm data - unrelated to the primary failure mode. console was received in danvers and tested - it passed ¿5s¿ specification with an engineering test cavity (cl = 14340nm) but failed it with a pump sensor (cl = 15910). monitoring optical bench charge-coupled device (ccd) analog output on an oscilloscope revealed a defect that was affecting placement signal calculation, due to low snr. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4164 | 1220648-2025-47623 | 23335253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23335253&pc=OZD | 09/22/2025 | 11/20/2025 |  | 10/20/2025 |  |  | 09/22/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Pumping Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient was prepped for impella support due to asystole. during preparation, the impella was started outside of the patient¿s body. an impella defective alarm was displayed. a second pump was opened and the same message was displayed. a new aic was turned on, the second pump was connected with no alarms, and the case proceeded. | [5 empty] | The impella device was received from the customer. upon completion of the investigation, a supplemental mdr will be filed. | Correction: d3, e4 the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: console logs from the reported day of event are mostly consistent with complaint and show ¿impella defective¿ alarm (#3, due to errors reading pump eeprom). device analysis: console was tested in the lab, and intermittent errors during pump eeprom reading were reproduced with a known-good engineering pump. controller error alarm related to optical bench was not reproduced, however corresponding data was consistent with loss of light in the optical system, not necessarily due to defective lamp. console inspection revealed that optical pump connector was heavily contaminated. root cause : ¿impella defective¿ alarm with pump, was due to pump eeprom communication issues caused by malfunction of the impellatronic board (epm sub circuitry). controller error alarm was due to momentary loss of light in the optical system was most likely caused by optical bench malfunction (exact component could not be identified). device history review: the device passed all post sterile inspection checks.
4165 | 1220648-2025-47625 | 23336421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23336421&pc=OZD | 08/26/2025 |  |  | 10/20/2025 |  |  | 09/26/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692545 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female | 70 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | Renal Failure | Thrombocytopenia | Thrombosis/Thrombus | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp for the support of a patient admitted in with acute myocardial infarction/cardiogenic shock. during support the patient subsequently had severe hemolysis with acute renal failure and the need for continuous renal replacement therapy. starting with impella implantation intermittent signs of malperfusion of the left leg (cooler than right leg, intermittently marbled). concomitant with bleeding tendency: massive transfusion. evidence of heparin-induced thrombocytopenia type ii. subsequently, pronounced malperfusion of the left leg under agatroban medication. malperfusion persisted and thrombotic closure of the pelvic vasculature and the femoral fork on the left diagnosed: removal of the impella and thrombectomy of the pelvic vasculature and the femoral fork as well as patchplastic of the left arteria femoralis communis. in the course of the next day, resuscitation was required again and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia/thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure/thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4166 | 1220648-2025-47626 | 23336441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23336441&pc=OZD | 09/24/2025 | 02/13/2026 |  | 10/20/2025 |  |  | 09/24/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703544 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 104 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | Ischemia | Tachycardia | Hemodynamic instability | [8 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). it was reported that the patient presented to the emergency department with complaints of shortness of breath and excruciating chest pain since having a colonoscopy the prior day. the patient was taken emergently to the cardiac catheterization laboratory for a left heart catheterization which revealed multi-vessel disease involved the left anterior descending (lad), diagonal, and right coronary artery. an impella cp was successfully placed for mcs, and the patient developed ventricular tachycardia that caused profound hypotension, which was successfully cardioverted after the first percutaneous transluminal coronary angioplasty of the lad. multiple attempts were made unsuccessfully to open the coronary arteries, but all attempts were unsuccessful. the patient was taken urgently to the operating room for a coronary artery bypass graft (cabg) procedure. after a successful cabg with three grafts placed, the patient was released to the intensive care unit for rest and recovery. on the third day of support, the decision was made to transfer the patient to another hospital for further treatment and an escalation of support to an impella 5.5. it was reported that during transport, the patient decompensated and there were suction alarms from the impella. upon arrival to the hospital, the patient was on power level 2 of support on the impella cp with very low mean arterial pressure and the patient coded. the patient was placed on veno-arterial extracorporeal membrane oxygenation for additional hemodynamic support. it was noted that the patient had no neuro response when off sedation and there were no pulses in either foot. the patient¿s care was withdrawn, and the patient expired. | The complainant reported additional information upon request: "they ended up withdrawing care. pump unavailable." |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | B5 and h6 updated. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: usb returned for evaluation. pump suction: the clinical details state that there was suction alarms and the patient was decompensating after start of ground transport between hospitals. additional information in the outreach calls presented that they gave 250 cc of fluid during the transport. versed, ketamine, fentanyl were all administered during the transport, as the patient had low blood pressure before they left. the data logs show that during the case the optical sensor was functioning appropriately. there were multiple suction alarms throughout support before the transport began with most being suction alarm #14 appearing to be triggered due to a potentially lower preload as it was thought that the patient was potentially dry and the patient was experiencing diastolic suction. there were no suction alarms noted for about 7 hours until transport began. around the suspected time of transport, the placement signal (ps) begins to decline less than 65mmhg and then suction alarms began to trigger likely from patient instability based on the ps and motor current (mc) values. throughout the time of transport there appeared to likely be volume intervention which was confirmed by the field team to maintain map > 65mmhg as the few suction events would temporarily resolve, ps would increase and pump flow would temporality increase after the suspected bolus. around the end of the transport, the ps decompensates and low pulstaility suction alarms are triggered during that time. based on the clinical details provided from the field team and physicians, as well as the trends seen in the data logs, the root cause of the pump suction is due to the patients condition. ischemia/tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest/hypotension/hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1947602. device history review: impella 613032 passed all post sterile inspection checks.
4167 | 1220648-2025-31874 | 23337002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23337002&pc=PYX | 09/13/2024 | 10/20/2025 |  | 10/20/2025 |  |  | 09/13/2024 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025536974 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after insertion on impella rp flex post tricuspid ring, the placement signal and motor current became unreliable and would not read. motor current showing appropriate waveform, placement checked several times under x-ray, and echocardiography images show impella is in good position. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the clinical details state that there was a mitral valve and tricuspid ring procedure right before implant. right after insertion it was noted that the placement signal and motor current became unreliable and would not accurately read. position was noted to be good. the data logs show alarm #1052 epm sensor failure and #1051 dp signal out of range. the logs also show placement signal trending upwards before going out of range to 3000mmhg. the cvp also goes out of range and flow calculation was disabled. there were continuous suction alarms and motor current spikes the product was returned and the pump passed laser continuity test and the 5s parameters were within specification, indicating the optical sensor was appropriately functioning. psnr did not produce when connected to the console. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4168 | 1220648-2025-31875 | 23337009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23337009&pc=OZD | 03/24/2025 | 10/20/2025 |  | 10/20/2025 |  |  | 05/29/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm was disabled during impella cp support with an automated impella controller (aic). the performance of the pump remained unaffected, and support continued with the implanted device. the decision was made to withdraw care, and the patient expired. the patient¿s death is captured in both the complaint and the report for the impella cp. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: note: the reported failure mode psnr was associated with the (b)(6), which is the second console associated with the reported case, whereas the first console is (b)(6). ¿ the case associated with the reported complaint was initiated on (b)(6) 2025, with the pump sn: (b)(6). ¿ on the complaint date of (b)(6) 2025, during support, the console displayed ¿placement signal not reliable (opm invalid signal, optical pump connected) ¿ alarm,¿ event #1036. ¿ the rt log shows the contrast oscillating between -3276 and 3274, and the signal-to-noise ratio (snr) dropped to zero. the lt log shows that the oscillating contrast was never resolved. this confirms the reported failure mode. ¿ additionally, there was a ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm,¿ event #1045. ¿ the rt logs confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) were represented as -16384 values due to the crc error. device analysis: this console was tested after arriving at field service. visually confirmed that the lamp was intermittently turning off by itself. when the lamp was off, 4.9 v was measured at the pbm j3 connector, an average of 4.8 v was measured at the p4 connector of the optical bench, and the lamp resistance measured 5.1 ohms. no issues were found with the analog output from the optical bench. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). root cause: ¿ the root cause of the psnr alarm was a lamp malfunction. ¿ the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4169 | 1220648-2025-31876 | 23337139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23337139&pc=OXD | 01/06/2025 | 10/20/2025 |  | 10/20/2025 |  |  | 01/06/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OXD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555795 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support. while on support, placement signal unknown alarms occurred. echocardiography performed showed proper placement, mid-left ventricle pointed towards apex, inlet measuring 5.5cm from atrioventricular (av). the physician did not feel the need to reposition or make any changes. furthermore, the patient had a slight ooze at the access site. the staff was aware and monitored. moreover, the anticontamination sleeve broke during support. the site was no longer sterile for repositioning. support continued. the abiomed medical safety officer reviewed and determined there was no console exchange or interruption in therapy or information that suggests such for the optical signal/placement signal issues during the ~6-month support duration. the patient expired for causes non-related to the impella devices, and impella devices (pump and automated impella controller) are not an associated factor in the demise outcome. the patient was waiting on a transplant and subsequently had been determined unable to get the transplant due to their condition. the patient declined any further line placement or advanced treatment. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cut pump was returned for investigation. no investigation could be performed due to the cut catheter on the returned product. the data log shows downward placement signal drift followed by the spike to 3000 with a placement signal not reliable (psnr) alarm on 75th day of use. no abnormalities were noted in the optical signal parameter. the cause of the placement signal issue was sensor silicone wear beyond the intended design life of 60 days, based on data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4170 | 1220648-2025-31877 | 23337148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23337148&pc=OZD | 01/05/2023 | 10/20/2025 |  | 10/20/2025 |  |  | 01/05/2023 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1292324 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal was unreliable. the protected percutaneous coronary intervention was successfully supported until patient could be weaned from the impella. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs from the reported day of event are consistent with complaint and show multiple error messages related to invalid placement signal samples due to low snr (ossiopsens set invalid_bad_snr_sample_mask), followed by a ¿placement signal not reliable (optical signal-to-noise ratio) - alarm issued¿, event #190. rtlog data from ic5199 shows that the raw optical signal oscillates between 0 and an unrealistic value of -3000. similarly, the placement signal oscillates between 0 and an unrealistic value of 3000. contrast levels also dropped below 20%, and the signal-to-noise ratio (snr) fell below 150. device analysis: console ic5199 was sent back to fs for further investigation. the root cause of the placement signal not reliable was a malfunction of the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4171 | 1220648-2025-47628 | 23339807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23339807&pc=OZD | 09/20/2025 |  |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support was presented with hemolysis and leg ischemia. additionally, the patient failed a wean attempt due to low oxygen saturation, requiring the addition of milrinone to support cardiac output. the patient was re-intubated due to agitation, and aspartate aminotransferase levels exceeded 1000, indicating possible hepatic injury. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation was completed. the device was discarded by the customer. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
4172 | 1220648-2025-47629 | 23340433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23340433&pc=OZD | 09/20/2025 |  |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719347 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 108 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with chest pain. the patient was found to have multi vessel disease. a decision was made for the impella cp to be placed. percutaneous coronary intervention was performed. however, it was noted that prior to leaving the catheterization lab the patient coded twice. the patient was taken to the intensive care unit for heart rest and recovery. a hematoma was reported at the insertion site. manual pressure was held and per the nurse the groin seemed soft. the patient was transfused with eight units of blood. sutures were reapplied around the access site and oozing had improved. ultimately, the patient was made do not resuscitate. per the family wishes all inotropes, ventilator, and impella support were turned off and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation was completed and no product was returned. hematoma : the cause of the injury is most likely use related since sutures were reapplied by md around access site and oozing has improved. cardiac arrest : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
4173 | 1220648-2025-47630 | 23341579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23341579&pc=OZD | 09/20/2025 |  |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628386 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was brought to the emergency department with chest pain. the impella cp was placed for cardiogenic shock and percutaneous coronary intervention was completed. in the intensive care unit, the patient suffered multiple ventricular fibrillation arrest episodes and the staff was unable to achieve return of spontaneous circulation. therefore, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The device was discarded and the investigation was completed. investigation summary: ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks.
4174 | 1220648-2025-47631 | 23341980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23341980&pc=OZD | 09/20/2025 |  |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026652520 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female | 63 KG | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that there was a pump flow issue with the impella cp once it had been inserted. it was noted as the unload was not possible, flow 0. the decision was made to explant and use a new device. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | D9 has been updated to reflect that the product was returned for investigation. h6 type of investigation, investigation findings, and investigation conclusions codes and h11 have been updated to reflect that the product was returned and the investigation is underway. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | Product complaint #(b)(4). investigation summary: impella cp sn (b)(6) + purge cassette returned for investigation. data review: data logs confirmed the failure mode & clinical narrative. logs showed after pump started & run for 2 minutes, mc spiked followed low flow that triggered auto suction response & suction alarms. this event remained to the rest of the case. product analysis: visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. low or blocked pump flow: the cause of the low or blocked pump flow was biomaterial ingestion. device history lot: device lot #: 1912976. device history batch: subcomponent lot#: n/a. device history review: pump sn: (b)(6) passed all post sterile inspection checks.
4175 | 1220648-2025-47632 | 23343508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23343508&pc=OZD | 09/20/2025 | 10/20/2025 |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026718418 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 104 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 while supporting a patient had placement signal not reliable alarm with absence of aortic and left ventricle placement signals. the alarm was silent. the patient continued support and was bridge to left ventricular assist device. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: data log review confirmed the reported issue. the cause of the optical signal issue can be traced to the console based on data log review. no defect was found with the pump based on data logs. device history review: the device passed all post sterile inspection checks
4176 | 1220648-2025-47633 | 23343514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23343514&pc=OZD | 09/20/2025 |  |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720050 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombocytopenia | Heart Block | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient on had questionable complete heart block and amio and precedex gtts were discontinued. the patient experienced torsades. cpr started until defibrillation converted back to sinus rhythm (sr). additionally, platelets were dropping possible heparin-induced thrombocytopenia (hit). the patient was switched to argatroban. patient was supported by device for almost 10 days , and patient is stable post issue. | [5 empty] | Section 6b: the explant date is unknown at this time. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia :the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia/ heart block : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4177 | 1220648-2025-47634 | 23343517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23343517&pc=OZD | 09/20/2025 | 02/18/2026 |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654227 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient (pt) was on impella 5.5 for heart recovery. during support it was noted of neuro deficit on pt assessment. observed decreased movement on patient's entire right side of body. a computed tomography (ct) scan of the head was conducted and revealed focal hypodensity in the left side of the brain. neurologist was consulted and decided that area affected was minimal and did not require neuro intervention. in addition, it was noted that ct head was in progress due to new neuro deficit. patient returned to operating room for intercostal bleed that needed repair, massive transfusions necessary. they noted that cerebrovascular accident on ct was stable on serial imaging. however, not moving right side, left side moving, left eye opening. the 5.5 was successfully weaned and explanted. | The complainant reported additional information upon request: "the pump was weaned and explanted. the team did not provide any further clinical information beyond device explant. the pump was discarded by the team and is not available for return." |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 and h6 updated. | The investigation for the reported major bleed and stroke has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of the major bleed and stroke could not be determined.
4178 | 1220648-2025-47635 | 23343636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23343636&pc=OZD | 09/20/2025 | 12/09/2025 |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026741325 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 70 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Inflammation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support was having episodes of profound hypotension. epinephrine, and sodium bicarbonate (bicarb) were given. additionally, the patient developed hematoma on the impella access site that was growing. the physician believed that the impella site could have been bleeding. two units of blood were administered. a couple of days later, the groin insertion site was tender and stable, but scrotal edema remained severe. the patient continued support and was explanted as planned. | [5 empty] | The investigation for the reported hematoma, major bleed and hypotension/ inflammation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma, major bleed and hypotension/ inflammation could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4179 | 1220648-2025-47637 | 23343720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23343720&pc=OZD | 09/20/2025 | 12/09/2025 |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support lost post tibial pulses. the patient remained on anticoagulation throughout device placement. cat scan angiography of groins and ultrasound was completed. | [5 empty] | The investigation for the reported ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ischemia could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4180 | 1220648-2025-47638 | 23343997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23343997&pc=OZD | 09/20/2025 | 12/04/2025 |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722336 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 87 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, the patient had a run of atrial fibrillation with heart rate in 180's. lidocaine was provided and patient was cardioverted. the next day, patient experienced hemineglect with head cat scan showing right side infarct. care was ultimately withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported paroxysmal atrial fibrillation has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of the paroxysmal atrial fibrillation could not be determined.
4181 | 1220648-2025-47639 | 23344005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23344005&pc=OZD | 09/20/2025 | 12/26/2025 |  | 10/20/2025 |  |  | 09/20/2025 | 10/20/2025 |  |  |  |  | 10/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026728966 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 78 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Cardiac Arrest | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Pain | Tachycardia | Hematuria | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the patient experienced ventricular tachycardia which required chest compressions lasting three minutes, however did not require intubation or medication during the code. a few days later the patient was successfully weaned and received a left ventricular assist device. | The user facility reported patient was placed onto impella support due to cardiomyopathy. while on support, patient presented hematuria. heparin infusion was paused. impella was repositioned due to concerns for hemolysis and persistent hematuria. right foot pain limited patient from ambulation; gout is suspected and patient was put on colchicine. pump presented suction alarms for which patient received albumin and one unit packed red blood cells. patient experienced gi bleed and was taken off anticoagulant. imaging was taken the next day and no evidence of gastrointestinal bleed was found. patient was placed back onto low dose heparin. patient experienced additional suction and fluid was ordered. |  |  |  |  | Additional information received from the clinical site and the following fields were updated based on the new information. b.7 added information that was omitted from the initial report that was submitted. d.4 revised udi number as it was previously entered incorrectly on the initial report that was submitted. e.1 added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported optical sensor issue, pump suction, hemolysis, tachycardia/ cardiac arrest, and hematuria/ major bleed has been completed. no device was returned for review. data log review showed a motor current shift upward (no spike) 30 minutes prior to impella position unknown and in aorta alarms. at the time of the mc shift, there was spiking in the raw optical signal that continued at the time of the alarms. the 5s parameters were stable otherwise. the returned product did not have any abnormalities related to the optical sensor. the root cause of the pump suction alarms were patient condition related as clinical details note low volume status. the root cause of the optical sensor issue, hemolysis, tachycardia/ cardiac arrest, and hematuria / major bleed could not be determined
4182 | 1220648-2025-31870 | 23331509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23331509&pc=OZD | 02/07/2025 | 10/17/2025 |  | 10/18/2025 |  |  | 02/07/2025 | 10/18/2025 |  |  |  |  | 10/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383536 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Obstruction of Flow | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that upon attempting to prime the setup, no purge fluid was observed flowing out of the pump assembly. it was noted that the purge line had been successfully primed prior to the pump being attached. the pump was replaced to resolve the issue. there was no patient harm. automated impella controller replacement was not required. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: there were no signs in the imc or lt/rt logs of no purge fluid flowing as described in the complaint. (b)(4) started with purge cassette 172319 and cp pump 561659. the console successfully primes but the user does not complete step 7 starting the case. the console is power cycled, and the same pump and cassette skip case wizard since the pump is already initialized but the user does not turn on the pump. the console is power cycled again and new purge cassette (161940) and pump (545263) are inserted and again the console primes successfully but this time the user completes the purge wizard and starts the pump. the rt log confirms the console successfully primed with purge flowing with the first pump. device analysis: the complaint was not reproduced during testing. purge and pump were primed successfully and fluid was flowing as expected. root cause: there was no issue found with purge fluid not flowing during the case. there was no issue found with the optical signal during the case. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4183 | 1220648-2025-31871 | 23331528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23331528&pc=OZD | 03/08/2025 | 10/17/2025 |  | 10/18/2025 |  |  | 03/08/2025 | 10/18/2025 |  |  |  |  | 10/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614262 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that impella 5.5 had placement signal not reliable alarm. impella flows remained the same and positioned was verified by transesophageal echocardiogram. patient expired on support - the death is captured in the sister record. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed on 03/08. at this time, signal-to-noise ratio (snr) drops to zero, lamp maxes out at 100%, light decreases, gain decreases, and contrast oscillates between +/-3,000. the rt logs were closely examined during the time of the psnr event, and the oscillating contrast trend was seen to have recovered intermittently. the pump was used 15 hours after the first psnr was reported. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4184 | 1220648-2025-47573 | 23322399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23322399&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026664364 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Cardiac Tamponade | Thrombosis/Thrombus | [6 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella 5.5 for mechanical circulatory support. while on impella 5.5 support the patient experienced arrhythmia, thrombus, tachycardia, cardiac tamponade, bleeding, hypotension, and cardiac arrest. the patient was a paced at 115. the patient went into atrial fibrillation with rapid ventricular response. an amiodarone bolus was administered and the patient's blood pressure dropped. with further decompensation, there was an escalation of vasopressors. levophed and phenylephrine were added. an echocardiogram was suggested to rule out tamponade or bleeding. the echocardiogram showed a posterior wall clot and concern for tamponade. 500 cubic centimeters bolus and one unit of packed red blood cells were administered. another echocardiogram confirmed a large clot and tamponade. the patient was taken to cardiovascular operating room with the doctor to attempt to open the sternum to relieve the tamponade. there was a large amount of blood and a clot. despite the clot removal, the patient progressed into pulseless electrical activity arrest. despite multiple rounds of epinephrine, the patient was unable to get return of spontaneous circulation and patient expired in the operating room. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. arrhythmia/ tachycardia/thrombosis/ cardiac arrest/ cardiac tamponade: the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
4185 | 1220648-2025-47574 | 23322718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23322718&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720778 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was explanted in routine fashion. upon explant a clot was expressed from the access site. the patient was taken to cardiac catheterization lab and contralateral access was obtained. an angiogram was performed which revealed a thrombus in the profunda. a thrombectomy was performed with excellent results. distal perfusion was assessed and were adequate post procedure. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation was completed and no device was received for evaluation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device history review was completed and passed all post sterile inspection checks.
4186 | 1220648-2025-47575 | 23322812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23322812&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722333 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 105 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient noted as receiving a high-risk coronary artery bypass graft was implanted with an impella 5.5 device for mechanical circulatory support post-surgery. the patient presented to the hospital with chest pain and elevated troponins. the patient was brought to the cardiac catheterization laboratory and was found to have triple-vessel disease, and the left ventricle was severely depressed. an intra-aortic balloon pump (iabp) was placed, and the patient was transferred to another hospital for further treatment and possible coronary artery bypass graft surgery. the patient underwent three coronary artery bypass grafts, and following the procedure, the impella 5.5 was placed directly through the aorta, and the iabp was weaned and explanted. on the sixth day of support, it was reported that the patient experienced ventricular fibrillation the previous night and was defibrillated three times and magnesium was given to resolve the arrhythmia. the patient¿s pacemaker was reviewed via echocardiography, and the medical team determined that the patient would require an automatic implantable cardioverter defibrillator. there were no further reports of arrythmia, and the device was subsequently weaned and explanted successfully. the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: the investigation has been completed. ventricular fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4187 | 1220648-2025-47577 | 23323801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23323801&pc=PYX | 09/16/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025571370 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. while on support, there was a placement signal not reliable alarm. motor current is unchanged so no action taken to address the issue and support continued without any further reported issues. . patient feels great, and his breathing is much better. no patient harm reported due to the issue. later after few days of support there was purge issue for which purge cassette was changed. the purge cassette is reported under separate report. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. placement signal issue: data logs were reviewed and the issue was confirmed. based on the patterns noted in data logs, the cause of the placement signal issue was determined to be dp sensor drift. device history review: the device passed all the post sterile inspection checks.
4188 | 1220648-2025-47580 | 23323836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23323836&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725427 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a delivery issue during impella cp implant. due to tortuous iliacs, the case was aborted. it was noted that the impella catheter couldn¿t get past the femoral artery due to vessel tortuosity and the sheath got kinked. there was no reported patient harm. the kinked sheath (lot number s9596062) will be reported on an additional manufacture report. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 added code b13 as it was omitted from the initial report that was submitted. update b17 code to b18 as new information was received that the pump was disposed of as the patient was questionable for a blood borne illness. h10 added related report number. investigation summary: the device was discarded and the investigation has been completed. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had iliacs that were too tortuous preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
4189 | 1220648-2025-47579 | 23323956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23323956&pc=OZD | 09/12/2025 | 10/16/2025 |  | 10/17/2025 |  |  | 09/18/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1803713 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 94 KG |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support the patient was ambulating up to chair, on return to bed impella console unexpectedly shut down. the nurse was able to turn it back on. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4190 | 1220648-2025-47587 | 23324881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23324881&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716951 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp insertion was aborted due to the patient's anatomy. no patient harm was reported. | [5 empty] | The investigation was completed and no product was returned for investigation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4191 | 1220648-2025-47588 | 23324971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23324971&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702590 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male | 100 KG |  |  | Malfunction | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced position issues when the impella slipped into the aorta. additionally, there was bleeding from the right groin and behind the peritoneum. the complication was managed through surgical intervention and three units of red blood cells were transfused. the pump was explanted and exchanged for a new impella. | [5 empty] | No product was returned. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position during the transport. the device passed all post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4192 | 1220648-2025-47589 | 23325132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23325132&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639888 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Cardiac Arrest | Hemolysis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient who had diagnostic left heart catheterization with severe disease noted in his coronary tree. the team decided to implant an impella cp for support as the patient was in cardiogenetic shock. during support, labs returned hemolyzed with dark red urine and multiple suction alarms were noted. the impella appeared too deep with "placement signal low" alarms. the physician attempted bedside repositioning, and despite cautiously backing out the catheter as it was pulled too far back into the aorta. while repositioning the alarms flipped from "in aorta" to "in ventricle". repositioning was paused but the alarms continued back and forth. the pump appeared to be in aorta rather than the ventricle with subtle pulsatility noted on the motor current. an echocardiogram was called to guide advancement, however, due to poor imaging windows availability the canula could not be visualized well/at all. after several attempts and advancing the catheter seven centimeters, the team decided to reposition in the catheterization lab. the patient continued to decompensate and was maxed out on pressor and respiratory support. the patient arrived to the lab in grim condition with mottled extremities, and oxygen stats were in the 70's to low 80's and a blood pressure was in the 40's/30's, and only detectable by impella. the pump was visualized with inlet well above the aortic valve but was easily repositioned. unfortunately, the patient coded while the lab and did not respond well to resuscitation. after a short period of time all supportive measures including impella were discontinued and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation was completed. data review: data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in ventricle/aorta. logs also showed multiple ps low and suction alarms triggered when pump was in improper position. no other issues were found on the logs. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was in improper position since the pump was pulled back too far to aortic area. device in wrong position: the cause of device in wrong position was most likely use related since the pump was pulled back too far to aortic area. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
4193 | 1220648-2025-31867 | 23325204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23325204&pc=OZD | 05/17/2025 | 10/17/2025 |  | 10/17/2025 |  |  | 05/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606213 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 89 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp suffered from ventricular tachycardia and limb ischemia. the limb was treated by perfusion and fasciotomy, but the patient ultimately expired. | [5 empty] | The root cause of limb ischemia could not be determined since the product was not returned and insufficient clinical details were provided. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4194 | 1220648-2025-47590 | 23325312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23325312&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1434325 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a console with a controller error upon start up. the yellow alarm caused no harm. the automated impella controller was replaced and no patient harm was reported since no patient was involved. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | In section g (manufacturing site name), the abiomed europe gmbh site address and required site information were inadvertently omitted from the initial report.
4195 | 1220648-2025-47591 | 23325606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23325606&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026720134 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 54 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that implantation of an impella 5.5 was aborted due to the patient's anatomy. no patient harm was reported. | [5 empty] | The investigation was completed. no product was returned for investigation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4196 | 1220648-2025-47593 | 23325725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23325725&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725979 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced drops in purge flow and high purge pressure alarms. troubleshooting was unsuccessful so the pump was explanted. no patient harm was reported. | [5 empty] | Product complaint (b)(4). the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The product was returned and evaluated. there was no evidence of a kink or damage found on the returned patient cable or catheter. high purge pressure was reproduced and resolved when cutting the catheter above the motor housing. there was some biomaterial observed in the purge gap. the cause of the high purge pressure was a kink since data logs showed an abrupt purge pressure rise with no motor current spike/trend and clinical details noted the patient cable being pinched in the bed railing but the exact location is unknown due to a lack of evidence on the returned product. the cause of the product damage was use related since clinical details note that the patient cable was pinched in the bed railing. the device passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4197 | 1220648-2025-47595 | 23325823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23325823&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612030 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that the patient presented to the hospital with worsening heart failure. an intra-aortic balloon pump was placed via axillary artery during a left ventricular assist (lvad) work up, however the following days the patient¿s hemodynamics worsened. the patient¿s ejection fraction was calculated to be 10% with a decreased mixed venous oxygen level of 47%, therefore the decision was made to escalate the patient to an impella 5.5 device. during implantation of the impella 5.5, it was reported that the impella was unable to advance past the axillary graft. the pump was removed to un-tunnel the graft, and it was noted that the pump had a white clot-like substance on the outlet. the pump was replaced and the second impella 5.5 was inserted via axillary artery and support was started with no issues, however within the first 7 hours of support the pump abruptly stopped functioning twice. each time, the pump was able to be restarted following the high motor current alarm with pump stop, however the p level would not go above p1 or p2 before stopping again. additionally, it was reported that the patient experienced hematuria. a decision was made to remove and replace the pump. the patient was brought back to the operating room for an impella 5.5 replacement. upon removal of the second pump, a clot was observed on the inlet of the pump. additionally, the axillary artery 8mm hemashield graft was found to have clots within it. the physician suspected the reason of the clots was due to suspected heparin induced thrombocytopenia of the patient. the graft was removed and replaced with a new 10 mm hemashield platinum graft which was anastomosed to artery. the new third impella 5.5 was implanted without issue, and the patient was released back to the intensive care unit (icu). upon returning to the icu, the patient was noted to have a facial droop and left side weakness. the patient was taken for an emergency computed tomography scan where a clot was found, and the patient was taken for an emergent thrombectomy. the following day the patient was noted to have no deficits from the thrombotic event. additionally, it was noted that the impella 5.5 was not well placed, and the surgeon was notified that the impella would need to be repositioned. the following day, the impella positioned was discussed. it was recommended to the surgeon to reposition the impella as the inlet was facing the mitral valve, and the pump was slightly deep at 6.4 cm. although the concern that the malposition could contribute to ectopy, the physician declined to reposition stating that the ventricle was large and was satisfied with the pump position. two days later, the pump was repositioned and pulled back to 5.1 cm and the power level was able to be increased from p-7 to p-8. the patient was placed on a two-week hold for additional surgery due to the stroke. the following day, it was reported that there were several impella in aorta and impella position unknown alarms which occurred during the previous night. echocardiography was performed to verify impella position, and impella position was confirmed to be well placed. on the seventh day of impella support, it was reported that the automated impella controller (aic) experienced an unexpected shut down. no alarms had been observed, and the medical team replaced the aic to resolve the issue. on the fourteenth day of support, it was reported that there were ongoing false placement signal alarms, and the medical team disabled the alarm. the following day it was reported that the patient was experiencing gastrointestinal bleeding. anticoagulation was withheld due to the ongoing bleeding and low hemoglobin. a unit of packed red blood cells were transfused as result of blood loss. on the eighteenth day of support, it was reported that there were placement signal and suction alarms. it was noted that there were no signs or symptoms of hemolysis due to the suction alarms, and that the patient¿s plasma free hemoglobin was <30 with adequate yellow urine output. on the twenty-fifth day of support, it was reported that the impella pump was repositioned due to the patient experiencing several runs of ventricular tachycardia. the patient remains on impella 5.5 support, with a plan for surgery for a durable lvad in the next week. this complaint involves four impella devices: pump 1: impella 5.5 with serial number: (b)6). pump 2: impella 5.5 with serial number: (b)(6). pump 3: impella 5.5 with serial number: (b)(6). automated impella controller with serial number: (b)(6). this report specifically addresses pump 1: impella 5.5, with serial number: (b)(6). separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | A4 and a6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary 5.5 pump, sn: (b)(6) was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had small and deep axillary artery preventing successful delivery of impella. device history review device (sn: (b)(6)) passed all post sterile inspection checks. | This follow up report is being submitted to report information for device return and the device evaluation. d9 updated for product return. h6 type of investigation updated to include testing of actual device. investigation summary: 5.5 pump, sn: (b)(6) was returned. unable to deliver or advance (delivery issue): cannula kink was observed on the returned product. the cause of the deliver issue was determined to be patient condition as the patient had small and deep axillary artery preventing successful delivery of impella.
4198 | 1220648-2025-47597 | 23326753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23326753&pc=PYX | 08/21/2025 | 10/17/2025 |  | 10/17/2025 |  |  | 08/21/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661955 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 123 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp flex and impella cp (bipella) for mechanical circulatory support following mitral valve and tricuspid valve surgery. due to the patient¿s severe pulmonary hypertension and poor right ventricular function, the decision was made to place the impella rp flex status post surgery. during chest closure, prior to impella rp flex implant, the patient¿s hemodynamics worsened and the decision was made to implant the impella cp after the impella flex was inserted and support initiated. the right internal jugular vein was accessed for the impella rp flex, and the impella rp flex was implanted without issues. the impella rp flex position was noted to be optimal when imaged with fluoroscopy, however it was noted that the placement signal of the impella rp flex was unreasonably high with a pulmonary artery catheter pressure higher than expected and pump flow was lower than expected for the p level of the pump. the placement signal issues did not resolve with adjusting the y-axis of the pump, and the medical team was advised by abiomed support that there was a potential issue with the optical sensor which was causing an inaccurate flow calculation. the impella cp was placed in the right femoral artery without issue, and the patient was brought back to the intensive care unit for rest and recovery. the patient then received volume resuscitation, after which the placement signal and flows returned to normal and expected levels. it was advised to the medical team by abiomed support that the optical sensor had likely been impeded by a vessel wall and correlated with hypovolemia. the following day, the patient was transferred to another hospital for advanced care. it was noted that the patient¿s right leg was cold, and pulses were unable to be found with doppler. additionally, a transesophageal echocardiogram was performed and revealed blood in chest causing right atrial compression. a plan was made to obtain computed tomography of the head, chest, and pelvis and then take the patient to the operating room for a chest washout, right femoral artery cutdown, and remove both impella pumps. if needed, the patient would then be cannulated for extracorporeal membrane oxygenation (ecmo). after the chest washout was performed, the patient¿s hemodynamics improved, and it was determined the patient did not require ecmo after impella pumps were explanted. the impella cp was emergently removed due to critical limb ischemia. ultimately, the patient required above knee amputation of the right leg due to complications of the limb ischemia. after both the impella rp flex and the impella cp were explanted, it was decided that the patient did not need to be cannulated for ecmo as the patient was hemodynamically stable and heart function had recovered enough for the patient to be stabilized with pressors and intubation. this complaint involves two impella devices: pump 1: impella rp flex with serial number (b)(6). pump 2: impella cp with serial number (b)(6). this report specifically addresses the impella rp flex. another report will be submitted for the other device associated with this complaint. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary major bleed: the cause of the injury was not determined due to insufficient clinical details. optical signal issue: the data logs confirm high placement signal and show placement signal not reliable (optical sensor out of range) alarms. per data log review, optical signal issue occurred 15:40 on (b)(6) 25 but resolved within an hour. at 20:00 the differential pressure (dp) sensor issue with psnr alarm occurred and did not resolve. the logs show there was negative cvp and suction alarms with decreased pump flows. clinical information mentions patient was hypovolemic indicating reduced volume and increased risk and that pump was likely impeded by a vessel wall as a result of suction. volume was added which resolved the suction and unreasonably high placement signal issue. the cause of the optical sensor issue was determined to be due to patient condition. placement signal issue: the logs show a sudden loss of placement signal and flow calculation at 20:00 on (b)(6) 25. placement signal not reliable (psnr) alarms were triggered. the 5s parameters were stable and issue did not resolve for the remainder of support. the log pattern identified did not match previously identified failures. the cause of the issue was not established as no product was returned for analysis. device history lot ==> device lot number: 1910086 device history review ==> (b)(4) revealed a particle larger than 0.2mm was present in the pouch for pump (sn: (b)(6). the affected pump had a rework order (1936187). this is unrelated to the reported failure mode.
4199 | 1220648-2025-47598 | 23327286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327286&pc=PYX | 09/17/2025 | 12/04/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661859 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 99 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of an impella rp flex that the patient had a hematoma at the impella rp flex access site of the right femoral vein. repositioning unit was adjusted and a suture was added for hemostasis. the patient was weaned off of the impella rp flex support. the medical doctor was happy with the improvement. post-operative surgery on the right-side support was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported hematoma has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of the hematoma could not be determined.
4200 | 1220648-2025-47599 | 23327587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327587&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710423 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. following insertion, the patient was noted to have an issue with distal perfusion in the right leg so a reperfusion line was utilized via the right-radial artery back towards an antegrade arterial stick in the right femoral artery. additionally, the patient was still on vasoactives and discussed this with the physician, but there were concerns over thrombocytopenia, neutropenia, possible hemolysis, and distal perfusion driving a more rapid wean. red blood cells and platelets were ordered. the patient was weaned from support and survived. | [5 empty] | H6 code e0506 was erroneously entered on the initial manufacturer device report number. code a24 was erroneously entered on the initial report and updated to a01. investigation summary: the investigation was completed. the product was not returned for review. data review: data logs showed the pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current spikes, abnormal placement signal or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
4201 | 1220648-2025-47600 | 23327730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327730&pc=OZD | 09/16/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697044 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 144 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during the course of support, the patient experienced a gastrointestinal bleed. consequently, heparin was temporarily withheld, and an esophagogastroduodenoscopy was performed, during which the bleeding site was identified and clipped. additionally, it was noted that the impella device was not securely positioned and migrated into the aorta when the patient was repositioned in bed. the impella was subsequently adjusted to p2, as verified through transesophageal echocardiogram and was explanted on the same day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was returned. major bleed: the cause of retroperitoneal hemorrhage (gi bleed) cannot be determined since insufficient clinical details were provided. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was not secured and upon moving in the bed impella migrated into aorta. device history review: the device passed all the post sterile inspection checks.
4202 | 1220648-2025-47592 | 23327747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327747&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  | 01/31/2026 |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460566 |  | 000000000001000080 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost distal flows past the repositioning sheath. in response, the pump was explanted. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4203 | 1220648-2025-47601 | 23327771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327771&pc=PYX | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026725991 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Fever | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient designated as a high-risk surgical case was implanted with an impella rp flex for mechanical circulatory support. during the support period, an alarm indicating an unreliable placement signal was triggered, and the patient subsequently developed a fever. it was noted that the placement signal not reliable alarm was activated, leading to the audio being silenced. furthermore, the patient's fever occurred during the period of support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4204 | 1220648-2025-47602 | 23327824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327824&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722308 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 99 KG | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the cassette and bag were changed but it remained high with low purge flow. tissue plasminogen activator was ordered. the team monitored closely. additionally, there was a concern for hemolysis following increase in extracorporeal membrane oxygenation speed and diuresis. the team was aware. also, bivalirudin was stopped due to bleeding. the patient was bridged to heart transplant. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. purge pressure high/purge system blocked: data log review showed a gradual rise in purge pressure (pp) over almost two hours. pp was near 1000 mmhg for ~1 hour before dropping to around 900 after the cassette change for about two more hours. pp then gradually decreased to 600 mmhg over two more hours. and remained stable for the rest of the case. the cause of the high purge pressure and blocked purge system was biomaterial buildup since data logs showed a gradual pp increase with no motor current (mc) spike/trend. hemolysis: data logs showed one suction alarm on 15-sep-2025 and no positioning issues. there were no mc spikes. the cause of the hemolysis was not determined since insufficient clinical details were provided and product was not returned. major bleed: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4205 | 1220648-2025-47603 | 23327981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23327981&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740381 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of an impella cp a hematoma was developing in the groin and vascular was to be consulted. arterial blood gas was acidotic with low hemoglobin and elevated potassium were noted. orders for two units of packed red blood cells, calcium, bicarb and d50. the plan is to wean pressors as tolerated and assess pulses via doppler study and recheck labs after transfusion. there are no impella related concerns or questions currently. about eight hours later, care was withdrawn and the patient expired. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned for investigation. major bleeding: a new hematoma was developing in the groin and vascular had to be consulted. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4206 | 1220648-2025-47604 | 23328150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23328150&pc=OZD | 09/09/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653075 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | False Alarm | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella cp device for mechanical circulatory support. on day 6 of an 8-day period of support for a patient with mitral valve regurgitation (mvr), a complaint was raised in japan regarding the cp pump used to assist the patient receiving peripheral cardiac circulation support (pccs). during this time, a low purge volume alarm was triggered. the medical team initiated troubleshooting by replacing the pump controller; however, this action did not resolve the alarm issue. ultimately, the patient successfully weaned off support on day 8. | [5 empty] | In section g (manufacturing site name), the abiomed europe gmbh site address and required site information were inadvertently included in the initial report, although this information was not required. | H6 revised medical device problem code since the initial report was submitted. type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. no product returned. false alarm - after six days on impella support, purge volume low alarm occurred with 100 milliliters (ml) left remaining in the purge bag. no purge leaks were observed. data log review showed purge volume low alarms occurring intermittently throughout case. purge pressure noted to be on the lower end of expected range and purge flows are elevated to 25 milliliters per hour plus. the purge volume low alarm will trigger when less than 121 ml remains in the bag if the bag size was 1000 ml. since the bag size is unknown, it was unable to be determined if the alarm triggered correctly or falsely. the cause of the false alarm issue with purge volume low alarms was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was discarded by the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4207 | 1220648-2025-47605 | 23330155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330155&pc=OZD | 09/07/2025 |  |  | 10/17/2025 |  |  | 09/07/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690359 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female | 97 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | Hematoma | Pain | Tachycardia | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced complications requiring intervention. the patient originally presented to emergency department with progressive shortness of breath and orthopnea worsening over time. the patient had a workup and was transferred to an outside hospital for further evaluation. the transthoracic echocardiography (tte) was indicative of newly diagnosed acute systolic heart failure secondary to peripartum cardiomyopathy. during hospitalization, rapid response was called for level of consciousness and brief tonic-clonic seizure activity attributed to decrease venous return from lying flat. a tte showed worsening left ventricular dysfunction with an ejection fraction of 15%, left ventricular end-diastolic dimensions were 7.7 cm, mild to moderate mitral regurgitation. the patient was started on diuretic therapy and transferred to advanced maternal care and fetal medicine. ongoing intra-disciplinary discussions regarding maternal decompensation due to high risk of acute cardiac events with significant morbidity and mortality with risk of intrauterine fetal demise. a right heart catheterization was performed, and the patient was transferred to unit and started on systemic epinephrine. cardiothoracic surgery was consulted. according to the team, after careful consideration between high-risk fetal medicine, cardiothoracic surgery, heart failure, and cardiac anesthesia, the decision was made to support with axillary impella 5.5 support until neonatal viability pending high risk obstetrics decision to continue with cesarean section possibly four days later. the patient would be evaluated for advanced heart failure therapies postpartum pending course and delivery. the patient arrived at the operating room on room air and systemic epinephrine at 2mcg/min. extracorporeal membrane oxygenation (ecmo) team was on standby for decompensation upon induction. induction was performed. esmolol was given after intubation. epinephrine was titrated to 4mcg/min and initiation of vasopressin at 0.03 and norepinephrine at 4mcg. transesophageal echocardiogram showed severely dilated left ventricular with an ejection fraction of 15%. the right ventricular function reduced. a right axillary incision was performed; the artery was identified and isolated. it was noted 5,000 units of heparin were given. baseline activated clotting time (act) was 154 seconds. the graft was sewn, and hemostasis was assured with adequate systemic anticoagulation. the introducer sheath was placed, and a .035 guidewire was advanced through the introducer. the impella 5.5 was eventually backloaded onto the .035 guidewire and advanced and good position was noted with inflow 5cm below aortic valve area. the guidewire was removed, and support was initiated and titrated to optimal support at performance level p-8 with flow at 4.8lpm. the patient was returned to unit for post operative management. high-risk obstetrics team was on the unit to manage maternal and fetal care. the next day after start of the support, the patient went into supraventricular tachycardia (svts) 140-160s, unresponsive to adenosine x3, converted after third dose of digoxin. the patient remains on low dose precedex. the plan was for tentative cesarean section at 23 weeks with hopes to extend out. however, it was noted that if at any time the patient decompensates, the delivery will be emergently completed in the unit. two days later, a ¿small¿ hematoma to site was noted. pressure dressing was applied, and systemic anticoagulation was on hold for 12 hours which resolved the hematoma. three days later, suction and position unknown alarms. the patient required one shock and was placed on lidocaine drip. teams have decided this day to proceed with cesarean section and prep for emergent ecmo cannulation. transesophageal echocardiogram showed the impella pump in good position and depth. left ventricular function was marginally better than on 5.5 implant. the patient experienced svt and adenosine x2 was administered. the incision was closed, and the patient was returned to the unit intubated with plans to extubate in a few hours. the baby was delivered at 22 weeks; however, the baby expired the following day. two days later an abdominal scan showed an anterior hematoma on the uterus but no active extravasation. one unit of packed red blood cells was given. the following two days, echocardiogram was completed and showed the impella measuring 6.3cm from the aortic valve. physicians were aware and made the decision not to reposition. the patient had some ectopy while ambulating but was not symptomatic with it. through the continued days of support, the patient experienced arrhythmic events, some self-resolved and some treated with medication. additionally, the patient had mild chest pain and neck pain at pulmonary artery catheter site as well as flank and abdominal pain. the latest update received notes the impella support continues for the patient now at day 41 (from implant) noting the patient remains listed unos (united network for organ sharing) status ii. of note, the patient experienced a pelvic hematoma. however, this is due to the cesarean section. | [5 empty] | A medical safety review was completed. arrhythmia experience by the patient cannot be ruled out as impella related given the noted suction event. however, in this case, the arrhythmia experienced by the patient is more likely due to the patient¿s heart failure and low ejection fraction. the baby was delivered because of arrhythmia and additionally, the supraventricular tachycardia could have been brought on by the pregnancy. the baby delivered alive at only 22 weeks and, however, expired the following day. the medical team's goal was to deliver at 24/28 weeks but due to the mother's condition, the medical team decided to deliver. the mother remained on impella support post-delivery. this is a serious injury type of reportable event. device status: the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added zip code extension as it was inadvertently omitted from the initial report that was submitted. additional information was received and noted on day 43 of impella mechanical circulatory support (mcs), the patient had runs of ventricular tachycardia on the night shift that self-resolved. the following morning report noted there were no impella alarms or concerns. the patient has been ambulating daily, and there are no changes in care plans. the patient remains on impella 5.5 mcs awaiting heart for transplant. | Additional information was received and note the patient underwent an uneventful orthotopic heart transplant. as a result, the impella 5.5 device was explanted (after approximately 64 days of support) with noted reason as a bridge to transplant and survival outcome at explant is patient survived. d6b was updated to reflect the date of explant accordingly. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the tachycardia, arrhythmia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. for the pain, the root cause of the injury was unable to be determined due to insufficient clinical details and hematoma, the cause of hematoma was use issue related since hemostasis was achieved by holding the heparin. pertaining to the pump suction, no product was returned. the patient had alarming suction and position unknown briefly. after reviewing the logs, multiple suction alarms were observed. no motor current or placement signal trend was observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. device in wrong position (positioning issue): the cause of positioning issue cannot be determined since insufficient clinical details were provided. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation.
4208 | 1220648-2025-47594 | 23330549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330549&pc=PYX | 09/17/2025 | 10/17/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025575268 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | Pulmonary Edema | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the patient had durable left ventricular assist device ( lvad) placed in (b)(6) 2025. the patient was admitted recently with infection. the patient developed right ventricular failure (rvf) with noted suction events on lvad. the physician elected to bring patient for rhc and impella rp insertion. in icu, the placement signal increased and impella flows decreased. motor current remained stable, and no acute spikes were observed. flows remained lower than expected throughout patient support. additionally, the impella inaccurate flow calculations were noted most likely due to a damaged differential pressure sensor. the plan was to remove the impella due to sustained low flows but when weaning impella the patient became unstable with pulmonary congestion/edema. the patient care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. low or blocked pump flow/ placement signal issue: the data logs confirm reduced pump flow with increased placement signal. the cause of the issue was determined to be due to a sensor drift of the differential pressure sensor as evidenced by log pattern. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4209 | 1220648-2025-47606 | 23330554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330554&pc=OZD | 09/17/2025 | 10/17/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700774 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Respiratory Failure | Thrombocytopenia | Thrombosis/Thrombus | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support on 16 sep 2025 for optimization prior to aortic valve replacement. patient was taken to operating room due to continued mediastinal hemorrhage. clot was evacuated, inferior mesenteric artery was ligated with sutures, and blood products were given, followed by washout and chest closure. the patient experienced acute hypoxic hypercapnic respiratory failure, associated with volume overload, for which he was reintubated. the patient went into atrial fibrillation, treated with amiodarone. patient experienced ventricular tachycardia following coughing spells. additional antiarrhythmic medications were given. impella support was increased due to hypotension. patient was also placed onto continuous renal replacement therapy for aggressive fluid removal. heparin was paused due to trach site bleeding. heparin was paused due to suspected thrombocytopenia. impella was explanted with noted difficulty, including moderate bleeding from the graft and requiring one unit packed red blood cells. explanting physician suspects device adherence to the vessel wall. | [5 empty] | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product returned for evaluation. major bleed, respiratory failure, hypotension, thrombocytopenia: clinical details noted that patient had bleeding from chest that required packing and additional blood products. additionally, patient experienced respiratory failure and thrombocytopenia while on support. and support of pump was increased due to patient experiencing hypotension. the cause of the injuries was not determined due to insufficient clinical details. thrombosis, atrial fibrillation, tachycardia: clinical details noted patient had clot in r/l chest. additionally, patient experiencing atrial fibrillation (afib) and ventricular tachycardia (vt) while on support. in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. difficulty to remove: clinical details noted that when pump was being explanted, moderate bleeding from graft was observed. physician noted possibility of impella had made adherence to vessel wall. no additional details were provided. the cause of the delivery issues could not be determined as limited clinical details were provided. device history review: pump passed all post-sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4210 | 1220648-2025-47607 | 23330556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330556&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740464 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the support of an impella cp, the patient presented with a hematoma at the site prior to preparation for surgery. the site was angle-matched with the suture hub positioned at the skin level. pressure was applied, and the site remained stable before the drape was placed. throughout the six-hour procedure, there was continuous bleeding from the impella groin site. upon completion of the case, the physician sutured down the blue suture hub, which appeared to improve the situation. additionally, the patient received one liter of plasmolyte. it was noted that the patient was oozing from multiple sites, with a knee immobilizer in place at the site, which was angle matched. pulses were detectable on doppler examination. the following day, the patient had received one unit of packed red blood cells (prbcs) overnight. a central line spot check of the central venous pressure (cvp) was performed, which revealed a cvp of 16. the site was assessed to be stable, locked, and tight, with the angle matched and the leg immobilizer still in place, and pulses were again present on doppler. a thorough washout of the chest cavity yielded no likely sources of bleeding. the sternum was closed, but the top layer of skin was left open with a wound vacuum in place. the patient maintained a stable condition until the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: impella groin site bled entire 6hr case. the patient is oozing from every site possible. knee immobilizer in place and site oozing but angle matched. the cause of the bleeding is determined to be patient condition because patient was bleeding from multiple locations. device history review: the device passed all the post sterile inspection checks.
4211 | 1220648-2025-47608 | 23330561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330561&pc=OZD | 09/17/2025 | 12/04/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574135 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 102 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Swelling/ Edema | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient on impella cp support for cardiac arrest. that the patient had groin site with some bleeding due to high international normalized ratio. md at bedside reinforced forward tension sutures which helped to resolve bleeding. oozing. pressure at site, surgicell applied. crrt for net even status and clearance. lfts remain very high but have been trending down slightly, patient inr 5 due to shock liver. 22- patients remain critically ill. ct head at bedside with no significant changes from prior, still showing edema. systemic ac infusing, no presser/inotrope requirement. plan to keep support in place, if mental status improves md does plan to escalate device. the device was explanted and a 5.5 was implanted. | [5 empty] | The investigation for the reported major bleed and renal failure/edema has been completed.the device was not returned for review. the root cause of the injury is most likely use related since reinforced forward tension sutures helped to resolve bleeding. the root cause of the renal failure/edema was not determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4212 | 1220648-2025-47609 | 23330563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330563&pc=OZD | 09/17/2025 | 10/27/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717553 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Fever | Hemolysis | Sepsis | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that the patient on impella cp for cardiac arrest. treatment for pneumonia and possible sepsis. decision was made to transfer to outside facility for escalation. the patient just arrived from outside hospital with plan to escalate to 5.5. the patient has hemolysis, low urine output, increased lactate and increased pressor support. will continue to monitor. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B1 product problem was selected incorrectly on the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. e4 added selection as it was omitted from the initial report that was submitted. h6 the first follow-up report that was submitted was to add clinical codes. this follow-up report was reported in error and codes were not added for investigation. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure." section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ investigation summary: the investigation for has been completed. data review: data logs showed the pump ran without issue during support. there were positioning alarms triggered in early of the case but resolved quickly. there were no motor current spikes, abnormal placement signal or purge issues observed during support. the product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. sepsis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | Corrections to the initial mfr report: h6 clinical codes e1302 and e230101 have been added. | H6 added codes as they were omitted from the second follow-up report that was submitted.
4213 | 1220648-2025-47610 | 23330572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330572&pc=OZD | 09/17/2025 | 10/17/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026729951 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 75 KG | Non Hispanic | White | Injury | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that patient impella 5.5 for support for shortness of breath and weakness. the physician was unable to advance beyond proximal subclavian. the vessel was evaluated with contrast - significant tortuosity with areas less than 5 mm. 8 mm and stent placed in subclavian by vascular surgeon. it was discussed to possibly use impella cp and the decision was made to abort due to extremely poor prognosis. | [5 empty] | Additional information was received, after review of the associated complaint by the investigation team and noted the automated impella controller (aic) as the potential cause of the pump not detected issue for the defective impella catheter event which was initially reported under the impella 5.5 sn: (b)(6) (mrn 1220648-2025-47619). note: this additional information is new, relevant and requires a supplemental report to depict the complete event story with all associated devices. however. there is no impact to the reportability assessment outcome of this complaint and the investigation remains unchanged. related medical device reports: this second impella 5.5 serial number: (b)(6) device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller and first impella 5.5 device serial number: (b)(6). | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and noted the following: [alarm of impella defective] the defective impella 5.5 device serial number (sn) 521818 was replaced with impella 5.5 device sn 620420. the physician met resistance in passing the impella 5.5 device sn 620420. this impella 5.5 device was removed, and the decision was made to perform contrast examination of the axillar / subclavian artery to determine why the impella 5.5 device sn 620420 was not advancing. significant stenosis of the distal subclavian artery was discovered and recognized as the reason for not being able to advance the impella sn 620420. a stent was subsequently placed and a second attempt to re-place the impella 5.5 sn 620420 failed as the pump was unable to pass through the stent. discussion was made regarding placing an impella cp from the prepared axilla; however, the team opted to not escalate care as notification was received intra-procedure that family wanted to go route of comfort care. the patient expired after care was withdrawn. medical safety review. medical safety reviewed the event and noted the following: the patient presented with worsening cardiogenic shock and experienced cardiac arrest prior to the procedure. an impella 5.5 device was initially prepared, but three ¿defective impella catheter¿ alarms occurred during setup. a second impella 5.5 catheter was prepared without error; however, insertion was unsuccessful due to axillary artery stenosis confirmed by angiography. the physician placed a stent in the stenotic segment, but a subsequent attempt to insert the device also failed. the implant procedure was aborted, and care was withdrawn in accordance with the family¿s wishes. the death is conservatively reported under the second impella 5.5 as a failed device delivery. however, the patient¿s critical condition requiring cardiopulmonary resuscitation and poor prognosis were the most likely contributors to the patient's outcome. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had significant tortuosity with areas less than 5 mm. preventing successful delivery of impella. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device serial number 620420 is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the first impella 5.5 device serial number 521818.
4214 | 1220648-2025-47612 | 23330589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330589&pc=OZD | 05/05/2023 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  | 01/31/2025 |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322669 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Rupture | Cardiac Tamponade | Hemodynamic instability | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 43-years-old male presented at a community hospital with an acute heart failure symptomatic for worsening dyspnea. a transthoracic echocardiography (tte) performed at admission revealed severe left ventricular dysfunction (left ventricular ejection fraction (lvef)= 25%) with global hypokinesia and secondary moderate mitral regurgitation (mr) due to bileaflet tethering. despite intensive pharmacological treatment for acute heart failure, the hemodynamic status worsened, necessitating mechanical circulatory support (mcs) with an impella cp implantation. a myocardial biopsy for suspected acute myocarditis yielded negative results but was complicated by cardiac tamponade, requiring surgical pericardial drainage; no active bleeding sources were identified. due to ongoing hemodynamic instability, a peripheral extracorporeal membrane oxygenation (ecmo) was implanted. the impella cp was initially left in place for venting, and continuous renal replacement therapy (crrt) was begun. however, due to severe bleeding from vascular access, along with the difficulty to maintain adequate anticoagulation, the impella cp had to be removed. a tte was repeated, showing a severe mr with posterior regurgitation jet secondary to medial a2 scallop flail, later confirmed by transesophageal echocardiography (toe). severe left ventricle dysfunction was also confirmed (lvef =25%). we hypothesized that the worsening of mr could have been generated by a mechanical trauma to a tethered mitral valve apparatus during impella removal. after heart team discussion and exclusion of heart transplantation and lvad implantation because of prohibitive surgical risk, a percutaneous mitral valve transcatheter edge-to-edge repair (m-teer) was planned as a potential bridge to ecmo weaning and recovery or to heart transplantation in a more stable condition. the procedure was performed, as usual, from the right femoral vein. transeptal puncture was located posteriorly and superiorly. a mitraclip g4 xtw was opened in the left ventricle and effective grasping was achieved on the second attempt. a mild residual regurgitation with disappearance of pulmonary venous systolic flow reversal was observed immediately; the final mean transmitral gradient was 2 mmhg. two days after the procedure ecmo could be successfully weaned and removed. an intra-aortic balloon pump (iabp) was then implanted to facilitate weaning from vasopressors and inotropic support. a few days later the iabp was removed, and the patient remained hemodynamically stable without inotropes or vasopressors. a tte performed at discharge confirmed a mild residual mr with reduced ventricular volumes and estimated pulmonary artery systolic pressure compared to the pre-procedural values. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. cardiac valve rupture, hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the bleed was not determined due to insufficient clinical details. device history lot: device lot: 1699216. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | . . this report is based on a systematic literature review and reports an adverse event associated with the impella cp. the incident is described in the article titled ¿transcatheter edge to edge repair as a bridge to recovery in patient with severe mitral regurgitation, advanced left ventricular dysfunction and ecmo support:: a case report' by marco licciardi, francesco sanna, marco melis, and bruno loi published in catheterization and cardiovascular interventions vol. 106, on august 1, 2025 pages 903-907. the literature reports that a 43-year-old male experienced bleeding, cardiac tamponade, hemodynamic instability and mitral valve damage following the use of the device. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4215 | 1220648-2025-47611 | 23330591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330591&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711166 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient brought to cardiac cath lab was found to have occluded left anterior descending artery- impella cp inserted via right fem without complication. patient decompensated had ventricular fibrillation arrest was shocked twicw and return of spontaneous circulation achieved. decision made to remove cp and place patient on extracorporeal membrane oxygenation (ecmo) | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4216 | 1220648-2025-47613 | 23330593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330593&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713153 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 94 KG | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that impella 5.5 continued with suction events over p-4 level. the impella had multiple repositions over a week span. most recent echocardiogram now showed a well positioned 5.5 that measured 5.25cm to atrioventricular annulus. the patient was bridged to heart transplant and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product was returned therefore d9 was updated. | The investigation was completed. impella 5.5 sn 611535 returned for evaluation. positioning issues: clinical reported that multiple repositions were made during the last week of the case. the product was returned for investigation, and there were no kinks or abnormalities observed. the cause of the positioning issues could not be determined due to insufficient clinical details. pump suction: clinical reported that suction was triggering above p-4. the product was returned for investigation, but the catheter was cut distal of the motor housing, so the pump could not be tested to try reproducing suction. there was no biomaterial observed inside the cannula and no kinks or abnormalities were observed. data log review confirmed suction events occurring during the case. there were no motor current spikes observed throughout the case. at the time of the first suction event, placement signal from ~ 100mmhg to 70mmhg, but began to increase again. no troubleshooting methods were reported in the clinical details at this time. suction also occurred over the last 5 days of the case, but at this time placement signal did not show any distinct trending or any signs of positioning issues, and clinical also stated that the pump was in proper positioning. clinical reported multiple repositions during this time periods, and they also reported plans to give fluids, but based on clinical details and data logs, it is not clear whether suction was triggering due to positioning issues or volume issues. the cause of the suction could not be determined due to inconclusive data log review and clinical details. device history review was completed and passed all post sterile inspection criteria.
4217 | 1220648-2025-47615 | 23330600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330600&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204293 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced battery problems. there was no reported patient harm. | [5 empty] | The impella device has been received from the customer. investigation is currently ongoing. a supplemental report will be filed once the investigation is complete.
4218 | 1220648-2025-47614 | 23330603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330603&pc=OZD | 09/13/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026652237 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an arch replacement, when attempted to insert a impella cp via the femoral artery the impella bent at the graft and could not be advanced. insertion was abandoned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. product damage (catheter kink): the cause of product damage cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks.
4219 | 1220648-2025-47616 | 23330713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330713&pc=OZD | 09/17/2025 | 10/30/2025 |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702652 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant from the united states reported an incident involving a continuous pump (cp) that was utilized for a patient admitted with a st-elevation myocardial infarction (stemi). the cp was implanted following the placement of a femoral-femoral bypass due to existing peripheral vascular disease. after the pump was implanted, a hematoma developed at the site. the clinical team monitored the hematoma as it increased in size and administered one unit of blood along with a fluid bolus, as the patient exhibited hypotension and experienced pump suction alarms. the cp was explanted after three hours, and the site was subsequently treated with a covered stent. the team later determined that the vessel damage and associated blood loss might have resulted from the femoral-femoral bypass placement rather than the use of the pump. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for investigation. hypotenstion: the cause of the injury was not determined due to insufficient clinical details. hematoma/asae: the cause of the hematoma was most likely use related, as clinical details state the cause of the hematoma was the antegrade access wire from the fem fem bypass rather than the impella. pump suction: the cause of the suction was most likely due to the patients condition, as the patient was hypotensive and bleeding at the time of suction, which can be seen by the trend down the placement signal during the suction events. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4220 | 1220648-2025-47618 | 23330718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330718&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2026654897 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | The complainant reported that the patient was placed on impella support for a coronary artery bypass graft. during the support period, the patient needed one unit of packed red blood cells and 1 mg of factor vii due to bleeding. the patient was taken to the operating room the next day for a washout. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
4221 | 1220648-2025-47619 | 23330732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330732&pc=OZD | 09/17/2025 |  |  | 10/17/2025 |  |  | 09/17/2025 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 SMARTASSIST |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2025505964 |  | 1000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. after connecting to the automated impella controller (aic), "defective impella catheter" alert was presented along with the instruction to "disconnect and replace impella catheter." medical team tried to disconnect and reconnect three times with the same results. the physician agreed to proceed with a new catheter. | [5 empty] | Additional information was received, after review of the complaint by the investigation team, and noted the automated impella controller (aic) as the potential cause of the pump not detected issue for the defective impella catheter event. related medical device reports: this first impella 5.5 device serial number: (b)(6) device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller and second impella 5.5 serial number: (b)(6), which is associated to the demise outcome. | Additional information was received and noted the following: [alarm of impella defective] the defective impella 5.5 device serial number (sn) (b)(6) was replaced with impella 5.5 device sn (b)(6). the physician met resistance in passing the impella 5.5 device sn (b)(6). this impella 5.5 device was removed, and the decision was made to perform contrast examination of the axillar / subclavian artery to determine why the impella 5.5 device sn (b)(6)was not advancing. significant stenosis of the distal subclavian artery was discovered and recognized as the reason for not being able to advance the impella sn (b)(6). a stent was subsequently placed and a second attempt to re-place the impella 5.5 sn (b)(6) failed as the pump was unable to pass through the stent. discussion was made regarding placing an impella cp from the prepared axilla; however, the team opted to not escalate care as notification was received intra-procedure that family wanted to go route of comfort care. the patient expired after care was withdrawn. medical safety review. medical safety reviewed the event and noted the following: the patient presented with worsening cardiogenic shock and experienced cardiac arrest prior to the procedure. an impella 5.5 device was initially prepared, but three ¿defective impella catheter¿ alarms occurred during setup. a second impella 5.5 catheter was prepared without error; however, insertion was unsuccessful due to axillary artery stenosis confirmed by angiography. the physician placed a stent in the stenotic segment, but a subsequent attempt to insert the device also failed. the implant procedure was aborted, and care was withdrawn in accordance with the family¿s wishes. the death is conservatively reported under the second impella 5.5 as a failed device delivery. however, the patient¿s critical condition requiring cardiopulmonary resuscitation and poor prognosis were the most likely contributors to the patient's outcome. device status. additional information also noted the impella 5.5 device (sn (b)(6) was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this impella 5.5 device serial number 521818 is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the second impella 5.5 serial number (b)(6), which is associated to the demise outcome. correction: in the initial report the section d4 ( udi ) was inadvertently reported incorrectly, which is corrected in this report. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4222 | 1220648-2025-31868 | 23330886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23330886&pc=OZD | 05/23/2023 | 10/17/2025 |  | 10/17/2025 |  |  | 05/25/2023 | 10/17/2025 |  |  | 12/31/2024 |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308792 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an intermittent placement signal coinciding with a placement signal not reliable alarm during impella support. patient support was able to continue with the implanted pump despite the intermittent issue. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. without returned product or data logs at the time of the event, the cause of the optical signal issue could not be determined. all pertinent information available to abiomed has been submitted at this time.
4223 | 1220648-2025-31869 | 23331190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23331190&pc=OZD | 10/17/2024 | 10/17/2025 |  | 10/17/2025 |  |  | 10/17/2024 | 10/17/2025 |  |  |  |  | 10/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) had optical signal issues. this issue was discovered during prep and was not related to a clinical procedure. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console imc logs confirm that optical system sensor error was issued as well as placement signal not reliable alarm. device analysis: the console's internal circuitry was inspected with no observable anomaly. the failure was not reproducible upon being plugged in with a similar known good pump. section 23 of the preventative maintenance procedure was performed after cleaning the front optical connector and found to be within specification. conclusion: the root cause for optical sensor system error was most likely contamination of the front optical connector. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4224 | 1220648-2025-47544 | 23309761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23309761&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725958 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed bleeding with a small hematoma at the impella access site. a washout was performed to resolve the issues. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary ==> hematoma/ major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history lot: device lot : 1960433. device history batch ==> sub-component lot : n/a. device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.
4225 | 1220648-2025-47545 | 23309777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23309777&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703705 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in post cardiotomy cardiogenic shock with low cardiac output syndrome was being implanted with an impella 5.5 for mechanical circulatory support. it was reported that during the insertion of the pump, the physician applied excessive tension as the device passed from the subclavian artery to the bifurcation which caused the catheter to kink. the physician decided to abort the insertion of the pump and replaced the pump with a new device. the new impella 5.5 was successfully inserted, and support was initiated. it was noted that prior to impella 5.5 support, the patient had been cannulated for extracorporeal membrane oxygenation and was intubated for respiratory support. the patient remained on impella 5.5 support with the new device for three days, and it was noted that the patient¿s outcome was expired while on support. there were no allegations of product malfunction or patient harm reported with the subsequent pump. | [5 empty] | A2, a4 is unknown. given the limited information surrounding the details the patient's condition and cause of death the initial impella 5.5 cannot be excluded as a possible contributing factor. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined as the use issue related since excessive tension was applied as the device passed from the subclavian artery to the bifurcation preventing successful delivery of impella. product damage (catheter kink): the cause of product damage was most likely use issue related since during insertion of the impella 5.5, excessive tension was applied as the device passed from the subclavian artery to the bifurcation, causing a kink in the catheter shaft device history lot: device lot: 1926437. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
4226 | 1220648-2025-47546 | 23309828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23309828&pc=DYB | 08/27/2025 |  |  | 10/16/2025 |  |  | 08/27/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DYB |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp insertion, the medical doctor placed a sentrant sheath and accessed the left ventricle with a pigtail. 018 wire was then placed but they were unable to advance the impella catheter around the aortic arch due to heavy resistance. the medical doctor had to remove the impella cp and the guidewire together. a new impella cp was placed successfully. this report represents the guidewire. a separate report was submitted for the pump. refer to section h.10 for the related report number. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The guidewire was not returned for investigation. the pump was returned for investigation with no physical damage noted. it was reported that the pump encountered heavy resistance at the aortic arch while being loaded on the 0.018 guidewire. the physician reported being unable to remove the wire, so both the device and the wire were pulled together. the cause of the removal issue was not determined, as the product was not returned and sufficient clinical information was not provided.
4227 | 1220648-2025-47547 | 23310589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23310589&pc=OZD | 09/15/2025 |  |  | 10/16/2025 |  |  | 09/20/2025 | 10/16/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146388 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Display or Visual Feedback Problem | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) screen froze and the "complete the procedure" alarm was observed. changing the purge cassette failed to resolve the issue. a new aic was used to successfully support the patient. | [5 empty] | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. a 70-year-old female with a medical history of persistent atrial flutter, paroxysmal atrial fibrillation, mixed hyperlipidemia, hypertension, and morbid obesity presented to the emergency department with several weeks of fatigue, dizziness, palpitations, and associated shortness of breath. the patient was newly diagnosed with heart failure with reduced ejection fraction (hfref) with a severely reduced ejection fraction of 10¿15% and severe global hypokinesis. the patient underwent left heart catheterization for ischemic evaluation, which demonstrated two-vessel coronary artery disease involving the mid left anterior descending artery and right coronary artery. right heart catheterization demonstrated worsening lactic acidosis, with lactate levels trending up to 12.9 mmol/l. the advanced heart failure team was consulted, and the patient was initiated on left-sided impella support, which was subsequently escalated to biventricular (bipella) support with impella devices: impella 5.5 and rp flex. during impella cp support, the patient developed acute renal failure requiring continuous renal replacement therapy and was treated with multiple medications. paroxysmal junctional rhythm was also reported. the patient was escalated to an impella 5.5, and bipella support was initiated with the addition of an impella rp flex. while on bipella support, the patient experienced multiple adverse clinical events. additionally, the automated impella controller (aic) connected to the impella rp flex displayed a frozen screen with a persistent ¿complete the procedure¿ alarm. replacement of the purge cassette did not resolve the issue. the aic was exchanged with no noted consequence to the patient, and the patient continued to receive support without further interruption. after approximately eight days of mechanical circulatory support, care was withdrawn, and the patient expired. after review of the case by medical safety, it was determined the automated impella controller is a serious injury type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this is one of four devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
4228 | 1220648-2025-47548 | 23310970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23310970&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740547 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Ischemia | Renal Failure | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient was placed onto impella support due to acute myocardial infarction/cardiogenic shock. after impella placement, distal perfusion cannulas were placed due to diminished distal perfusion. the patient experienced multiple runs of non-sustained ventricular tachycardia and was re-bolused with amiodarone and a lidocaine drip was added. the patient was placed on antibiotics due to elevated white blood cells. the patient eventually experienced three sustained runs of ventricular tachycardia requiring three defibrillations. additionally, the patient was placed on continuous venovenous hemodialysis. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Investigation summary: the investigation has been completed. no product was returned for investigation. ischemia/ tachycardia/ infection: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4229 | 1220648-2025-47549 | 23312467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23312467&pc=DQX | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018" PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 77 KG | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician encountered delivery issue when attempting to provide impella cp support. the physician (md) obtained vascular access using best practices and followed step by step as per ifu's. pump was backloaded over abiomed 0.18 wire into left ventricular and proper placement was obtained. md attempted to remove wire and start support; however, they were unable to remove wire. md stared this had never happened to them before. md then removed wire and catheter as a unit. md regained ventricular access with 5 fr pigtail over a v-18 0.18, pigtail removed, and impella catheter was backloaded over v-18. impella was properly positioned and md attempted once again to remove wire and once again they were unable to remove the guidewire. wire and catheter were removed as a unit and md asked for a new impella pump. no patient harm was reported due to the issue. | [5 empty] | Device returned for investigation. d9 updated. investigation results are pending. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. unable to deliver or advance (delivery issue): no issues were found with the returned product. the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | D4: the lot, serial and udi is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4230 | 1220648-2025-47551 | 23312842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23312842&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the access site which was managed with redressing. the patient also suffered from pulseless electrical activity and cardiopulmonary resuscitation was given until return of spontaneous circulation was achieved. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was omitted on the initial report that was submitted. d.4 added primary udi number as it was omitted from the initial report that was submitted. h.4 added device manufacture date as it was omitted from the initial report that was submitted. h.6 code e0506 was reported incorrectly on the initial report that was submitted. investigation summary: cardiac arrest : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
4231 | 1220648-2025-47552 | 23314004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23314004&pc=OZD | 09/16/2025 | 10/16/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686387 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient experienced a subarachnoid hemorrhage while on impella cp support. in addition to bleeding at the puncture site. the bleeding was controlled by restraining his lower limbs and adjusting the angle. the cp was successfully weaned and explanted. | [5 empty] | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4232 | 1220648-2025-47553 | 23315057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23315057&pc=OZD | 09/16/2025 | 10/22/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1940063 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user reported that during preparation of a impella cp procedure the console had a connection cable that was connected to the aic, the previous product preparation screen was displayed. (the process of connecting the connection cable to the catheter plug.) pressing the cancel button and reconnecting the connection cable returned the screen to normal. there was a delay of approximately two minutes before assistance began, but this did not affect the patient's condition. in addition to a impella connect connection failure that occurred during use in the cath lab. it was noted that after connecting the pump catheter, an impella connect alert was not sent. when checking impella connect, it was not displayed. they contacted the service department and attempted to check the usb port connection and restart the device, but the situation did not improve. however, the procedure was completed successfully with the same console. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4233 | 1220648-2025-47554 | 23315209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23315209&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719349 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced "groin issues" that prompted explantation of the pump. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
4234 | 1220648-2025-47556 | 23316572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316572&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701702 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that patient implanted with an impella 5.5 had hematuria and a hematoma at the insertion site. the heparin infusion was stopped. the patient was successfully bridged to a heart transplant, and the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
4235 | 1220648-2025-47557 | 23316912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316912&pc=OZD | 09/16/2025 | 12/01/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655528 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 65 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to cardiomyopathy. while on support, the patient experienced stroke, thrombus, and ultimately passed away. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. date of death unknown. | The investigation for the reported thrombosis and stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombosis and stroke could not be determined.
4236 | 1220648-2025-47558 | 23316934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316934&pc=PYX | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026671759 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Difficult to Open or Close | [5 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention |  |  |  |  | INTRA-AORTIC BALLOON PUMP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient post coronary artery bypass graft and ventricular septal defect repair had a right heart catheterization performed and volume given and the medical team decided to implant an impella rp flex for mechanical circulatory support via the right femoral vein. the patient was already supported with an intra-aortic balloon pump (iabp) that had been initially placed in the right femoral artery but moved to the left femoral artery due to limb ischemia associated with the iabp placement. the patient was released from the operating room to the intensive care unit, and upon arrival it was noted that the touhy lock was unable to be tightened and secured. a sterile dressing was applied that the catheter was sutured in place. it was additionally noted that there was oozing at the impella access site. to resolve the oozing, purse string sutures were deployed, and manual pressure was held. the following day, it was reported that the patient experienced episodes of atrial fibrillation overnight. on the ninth day of impella support, the patient was weaned from the iabp and the iabp was explanted. in preparation for the iabp explant, the patient¿s heparin was withheld. a few hours following the explant of the iabp, the impella rp flex alarmed for purge system blocked. the pump¿s position was verified, and the medical team decided to add tenecteplase to the purge fluid. the purge flow increased, the purge pressure decreased, and the alarms resolved. the medical team decided to replace the impella rp flex with a new device as there were concerns whether the pump was functioning properly after alarms were resolved. a new impella rp flex was prepped and inserted via the right jugular vein, and the first pump was removed and the access site sutured to achieve hemostasis. the patient remained on support with the replacement pump for approximately 1 hour and it was noted that the patient expired during support. there were no malfunctions or adverse events reported with the second pump. | [5 empty] | Given the limited information surrounding the details the patient's condition and cause of death the initial impella rp flex cannot be excluded as a possible contributing factor. a4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4237 | 1220648-2025-47559 | 23316940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316940&pc=OZD | 09/16/2025 | 10/16/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738832 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | High Blood Pressure/ Hypertension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with ectopy. the impella cp smartassist moved into the aorta and was repositioned successfully. prostate biopsy done and hypertension. received red blood cells. the patient survived and was bridged to support on an impella 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4238 | 1220648-2025-47561 | 23316943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316943&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683161 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 98 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a cardiac arrest and expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
4239 | 1220648-2025-47560 | 23316945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316945&pc=PYX | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026622747 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 76 KG | Hispanic | White | Injury | Yes | No | Restricted Flow rate | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Hypovolemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced bleeding from the access site so heparin was paused and two units of packed red blood cells and one unit of platelets were transfused. the patient appeared hypovolemic and had negative access pressure alarms on dialysis. also, there were high purge pressure alarms. the p level of the pump was adjusted without resolution. upon explantation a clot was visualized at inflow cage. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4240 | 1220648-2025-47562 | 23316957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316957&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026667462 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced aspiration alarms and pump in aorta alarms. pump flow stopped and a thrombus was identified at the pump inflow, | [5 empty] | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product and data logs were not returned for investigation. the cause of saturated pump flow issue was not determined due to product/logs not being returned for investigation. the cause of low pump flow issue was not determined due to product/logs not being returned for investigation. the cause of suction issue was not determined due to product/logs not being returned for investigation. the device passed all post sterile inspection checks.
4241 | 1220648-2025-47563 | 23316958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316958&pc=OZD | 08/26/2023 | 11/04/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  | 05/31/2025 |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363539 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry during use with an impella cp: "loqi reported a right femoral artery insertion site infection with relationship to impella listed as definitely related. infectious disease was consulted on (b)(6) 2023 and the patient started on doxycycline 100 mg bid for 14 days." | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the infection was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
4242 | 1220648-2025-47550 | 23316976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23316976&pc=OZD | 09/12/2025 |  |  | 10/16/2025 |  |  | 09/12/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700170 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 61 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | False Alarm | Malposition of Device | [5 empty] | Atrial Fibrillation | Unspecified Infection | Pleural Effusion | Great Vessel Perforation | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a perforation of the left ventricle. the perforation was surgically repaired and volume was administered. the patient also experienced atrial fibrillation that self resolved. the pump generated impella in aorta alarms. in response, the pump was repositioned and albumin was given. additionally, the patient was given antibiotics due to positive sputum culture results. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4243 | 1220648-2025-47564 | 23317005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317005&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696194 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 125 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Inflammation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. patient experienced inflammation, hematoma, and arrhythmia during impella 5.5 support. while on support, the patient had a cat scan of the right leg due to increased swelling and tenderness to the thigh. a right groin hematoma was noted. one unit of packed red blood cells was administered. an echocardiogram was completed, and the small hematoma was still evident. anticoagulation remained on hold due to groin and axillary site. there was a plan to wean the patient off dobutamine for supraventricular tachycardia (svt) and continue milrinone and bumex drops. there was an additional three-minute episode of svt and the patient converted to normal sinus rhythm. patient successfully weaned and explanted 2 day later. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. arrhythmia : the root cause of the injury was unable to be determined due to insufficient clinical details. inflammation : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4244 | 1220648-2025-47566 | 23317006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317006&pc=PYX | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719846 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with pulmonary embolism was implanted with an impella rp flex device for mechanical circulatory support. upon arrival to icu, patient began to get hypotensive and suction was noted on rp flex. filling pressures were low. inotropes were restarted and patient was given lactated ringer's solution (lr) until blood products arrived. patient also had bleeding from nares and mouth since tpa (tissue plasminogen activator), so plan was made to do a full body catscan (ct) to assess for any possible source of bleeding. a few hours later, bleeding slowed, but patient had received 6 units of blood that day. ct also showed a small brain bleed as well as some potentially ischemic segments. the patient was on support for a total of two days before pump was explanted. ct showed new brain bleeds, decision to remove impella so heparin can be stopped and neuro status checked md and family aware of high chance of decompensation and may go comfort care. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed, epistaxis: the root cause of the bleed cannon be determined since insufficient clinical details were provided. hypotension: the root cause of the injury was unable to be determined due to insufficient clinical details. stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no logs are available for this device. the root cause of the pump suction was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4245 | 1220648-2025-47567 | 23317052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317052&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711584 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp escalated to the impella 5.5 pump. the 5.5 pump was supporting the patient bridging to transplant. during the lengthy support there was suction treated with albumin and fluids. additionally there was bleeding from multiple sites that prompted anticoagulation to hold and units of blood to be delivered. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B3 revised date of event as it was reported incorrectly on the initial report that was submitted. d3 added manufacturer fax number as it was omitted from the initial report that was submitted. e1 added initial reporter city , state, phone number, zip code and zip code extension as they were omitted from the initial report that was submitted. h6 added b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. major bleed/asae: clinical details reported blood oozing from multiple sites. the patient's activated clotting time was not provided. the product was not returned for investigation. the cause of the bleeding was most likely related to patient condition because it occurred at multiple access sites. pump suction: clinical details reported suction occurring but stated that it was likely volume related and albumin was given. the product was not returned. data log review confirmed suction events during the case. in the time leading up to suction, placement signal is trending down. placement signal slightly recovers after this drop, which aligns with clinical reporting albumin was given in response to suction. suction occurred later in the case as well, and placement signal is seen trending down again from approximately 100 to approximately 30 mmhg leading up to the suction. the cause of the suction was most likely related to the patient's condition, as placement signal is trending down before suction is triggered, and clinical reported giving albumin to resolve the issue. device history review: the pump passed all post sterile inspection criteria.
4246 | 1220648-2025-47569 | 23317093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317093&pc=OZD | 09/16/2025 | 12/04/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738825 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella cp support during percutaneous coronary intervention (pci) after the pump was inserted and about to start the pci, the patient lost all pulsatility while going into ventricular tachycardia. required pushes of epinephrine to regain pressure. had hematoma at the access site that resolved with manual pressure. the family elected to withdraw care and the patient ultimately expired. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. hematoma: in cath-lab, after removal, noticed a hematoma as they were dressing the impella site. scrubs held pressure on groin site and it went away. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4247 | 1220648-2025-47568 | 23317094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317094&pc=OZD | 09/17/2025 | 12/26/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  | 09/30/2025 |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399221 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male | 85 KG |  |  | Injury | Yes | Yes | Failure to Pump | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during support of impella cp on (b)(6) 2025, md reported, that he had to reposition the pump; afterwards he saw a spike in mc and had no flow anymore; despite several attempts the impella could not be started; recommendation given to exchange the impella the pump could be restarted and is running on p-1, plan to explant in the next few hours due to suspected clot; pt. stable with extracorporeal membrane oxygenation, mean arterial pressure 60mmhg impella was explanted because of pump stop. patient is stable. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9: device returned and date device returned. h3: changed device evaluated by manufacturer? to yes. h6: investigation findings: corrected code, type of investigation, corrected code. h11: added information regarding the results of the investigation. the investigation for the reported thrombosis, temporary pump stop, and pump stop has been completed. thrombosis: the cause of the injury was not determined due to insufficient clinical details. returned product analysis: the returned pump was inspected and there were no abnormalities on visual inspection. the pump passed electrical testing and no biomaterial was observed. the pump was run in the lab at p9 for about an hour and pump stop did not reproduce. no defects were found with returned product. temporary pump stop: md reported, that he had to reposition the pump; afterwards he saw a spike in mc and had no flow anymore. pump could be restarted and was running on p-1. the cause of the temporary pump stop could not be determined as no data logs were returned for analysis and no defects on returned product pump stop: impella was explanted because of pump stop after temporary pump stop earlier. the cause of the pump stop could not be determined as no data logs were returned for analysis and no defects on returned product.
4248 | 1220648-2025-47570 | 23317103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317103&pc=OZD | 09/16/2025 | 12/04/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026739471 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 53 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed using best practices, flows 3.6l in auto. single access percutaneous coronary intervention(pci) performed, peel away was removed and repo was placed, and impella iliac angiogram reviewed no flow past repo sheath. 5fr ante grade sheath placed and 7fr placed in left fem, patient transported to coronary care unit for further management. the next day, the impella was removed due to concern for right leg ischemia. the cp was removed in the cath lab but there was still concern for leg ischemia after removal. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ischemia has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of the ischemia could not be determined.
4249 | 1220648-2025-47565 | 23317592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317592&pc=OZD | 09/16/2025 | 10/16/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740250 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. after impella cp was placed and post peel away sheath was removal, hematoma noted. perclose sutures tightened to try to stop bleeding due to high active coagulation time, but hematoma kept growing. hematoma evacuation was done to the left femoral artery, and the patient was sent to icu. when support was completed and impella was explanted the physician repaired the groin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma/ major bleed: the cause of the hematoma and major bleeding were unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
4250 | 1220648-2025-47571 | 23317605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23317605&pc=OZD | 09/16/2025 | 10/16/2025 |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677285 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. upon arrival to pt room, hematoma, with blood oozing into jp drain noted. manual compression and sandbag added to site. additionally, patient has spiked a fever overnight and more confused this am. on (b)(6) am still unable to find source of bleeding. going for endoscopy today. patient has received 5 units of blood in past 24 hours. dr. (b)(6) would like to explant impella within the next couple days. multiple suction events overnight, dropped to p7, from p8. hemoglobin dropped to 6.9, receiving second unit of packed red blood cells now (hgb goal >8) further maroon bm overnight. remains off anticoagulation. esophagogastroduodenoscopy showed no acute areas of bleeding. echo wean this am, results pending, dropped to p3 briefly. | [5 empty] | Correction: b1, d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed, hematoma: the cause of the bleed was most likely use issue related since hematoma and bleeding started occurring after the arrival of patient to icu from catheterization lab. fever: the cause of the injury was not determined due to insufficient clinical details. pump suction: no pump was returned. patient had multiple suction alarms. no logs are available for this device. the cause of the pump suction was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4251 | 1220648-2025-31866 | 23318659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23318659&pc=OZD | 02/25/2020 | 10/16/2025 |  | 10/16/2025 |  |  | 02/25/2020 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1126567 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R | Recall |  |  | Initial submission | [9 empty] | The us complainant reported to have an automated impella controller (aic) failure. the aic displayed a failure warning, then shut down. the aic was replaced per instructions for use recommendations. there was no patient harm or injury reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data review: data confirmed the failure mode as imc log showed multiple errors "imcadmin: can't open shared memory when console was powering on. console analysis: console was tuned on and revealed no issue. manually restarted the console couple times could not reproduce the failure mode. the console then was power cycled 500 times to observe if any issue occur. after couple times of power cycling, the failure mode was reproduced as the console's boot up process was crashed and the screen showed system self check failure. additional test revealed the pump motor driver (pmd) communication issue of impellatronic. the issue was resolved when impellatronic was replaced. the ribbon cable between 4 in 1 assembly and impellatronic was bent but not affect to the function of the console. the cause of the system start up issue was a defective impellatonic pcb. this report is related to recall: rcl250612/res#97440.
4252 | 1220648-2025-47572 | 23319201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23319201&pc=OZD | 09/16/2025 |  |  | 10/16/2025 |  |  | 09/16/2025 | 10/16/2025 |  |  |  |  | 10/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719197 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 111 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp for approximately 17 hours for pre-percutaneous coronary intervention support. the patient was unable to successfully wean drips overnight and the heart team decided to escalate to an impella 5.5 for further mechanical circulatory support (mcs) and during this support, the patient experienced device sensing issue, bleeding, device malpositioning, and ventricular tachycardia requiring intervention. the patient ultimately expired. the impella was successfully placed and impella 5.5 immediately placed on p-2 support level. impella cp was decreased to p-1 and pulled back into the aworta. impella 5.5 was slowly increased to p-7 support level. impella cp was then turned off and removed. perclose x2 were tied down and manual pressure used for hemostasis. transesophageal echocardiogram confirmed measurement of 4.7cm. the patient was transferred to the unit in critical condition. post implant, same day, the optical sensor measured ¿significantly¿ lower than the blood pressure and remained low. support continued. two days later the patient remained intubated and still not following commands. the patient was weaned from epinephrine and remained on dobutamine. the patient had good diuresis over the last 24 hours. the were no plans to explant the impella any time soon but the team was hopeful for recovery. it was further noted there were no systemic anticoagulations due to gastrointestinal bleed concerns. placement signal and blood pressure still did not match up. the placement signal was noted approximately 400 points lower. as the patient continued on support now a day later, it was noted the patient continued to have bouts of ventricular tachycardia. the physician pulled the impella back 1cm. the patient continued to have black stools and received one unit of blood daily. systemic bivalirudin was placed on hold. colonoscopy and ganglion block procedures were scheduled. however, completion and results were unnoted. two days later, the patient was transferred for palliative care, and the following day the patient expired; care was withdrawn. | [5 empty] | As the impella device had associated events, the device cannot be dissociated and therefore is being reported as death type of reportable event. however, in this case, patient underlying condition likely contributed most to an outcome of death (cardiogenic shock stage e; patient remained critically ill after intervention and unable to fully recovery). the patient was transitioned into palliative care and care was withdrawn leading to the patient expiring. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined due to insufficient clinical details. device in wrong position: due to lack of clinical details, the root cause of the device in wrong position cannot be established. optical signal issue: due to a lack of clinical details about what was occurring with the patient during this and no product returned, the root cause of the optical sensor issue cannot be established. device history review: the device passed all post sterile inspection checks.
4253 | 1220648-2025-31865 | 23278333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23278333&pc=OZD | 09/10/2025 | 10/13/2025 |  | 10/15/2025 |  |  | 09/10/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683477 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Disability | Required Intervention |  |  |  | ABM14FR, 14F COOK 45CM AND 16F COOK 45CM SHEATHS. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced a stroke. an angiogram showed severe triple vessel disease. an impella cp and swan-ganz catheter were placed and the patient was referred for bypass. the patient had prior stent grafts placed and severe iliac disease. the impella 14f 25cm peel away sheath kinked and was unable to pass impella sheath. the physician attempted with a 14f cook 45cm sheath and still was unable to pass. the physician then placed 16f cook 45cm sheath and inserted the impella cp through 16f cook 45cm sheath and the pump was started on performance level p-6. the right femoral vein swan was placed for monitoring with plans to refer for bypass surgery. the patient was taken to the operating room for coronary artery bypass graft surgery (2 vessels) and direct impella 5.5 placement. the patient was cannulated for bypass, impella cp was removed, case continued, and the physician placed the impella 5.5 at end to come off bypass. the patient was transferred to the unit and later this day the patient was noted to have a right middle cerebral artery infarct overnight. no intervention details if any were available. it was noted, however, the stroke event was not successfully managed. care was withdrawn and the patient expired. additional information was subsequently received and noted after review with team, the stroke is believed to be from removal of impella cp before escalation to impella 5.5. | [5 empty] | Medical safety review was completed and noted the following: the stroke with the demise outcome was reported on the impella 5.5. additional information was received and noted it is believed the stroke occurred after removal of the impella cp before escalation to the impella 5.5. as the impella 5.5 device was in use at time of expiration and timing of the stroke is indeterminate, the impella 5.5 association with stroke and demise outcome will remain unchanged. the associate director of post market clinical review noted without information on when the stroke occurred this is hard to determine definitely. however, it should be noted that the impella cp was placed through a 16fr cook sheath which is off label and could lead to increased risk of clot formation in the sheath. the impella cp device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of three devices associated with the event. refer to the following related report numbers: manufacturing report number 1220648-2025-31138 for the abiomed kit, 14fr introducer manufacturing report number 1220648-2025-45743 for the impella cp instructions for use for the related event are as follows: impella cp with smartassist system ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
4254 | 1220648-2025-47446 | 23278599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23278599&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764459 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 68 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The patient was placed onto impella support for intervention for abdominal aortic aneurysm. while on support, during a routine purge cassette change, the automated impella controller (aic) failed to recognize the new purge disc. the patient had to be transferred to a new aic. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4255 | 1220648-2025-47447 | 23278877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23278877&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724338 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp had hematoma around the impella site. femstop was initially applied. then manual pressure with medical doctor placing a stitch at site. the perclose was tightened around the repositioning sheath. additionally, the patient had gi bleed with orogastric tube placement. two units of blood products were administered. the patient continued support and was explanted as planned. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma is determined use issue because the hematoma is resolved by tightening of perclose sutures. major bleed: the root cause of the clinical symptom(bleeding) was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4256 | 1220648-2025-47448 | 23279506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23279506&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719074 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 67 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was placed but suction occurred on p-2 level and above. there was an apparent kink in the cannula as visualized from fluoroscopy. a new impella cp was placed successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. product damage (cannula kink): the cause of product damage cannot be determined since insufficient clinical details were provided. pump suction: no logs are available for this product. the cause of pump suction cannot be determined since insufficient clinical details, no logs or product was provided. device history review: the device passed all the post sterile inspection checks.
4257 | 1220648-2025-47449 | 23279747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23279747&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701732 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 63 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella 5.5 for the patient admitted in with cardiogenic shock and cardiomyopathy. the patient was having atrial fibrillation/ventricular tachycardia and being evaluated for advanced care. the pump was placed along with prior ecmo but, after 3 days there were "in aorta" alarms and there was pulsatility issues observed. the family decided to withdraw care the following day and the patient expired. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the optical sensor issue was not determined since product/logs were not returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4258 | 1220648-2025-47445 | 23279966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23279966&pc=OZD | 09/12/2025 |  |  | 10/15/2025 |  |  | 09/12/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709732 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced insufficient blood circulation in punctuated leg. explantation impella and vascular surgery. the patient outcome is reported as stable. | [5 empty] | Correction: g4. the investigation of the reported failure mode has been completed. no product was received for evaluation. vascular dissection: the cause of the vascular dissection was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4259 | 1220648-2025-47452 | 23280930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23280930&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676683 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 70 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The eu/ swizz complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support on day 3 patient experienced ischemia of the left leg at the impella cp position side. no perfusion and elevation of lactate, decision to remove the pump and escalate to 5.5 with aortic access. after implanted with 5.5 for about 16 days the physicians decided to withdraw the therapy because they had no other therapeutic option since the patient had a very low left ventricular ejection fraction (lvef) (5%) and an atrial septal defect that they could not treat. a definitive left ventricular assist device (lvad) was not possible because of the atrial septal defect and the patient was no candidate for a heart transplant. the death of the patient is reported under the 5.5 under a separate report. | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp. three days after start of the impella cp support, the patient developed limb ischemia (impella leg). there were no perfusion and elevation of lactate. additionally, during the impella cp support there was bleeding at the extracorporeal membrane oxygenation (ecmo) site, but no noted intervention. the decision was made to remove the impella cp and escalate to an impella 5.5 with aortic access. when escalated to the 5.5 the patient experienced a shunt closing failure. approximately 16 days after start of the impella 5.5 support, it was reported the placement signal was found to be unreliable. position signal was recalibrated without repositioning of the device. the physician reported a sudden drop in the ao position signal on the aic without any changes in patient status, medicines given or ecmo flow/rpm. aortic position signal was recalibrated without repositioning of the device. instantly parameters improved and became similar as before the sudden change. there was no harm to the patient as a result of this event. additionally, it was noted the patient had oeasophagic bleeding and developed sepsis. there were no advance therapy options (left ventricular assist device, heart transplant) to treat the patient. a multidisciplinary decision for withdrawal of care was made, which was completed three days later. the patient expired. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d4, e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. minor bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | An update was made to section b5 event description to include complete details regarding the events. additional information was subsequently received and noted the following: the patient had very sick arteries, already before the implantation, the toes were blue before the puncture, and it was of course worse afterwards. the axillary artery was also way too narrow for an impella 5.5, and this is why the physician opted for direct access. the physician's decided to withdraw the therapy because they had no other therapeutic option since the patient had a very low left ventricular ejection fraction (5%) and an atrial septal defect that they could not treat. a definitive left ventricular assist device (lvad) was not possible because of the asd, and the patient was no candidate for a heart transplant. medical safety review: medical safety reviewed the event. the pump was placed but the impella cp limb became ischemic, and this prompted pump explant and escalation to a 5.5 pump after 3 days of support. additionally, during the impella cp support there was bleeding at the extracorporeal membrane oxygenation (ecmo) site, but no noted intervention. the patient was successfully escalated to an impella 5.5. this is a serious injury type of reportable event. while on impella 5.5 the patient had oeasophagic bleeding. it cannot be determined if the use of anticoagulant for use of the impella devices impacted/exacerbated the out of the bleeds for the patient. following, the patient developed sepsis. there were no advance therapy options (left ventricular assist device, heart transplant) to treat the patient. ultimately the patient expired. a multidisciplinary decision was made to withdraw care leading to the patient's demise. it is unknown how or if the impella 5.5 pump played a role involving the sepsis and given this information at hand there is not enough evidence to dissociate the impella 5.5 device from the demise outcome. note: placement signal issue or noted drop in the aortic position signal is not association to the demise outcome. there were no changes in patient status, medicines given or ecmo flow/rpm. related medical device reports: this impella cp device is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
4260 | 1220648-2025-47453 | 23281012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23281012&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717062 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the case was aborted case after initial pump insertion, per md, patient is hemodynamically unstable with worsening blood pressure and aortic insufficiency and placement signal loss. pump removed & re-insertion attempted. the pump was explanted and new cp placed. the 2nd cp (sn: (b)(6)) is captured in another complaint. | [5 empty] | Correction d4. the investigation of the reported failure mode has been completed. no product returned. optical sensor issue - data log review shows placement signal (ps) railing to 3000mmhg after 3 minutes on support, and never recovering. pump positioning noted to be in the ventricle during ps failure. at that time, signal not reliable (snr) drops to 0, contrast decreases, lamp increases, and gain increases. the cause of the optical signal issue could not be determined to no product returned, inconclusive data log review, and insufficient clinical details. patient code - heart valve incompetence: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event and noted the following: the patient presented in cardiogenic shock and an impella cp was placed for hemodynamic support. after initial positioning and activation, the patient became increasingly unstable, and the physician suspected device-related aortic insufficiency. the first impella cp was removed and reinserted; however, a ¿placement signal not reliable¿ alarm occurred with no placement signals. the device was subsequently removed. a second impella cp was inserted successfully, and support was initiated without issue. upon icu arrival, a hematoma was observed at the access site. the device was re sutured and redressed to optimize angle and tension, and a femoral compression device was applied. the hematoma softened, but the following day the patient required transfusion of two units of packed red blood cells due to decreased hemoglobin and hematocrit, despite a stable access site. the patient expired the next day. the first impella cp is reportable for device malfunction and serious injury due to suspected aortic insufficiency and placement signal failure. the second impella cp is related to bleeding and hematoma. the patient¿s death occurred while on support; however, there is insufficient evidence to conclude that the second impella cp device contributed to the fatal outcome so will be conservatively reported. note: h6 heath effect ¿ clinical/impact coding, medical device problem, component, and investigation type/findings/conclusion coding remain unchanged as previously reported. related medical device reports: this impella cp pump 1 device is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 2 device, which is associated with the demise outcome.
4261 | 1220648-2025-47455 | 23285976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23285976&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716184 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a intra-aortic balloon pump placed for coronary perfusion and cardiogenic shock with the plan to escalate to impella 5.5 to prehab patient for coronary artery bypass graft (cabg). during support a hematoma was noted at insertion site with no progression, plan to evacuate during surgery monday. later it was reported that the patient had increasing white blood cell (up to 21 from 16) and low grade fever overnight, so cabg postponed. some confusion over the weekend, likely combined icu delirium, age, and lack of sleep. hemodynamics stable, dobutamine down to 1.5. r axillary impella pocket with stable hematoma, no drainage. going for cabg this am. plan to keep impella in place, explore and washout axillary pocket. slight facial droop noted, considering ct. dobutamine also restarted. later the patient still with right sided facial droop and mumbled speech-clinically suspicious for stroke. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed and no product was returned for investigation. hematoma: hematoma at insertion site present with no progression, plan to evacuate during surgery monday. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. stroke/ infection: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
4262 | 1220648-2025-47457 | 23286052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286052&pc=PYX | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026647401 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 67 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced possible sepsis and was started on antibiotics. additionally, the patient experienced bleeding. the patient was treated with pressors and blood products. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the sepsis and major bleed were unable to be determined due to insufficient clinical details.
4263 | 1220648-2025-47458 | 23286144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286144&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728340 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding issues, and the pump came out of position into the aorta. the pump was explanted and the patient was escalated to an impella 5.5. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the bleed was not determined due to insufficient clinical details. the cause of the positioning issue was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4264 | 1220648-2025-47460 | 23286269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286269&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714091 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp had bleeding at the impella access site. the site was re-sutured using forward tension sutures. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1. added initial reporter address line 1 as it was omitted from the initial report that was submitted. h6. added code b13 as it was omitted from the initial report that was submitted. h11. added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. investigation summary: the investigation has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be use issue because the bleeding was resolved by resuturing using forward tension sutures. device history review: the device passed all post sterile inspection checks.
4265 | 1220648-2025-47466 | 23286420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286420&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693629 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 120 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a gastrointestinal (gi) bleed. the patient received a transfusion of red blood cells. later, the patient expired. | [5 empty] | The cause of the major bleed was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4266 | 1220648-2025-47461 | 23286437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286437&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725335 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. patient have site oozing. nurse changed dressing earlier this morning. blood products continue to be given as needed. heparin in hold per physician platelet transfused, 1 unit red blood cells given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d4. the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding was determined to be anticoagulation management. device history review: the device passed all the post sterile inspection checks.
4267 | 1220648-2025-47468 | 23286483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286483&pc=OZD | 09/13/2025 | 01/22/2026 |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618442 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 129 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support post percutaneous coronary intervention (pci) was shocked in the cath lab twice for ventricular tachycardia and ventricular fibrillation. cpr was done for 5 mins and achieved return of spontaneous circulation. on day five impella support dressing on the groin was oozing. dressing was changed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4268 | 1220648-2025-47462 | 23286557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286557&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700120 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion of the impella cp, the impella would not traverse the distal aorta & common iliac due to disease and tortuosity. the physician decided to abort the impella cp insertion and placed an intra-aortic balloon pump through the impella sheath without difficulty. there was no patient harm reported. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the delivery issue was determined to be patient condition as the patient had distal aorta & common iliac due to disease and tortuosity preventing successful delivery of impella. device history lot: device lot: 1947623. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4269 | 1220648-2025-47463 | 23286647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286647&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689454 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 68 KG | Non Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | Stroke/CVA | Fever | Low Blood Pressure/ Hypotension | Unspecified Infection | Ventricular Fibrillation | Hematuria | Valvular Insufficiency/ Regurgitation | [6 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. the pump was repositioned to resolve the issue. the patient then developed pneumonia with a fever, ventricular tachycardia, and a stroke. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | It was reported that care was withdrawn and the patient expired. b2 and h1 have been updated. | No product was returned for investigation. the root cause of the hematuria, fever, and hypotension was unable to be determined due to insufficient clinical details. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the root cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | D6b, date of explant, has been added.
4270 | 1220648-2025-47470 | 23286704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286704&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682534 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered continuous ventricular tachycardia (vt) during impella 5.5 support. the pump was repositioned due to vt and the patient was given 1 shock, amiodarone bolus and drip. the patient was shocked again the following night. the night after, patient encountered vt again requiring cardioversion. the pump was later explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the arrhythmia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4271 | 1220648-2025-47456 | 23286808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23286808&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026627726 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp was inserted and there was a hematoma at the site of access. no reported issues during sheath exchange in cardiac cath lab. manual pressure held, site re-sutured for increased forward tension, angle adjusted with gauze & dressed per best practices. fem-stop applied & hematoma softer, stable. the patient later expired after 2 days of support when care was withdrawn. the 1st implanted cp (sn: (b)(6) was explanted and replaced with this cp due to placement signal issue and is reported under separate report. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the returned product was investigated. some dried blood was seen on the inflow cage and suture hub. no other abnormalities were seen. access site adverse event (hematoma): clinical details mention that a hematoma was noted at the access site. the cause of the access site adverse event is use related since after arrival to intensive care unit (icu) the activated clotting time (act) were above 250 and the issue was improved by re-suturing for increased forward tension and adjusting the angle with gauze. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event and noted the following: the patient presented in cardiogenic shock and an impella cp was placed for hemodynamic support. after initial positioning and activation, the patient became increasingly unstable, and the physician suspected device-related aortic insufficiency. the first impella cp was removed and reinserted; however, a ¿placement signal not reliable¿ alarm occurred with no placement signals. the device was subsequently removed. a second impella cp was inserted successfully, and support was initiated without issue. upon icu arrival, a hematoma was observed at the access site. the device was re sutured and redressed to optimize angle and tension, and a femoral compression device was applied. the hematoma softened, but the following day the patient required transfusion of two units of packed red blood cells due to decreased hemoglobin and hematocrit, despite a stable access site. the patient expired the next day. the first impella cp is reportable for device malfunction and serious injury due to suspected aortic insufficiency and placement signal failure. the second impella cp is related to bleeding and hematoma. the patient¿s death occurred while on support; however, there is insufficient evidence to conclude that the second impella cp device contributed to the fatal outcome so will be conservatively reported. note: h6 heath effect: clinical/impact coding, medical device problem. component, and investigation type/findings/conclusion coding remain unchanged as previously reported. related medical device reports: this impella cp pump 2 device is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 1 device. | The product was received for investigation. d9 updated with receipt date. once investigation is complete a supplemental report will be submitted.
4272 | 1220648-2025-47467 | 23287156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287156&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724422 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 88 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia so amiodarone was given. the patient continued support until the pump was successfully weaned and explanted as planned. | [5 empty] | The root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4273 | 1220648-2025-47459 | 23287253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287253&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621528 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure and purge system blocked alarms. tissue plasminogen activator was administered. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the high purge pressure was a purge tubing kink but the location is unknown since product was not returned. the device passed all post-sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4274 | 1220648-2025-47465 | 23287288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287288&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672830 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed an access site bleed. as a result, the patient received blood transfusions. the patient was later weaned from the pump and survived. | [5 empty] | B.7 added information was that omitted on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name should of been left blank on the initial report that was submitted. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary no product was returned for investigation. major bleed: the impella site has a small ooze, and no hematoma. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4275 | 1220648-2025-47469 | 23287324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287324&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712189 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 80 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient encountered ventricular tachycardia (vt) during impella cp support. patient arrived at icu and went into vt. as a result, cpr was started, 12 min downtime, and return of spontaneous circulation was achieved. however, after a few hours of support, care was withdrawn and patient ultimately expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. arrhythmia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4276 | 1220648-2025-47464 | 23287515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287515&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701670 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was presented to the hospital with complaints of shortness of breath that had been ongoing for three days. an electrocardiogram was performed, and the results were abnormal. the patient¿s troponin results were elevated and increasing. the patient was then transferred to another hospital for left heart catheterization. the left heart catheterization revealed severe multi-vessel disease with an elevated end-diastolic pressure. the decision was made to place the impella cp in a pre-operative setting and transfer the patient to another hospital for a high-risk coronary artery bypass graft procedure. the patient was successfully transferred to the hospital, and the following day the patient was taken to the cardiac catheterization laboratory (ccl) for a percutaneous coronary intervention (pci). two drug-eluting stents were placed to the left circumflex coronary artery, however the attempt to intervene on the left anterior descending coronary artery (lad) was not successful, and the medical team determined that the patient was no longer a surgical candidate for coronary artery bypass graft. additionally, it was reported that the automated impella controller stopped displaying the placement signal and shortly after the loss of placement signal, the aic had a yellow controller error alarm. the aic was exchanged for another unit and the yellow alarm resolved. the following day, it was noted that the patient¿s urine output was slightly pink and was decreased since admission to the hospital. the patient was given fluid volume, and it was noted the urine color cleared throughout the day. the following day, the patient was taken back to the ccl to address the lad with another pci. the pci was successful, and the patient was implanted with two stents to the lad and two stents to the right coronary artery. following the procedure, the patient was successfully weaned and the impella cp was explanted. this complaint involves two impella devices: impella cp with serial number (b)(6). automated impella controller with serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 added as it was omitted from the initial report that was submitted. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. h10 related report number was added. investigation summary: no product was returned for investigation. hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4277 | 1220648-2025-47471 | 23287556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287556&pc=OZD | 09/13/2025 | 10/13/2025 |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569908 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hematoma | Sepsis | Renal Impairment | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a cp pump placed to support the patient admitted in with a st-segment elevation myocardial infarction. the pump was placed and was supporting when there was a hematoma observed on the site of the pump. the team did not explant the pump, but the pump remained on for support as the team assessed for renal failure, and started continuous renal replacement therapy, and for possible sepsis. the patient was in multiple organ failure and decision made not to escalate but rather withdraw care. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4278 | 1220648-2025-47473 | 23287577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23287577&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/13/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714579 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with chest pains and eventually coded requiring cardiopulmonary resuscitation and intubation. decision for impella cp was made. the cp was successfully placed in right femoral artery. the patient was on cp support for 9.62 hours and was escalated to 5.5 in right axillary/subclavian artery. it was reported that there was a loss of pulse on right foot noted so computed tomography angiography performed showing thrombus in common femoral artery. vascular surgeon to bedside to evaluate and planning for open debridement and vascular repair. according to the note, upon open assessment, the posterior wall of the common femoral artery was held together with only the adventitia layer as the intima was dissected. there was speculation that the perclose closure could have damaged the artery wall during closure attempts. the artery was repaired with a synthetic graft. patient was recovering well and pulses intact. the patient tolerated the vascular repair of the common femoral artery well. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia/ thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. vascular dissection: on open assessment, the posterior wall of the common femoral artery was held together with only the adventitia layer as the intima was dissected. there was speculation that the perclose closure could have damaged the artery wall during closure attempts. the artery was repaired with a synthetic graft. the cause of the vascular dissection was determined to be use issue because perclose closure could have damaged the artery wall during closure attempts. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation.
4279 | 1220648-2025-47475 | 23289802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23289802&pc=OZD | 09/13/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090366 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was presented to the hospital with complaints of shortness of breath that had been ongoing for three days. an electrocardiogram was performed, and the results were abnormal. the patient¿s troponin results were elevated and increasing. the patient was then transferred to another hospital for left heart catheterization. the left heart catheterization revealed severe multi-vessel disease with an elevated end-diastolic pressure. the decision was made to place the impella cp in a pre-operative setting and transfer the patient to another hospital for a high-risk coronary artery bypass graft procedure. the patient was successfully transferred to the hospital, and the following day the patient was taken to the cardiac catheterization laboratory (ccl) for a percutaneous coronary intervention (pci). two drug-eluting stents were placed to the left circumflex coronary artery, however the attempt to intervene on the left anterior descending coronary artery (lad) was not successful, and the medical team determined that the patient was no longer a surgical candidate for coronary artery bypass graft. additionally, it was reported that the automated impella controller stopped displaying the placement signal and shortly after the loss of placement signal, the aic had a yellow controller error alarm. the aic was exchanged for another unit and the yellow alarm resolved. the following day, it was noted that the patient¿s urine output was slightly pink and was decreased since admission to the hospital. the patient was given fluid volume, and it was noted the urine color cleared throughout the day. the following day, the patient was taken back to the ccl to address the lad with another pci. the pci was successful, and the patient was implanted with two stents to the lad and two stents to the right coronary artery. following the procedure, the patient was successfully weaned and the impella cp was explanted. this complaint involves two impella devices: impella cp with serial number (b)(6), automated impella controller with serial number (b)(6). this mdr specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
4280 | 1220648-2025-47474 | 23289850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23289850&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703366 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 49 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed a gastrointestinal (gi) bleed and heparin was held. the gi bleed was resolved. the patient was escalated to an impella 5.5 where support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. ppae: major bleed. the cause of the injury was not determined due to insufficient clinical details. device history review: this pump passed all post sterile inspection checks.
4281 | 1220648-2025-47476 | 23290476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23290476&pc=OZD | 09/18/2025 |  |  | 10/15/2025 |  |  | 09/18/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1762746 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported while on a patient was on impella 5.5 support, there was an aortic signal on the automated impella controller used that was significantly different from cuff blood pressure. the patient remains on support. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the automated impella controller. another report was submitted to represent the impella 5.5. refer to section h.10 for the related report number.
4282 | 1220648-2025-47405 | 23290987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23290987&pc=OZD | 08/20/2025 | 10/14/2025 |  | 10/15/2025 |  |  | 08/20/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699710 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 114 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [8 empty] | Required Intervention | Death |  |  |  |  | JACKSON-PRATT (JP) DRAIN | VENTILATOR | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed on impella cp for mechanical circulatory support. the patient had a previous retroperitoneal bleed prior to impella placement and was found to have a hematoma. anticoagulation was held. the impella site was slightly oozy with a jp drain in place. the patient went into ventricular tachycardia requiring one round of cardiopulmonary resuscitation. the size is slightly oozy but soft with distal pulses present. the patient went into atrial fibrillation and amiodarone was started. the patient returned to normal sinus rhythm. the patient remains on a ventilator and has a vascular catheter placed. pulses were lost in the bilateral lower extremities. the family ultimately withdrew care and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. tachycardia/ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma/ major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. paroxysmal atrial fibrillation: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
4283 | 1220648-2025-47481 | 23291211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23291211&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691945 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Heart Block | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there were positioning issues, oozing, heart block, and infection during impella cp patient support. while on support and during transfer to a bed, impella was dislodged. the patient was placed back on table and the left ventricle was re-engaged. the impella sheath was removed, and the repositioning sheath was advanced. the staff reports the site was redressed and figure eight suture applied due to oozing at site. later it was noted that the site looked good with angle matching and no bleeding. furthermore, the patient was in complete heart block. phenylephrine was pushed. the patient was transferred to an impella 5.5 pump for continued support. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed: the cause of the access site bleed was use related since re-dressing and maintaining angle resolved bleed. infection: the root cause of the injury was unable to be determined due to insufficient clinical details. heart block: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the user/patient during the transport to bed. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the bleed was not determined due to insufficient clinical details. the root cause of the infection was unable to be determined due to insufficient clinical details. the root cause of the heart block was unable to be determined due to insufficient clinical details. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the user/patient during the transport to bed. the device passed all post sterile inspection checks.
4284 | 1220648-2025-47477 | 23291662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23291662&pc=OZD | 09/17/2025 |  |  | 10/15/2025 |  |  | 09/17/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728702 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hematuria | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that the patient presented to the hospital with worsening heart failure. an intro-aortic ballon pump was placed via axillary artery during a left ventricular assist (lvad) work up, however the following days the patient¿s hemodynamics worsened. the patient¿s ejection fraction was calculated to be 10% with a decreased mixed venous oxygen level of 47%, therefore the decision was made to escalate the patient to an impella 5.5 device. the impella 5.5 was inserted via axillary artery and support was started with no issues, however within the first 7 hours of support the pump abruptly stopped functioning twice. each time, the pump was able to be restarted following the high motor current alarm with pump stop, however the p level would not go above p1 or p2 before stopping again. additionally, it was reported that the patient experienced hematuria. a place was made to remove and replace the pump. the patient was brought back to the operating room for an impella 5.5 replacement. upon removal of the first pump, a clot was observed on the inlet of the pump. additionally, the axillary artery 8mm hemashield graft was found to have clots with in it. the physician suspected the reason of the clots was due to suspected heparin induced thrombocytopenia of the patient. the graft was removed and replaced with a new 10 mm hemashield platinum graft which was anastomosed to artery. the new impella 5.5 was implanted without issue, and the patient was released back to the intensive care unit (icu). upon returning to the icu, the patient was noted to have a facial droop and left side weakness. the patient was taken for an emergency computed tomography scan where a clot was found, and the patient was taken for an emergent thrombectomy. the following day the patient was noted to have no deficits from the thrombotic event. additionally, it was noted that the impella 5.5 was not well placed, and the surgeon was notified that the impella would need to be repositioned. the following day, the impella positioned was discussed. it was recommended to the surgeon to reposition the impella as the inlet was facing the mitral valve, and the pump was slightly deep at 6.4 cm. although the concern that the malposition could contribute to ectopy, the physician declined to reposition stating that the ventricle was large and was satisfied with the pump position. two days later, the pump was repositioned and pulled back to 5.1 cm and the power level was able to be increased from p-7 to p-8. the patient was placed on a two-week hold for additional surgery due to the stroke. the following day, it was reported that there were several impella in aorta and impella position unknown alarms which occurred during the previous night. echocardiography was performed to verify impella position, and impella position was confirmed to be well placed. on the seventh day of impella support, it was reported that the automated impella controller (aic) experienced an unexpected shut down. no alarms had been observed, and the medical team replaced the aic to resolve the issue. on the fourteenth day of support, it was reported that there were ongoing false placement signal alarms, and the medical team disabled the alarm. the following day it was reported that the patient was experiencing gastrointestinal bleeding. anticoagulation was withheld due to the ongoing bleeding and low hemoglobin. a unit of packed red blood cells were transfused as result of blood loss. on the eighteenth day of support, it was reported that there were placement signal and suction alarms. it was noted that there were no signs or symptoms of hemolysis due to the suction alarms, and that the patient¿s plasma free hemoglobin was <30 with adequate yellow urine output. on the twenty-fifth day of support, it was reported that the impella pump was repositioned due to the patient experiencing several runs of ventricular tachycardia. the patient remains on impella 5.5 support, with a plan for surgery for a durable lvad in the next week. this complaint involves three impella devices: pump 1: impella 5.5, with serial number (b)(6), pump 2: impella 5.5, with serial number (b)(6), automated impella controller with serial number (b)(6). this mdr specifically addresses pump 1: impella 5.5, with serial number (b)(6),, which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | A4, a6 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4285 | 1220648-2025-47484 | 23291905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23291905&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704218 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Atrial Fibrillation | Atrial Flutter | Fever | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced cardiac fibrillation, fever, and hematuria. an echocardiogram revealed the pump was in a deep position so it was pulled back into position. later, the patient was successfully weaned from the pump and survived. | [5 empty] | G1 manufacturing site name was incorrect and should of been blank on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for review but the data logs were received on 03-nov-2025. data review: data logs showed the pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current spikes, abnormal placement signal or purge issues observed during support. the product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to pump was not in the proper position since it improved after repositioning of the pump. the cause of positioning issue was unknown. atrial flutter/paroxysmal atrial fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. access site bleeding: the cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4286 | 1220648-2025-47480 | 23292095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23292095&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026661726 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Decrease in Pressure | Increase in Pressure | [5 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a declining purge rate so tissue plasminogen activator (tpa) was added to the purge. the patient became septic and was treated with antibiotics. out of concern for possible heparin-induced thrombocytopenia (hit) the patient was treated with therapeutic heparin. a thrombectomy was performed and the next day, the pump was explanted. | [5 empty] | No product returned for investigation. the cause of the high purge pressure was biomaterial build up due to tpa resolving the purge issues with no mc spikes noted outside of p-level changes. the root cause of the thrombus and sepsis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4287 | 1220648-2025-47483 | 23292334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23292334&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657745 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 151 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 had suction while supporting a patient. the impella p-level was decreased, and one unit of albumin and blood was given. additionally, the patient was going to be evaluated for heparin-induced thrombocytopenia due to low platelets. two units of blood were given the day before. it was further reported that a ct showed subacute nonhemorrhagic infarct in the right frontal lobe. the patient continued impella support and was explanted as planned. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was discarded. the investigation was completed. thrombocytopenia/stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: clinical reported suction occurring a number of times during the case, but it would resolve after fluid administration. the product was not returned for investigation. data log review confirms suction triggering throughout the case. at the time of suction, ps trends down to low magnitudes, triggering suction and "impella position unknown" alarms. clinical reported at each suction event, albumin and/or blood was administered, and suction resolved. the cause of the suction was most likely related to the patients condition, as the downward trend in ps indicates the patient had volume issues, which was confirmed by clinical details stating suction alarms resolved when fluid was given.
4288 | 1220648-2025-47478 | 23292362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23292362&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719149 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Atrial Fibrillation | Ischemia | Respiratory Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that the patient experienced atrial fibrillation with rapid ventricular response and was treated with amiodarone. spontaneously converted to sinus rhythm. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during the support, the patient's left lower extremity had become ischemic because of the impella being in place. the patient had severe peripheral artery disease (pad) prior to placing the impella. additionally, there was a disparity between the patient's radial arterial line and placement signal (70/60 versus 173/72) potentially mediated by the patient's severe pad which had made intervention planning challenging for the team. the impella position was stable on echocardiogram. additionally on day two of support, the patient was intubated for respiratory failure but its unknown if this is related to the device. the impella was removed due to ischemia. the patient's outcome is unknown. |  |  |  |  | Correction h6. the investigation of the reported failure mode has been completed. no product was received for evaluation. atrial fibrillation/ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. respiratory failure: the patient had respiratory failure on (b)(6). the cause of the injury was not determined due to insufficient clinical details. optical signal issue: data logs show that upon starting the pump the placement signal (ps) was 130mmhg and then the patient would decompensate causing the ps to go from ~130mmhg to 60mmhg and then it would increase again until the ps would lower again with slight changes in flow. motor current was relatively flat. the ps dropping most likely relates to the patient decompensating and then receiving boluses of drugs which then causes it to increase after. this occurred throughout the duration of the case. no alarms were triggered. the logs do not show an error with the calibration of the optical sensor. due to lack of product returned and insufficient clinical details that confirm the timing of drug boluses as well as lack of information regarding the a-line readings, the root cause of the optical signal issue cannot be established. device history review: the device passed all the post sterile inspection checks. | Additional information/correction: b5, h6. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
4289 | 1220648-2025-47485 | 23292639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23292639&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594879 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered pump stop during impella cp support. after around a month of pump support, the pump suddenly stopped. the medical team did not know the cause of the pump stop. recommendations were given, including to removed or exchange the pump, and restart of the pump was not possible. no further information is noted on the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump stop: the cause of the pump stop could not be determined due to insufficient clinical details, and lack of product or data logs returned. device history review: the device passed all post sterile inspection checks.
4290 | 1220648-2025-47482 | 23292743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23292743&pc=OZD | 09/14/2025 | 11/04/2025 |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682573 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 142 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Fever | Thrombocytopenia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was fever, pulmonary embolism, and thrombocytopenia during impella 5.5 patient support. while on support, it was noted that the patient remained febrile on multiple antibiotics. a head computed tomography was negative and showed multiple pulmonary emboli. the patient had low platelets with system heparin off. the patient was getting platelet infusions. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The cause of the thrombocytopenia, fever, and pulmonary embolism was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
4291 | 1220648-2025-47486 | 23293914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23293914&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734403 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 71 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, chest x-ray was unchanged and stable pulmonary edema was noted. | [5 empty] | Investigation summary: the investigation has been completed. no product was returned for investigation. pulmonary edema: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4292 | 1220648-2025-47487 | 23294143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23294143&pc=OZD | 09/14/2025 | 12/09/2025 |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026738332 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 100 KG | Hispanic |  | Death | Yes | Yes | Malposition of Device | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp smart assist per protocol. however, the patient had moderate to severe aortic stenosis (as), making the positioning difficult. single access was obtained and percutaneous coronary intervention (pci) was performed with stent placement. the team was able to wean levophed and removed the peel away sheath at the end of the procedure. the reposition sheath was inserted and the impella was repositioned. the pump was deep; however it was the best placement that had minimal ectopy. the family made the decision to make the patient do not resuscitate and care was withdrawn. the patient ultimately expired. | [5 empty] | The investigation for the reported arrhythmia, delivery issue, positioning issue has been completed. no product wa returned for review. the root cause of tharrhythmia was unable to be determined due to insufficient clinical details. the cause of the delivery issue and positioning issue was determined to be patient condition as the patient had moderate to severe aortic stenosis (as) preventing successful delivery of impella. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4293 | 1220648-2025-47488 | 23294274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23294274&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 67 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient on impella cp support developed bleeding from the impella access site. three units of blood were administered. two mattress sutures were placed, and bleeding was resolved. it was further reported that due to the patient¿s cancer that had metastasized care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product received for investigation. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4294 | 1220648-2025-47491 | 23294471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23294471&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026715031 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Death | Yes | No | Increase in Pressure | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered high purge pressure, placement signal issue and thrombus during impella 5.5 support, which was second pump for the patient after an impella cp. it was reported that high purge pressure alarm and impella in aorta alarm was on. however, the pump was found to be freely positioned in the ventricle at approximately 4.9 cm. additionally, due the possibility of thrombus, the decision was made on starting lysis therapy. the purge was changed to lysis, which resolved the issues and made purge values go back within normal range. flow data improved after repositioning the pump. however, after 18 days of support, care was withdrawn and the patient expired. | [5 empty] | Product complaint # (b)(4). investigation summary: no product returned. only the imc logs were recovered, no lt or rt logs. purge pressure high - after 2 days on support, pp high alarms after repositioning pt. no kinks seen. lysis executed and purge issues said to have resolved within 24 hours. the cause of the purge pressure high was not determined due to no product returned, no data logs recovered, and insufficient clinical details. optical sensor issue - after 2 days on support, impella in aorta alarms after repositioning pt; however, pump noted to be positioned in the ventricle. the cause of the optical signal issue was not determined due to no product returned, no data logs recovered, and insufficient clinical details. patient codes - thrombosis: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the cause of the thrombus was unable to be determined due to insufficient clinical details. device history lot: device lot: 1947648. device history batch: subcomponent lot: n/a. device history review: pump s/n (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4295 | 1220648-2025-47489 | 23295521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23295521&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706521 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during the support the patient experienced toresades which requires 1 defib. physician planned to place an implantable cardioverter defibrillator (icd). the physician does not believe it was device related, but due to electrolyte imbalance. patient was successfully weaned and explanted after 7 days of support post issue. | [5 empty] | Correction: d3, h3. the investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4296 | 1220648-2025-47490 | 23295540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23295540&pc=OZD | 09/14/2025 | 10/14/2025 |  | 10/15/2025 |  |  | 09/14/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719722 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was a purge system blocked alarm and high purge pressure during impella 5.5 patient support. while on support, the perfusionist reported an alarm for purge system blocked. the physician decided to start lysis with the purge flow under 1ml per hour and the purge pressure 1056mmhg. the perfusionist reported that lysis was successful; the purge flow and purge pressure was in range (7ml per hour and 500mmhg). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. purge pressure high/purge system blocked: the cause of the high purge pressure/purge system blocked could not be determined as no product or logs were returned for evaluation. device history review: the device passed all post sterile inspection checks.
4297 | 1220648-2025-31495 | 23296681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23296681&pc=OZD | 01/27/2025 | 10/14/2025 |  | 10/15/2025 |  |  | 01/27/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601306 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. while on support there were placement signal issues. they found that the placement signal was not reliable and was failing. however, the flow and motor current were stable and support continued with the pump without any further intervention for the issue. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were reviewed, and the out of box placement signal not reliable (psnr) alarm was confirmed. this pattern matches an air gap between the patient cable plug and console receptacle. returned product was investigated and there were no abnormalities. the pump passed laser continuity, 5s parameters were all within specification when connected to a console, and psnr did not reproduce. based on data logs and returned product, the root cause of the optical signal issue was determined to most likely be an air gap between the aic and patient cable plug. since clinical details did not provide any details regarding when/how the psnr alarm resolved, it can not be verified. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4298 | 1220648-2025-47492 | 23299123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23299123&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/17/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710687 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male | 80 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and following a ¿successful¿ implant of the impella device a leak in the purge cassette was detected approximately two hours later. the cassette was successfully replaced. however, three days later the patient expired noted to be ¿not impella related.¿ the patient arrived via ambulance from home and quickly deteriorated. the patient was in cardiogenic shock, and the impella cp device was implanted prior to intervention. there were no complications with placement and none afterwards. the impella is running well. however, approximately two hours later a purge cassette leak was detected and it was changed. the following day, a call was made to the emergency call center to discuss the case, and it was noted that the pump runs well but the patient has a ¿bad¿ prognosis; catecholamines needed to be increased, and escalation was not an option. the next day it was noted the patient¿s overall condition was now ¿very bad.¿ the pump runs well and there were no signs of heart recovery. further higher doses catecholamines were needed; again, escalation was not an option. thereafter, next day, the patient remained critically ill and expired. | [5 empty] | Medical safety reviewed the event. the event describes a patient who presented in critical condition in cardiogenic shock, decompensating on arrival with no other therapy options and subsequently expired. there was no report or indication of pump malfunction and/or interruption, impact to the patient¿s impella mechanical circulatory support. it was noted that the impella ran well, and the adverse outcome is ¿not impella related.¿ in this case, the patient¿s underlying condition likely contributed most to an outcome of death (cardiogenic shock stage extremis). the patient¿s history/comorbidities details are unknown. however, conservatively, the death should be reported (against the impella pump) as the device was in use at the time of the expiration. note: purge cassette is not reportable; the chance of death or serious injury is remote. there is very limited; no potential for adverse impact to the patient given the details of this event. | The event describes a patient who presented in critical condition in cardiogenic shock, decompensating on arrival with no other therapy options and subsequently expired. there was no report or indication of pump malfunction and/or interruption, impact to the patient¿s impella mechanical circulatory support. it was noted that the impella ran well, and the adverse outcome is ¿not impella related.¿ in this case, the patient¿s underlying condition likely contributed most to an outcome of death (cardiogenic shock stage e). however, patient¿s history/comorbidities details are unknown. as the device was in at the time of the demise outcome and cannot be dissociated, conservatively the report for death is being submitted for the impella pump (not including the cassette leak). patient-specific information a5: ethnicity and a6: race is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4299 | 1220648-2025-47497 | 23300118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300118&pc=OZD | 09/05/2025 | 10/08/2025 |  | 10/15/2025 |  |  | 09/18/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1631552 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male | 65 KG |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that there was a battery failure with the automated impella controller (aic). the biomedical engineer onsite reported that there was a battery failure with the console. battery issue on console ic9327 was that when it was plugged in, the battery did not charge past 75% and when unplugged the battery did not drop past 75%. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4300 | 1220648-2025-47493 | 23300154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300154&pc=OZD | 09/18/2025 |  |  | 10/15/2025 |  |  | 09/18/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026728688 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that the patient presented to the hospital with worsening heart failure. an intro-aortic balloon pump was placed via axillary artery during a left ventricular assist (lvad) work up, however the following days the patient¿s hemodynamics worsened. the patient¿s ejection fraction was calculated to be 10% with a decreased mixed venous oxygen level of 47%, therefore the decision was made to escalate the patient to an impella 5.5 device. the impella 5.5 was inserted via axillary artery and support was started with no issues, however within the first 7 hours of support the pump abruptly stopped functioning twice. each time, the pump was able to be restarted following the high motor current alarm with pump stop, however the p level would not go above p1 or p2 before stopping again. additionally, it was reported that the patient experienced hematuria. a place was made to remove and replace the pump. the patient was brought back to the operating room for an impella 5.5 replacement. upon removal of the first pump, a clot was observed on the inlet of the pump. additionally, the axillary artery 8mm hemashield graft was found to have clots with in it. the physician suspected the reason of the clots was due to suspected heparin induced thrombocytopenia of the patient. the graft was removed and replaced with a new 10 mm hemashield platinum graft which was anastomosed to artery. the new impella 5.5 was implanted without issue, and the patient was released back to the intensive care unit (icu). upon returning to the icu, the patient was noted to have a facial droop and left side weakness. the patient was taken for an emergency computed tomography scan where a clot was found, and the patient was taken for an emergent thrombectomy. the following day the patient was noted to have no deficits from the thrombotic event. additionally, it was noted that the impella 5.5 was not well placed, and the surgeon was notified that the impella would need to be repositioned. the following day, the impella positioned was discussed. it was recommended to the surgeon to reposition the impella as the inlet was facing the mitral valve, and the pump was slightly deep at 6.4 cm. although the concern that the malposition could contribute to ectopy, the physician declined to reposition stating that the ventricle was large and was satisfied with the pump position. two days later, the pump was repositioned and pulled back to 5.1 cm and the power level was able to be increased from p-7 to p-8. the patient was placed on a two-week hold for additional surgery due to the stroke. the following day, it was reported that there were several impella in aorta and impella position unknown alarms which occurred during the previous night. echocardiography was performed to verify impella position, and impella position was confirmed to be well placed. on the seventh day of impella support, it was reported that the automated impella controller (aic) experienced an unexpected shut down. no alarms had been observed, and the medical team replaced the aic to resolve the issue. on the fourteenth day of support, it was reported that there were ongoing false placement signal alarms, and the medical team disabled the alarm. the following day it was reported that the patient was experiencing gastrointestinal bleeding. anticoagulation was withheld due to the ongoing bleeding and low hemoglobin. a unit of packed red blood cells were transfused as result of blood loss. on the eighteenth day of support, it was reported that there were placement signal and suction alarms. it was noted that there were no signs or symptoms of hemolysis due to the suction alarms, and that the patient¿s plasma free hemoglobin was <30 with adequate yellow urine output. on the twenty-fifth day of support, it was reported that the impella pump was repositioned due to the patient experiencing several runs of ventricular tachycardia. the patient remains on impella 5.5 support, with a plan for surgery for a durable lvad in the next week. this complaint involves three impella devices: pump 1: impella 5.5, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). automated impella controller with serial number (b)(6). this mdr specifically addresses pump 2: impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The cause of the false alarms could not be determined as insufficient clinical details were provided and product wasn't returned for investigation. the cause of the alarms and optical signal issue was determined to be due to sensor wear. the cause of the bleeding cannot be determined as insufficient clinical details were provided. the root cause of the tachycardia and stroke was unable to be determined due to insufficient clinical details. the device passed all post-sterile inspection checks. | Additional information received for d6 explant. | A4,a6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4301 | 1220648-2025-47495 | 23300220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300220&pc=OZD | 08/21/2025 |  |  | 10/15/2025 |  |  | 08/21/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716054 |  | 0048-0045 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a gastrointestinal (gi) bleed. systemic anticoagulation was withheld to obtain hemostasis. several days later, the patient had an upper gi bleed and had to get reintubated for an endoscopy. the source of bleeding was found and intervention was done. the patient was escalated to impella 5.5 to support coronary artery bypass graft (cabg) and ventricular septal defect (vsd) repair. after the patient had cabg and vsd repair, the patient soon expired while on impella support. | [5 empty] | The investigation has been completed. major bleed : the cause of the injury was not determined since the product was not returned and insufficient clinical details were provided. device history review: the pump passed all the post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
4302 | 1220648-2025-31838 | 23300328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300328&pc=OZD | 08/27/2024 | 10/15/2025 |  | 10/15/2025 |  |  | 08/27/2024 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507182 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella cp support. during support, a 'placement signal unreliable' alarm was present. the alarm was disabled. there were no kinks observed, and the cable was securely connected. the device was not explanted due to this issue. a decision was made to withdraw care, and the patient subsequently expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the failure mode was changed from placement signal issue to optical signal issue because the impella cp only has an optical sensor and, therefore, can only have optical signal issues. data logs were reviewed, and the "placement signal not reliable" alarm was confirmed. product was not returned for investigation. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4303 | 1220648-2025-47013 | 23300584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300584&pc=PYX | 08/21/2025 |  |  | 10/15/2025 |  |  | 08/21/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661955 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 123 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp flex and impella cp (bipella) for mechanical circulatory support following mitral valve and tricuspid valve surgery. due to the patient¿s severe pulmonary hypertension and poor right ventricular function, the decision was made to place the impella rp flex status post surgery. during chest closure, prior to impella rp flex implant, the patient¿s hemodynamics worsened and the decision was made to implant the impella cp after the impella flex was inserted and support initiated. the right internal jugular vein was accessed for the impella rp flex, and the impella rp flex was implanted without issues. the impella rp flex position was noted to be optimal when imaged with fluoroscopy, however it was noted that the placement signal of the impella rp flex was unreasonably high with a pulmonary artery catheter pressure higher than expected and pump flow was lower than expected for the p level of the pump. the placement signal issues did not resolve with adjusting the y-axis of the pump, and the medical team was advised by abiomed support that there was a potential issue with the optical sensor which was causing an inaccurate flow calculation. the impella cp was placed in the right femoral artery without issue, and the patient was brought back to the intensive care unit for rest and recovery. the patient then received volume resuscitation, after which the placement signal and flows returned to normal and expected levels. it was advised to the medical team by abiomed support that the optical sensor had likely been impeded by a vessel wall and correlated with hypovolemia. the following day, the patient was transferred to another hospital for advanced care. it was noted that the patient¿s right leg was cold, and pulses were unable to be found with doppler. additionally, a transesophageal echocardiogram was performed and revealed blood in chest causing right atrial compression. a plan was made to obtain computed tomography of the head, chest, and pelvis and then take the patient to the operating room for a chest washout, right femoral artery cutdown, and remove both impella pumps. if needed, the patient would then be cannulated for extracorporeal membrane oxygenation (ecmo). after the chest washout was performed, the patient¿s hemodynamics improved, and it was determined the patient did not require ecmo after impella pumps were explanted. the impella cp was emergently removed due to critical limb ischemia. ultimately, the patient required above knee amputation of the right leg due to complications of the limb ischemia. after both the impella rp flex and the impella cp were explanted, it was decided that the patient did not need to be cannulated for ecmo as the patient was hemodynamically stable and heart function had recovered enough for the patient to be stabilized with pressors and intubation. this complaint involves two impella devices: pump 1: impella rp flex with serial number (b)(6). pump 2: impella cp with serial number (b)(6). this report specifically addresses the impella rp flex. another report will be submitted for the other device associated with this complaint. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
4304 | 1220648-2025-47494 | 23300614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300614&pc=OZD | 09/24/2025 |  |  | 10/15/2025 |  |  | 09/24/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1076156 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Battery Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that the patient presented to the hospital with worsening heart failure. an intro-aortic balloon pump was placed via axillary artery during a left ventricular assist (lvad) work up, however the following days the patient¿s hemodynamics worsened. the patient¿s ejection fraction was calculated to be 10% with a decreased mixed venous oxygen level of 47%, therefore the decision was made to escalate the patient to an impella 5.5 device. the impella 5.5 was inserted via axillary artery and support was started with no issues, however within the first 7 hours of support the pump abruptly stopped functioning twice. each time, the pump was able to be restarted following the high motor current alarm with pump stop, however the p level would not go above p1 or p2 before stopping again. additionally, it was reported that the patient experienced hematuria. a place was made to remove and replace the pump. the patient was brought back to the operating room for an impella 5.5 replacement. upon removal of the first pump, a clot was observed on the inlet of the pump. additionally, the axillary artery 8mm hemashield graft was found to have clots with in it. the physician suspected the reason of the clots was due to suspected heparin induced thrombocytopenia of the patient. the graft was removed and replaced with a new 10 mm hemashield platinum graft which was anastomosed to artery. the new impella 5.5 was implanted without issue, and the patient was released back to the intensive care unit (icu). upon returning to the icu, the patient was noted to have a facial droop and left side weakness. the patient was taken for an emergency computed tomography scan where a clot was found, and the patient was taken for an emergent thrombectomy. the following day the patient was noted to have no deficits from the thrombotic event. additionally, it was noted that the impella 5.5 was not well placed, and the surgeon was notified that the impella would need to be repositioned. the following day, the impella positioned was discussed. it was recommended to the surgeon to reposition the impella as the inlet was facing the mitral valve, and the pump was slightly deep at 6.4 cm. although the concern that the malposition could contribute to ectopy, the physician declined to reposition stating that the ventricle was large and was satisfied with the pump position. two days later, the pump was repositioned and pulled back to 5.1 cm and the power level was able to be increased from p-7 to p-8. the patient was placed on a two-week hold for additional surgery due to the stroke. the following day, it was reported that there were several impella in aorta and impella position unknown alarms which occurred during the previous night. echocardiography was performed to verify impella position, and impella position was confirmed to be well placed. on the seventh day of impella support, it was reported that the automated impella controller (aic) experienced an unexpected shut down. no alarms had been observed, and the medical team replaced the aic to resolve the issue. on the fourteenth day of support, it was reported that there were ongoing false placement signal alarms, and the medical team disabled the alarm. the following day it was reported that the patient was experiencing gastrointestinal bleeding. anticoagulation was withheld due to the ongoing bleeding and low hemoglobin. a unit of packed red blood cells were transfused as result of blood loss. on the eighteenth day of support, it was reported that there were placement signal and suction alarms. it was noted that there were no signs or symptoms of hemolysis due to the suction alarms, and that the patient¿s plasma free hemoglobin was <30 with adequate yellow urine output. on the twenty-fifth day of support, it was reported that the impella pump was repositioned due to the patient experiencing several runs of ventricular tachycardia. the patient remains on impella 5.5 support, with a plan for surgery for a durable lvad in the next week. this complaint involves three impella devices: pump 1: impella 5.5, with serial number: (b)(6). pump 2: impella 5.5, with serial number: (b)(6). automated impella controller with serial number: (b)(6). this mdr specifically addresses automated impella controller with serial number: (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | This report is to correct the h6 medical device problem code. h6. medical device problem code a0705 removed battery problem was reported in the initial report and has been corrected to (b)(6) unintended application program shut down. | A4. a6 is unknown. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
4305 | 1220648-2025-47503 | 23300791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300791&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714554 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had positive blood cultures. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Ppae(infection): the root cause of the infection was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4306 | 1220648-2025-47505 | 23300817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300817&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026724411 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impella cp was removed due to a unknown left ventricular thrombus. the patient survived. | [5 empty] | The investigation was completed and no product was returned for investigation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4307 | 1220648-2025-47501 | 23300838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300838&pc=OZD | 09/26/2025 |  |  | 10/15/2025 |  |  | 09/26/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1887775 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) failed to recognize the purge disc. multiple attempts were unsuccessful. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4308 | 1220648-2025-47509 | 23300939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23300939&pc=OZD | 08/25/2025 |  |  | 10/15/2025 |  |  | 08/25/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705171 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | False Alarm | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Thrombocytopenia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced heparin induced thrombocytopenia (hit), bleeding, tachycardia, ectopy, hypotension, and false alarms. proper pump positioning was confirmed via an echocardiogram. the patient was treated with bivalirudin for the hit. the patient was started on amiodarone and was cardioverted in response to the ectopy. one unit of packed red blood cells was transfused in response to oral bleeding. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4309 | 1220648-2025-46767 | 23301202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301202&pc=OZD | 08/18/2025 |  |  | 10/15/2025 |  |  | 08/18/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696501 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 132 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient admitted in with known cardiac history and cardiogenic shock. the pump was placed but the delivery was ischemia inducing so the team placed an antegrade sheath to perfuse the limb. the pump was explanted and escalated to an impella 5.5 after 11 hours of support. this report represents the impella cp. another report was submitted for the impella 5.5. refer to section h.10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
4310 | 1220648-2025-47498 | 23301297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301297&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569592 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Death | Yes | Yes | False Alarm | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Asystole | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | A 56-year-old male patient was already in the hospital with acute myocardial infarction cardiogenic shock with a mean atrial pressure (map) of 102 mmhg (under 5 catecholamines and respiratory support) with increasing arterial lactate of 9.2 mmol/l with a left ventricular ejection fraction of 15 mmhg in society for cardiovascular angiography and interventions shock stage e on impella cp support. for escalation of care, right axillary/subclavian access for impella 5.5 placement was successfully obtained without any complications. the change of the heart pumps was done by general recommendations and vessel closure was also without issues done by manual compression. the patient was adequately supported after pump change by the impella 5.5 with 5.3 liter/ minute and lost pulsatility, so, he was completely dependent on the heart pump completely took over the heart function. protamine was given to reverse heparin administration. moreover, as the hemoglobin dropped, the patient received 2 units of packed red blood cells per anesthesia. the first complaint was filed after initiation of 5.5 therapy as the impella automatic controller (aic) showed a false "impella in aorta alarm." the correct position was confirmed via transesophageal echocardiogram (tee) as best practice. as the alarm did not resolve (because no change in position was done) the placement signal has been disabled from now on. a pulmonary artery catheter was in place, and the patient was sent to the intensive care unit (icu). the second complaint was documented on the same day on the icu. as the patient remained critically ill, he required medication during impella 5.5 support (epinephrin was removed, while norepinephrine and dobutamine was increased or added). patient remained non-pulsatile on impella 5.5 and a map of greater than 65. one suction alarm occurred but resolved by reducing the performance level to p8. the patient was bleeding from all sites except for the impella site so, he received 2 units o fresh frozen plasma (ffp). as planned when on impella 5.5, continuous renal replacement therapy (crrt) was started at 1700 and pulling 50cc/hour. on the second day of impella 5.5 support, positive blood cultures from methicillin-resistant staphylococcus aureus (mrsa), so penicillin was started with vancomycin (unasyn). on the third day of impella 5.5 support, the patient remained critically ill on high vasopressor doses and has bleedings from nose, stomach, artery line site, and also impella 5.5 site. two ffps and one packed red blood cell unit was given. the impella 5.5 access site was free of hematoma. the placement monitoring on the aic was still disabled and the patient had no native pulsatility and the console showed false ¿impella in aorta alarm (see complaint 1). eventually, the patient was with no neurological reflexes and a lengthy discussion held with family led to the decision to withdraw care. the patient died on impella 5.5 crrt, and vasoactive medication when the machines were turned off. | It was reported that a patient on impella cp with worsening cardiogenic shock, increasing lactate, and maxed on 5 vasopressors was being escalate to impella 5.5 and start continuous renal replacement therapy (crrt). the impella cp pulled back into descending aorta. 5.5 advanced into left ventricle and flows slowly ramped up. however, the patient lost all pulsatility when put on 5.5 support. 50 of protamine given and pocket closed. cp explanted and manual pressure held for 30 mins with excellent hemostasis. it was reported of a false impella in aorta alarm after initiation of 5.5 therapy. position confirmed. placement signal was disabled since implant of 5.5. during support the patient remained critically ill. there was a suction alarm with resolution by dropping to p8. patient was bleeding from all sites except impella site and received (b)(4) units fresh frozen plasma. crrt was started. there was also positive blood cultures and methicillin-resistant staphylococcus aureus in the nares so starting vancomycin and unasyn. it was noted that the patient remained intubated and unresponsive. remained critically ill on high vasopressor doses. pt bleeding from nose, stomach, art line site, and impella site. received more fresh frozen plasma and one packed red blood cells. impella site with no hematoma and dressing changed. placement monitoring disabled on console as patient has no native pulsatility and false ¿impella in aorta alarm¿. ultimately care was withdrawn and the patient expired. |  |  |  |  | H1. type of reportable event was incorrectly selected on the original report. h6. health effect- impact code death was inadvertently missed on the original report. | Additional information was received and noted the automated impella controller may be the potential cause of the false alarm/buzzer issue. medical safety review. medical safety/clinical reviewed the event. a 56-year-old male presented to the hospital with an electrocardiogram demonstrating an anterolateral st-elevation myocardial infarction (stemi) and was emergently taken to the cardiac catheterization laboratory. left heart catheterization revealed occlusion of the left anterior descending artery with thrombus extending into the left main coronary artery and the ostium of the circumflex artery. following the procedure, the patient became increasingly agitated and required urgent intubation. in the setting of cardiogenic shock due to acute myocardial infarction, an impella cp was placed prior to percutaneous coronary intervention. after device placement, the patient developed gross hematuria and a groin hematoma at the access site, which was described as soft and non-expanding. this event is considered a serious injury attributable to the impella cp. despite support, the patient¿s cardiogenic shock worsened, with rising lactate levels and the need for five vasopressors. care was escalated to an impella 5.5. following initiation of impella 5.5 support, a ¿false impella in aorta¿ alarm occurred. device position was confirmed by transesophageal echocardiography, and the placement signal was subsequently disabled. the patient developed acute renal failure with minimal urine output and was started on continuous renal replacement therapy. the patient experienced bleeding from multiple sites, excluding the impella access site, and received transfusion support including fresh frozen plasma and packed red blood cells. the patient remained intubated and unresponsive, and neurological assessment was limited due to hemodynamic instability. on the following day, blood cultures were positive, and methicillin-resistant staphylococcus aureus was identified in the nares; antibiotic therapy was initiated. the patient remained critically ill on high-dose vasopressors and exhibited absence of neurological reflexes. the family elected to withdraw life-sustaining therapy and the patient expired. the death is most likely due to the patients underlying critical condition. related medial device reports. this is now one of three devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. following the cross-functional review by the investigation team, the automated impella controller was identified as a potential contributor to the false alarm/buzzer activation. there is no report or indication of any interruption to patient support; the automated impella controller appeared to remain fully operational throughout the case. based on this assessment, the event has been classified as a reportable malfunction accordingly. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. | This is one of two reports and this one reflects the pump and the other reflects the controller h10 was updated with the related report number 1220648-2025-48895. | Event description details and complaint/patient coding was updated/corrected following further clinical review. therefore, sections b5 event description has been updated and h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the coding that accurately reflects the reported event. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4311 | 1220648-2025-47500 | 23301339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301339&pc=OZD | 09/22/2025 | 11/20/2025 |  | 10/15/2025 |  |  | 09/22/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636854 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pumping Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient was placed onto impella support due to asystole. while on support, the patient experienced an access site bleed and hematoma that was resolved by removal of the pump. patient expired shortly thereafter. | [5 empty] | This complaint was incorrectly assessed as a death. this pump was never used on a patient and has been updated to reflect a malfunction. refer to mwr-19112025-0001976792 for details on the reported patient death. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4312 | 1220648-2025-47450 | 23301365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301365&pc=OZD | 09/15/2025 | 10/13/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1635044 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during a prep when powering on the automated impella controller (aic), a system check failure alert was present. the buttons on the bottom of the console were checked and the appropriate button was on. the console was unplugged and swapped out. no patient was involved or was harmed due to the issue. | [5 empty] | H6 medical device problem code has been corrected from a1102 to a1104. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
4313 | 1220648-2025-47499 | 23301410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301410&pc=OZD | 08/21/2025 |  |  | 10/15/2025 |  |  | 08/21/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698058 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that upon starting the device, flows would not go above 1.5 liters per minute (l/min). the motor current was lower than expected and a suction alarm was present. it was explained to the surgeon that there was a possibility that the pump ingested a clot ingestion as the patient had suspected heparin inducted thrombocytopenia and a low platelet count of 60,000/l prior to arriving at the cardiovascular operating room. the surgeon elected to remove the impella cp and insert new impella cp device. a clot was noted on the inlet once device was removed. a new impella cp was prepped and the sheath sidearm was flushed multiple times, however bleed back was noted upon impella cp insertion. placement in the left ventricle was confirmed with imaging, and the impella was started with flows as expected at 3 l/min. p level was decreased to p7 and suction was noted. to resolve the suction issues, the patient was given albumin. it was noted that after support was initiated, urine output increased, and the patient was able to be weaned off of levophed. the patient was released to the intensive care unit for recovery. four days later, the patient was no longer making urine and started continuous renal replacement therapy (crrt). it was noted the patient had poor prognosis at that time. the patient remained on crrt, and although very ill, the left heart function recovered for impella cp to be slowly weaned. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. this complaint involves two impella cp pumps: pump 1: impella cp with serial number (b)(6). pump 2: impella cp with serial number (b)(6). this report specifically addresses pump 2, impella cp with serial number (b)(6). a separate report will be submitted for the other device associated with this complaint. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product returned. pump suction - after implant, pump positioning was confirmed by fluoroscopy and transesophageal echocardiogram, but suction noted at p-7. albumin was given. data log review shows the majority of suction alarms occur above p-6 and resolve with a reduction in p-level. no motor current spikes outside of p-level changes. the cause of the suction alarms was patient condition related due to lack of volume that triggered suction at elevated p-levels. renal failure - the cause of the renal failure was not determined due to insufficient clinical details. device history lot: device lot: 1945887 device history review: pump (b)(6) passed all post sterile inspection checks.
4314 | 1220648-2025-47506 | 23301488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301488&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026715430 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male | 85 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Multiple Organ Dysfunction Syndrome | No Clinical Signs, Symptoms or Conditions | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and expired on support two days later. the patient presented for a high-risk pci for a blocked left internal mammary artery to the left anterior descending artery (lima-lad). the impella cp was implanted pre-pci. pci was coronary angioplasty, and it was unknown if untreated diseased was greater or equal to 50%. no further information regarding pre and post procedure was noted. in the case wrap up notes it was noted the patient experienced multiorgan failure and sepsis (onset of these adverse events were unnoted). the patient expired on support 2 days after start of the impella mcs. additional information was received and noted the patient came in a "really bad" shape overall with confirmed sepsis. the patient had coronary artery bypass graft (cabg) surgery and was in the intensive care unit for a prolonged time before cardiology took over and tried to resolve blocked lima-lad bypass. sepsis and liver and kidney failure were confirmed before impella insertion. additionally, the interventional cardiologist confirmed that the patient's health status was already compromised, and it is not related to the impella mcs. | It was reported that a patient implanted with an impella cp experienced organ failure and sepsis. ultimately, the patient expired. |  |  |  |  | The cause of the organ failure and sepsis was not determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | Sections a4 weight, b5 event description, and b7 medical history/preexisting condition were updated to include clinical details at hand and subsequently received additional information. b2 revised date of death as it was reported incorrectly on the initial report that was submitted. medical safety review. medical safety reviewed the event. the event describes patient post coronary artery bypass graft surgery with a block bypass (noted during post-surgical care). the patient was in a decompensating state and had developed sepsis and went into liver and kidney failure. a decision was made to place the patient on an impella cp for mechanical circulatory support (mcs) for a planned intervention to the blocked left internal mammary artery to the left anterior descending artery (lima-lad). the patient expired 2 days later on support. there is no report or indication that the impella cp malfunctioned or was associated with an adverse event during support. the impella cp was in use during patient expiration but would not have materially contributed to the patient death given patient underlying condition compromised with sepsis and multiorgan failure that was confirmed prior to impella support. however, conservatively, the death should be reported as device was in use at time of expiration. as a result of medical safety review outcome, the sepsis and multi-organ failure adverse events have been dissociated from the impella cp device used. therefore, h6 health effect - clinical and impact coding have been updated accordingly. note: h6 medical device problem and investigation type/findings/conclusion remain unchanged as previously reported. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4315 | 1220648-2025-47504 | 23301584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301584&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699471 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient¿s urine was showing some hemolysis while on impella cp support. the impella was repositioned. the pump was later explanted and the patient survived. | [5 empty] | The investigation was completed. data review: data logs showed pump ran without issue during support. there were no alarms triggered during the case. there were no mc spikes, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. the device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4316 | 1220648-2025-47502 | 23301624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301624&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674417 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hematuria | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hematuria. an echocardiogram showed the pigtail up against the ventricle wall. additionally, the patient had a hematoma at the groin site. the patient stopped following commands and was twitching. the patient was escalated to an impella 5.5 pump. | [5 empty] | Investigation summary: no product was returned. hematoma, bleed: the cause of the hematoma and bleed was not determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4317 | 1220648-2025-47513 | 23301743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301743&pc=OZD | 08/28/2025 |  |  | 10/15/2025 |  |  | 08/28/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714024 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Difficult or Delayed Positioning | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was product damage during impella 5.5 support. during repositioning at the bedside, the device¿s anti-contamination sleeve and top cap separated causing the physician to pull the impella 5.5 back into the ascending aorta. the patient returned to the operating room for emergent impella 5.5 replacement and survived. | [5 empty] | A.4 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H.6 added medical device problem code as it was inadvertently omitted from the initial report that was submitted. investigation summary: no products were returned for evaluation. product damage: clinical details noted that sterile sleeve and top cap became separated during pump repositioning at bedside. the cause of the product damage could not be determined as no product or images were returned for analysis. positioning issues: clinical details noted that during repositioning, the sterile sleeve and top cape became separated and caused the physician to pull the pump back into the ascending aorta. the cause of the positioning issues was a user related issue as the pump was pulled back into the ascending aorta by the physician per clinical details. device history review: the pump passed all post-sterile inspection checks.
4318 | 1220648-2025-47507 | 23301930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23301930&pc=OZD | 08/25/2025 | 10/29/2025 |  | 10/15/2025 |  |  | 08/25/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639912 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Myocardial Infarction | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient who presented as st-segment elevation myocardial infarction was taken to the cardiac catheterization lab for emergent intervention. while on the table, the patient coded and required cardiopulmonary resuscitation. the medical team made every effort to access the femoral artery, but had difficulty for several minutes. upon gaining access, an image was taken of the femoral artery, which was extremely small. the pump had already been opened and prepped, but the physician did not feel it was possible to place the sheath through the leg safely. the pump was discarded after being unable to pass through vasculature due to peripheral artery/vascular disease. the vessel diameter was less than four millimeters. an impella cp pump was opened and connected, but was never attempted as resuscitation efforts were stopped. the patient expired. | [5 empty] | No product was returned for investigation. the cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to patient having pad/pvd and had resistance in femoral artery and an image was taken of the femoral artery, which was extremely small. impella insertion was aborted. the cause of the clinical symptom was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A.2, a.4, and a.5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4319 | 1220648-2025-47508 | 23302030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302030&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705340 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male | 68 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had a impella 5.5 pump placed as care escalation from a impella cp pump. the 5.5 was supporting while awaiting a transplant. while on support it was reported that there placement signal and lv signal not displayed, beside this pump worked properly. the patient remained on support and there was no patient harm. | [5 empty] | The investigation was completed. optical sensor issue: clinical report missing left ventricle placement signals. the logs show placement signal not reliable alarms and the 5s parameters are of the pattern identified as an oscillating contrast issue. this issue has been characterized as a console-related issue. the device history review was completed and passed all post sterile inspection checks. | In section g (manufacturing site name), the abiomed europe gmbh site address and required site information were inadvertently omitted from the initial report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4320 | 1220648-2025-47512 | 23302056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302056&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662161 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 87 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support patient developed bleeding from the impella access site. manual pressure was held. two units of blood products were administered, and albumin bolus was given. patient is stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4321 | 1220648-2025-47510 | 23302058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302058&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705345 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. on the eighth day of support, it was reported that the side arm of the pump was leaking. the medical team reported that it was leaking intermittently, and that the purge flow and purge pressure were in expected ranges. there were no issues observed with the luer connections or the purge cassette. no patient harm was reported. it was noted that the pump was explanted 11 days later, however patient outcome is unknown. | Additional information was provided. the device was explanted successfully after weaning. the explant was not considered premature and was unrelated to the reported complication. the patient's outcome at the time of explant was survived. |  |  |  |  | A2, a4, a5, a6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 b5 additional information was added for received data logs. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. no issues with luer connections or purge cassette. data log review showed no significant changes in purge pressure/purge flow during the case. the cause of the purge leak was not determined since product was not returned. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | B5 additional information was added for patient outcome. d6b explant date was revised due to new information received.
4322 | 1220648-2025-47514 | 23302087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302087&pc=OZD | 09/14/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725952 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a placement signal not reliable alarm. the alarm lasted approximately 6 minutes, and self-resolved. proper pump placement was confirmed with echo. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned. the most likely cause of the optical sensor issue was the automated impella controller. the device passed all post sterile inspection checks.
4323 | 1220648-2025-47511 | 23302313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302313&pc=OZD | 08/25/2025 |  |  | 10/15/2025 |  |  | 08/25/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697888 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | White | Injury | Yes | Yes | False Alarm | Failure to Pump | Failure to Advance |  |  |  |  | Heart Failure/Congestive Heart Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that during the setup of the device, when the impella was connected to the automated impella controller (aic), the aic had alarms for high motor current and impella stop prior to the impella being started. it was then reported that there were alarms for purge flow high and impella stopped, retrograde flow. the aic was exchanged as a result of the alarms reported. additionally, it was reported that the patient was unstable and required defibrillation during the attempted implant of the device. it was reported that the physician felt that the patient¿s vessels were too vulnerable to implant the impella cp, and the insertion was aborted. the outcome of the patient is unknown. this complaint involves two impella devices: impella cp with serial number (b)(6) automated impella controller, with serial number (b)(6) this report specifically addresses the impella cp with serial number (b)(6). a separate report will be submitted for the other device associated with this complaint. | [5 empty] | Complaint product experience coding was updated to better describe the reported event. therefore, h6 medical device problem coding was revised accordingly. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, data logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. pertaining to the alarm, right after connecting the pump to the automated impella controller (aic) the alarm "motor current high, impella stops¿ occurred, even if it was not started before and the device was not inserted in the patient. then the alarm "purge fluid increased,¿ and "impella stopped. retrograde flow¿ occurred. it was noted the staff did not properly complete the de-air procedure process. the aic was exchanged for a new one which did not succeed due to delivery issues. regarding the pump did not start, data log review showed motor current spike to 1501 on impella start which triggered an impella stopped (rpm deviation) alarm. the pump never started after that. throughout the 12 minutes the impella was plugged into console, contrast is high at 37.5%, lamp low at 74, all 5s parameters nominal. no purge issues were noted. the cause of the pump not starting was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. the cause of the unable to deliver or advance issue was patient condition related due to the noted vulnerability of the vessels that prevented further implant attempts. regarding the cardiac failure, the cause was not determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller. | A.4 and a.5 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4324 | 1220648-2025-47515 | 23302673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302673&pc=OZD | 08/29/2025 |  |  | 10/15/2025 |  |  | 08/29/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630939 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Male | 150 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a clot and low pump flows. this impella was an exchange. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. low or blocked pump flow: the cause of low or blocked pump flow was unable to be determined as limited clinical details were provided and no product was returned for review. thrombus: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Additional product information received: d4, d6a and d6b updated.
4325 | 1220648-2025-47516 | 23302768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302768&pc=OZD | 09/01/2025 |  |  | 10/15/2025 |  |  | 09/01/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716930 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 52 KG | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Cardiac Tamponade | Pericardial Effusion | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was low pump flow, ischemia, cardiac tamponade, pericardial effusion and minor bleeding during impella cp patient support. while on support, there was low impella flow noted. multiple adjustments were made. an echocardiogram done in the catheterization lab and a large pericardial effusion was noted. an emergent pericardiocentesis was completed with approximately 200 milliliters of blood returned. immediately, mean arterial pressure and impella flows improved significantly. the doctor also did an angiogram to see lower extremity flow, and was only able to see a very small flow to legs. the patient¿s groin was bleeding. the team did not want to use a fem-stop due to low flows into the legs. pedal pulse on left (impella) site were not present. the right pedal pulse was faint. the abiomed representative recommended the suture being redone with forward tension sutures as the suture pad was not forward and could have been at the arteriotomy site. it was noted that the groin had stopped bleeding with re-dressing. the nurse was continually having to work with the pericardial drain due to tamponade that continued to come back. the nurse had decreased drips but impella p level remained the same. the access site was clean, dry and intact. it had not been re-dressed since arrival to the intensive care unit. flow to the lower extremity was reduced due to vessel size and impella sheath. the patient expired. | [5 empty] | No product was returned for investigation. the cause of the bleed was most likely use issue related since hemostasis was achieved by redoing the forward tension sutures. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the root cause of the cardiac tamponade and pericardial effusion was unable to be determined due to insufficient clinical details. the cause of low pump flow was patient related since pericardiocentesis procedure improved flows confirming there was low blood volume due to pericardial effusion. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4326 | 1220648-2025-47519 | 23302920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23302920&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738278 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon arrival at the unit the patient had bleeding from around the site. manual pressure was held as well as fem stop on the insertion site. heparin was held. the physician ultimately decided to return to cardiac cath lab for impella explant. the patient is stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product received for investigation. major bleed: the cause of the injury is most likely use related. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4327 | 1220648-2025-47521 | 23303063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23303063&pc=OZD | 09/05/2025 |  |  | 10/15/2025 |  |  | 09/05/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694951 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 56 KG | Non Hispanic | White | Injury | Yes | Yes | Priming Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered air in purge and the patient had gastrointestinal (gi) bleeding during impella 5.5 support. it was reported that the distal end of the purge tubing was found broken in the pump check valve and that there was an air in purge alarm. the staff was unable to remove from the tubing piece from the check valve. no further intervention was noted and support was continued. also, the patient had a gi bleed for which anticoagulant was withheld. | [5 empty] | No product returned. the cause of the air in purge alarm could not be determined as the product was not returned, the data log review was inconclusive, imc logs were not recovered, and insufficient clinical details were provided. the cause of the gastrointestinal bleed was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4328 | 1220648-2025-31674 | 23303087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23303087&pc=OZD | 07/14/2024 | 10/15/2025 |  | 10/15/2025 |  |  | 02/01/2024 | 10/15/2025 |  |  | 02/28/2026 |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025469459 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella 5.5 pump to support a patient. the pump lost placement signal, but the pump remained on for the support despite the optical signal loss. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product returned. data logs show that about one hour into the case, the placement signal goes out of range with placement signal not reliable (psnr) alarms. signal-to-noise ratio (snr) drops to 0, lamp maxes at 100, gain drops, light drops. contrast oscillates +/- 3200. op metrics do not recover. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4329 | 1220648-2025-47522 | 23303157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23303157&pc=OZD | 09/05/2025 |  |  | 10/15/2025 |  |  | 09/12/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704481 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 78 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Fever | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing from the central cannulation site. the patient received a blood transfusion. the patient then developed pneumonia with a fever and hemolysis. the patient was treated with antibiotics and continuous renal replacement therapy and the pump was repositioned. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4330 | 1220648-2025-47518 | 23303362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23303362&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688093 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. physician was concern of a noted bleed to posterior of heart. the bleed was very difficult to access. at the time of this update, it is believed to be from left appendage. once bleed was controlled to doctor¿s satisfaction, left ventricle access was attempted. multiple attempts were made and the catheter traversed descending aorta with each attempt. contrast was injected showing ¿uncommon vasculature¿ per physician. it was noted that patient began to bleed again. multiple units of blood products and been administered and continued to be administered. physicians agreed at this point that all efforts had been exhausted on placing 5.5 and aborted insertion. at this time, they both attempted to control bleeding but were unfortunately unsuccessful. patient expired in few hours. | [5 empty] | Correction b1, d4 the investigation of the reported failure mode has been completed. no product was returned. major bleed: the cause of bleed cannot be determined since insufficient clinical details were provided. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had uncommon vasculature preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4331 | 1220648-2025-47526 | 23303526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23303526&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722092 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support placed for the coronary artery bypass graft patient in need of support. during removal of the pump it was reported that there was slight resistance verbalized by physician but not enough to warrant concern. small bleed back & clamped to trim distal portion of graft. anesthesia reported significant drop in blood pressure moments after explant with suspicion of graft removal on explant. pulseless electrical activity arrest to follow and chest emergently reopened. 1l blood suctioned from chest into cell saver. bleed noted from right anterior lateral puncture site where graft was and hemostasis achieved rather quickly with heavy ties. brief episode of cardiac massage with return of spontaneous circulation and 5u packed red blood cells given. | [5 empty] | Additional information was given for the explant date therefore d6b was updated. the investigation was completed. the pump was not returned for investigation. after removal of the suture hub and opening the pocket, the graft was secured and the impella 5.5 was removed with only slight resistance. minor bleeding was controlled during graft trimming. immediately after explant, the patient experienced a significant drop in blood pressure and pea arrest, prompting emergency chest reopening. one liter of blood was evacuated, and bleeding from the graft site was quickly controlled with heavy ties. cardiac massage restored circulation, and 5 units of prbcs were given. the patient was stabilized, sent to cvicu, and remained in critical condition but was weaning off vasopressors. the physician believed the event was related to tension on the suture line during explant, due to the device¿s placement and exit site. access site adverse event (hemorrhage/bleeding): the cause of the access site bleeding issue was most likely related to unintentional use, based on the given clinical details. difficult/unable to remove: the cause of the removal issue was most likely related to unintentional use, based on the given clinical details. clinical symptoms (cardiac arrest): the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4332 | 1220648-2025-47524 | 23303933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23303933&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026740276 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella cp experienced purge fluid not detected. both the pump and the supporting automated impella controller (aic) were replaced to continue patient support. no patient harm was reported. | [5 empty] | After further information was received, this event was determined not to be a reportable event. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | After further review of the event, the complaint coding required updating to better describe the reported event to include hemodynamic instability. additionally, the type of reportable event was corrected to serious injury based on the complaint details. correspondingly, fields b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes and medical device problem codes) have been updated, corrected accordingly. note: h6: component/investigation (type/findings/conclusion coding remains unchanged. this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller. | B.7 information was added that was omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was omitted from the initial report that was submitted. h.10 added related report number. this is one of two related reports. this report represents the pump. a separate report was submitted to represent the automated impella controller. investigation summary: purge fluid not detected: the pump was not returned for investigation. data log was not provided or could not be retrieved from the remote server. the cause of the purge fluid not detected issue was not determined, as the product was not returned and sufficient clinical information was not provided. device history review: the pump passed all post sterile inspection checks.
4333 | 1220648-2025-47525 | 23304036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23304036&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711566 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 68 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced signal and suction alarms. the patient received a transplant and the pump was explanted. post transplant, the patient experienced a stroke. | [5 empty] | No product was returned for investigation. the cause of the optical signal issue was not determined due to no product returned, insufficient data logs returned, and insufficient clinical details. the cause of the pump suction was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. the root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4334 | 1220648-2025-47517 | 23304076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23304076&pc=OZD | 09/02/2025 |  |  | 10/15/2025 |  |  | 09/02/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515780 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on an impella cp the patient arrested for four minutes. return of spontaneous circulation was achieved and the patient was given two units of blood. unfortunately, after the procedure, the patient coded and is now intubated and on pressors. the impella cp was removed. | [5 empty] | Investigation summary: the investigation was completed. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
4335 | 1220648-2025-47523 | 23304104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23304104&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/16/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717695 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was placed onto impella support due to cardiogenic shock. while on support, the patient experienced ectopy that was resolved by repositioning. the patient remained on support until successful explant and survived. | [5 empty] | The investigation was completed and no product was returned. arrhythmia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4336 | 1220648-2025-47520 | 23304137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23304137&pc=OZD | 09/15/2025 | 10/27/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received. re-suturing was not needed but heparin was paused due to the bleeding. | The user facility reported that the patient on impella cp support developed bleeding from the impella access site. manual pressure was held, and dressing was changed multiple times. the patient was not on heparin drip. the bleeding started to slow down with angle matching and manipulation of the catheter. there was discussion of re-suturing the groin if it was needed. the patient continued impella support and bleeding was resolved. |  |  |  |  | Corrections to the initial mfr report: h6 impact code f01 has been added. | B.5 added information that was omitted from the first follow up report. e.1 added zip code extension as it was omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.2 additional information should of been the only selection on the first follow up report that was submitted. all information was additional and not a correction. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. no product was returned for investigation. major bleed: the patient was transferred with an impella cp in place. the nurse reported bleeding from the impella site and was having to change dressing multiple times. hemostasis was achieved by manual pressure. the cause of the bleeding was determined to be use issue due to patient movement. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4337 | 1220648-2025-47527 | 23306663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306663&pc=OZD | 09/15/2025 | 12/02/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657833 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 83 KG |  |  | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Arrhythmia | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a 5.5 pump placed for the support of the patient admitted into tertiary center after transfer from the community with extracorporeal membrane oxygenation (ecmo). the 5.5 was placed an ecmo explanted. the team had the pump supporting when there was a hematoma at the ecmo site, bleeding at the axillary site, a fever, and an arrhythmia of junctional escape rhythm. the pump supported additionally with sleeve damage. the pump was explanted after transplant and 28 days of support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported arrhythmia and product damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia and product damage could not be determined.
4338 | 1220648-2025-47528 | 23306664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306664&pc=PYX | 08/12/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  | 01/31/2026 |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025480905 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Pericardial Effusion | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp and impella rp devices for mechanical circulatory support for a percutaneous coronary intervention (pci). during the pci it was reported that the impella rp was not positioned optimally. the physician attempted to advance the rp, but the pigtail on the impella rp did not move forward. after the pci was completed, and prior to being released to the cardiac care unit (ccu), pericardial effusion was discovered with an echocardiogram. a drain was placed for the pericardial fluid. it was then noted that suction alarms occurred on both devices and position was reviewed and no impella position adjustments were needed. the source of the bleed could not be determined with echocardiogram or angiogram. the patient deteriorated and went into pulseless electrical activity. the patient expired upon arrival to the ccu. the physician who was the head of the medical team stated he did not believe the cause of the pericardial effusion to be related to the impella rp. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. cardiac arrest, pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: patient had suction alarms. no logs are available for review. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history lot: device lot: 1811679. device history review: device (sn: (b)(6)) passed all post sterile inspection checks. | Given the limited information surrounding the details the patient's condition and cause of death the impella rp cannot be excluded as a possible contributing factor. a.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4339 | 1220648-2025-47529 | 23306670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306670&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1815943 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | Yes | Initial submission | [9 empty] | The us reporter stated that, while in use during a procedure, the console cord was pulled away from the wall and the cord broke away from the plug. the automated impella controller (aic) was swapped out and there was reportedly no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be file.
4340 | 1220648-2025-47531 | 23306674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306674&pc=OZD | 09/15/2025 | 10/15/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026721970 |  | 000000000001000770 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella 5.5 on september 15th, 2025, that the bleeding; hemolysis; thrombus. thrombus noted on extracorporeal membrane oxygenation (ecmo) venous drainage, one unit of blood and one unit of platelets given over night. the medical team observed thrombus formation on venous drainage of ecmo circuit and are continuing to withhold anti-coagulation regardless of this observation. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint #: (b)(4). investigation summary: the pump was not returned for investigation. data log was not provided or could not be retrieved from the remote server. clinical symptoms (thrombosis/thrombus): in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (hemorrhage/bleeding): the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was not determined without returned product and sufficient clinical details. device history lot: device batch: 1958867. device history batch: subcomponent lot: na. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
4341 | 1220648-2025-47530 | 23306675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306675&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738708 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that approximately thirty minutes into the hrpci there were alarms for purge pressure high. the purge pressure was reported to be 1050 and the purge flow was lower than expected at 1.4ml/min. the medical team changed the purge cassette; however, this did not resolve the alarms. the medical team was advised by abiomed support to monitor motor current and silence purge pressure high alarms for the duration of the procedure, as the intent was to explant the pump at the end of the case. the patient was successfully weaned off support at the end of the case. no patient harm was reported as a result of the malfunction. | [5 empty] | D9 has been updated to reflect product was returned for investigation. h6 type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect that product was returned and investigation is underway. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary impella cp sn (b)(6) + cut pc + introducer returned for evaluation. high purge pressure: clinical details note high purge pressure (pp) alarms about 30 minutes into the case. changing the purge cassette did not resolve the alarm. data log review showed a 9 minute rise in pp with a motor current spike at the start of the rise. the purge pressure remained high for the rest of the case. the product was returned and evaluated. there was some dried dextrose in the purge gap and around the outlet cage before it was soaked in warm water. high purge pressure reproduced with the returned pump. the high purge pressure resolved when cutting the catheter near the motor housing. there was no sign of blood ingress in the purge tubing above the motor housing. the impeller was re-imaged after purging and there was biomaterial visible in the purge gap. the cause of the high purge pressure was biomaterial ingestion since data logs showed a mc spike followed by a rise in pp and hpp reproduced with the returned product. device history lot device lot: 1970179 device history review. the pump sn (b)(6) passed all post-sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4342 | 1220648-2025-47532 | 23306681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306681&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697931 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient during the support developed a small hematoma near the insertion site. during foley insertion as the patient moved the right lower extremity, hematoma noted. pressure was held and hematoma resolved. sheath angle matched and site dressed. | [5 empty] | Correction: d4 the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of hematoma is determined to be use issue due to patient movement because the hematoma was noted after patient movement. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4343 | 1220648-2025-47533 | 23306683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306683&pc=PYX | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026717913 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Fracture | [5 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post -op was implanted with an impella rp device for mechanical circulatory support. the patient went into atrial fibrialtion overnight, team unable to control rhythm despite multiple cardioversions and amio bolus and drip. overnight impella purge suction increased, motor current increased, and flows dropped to 0 over approximately 8 minutes time. pressor requirements increased during time of low flow. position unchanged. no acute signs or symptoms of hemolysis. on (b)(6) 2025 micron filter broke off purge side arm. side arm bypassed by blunt tip needle. patient successfully weaned and explanted this morning despite stable sidearm bypass in place. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4344 | 1220648-2025-47535 | 23306686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306686&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205603 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 141 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during impella cp support, there was ongoing placement signal low alarm. in addition, when they turned down from p2 to p1 and then got the "in aorta" alarms. they attempted to disable placement monitoring per ic but didn't 'accept". the impella cp (sn : (b)(6). is reported under separate report. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d9 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. when device received, investigation incomplete:
4345 | 1220648-2025-47534 | 23306687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306687&pc=OZD | 09/15/2025 | 12/01/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility had a cp-aic support being initiated for the patient admitted in with cgs and a positive stress test. the pump and console failed at the case start. the purge fluid was not detected and so the pump and aic were replaced as the team made choice to not troubleshoot. | [5 empty] | The investigation for the reported controller failure alarm has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the controller failure alarm could not be determined. | After further review of the event, the complaint coding required updating to better describe the reported event to include hemodynamic instability. additionally, the type of reportable event was corrected to serious injury based on the complaint details. correspondingly, fields b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes and medical device problem codes) have been updated, corrected accordingly. investigatioin summary: the product with logs were received and the investigation was completed. console logs were consistent with what was reported in the complaint. purge fluid not detected errors were observed in the imc logs. rtlog data showed that there was almost no purge pressure when the purge fluid not detected errors occurred during purge priming. device analysis: the console was tested, and the reported purge fluid not detected error was not able to be reproduced. the console operated as intended while running with a known good pump and purge cassette kit. conclusion: the root cause of the purge fluid not detected issue on ic6115 could not be determined due to the inability to reproduce during testing. the purge pressure was almost nonexistent when these errors occurred which could be indicative of a kink in the purge line or a possible leak but there was no conclusive evidence. product history review was completed and there were no issues related to the complaint discovered when reviewing the product history of console ic6115. this automated impella controller device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | After further information was received, this event was determined not to be a reportable event. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4346 | 1220648-2025-47536 | 23306692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306692&pc=OZD | 09/15/2025 | 11/28/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722892 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during support of impella 5.5 on (b)(6) 2025, that the patient was taken to cardiovascular operating room. 10mm graft was anastomosed to the right axillary artery, impella 5.5 attempted to be placed under fluoroscopy. aborted due to suspected calcium within axillary artery. impella cp was placed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted to provide medical safety/clinical review. no new information was identified that changes the previously submitted device investigation or its conclusions. medical safety/clinical review. medical safety/clinical reviewed the event. the patient was admitted for congestive heart failure exacerbation and supraventricular tachycardia, underwent cardioversion, and was initially transferred to telemetry. due to worsening cardiogenic shock, the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va-ecmo) on (b)(6) 2025 (4 days prior to impella support), subsequently weaned from vasopressor and inotropic support, had diuretics discontinued, and was supported with a femoral intra-aortic balloon pump (iabp). four days later, escalation to impella support was planned. on (b)(6) 2025, the patient was taken to the cardiovascular operating room for implantation of an impella 5.5 (pump 1). a 10-mm graft was anastomosed to the right axillary artery, and impella 5.5 advancement was attempted under fluoroscopic. the procedure was aborted due to suspected calcification of the axillary artery, and the impella 5.5 was not implanted. an impella cp (pump 2) was subsequently implanted and following this implant, the patient became hypotensive with low mixed venous oxygen saturation (svo2). dobutamine was initiated and escalated without improvement and was associated with arrhythmia, requiring multiple cardioversions. the right hand was noted to be dusky with low flow. the patient required multiple defibrillation shocks. plans were made to escalate support with a left axillary impella 5.5, and the impella cp was explanted on (b)(6) 2025. an impella 5.5 (pump 3) was implanted and maintained through (b)(6) 2025. on (b)(6) 2025, nursing staff reported increasing vasopressor requirements, administration of multiple blood products and fluids, and concern for possible sepsis. nephrology adjusted continuous renal replacement therapy (crrt). the following day on (b)(6) 2025, interrogation of the patient¿s implanted cardioverter-defibrillator revealed frequent episodes of atrial fibrillation with rapid ventricular response, requiring cardioversion with an external defibrillator. on (b)(6) 2025, the patient developed sustained ventricular tachycardia overnight, receiving one icd shock and one external defibrillation at 100 joules, resulting in conversion to a paced rhythm. the patient received an amiodarone bolus and remained on continuous infusion. labs revealed an elevated plasma-free hemoglobin of 189, increased from the prior day. central venous pressure was reported as 10 mmhg, and the left ventricular pressure on the automated impella controller was reported as ¿31 mmhg, raising concern for possible impella positioning. echocardiography was pending and possible impella pump repositioning (echo results and if the impella was repositioned are unknown; no further details were noted on this). impella 5.5 and va-ecmo support were continued. subsequently, impella support was withdrawn overnight. no additional details regarding the exact timing or circumstances of withdrawal or patient outcome were noted. based on the information at hand, the aborted implantation of impella 5.5 (pump 1) was attributed to patient anatomy, specifically suspected axillary artery calcification. this a malfunction type of reportable event. the adverse events observed during impella cp (pump 2) with no device malfunction is a serious injury. regarding impella 5.5 (pump 3) support, there were recurrent atrial and ventricular arrhythmias, rising plasma-free hemoglobin, and concern for possible device positioning were noted. there is not enough evidence to exclude the impella 5.5 pump 3 as an association factor to the patient's outcome. however, patient's underlying condition contributed most to the outcome (advanced heart failure progression). the event story involves three product complaints. mpella 5.5 - mdr 1220648-2025-47536: malfunction. impella cp pump 2 - mdr 1220648-2025-47828: serious injury. impella 5.5 pump 3 - mdr 1220648-2025-48483: death. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp pump 2 and impella 5.5 pump 3. | The investigation for the delivery issue has been completed.: the cause of the delivery issue was determined to be patient condition as the patient most likely had calcium within axillary artery preventing successful delivery of impella
4347 | 1220648-2025-47537 | 23306693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306693&pc=OZD | 09/15/2025 | 11/28/2025 |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported they were unable to get the impella to pass all the way through the sheath. attempted to place the long sheath without success. the team decided to abort the case and get vascular involved to proceed with the case at a different time | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported minor bleed and difficulty of insertion of introducer has been completed. no product was returned for review. the root cause of the difficulty of insertion of introducer was most likely patient condition related since pump was unable to pass through vasculature and met resistance at femoral artery. the root cause of the minor bleed cannot be determined since insufficient clinical details were provided
4348 | 1220648-2025-47538 | 23306696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306696&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732042 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient developed altered mental status while on impella cp support and was taken for computed tomography scan (ct). ct scan showed frontal brain bleed with midline shift. patient unable to be anticoagulated at this time due to brain bleed. the cp was successfully weaned and explanted. the patient survived. | [5 empty] | The investigation was completed and no device was returned. stroke : the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4349 | 1220648-2025-47539 | 23306697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306697&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704256 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 154 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while a patient was on impella cp support the patient became hypotensive and levo was started. the patient taken to the operating room with vascular surgery. explant of impella cp successfully completed via surgical cutdown to right groin. right femoral artery repaired and distal blood flow confirmed. the patient survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the hypotension and vascular dissection were unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4350 | 1220648-2025-47540 | 23306698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306698&pc=OZD | 09/15/2025 |  |  | 10/15/2025 |  |  | 09/15/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738865 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 154 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with altered mental status. he worsened over the day and was placed on multiple medications for blood pressure support. however the patient coded. electrocardiogram showed an inferior st-elevation myocardial infarction. his cath revealed clean coronaries and ejection fraction of 20%. decision too place impella cp was made. impella was prepared and inserted without issue. upon arrival to the intensive care unit he coded and was unable to be resuscitated. the patient expired on support. | [5 empty] | H6 added code b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been complete. peri-procedural adverse event (cardiac arrest): in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4351 | 1220648-2025-47541 | 23306865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306865&pc=OZD | 09/20/2025 | 02/24/2026 |  | 10/15/2025 |  |  | 09/20/2025 | 10/15/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725989 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Cardiac Tamponade | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information has been provided upon request. the pump was temporarily stopped by the computed tomography surgeon while the patient was on bypass. the team was unaware of the surgical mode at that time. the pump was plugged back into the console after the patient came off bypass and restarted. the pump is not available for return. | The complainant had placed a impella 5.5 to support a surgical patient. the prior 5.5 had been replaced due to loss of optical signal. during the support there were persistent suction and a defective pump alarm that was remedied by a pump re-start. additionally, there was an effusion during the support which prompted intervention/holding of the anticoagulation. the patient was later successfully weaned and survived. |  |  |  |  | B5 updated. the investigation is still ongoing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4352 | 1220648-2025-47542 | 23306869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23306869&pc=OZD | 09/16/2025 |  |  | 10/15/2025 |  |  | 09/16/2025 | 10/15/2025 |  |  |  |  | 10/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731858 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 146 KG | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted, and support was started. flows were noted to be suboptimal. confirmed with fluoro and a kink in the catheter was noted. impella will not flow above p4. the medical doctor made the decision to continue with the flow that was being achieved. there was no patient harm. | [5 empty] | The investigation was completed and no product was returned. product damage (catheter kink): the cause of product damage catheter kink cannot be determined since insufficient clinical details were provided. low pump flow: no product was returned. patient had low pump flow due to catheter kink. no logs are available for this product. the cause of low pump flow cannot be determined since insufficient clinical details were provided. device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4353 | 1220648-2025-47428 | 23275728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275728&pc=OZD | 09/11/2025 | 10/10/2025 |  | 10/12/2025 |  |  | 09/11/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685614 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Fever | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp had pink-tinged urine and hematuria was noticed. additionally, the patient spiked a temperature, and antibiotics started. the patient required more support and was escalated to impella 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the hematuria/ fever: the cause of the injury was not determined due to insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. b1 product problem was incorrectly selected in initial report e4 added selection as it was omitted from the initial report that was submitted.
4354 | 1220648-2025-47429 | 23275766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275766&pc=OZD | 09/07/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711549 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan had a cp pump placed for the support of a patient admitted in post ventricle fibrillation arrest in the community. the pump was placed along with extracorporeal membrane oxygenation (ecmo) for support. during the support the optical signal was not reliable but the pump was not replaced due to patient's condition was already severe and the device was still functioning, after 4.5 days the patient expired due to underlying multiple organ failure. no allegation made against the pump contributing to the death. | [5 empty] | Correction: b1. the investigation of the reported failure mode has been completed. no product was received for evaluation. optical sensor issue: the cause of the optical sensor issue was not determined without returned product information and sufficient clinical details, including data logs. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
4355 | 1220648-2025-47431 | 23275771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275771&pc=OZD | 09/12/2025 | 12/01/2025 |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713131 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a 5.5 being placed to support the patient admitted in with known coronary artery disease and in need of a high-risk percutaneous coronary intervention (hrpci). the pump was to placed to support for the hrpci and abdominal aortic aneurysm repair. the pump failed to be delivered due to native anatomy and so the team instead placed a cp pump. there was no patient injury. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. the patient is a 53-year-old female with a medical history significant for coronary artery disease status post coronary artery bypass grafting (cabg), hypertension, systemic lupus erythematosus, and kidney transplant (1998). the patient presented with chest pain and was diagnosed with a non-st elevation myocardial infarction (nstemi). cardiac catheterization demonstrated a patent left internal mammary artery (lima) to the left anterior descending artery (lad), with significant disease in the left circumflex and right coronary arteries. a subsequent abdominal ct scan identified a 5.3 cm infrarenal abdominal aortic aneurysm (aaa). cardiothoracic surgery was consulted for placement of an axillary impella 5.5 device to provide hemodynamic support during high-risk percutaneous coronary intervention (pci), with planned bilateral femoral artery cutdowns to facilitate simultaneous endovascular repair of the abdominal aortic aneurysm (evar). attempted placement of the impella 5.5 was unsuccessful due to small and tortuous vessel anatomy. the device was removed and replaced with an impella cp, which was successfully placed and initiated on auto mode with normal flows. pci was performed on the remaining conduit to the lad, and the abdominal aortic aneurysm was stented via femoral access. a sponge plug was placed in the graft, partially externalized at the axillary incision site, and secured with three-point fixation. the patient remained on impella support post-procedure. on post-implant day 1, generalized bleeding was noted to have improved; however, renal function worsened. on (b)(6) 2025, the patient experienced a cardiac arrest (rhythm unspecified) and received two units of packed red blood cells. the following day, the patient required dialysis and vasopressor support with norepinephrine (levophed) and epinephrine for hypotension. impella support was continued for hemodynamic recovery. despite ongoing support measures, the patient expired while on impella support. no additional information was provided. the patient had pre-existing comorbidities, including advanced coronary artery disease, prior cabg, systemic lupus erythematosus, renal dysfunction requiring dialysis, and a large abdominal aortic aneurysm requiring intervention. the initial inability to advance the impella 5.5 due to small and tortuous vessel anatomy is consistent with vascular access limitations described in the device instructions for use (ifu). the patient¿s death appears multifactorial in the setting of complex cardiovascular disease, high-risk pci, aaa repair, worsening renal failure, bleeding, and hemodynamic instability. however, the impella cp, which was in use at the time of expiration, will be reported for the death. the event involving the impella 5.5 is a malfunction type of reportable event. h6: investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation for the reported delivery issue has been completed. no product was returned for review. the root cause of the delivery is determined to be patient condition because the pump was unable to pass through vasculature due small and tortuous vessel.
4356 | 1220648-2025-47432 | 23275774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275774&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 2000154 |  |  | ABIOMED, INC. |  |  | US |  | 2026703531 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp pump placed to support the patient admitted in with acute myocardial infarction/cardiogenic shock. while receiving support, the patient experienced a groin hemorrhage. this bleeding was controlled for 45 minutes before it ceased. subsequently, the patient coughed and bore down, causing the bleeding to restart. the medical team once more applied manual pressure, utilized a femostop, administered 4 units of blood, and disconnected the pump. the vascular team was consulted and performed a cut down to seal the site of the bleed. the patient successfully survived the procedure. | Additional information was received stating that the device was in the incorrect position. the pump migrated into the aorta, leading to malposition and subsequent device removal. laboratory findings were consistent with hemolysis. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information was received and h6 codes updated.
4357 | 1220648-2025-47433 | 23275776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275776&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026694913 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. immediately upon implant, impella aortic waveform noted to be 20-30 pts lower than the a-line. of note, the impella tray became contaminated when the circulator was handing off the pump to the scrub. for this reason, the scrub had to manually remove the impella from the tray to maintain sterility. the impella was primed, but not while laying flat in the tray. upon implant, the impella flows were appropriate for the p-level. no abnormal or inappropriate alarms noted. additionally significant oozing and drain output from incision was noted. stat echo showed impella measuring 3.3 cm and left ventricle (lv ) thrombus in apex. noted not seen prior to impella implant. hemodynamic instability and amount of drain output determined need to return to operation room (or) . patient taken back to or with on-site support provided. bleeding controlled at this time. also there was suction issue. the pump remained on til pump explanted for left ventricular assist device ( lvad) placement on day 5. patient successfully bridged to the lvad and is stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. asae: there was significant oozing from an incision and drain output. it was unknown how the bleeding became controlled. the cause of the injury was not determined due to insufficient clinical details. optical sensor issue: only real time (rt) logs were returned. due to a lack of product returned and insufficient clinical information, the root cause cannot be determined. pump suction: based on the logs and the clinical information, the root cause is due to the patients condition as they were given boluses and epi to increase volume and contractility which seemed to improve suction. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4358 | 1220648-2025-47430 | 23275781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275781&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 2000154 |  |  | ABIOMED, INC. |  |  | US |  | 2026700947 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump provided support when it became dislodged and was subsequently redelivered. following the redelivery, it was observed that the pump was experiencing suction issues, leading to its removal and replacement. the second cp support was then implanted. there was no patient injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information: d9 the investigation of the reported failure mode has been completed. the pump was returned for investigation. pump suction: data log review confirmed suction events occurring. after suction events occurred, the motor current (mc) spiked and motor speed deviating down, which is indicative of biomaterial ingestion. this ingestion happened right before the drop in flow, when suction alarms triggered as well as a "position in aorta" alarm. the product was returned, and biomaterial was observed at the impeller inside the cannula. however, suction was occurring before the mc spike, and placement signal dropped significantly at this time as well, which means that positioning issues were occurring at the time of suction, and it's possible that biomaterial was picked up during repositioning, seen by the mc spike, as well as the saturated flow near the end of the case. the cause of the suction was most likely improper positioning, as clinical reported the pump was inadvertently pulled back into the aorta, which was verified by the placement signal drop at the time of suction in the data logs. positioning issues: the cause of the positioning issue was most likely use related, as clinical details state the pump came out of position because the patient was moving uncontrollably. device history review: the device passed all the post sterile inspection checks.
4359 | 1220648-2025-47435 | 23275801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275801&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1731823 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. automated impella controller (aic) made weird noises coming from the purge at the beginning of the case. then later, the aic experienced a controller, failure alarm along with the waveforms dropping. the waveforms eventually came back after about 2-3 minutes but the pump still ran normally without failure. no patient harm was reported during the issue. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4360 | 1220648-2025-47436 | 23275802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275802&pc=OZD | 09/11/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692884 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support the patient experienced ventricular tachycardia (vt) when sits up, in addition had some runs of vt which it is noted that it was related to low magnesium and was replaced and runs of vt resolved. the patient also stated some numbness the left leg. the patient is still on support awaiting transplant. | Additional information received. the numbness was never in relation to impella as patient never had a cp, only iabp to 5.5. numbness was not reported in upper extremity or mentioned lower extremity numbness to clinical rounding the following day. the oozing- was serous drainage, not bleeding. no further issue after dressing changed done with clinical present with two nurses in a sterile fashion per hospital policy. ventricular tachycardia was noted to resolve with electrolyte correction. in addition to the patient was successfully weaned and explanted as a bridge to transplant. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b1. the investigation of the reported failure mode has been completed. no product was received for evaluation. vt/ischemia: the cause of the injuries was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | Additional information received the following fields were updated: b5, d6b, h6. should the investigation results provide additional information a supplemental manufacturer device report will be filed.
4361 | 1220648-2025-47437 | 23275803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275803&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026711431 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during the support the patient experienced hematoma for which heparin was put on hold. patient is still on device support awaiting heart and kidney transplant. | Additional information received the patient developed aphasia and slurred speech yesterday. a computed tomography scan demonstrated a small cerebrovascular accident, or stroke. the patient was therapeutic on anticoagulation at the time of the event. per staff, there were no issues or alarms with the pump at the time of the neurological status changes. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. stroke: ct scan scan showed small cva. the root cause of the injury was unable to be determined due to insufficient clinical details. additional information was received and h6 codes updated. | The investigation of the reported failure mode has been completed. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | D6b explant date provided.
4362 | 1220648-2025-47438 | 23275804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275804&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697160 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 84 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | [9 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support patient experienced bleeding from tunnel site and hematoma in surgical pocket. additional suture placed and pressure bag applied on top of pocket. nursing staff addressed low central venous pressure (cvp) with provider and patient is receiving fluids. patient to receive 1 unit of packed red blood cells. (prbc). nursing weaning pressor supports as patient tolerates. impella noted to measure 4.3 cm on echo imaging. provider at bedside and repositioned. patient was started on precedex overnight for agitation but has now spiked a fever and precedex was discontinued. staff obtaining new cultures for septic work up. white blood count (wbc) increased significantly to 39. the patient is receiving fluid bolus and another unit of prbc. the patient has been decompensating throughout the night. patient¿s family at bedside and chose comfort measures. patient expired. | [5 empty] | Correction: b1, d4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma, bleed: the cause of the bleed and hematoma was most likely use issue related since hemostasis was achieved by placing additional suture. sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4363 | 1220648-2025-47439 | 23275805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275805&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710808 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support patient had a multiple runs of no sustained ventricular tachycardia (vt) and, magnesium was administered. patient had another episode of vt following night for which patient was shocked and amio, bolus administered, echo done and impella position confirmed. cat scan revealed small infarction, patient more alert today and able to follow commands. impella successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation was completed and no product was returned. tachycardia, thrombosis): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and passed all post sterile inspection checks.
4364 | 1220648-2025-47440 | 23275808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275808&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690791 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The eu complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient in critical state, sepsis ongoing therefore high dosages of vasopressors were administered. pump ran without any problem but patient was in critical condition and did not improve so care was withdrawn and patient passed away. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. sepsis: the root cause of the sepsis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4365 | 1220648-2025-47441 | 23275809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275809&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618685 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support the patient was bleeding around the repo sheath. the patient received 2 units of red blood cells (rbc). the patient taken to operation room for 5.5 escalation and cp removal. patient is stable post explant. | [5 empty] | The investigation of the reported failure mode has been completed. major bleed : the cause of the injury is most likely use related since bleeding around the repositioning sheath post transfer. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4366 | 1220648-2025-47426 | 23275815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275815&pc=OZD | 09/10/2025 | 10/10/2025 |  | 10/12/2025 |  |  | 09/10/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1108441 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prepping, the impella console had a controller error alarm immediately after boot-up. there are no patient involved in this event. | [5 empty] | The investigation of the reported failure mode has been completed. data review: console logs were consistent with what was reported in the complaint. device analysis: console was tested after and the controller error was unable to be reproduced during testing. no abnormalities were observed while inspecting the connection between the 4-in-1 and the glass keyboard. conclusion: the root cause of the controller error could not be determined due to the inability to be reproduced but was likely due to a communication breakdown during initialization of kbd to 4-in-1 communication. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
4367 | 1220648-2025-47442 | 23275817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275817&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133145 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us based biomed team observed a malfunction of the automated impella controller (aic) outside of patient use. there was "failed electrical test ground resistant" and purge cassette was not detected there was no harm reported as no patient is involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Manufacturing site name was incorrectly reported.
4368 | 1220648-2025-47443 | 23275818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275818&pc=OZD | 09/11/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738752 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. patient had bleeding from impella site requiring transfusions. impella was repositioned on arrival and angle matching was optimized. had some resolution in bleeding but overnight developed significant bleeding from impella site per rn requiring fem stop. had significant hemolysis with elevated plasma-free hemoglobin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of the bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all the post sterile inspection checks.
4369 | 1220648-2025-47444 | 23275819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275819&pc=OZD | 09/12/2025 |  |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705270 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support the patient was bleeding at impella access site, manual pressure held, and dressing changed. following this event, bleeding was uncontrolled. the patient brought back to cardiac cath lab where they replaced impella repo unit for 16fr sheath and consulted vascular surgery. vascular did surgical repair of iliac due to dissection. | [5 empty] | D7a revised as it was reported incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. vascular dissection: vascular did surgical repair of iliac due to dissection. the cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
4370 | 1220648-2025-47434 | 23275827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23275827&pc=OZD | 09/12/2025 | 10/12/2025 |  | 10/12/2025 |  |  | 09/12/2025 | 10/12/2025 |  |  |  |  | 10/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119121 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The facility reported that during training/in-service, ongoing issues were noted with the console failing at the start up and a controller error was observed on the automated impella controller (aic). there was no patient involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4371 | 1220648-2025-47384 | 23263162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263162&pc=OZD | 09/10/2025 | 11/26/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703704 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 74 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed to support the percutaneous coronary intervention (pci). the patient developed ventricular fibrillation, defibrillated x 1 and converted back to normal sinus rhythm. the physician proceeded to complete pci. during the procedure, the patient again experienced hypotension & bradycardia. additional adenosine given. hemodynamics returned to baseline. the final images revealed timi 3 flow in all vessels, so a decision was made to wean and explant impella as normal blood flow was seen throughout the coronary arteries. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ bradycardia and hypotension has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ventricular fibrillation/ bradycardia and hypotension could not be determined.
4372 | 1220648-2025-47385 | 23263163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263163&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702833 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 78 KG | Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing from the access site. in response, heparin was withheld. the oozing was resolved with manual compression. the patient also experienced ventricular tachycardia and ventricular fibrillation, and the device was repositioned. vasoactive medications were given in response to hypotension. the patient also had a fever that self-resolved. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the bleed cannot be determined since insufficient clinical details were provided. the root cause of the ventricular fibrillation and tachycardia was unable to be determined due to insufficient clinical details. the cause of the fever was not determined due to insufficient clinical details. the cause of the positioning issue was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4373 | 1220648-2025-47386 | 23263165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263165&pc=OZD | 09/10/2025 | 12/05/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738276 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 92 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp pump placed to support the patient admitted in with cardiogenic shock, cardiomyopathy, and other cardiac comorbidities. the pump was placed and peel away sheath removed. upon this, there was an access site bleed that prompted adjustment of the perclose, suture, and manual hold. the team achieved hemostasis. the pump was later showing suction and albumin was infused. the pump remained on until day 3 when care was withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported major bleed and pump suction has been completed. no product was returned for review. patient had multiple suction alarms and low flow. after reviewing the logs, multiple suction alarms were observed. no motor current spike or trend in placement signal was present. the root cause of pump suction and major bleed cannot be determined since insufficient clinical details were provided.
4374 | 1220648-2025-47387 | 23263168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263168&pc=OZD | 09/09/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398838 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient who decompensated during a percutaneous coronary intervention was decannulated with extracorporeal membrane oxygenation and implanted with an impella cp pump or mechanical circulatory support. it was reported an hour after support was initiated, the automated impella controller (aic) had an alarm for controller failure. the aic was exchanged, and the issue was resolved with the new aic unit. there was no patient harm reported. | [5 empty] | Correction: d6a and d6b should have been left blank. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
4375 | 1220648-2025-47388 | 23263177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263177&pc=OZD | 09/08/2025 |  |  | 10/10/2025 |  |  | 09/08/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026718423 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 121 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced suboptimal flows. it was suspected that the pump ingested a clot from a preexisting va ecmo graft. the patient suffered recurring sustained ventricular tachycardia and required defibrillation. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.1. added selection as it was inadvertently omitted from the initial report that was submitted. e.1. added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11. added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. investigation summary: no product was returned. thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: no product was returned. the pump ran for less than an hour. the patient had low flows. after reviewing the logs, multiple suction alarms were observed along with motor current spike. small motor current spikes/mimicking with stable revolutions per minute and drop in flows started there. this most likely occurred after impella being in aorta for 1 to 2 mins prior to motor current spike as something got pulled from the aorta likely. the cause of low pump flow most likely biomaterial ingestion per the logs. device history review: the device passed all post sterile inspection checks.
4376 | 1220648-2025-47357 | 23263184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263184&pc=OZD | 09/08/2025 | 11/06/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667110 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the support on day 4 of the support there was rising purge pressure that prompted the team to troubleshoot. the troubleshooting was inclusive of checking the tubing, changing purge bag, and changing the purge cassette. when these measures failed the pump was explanted and replaced by a 5.5 pump. the explant date and patient outcome is unknow. | [5 empty] | D6b updated with additional information the device was returned and d9 was updated as well as h6 type of investigation and investigation findings. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. purge pressure high: data logs were not returned. the product was returned and evaluated. biomaterial was found in the purge gap. high purge pressure reproduced and resolved when cutting the catheter just above the motor housing. there was no blood ingress observed in the purge line above the motor housing. no other abnormalities observed. one potential cause of the high purge pressure (hpp) was biomaterial based on returned product analysis but there is no conclusive evidence of whether a kink or other events could have contributed to the hpp since data logs are not available. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4377 | 1220648-2025-47389 | 23264052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264052&pc=OZD | 09/10/2025 | 10/10/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1393502 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 84 KG | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp - automated impella controller was being started up for the support of a patient admitted in with acute myocardial infarction with cardiogenic shock. the console had a startup failure and was replaced. no harm or injury was noted from any delay in support/console replacement. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | D6a and d6b should have been left blank.
4378 | 1220648-2025-47390 | 23264196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264196&pc=OZD | 09/12/2025 | 11/06/2025 |  | 10/10/2025 |  |  | 09/12/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299522 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed hemolysis and encountered pump detection issue with the automated impella controllers (aic). the patient would eventually expire. the patient required c-pap for over 24hrs but continued to decompensate overnight, was intubated and was sent to catheterization lab for an emergent catheterization. the impella cp was implanted for protected percutaneous coronary intervention, which was completed. however, during the pci, the physician was unable to cross the proximal left anterior descending occlusion. the pci was paused for shock call. the decision was made to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo). penumbra was used and the occlusion was ballooned, and one drug-eluting stent was implanted. after the pci, the peel-away sheath was removed, touhy-borst locked at 91cm and grin was dressed, dry and intact. the patient was sent to the unit on ecpella (ecmo and impella) support. the following day it was noted that the impella was repositioned with point-of-care ultrasound for an alarm in left ventricle. hemolysis complication was noted and consequently, the impella catheter was pulled back 1cm. of note, placement signal alarm was disabled for frequent placement signal low alarms occurring. however, the impella position still appeared to be malpositioned, caught in the mitral apparatus and sitting deep at 6 cm. patient care note unrelated to the impella, noted the patient continued to require blood products related to the at home fall injury. heparin was placed on hold and possible surgery was noted for the left flank hematoma/bleed. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. ecpella support continued and the next day ecmo was weaned to 3800rpm/3.7l/min. p level increased to 4, flow 2.5l. drips were weaned downed. surgical consultation did not require intervention for hematoma/bleed. the patient received one unit of red blood cells over the night. of further note, the placement signal alarm was enabled again. ecpella support continued. two days later echocardiogram showed ejection fraction improvement 5% to 10-20%. the impella was still angled towards the mitral valve but less than initial implant. inlet at 4.7 cm from aortic valve. plasma free hemoglobin at 10mg/dl and alanine aminotransferase down trending was noted. low dose heparin drip was started. intros adjustment was noted. also noted was that the patient experienced another vt storm overnight with torsades and received one shock from implantable cardioverter defibrillator. support continued with target for heart recovery versus transfer for potential left ventricular assist device (lvad). ecpella support continued and three days later [11-sep-2025] ventricular tachycardia and supraventricular tachycardia was noted through the day and over the night. now transplant versus lvad work up started. the following day, ganglion block was placed to help control ectopy. this day the aic had to be switched out and when the nurse attempted to switch out the aic from the outside hospital console, pump not detected was encountered. when white cable was plugged in to aic ic 9571, instead of restarting the pump on current setting on the new aic it went to ¿spike dextrose bag¿ screen. because it did this, the care team switched it back to the original aic ic 5269, and when the white cable was plugged back into aic ic 5269, it did the same thing. therefore, to avoid pump being off for any longer, the nurse continued through the spike/prime purge menu and impella restarted on auto as if it were a new impella cp case. the device was restored to previous p-level and patient stable. the following day palliative care was discussed; the patient was not a candidate for transplant or lvad. the patient would subsequently expire. | [5 empty] | The product received date was erroneously captured as 06-oct-2025 instead of 22-sep-2025. accordingly, correction has been made to section d9. date device returned to manufacturer. related medical device reports: this automated impella controller is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other aic and the impella cp, which is associated with the demise outcome. | Medical safety reviewed the event. the patient presented in a decompensating state, was shocked 10 times for ventricular tachycardia; the patient presented with non-st-segment elevation myocardial infarction (nstemi). the patient would eventually expire on support. the palliative care was discussed; the patient was not a candidate for transplant or left assist ventricular device (lvad). the patient had additional support in the form of veno-arterial extracorporeal membrane oxygenation (va ecmo), experienced complications related to an at-home fall injury and was in a poor state. the impella was there to support the patient. the patient expired and the impella may have contributed to the patient's demise given the hemolysis and positioning issues encountered (thus may have added to an already complex situation or compounded the situation to where the patient actually expired). regarding automated impella controllers (aics) the events describes failure of pump to be detected by the aic resulting in an alarm. the original aic was reconnected with the same outcome. however, to avoid the patient being off pump longer, the nurses continued through the spike/prime purge mean and restarted the pump on auto as if it was a new case. although interruption in support, it appears the patient expired for reasons unrelated to the aics (as noted above). the aics malfunctioned and contributed to the patient's implant experience event and should be reported as a malfunction type of reportable event. related medical device reports: this automated impella controller is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other aic and the impella cp, which is associated with the demise outcome. | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed. related medical device reports: this automated impella controller is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other aic and the impella cp, which is associated with the demise outcome. this event describes transfer failure of the automated impella controllers. there was no report or indication of significant hemodynamic decompensation. the patient was noted to have remained stable. the impella cp pump, its positioning was suboptimal, and it was noted patient developed hemolysis. given the details of the case dissociation cannot be made although the patient¿s underlying condition appears to have contributed more to the demise outcome (cardiogenic shock, hematoma/bleeding from fall injury and ejection fraction of 5%-10%).
4379 | 1220648-2025-47391 | 23264209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264209&pc=OZD | 09/12/2025 | 11/06/2025 |  | 10/10/2025 |  |  | 09/12/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1646194 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed hemolysis and encountered pump detection issue with the automated impella controllers (aic). the patient would eventually expire. the patient required c-pap for over 24hrs but continued to decompensate overnight, was intubated and was sent to catheterization lab for an emergent catheterization. the impella cp was implanted for protected percutaneous coronary intervention, which was completed. however, during the pci, the physician was unable to cross the proximal left anterior descending occlusion. the pci was paused for shock call. the decision was made to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo). penumbra was used and the occlusion was ballooned, and one drug-eluting stent was implanted. after the pci, the peel-away sheath was removed, touhy-borst locked at 91cm and grin was dressed, dry and intact. the patient was sent to the unit on ecpella (ecmo and impella) support. the following day it was noted that the impella was repositioned with point-of-care ultrasound for an alarm in left ventricle. hemolysis complication was noted and consequently, the impella catheter was pulled back 1cm. of note, placement signal alarm was disabled for frequent placement signal low alarms occurring. however, the impella position still appeared to be malpositioned, caught in the mitral apparatus and sitting deep at 6 cm. patient care note unrelated to the impella, noted the patient continued to require blood products related to the at home fall injury. heparin was placed on hold and possible surgery was noted for the left flank hematoma/bleed. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. ecpella support continued and the next day ecmo was weaned to 3800rpm/3.7l/min. p level increased to 4, flow 2.5l. drips were weaned downed. surgical consultation did not require intervention for hematoma/bleed. the patient received one unit of red blood cells over the night. of further note, the placement signal alarm was enabled again. ecpella support continued. two days later echocardiogram showed ejection fraction improvement 5% to 10-20%. the impella was still angled towards the mitral valve but less than initial implant. inlet at 4.7 cm from aortic valve. plasma free hemoglobin at 10mg/dl and alanine aminotransferase down trending was noted. low dose heparin drip was started. intros adjustment was noted. also noted was that the patient experienced another vt storm overnight with torsades and received one shock from implantable cardioverter defibrillator. support continued with target for heart recovery versus transfer for potential left ventricular assist device (lvad). ecpella support continued and three days later (b)(6) 2025 ventricular tachycardia and supraventricular tachycardia was noted through the day and over the night. now transplant versus lvad work up started. the following day, ganglion block was placed to help control ectopy. this day the aic had to be switched out and when the nurse attempted to switch out the aic from the outside hospital console, pump not detected was encountered. when white cable was plugged in to aic ic 9571, instead of restarting the pump on current setting on the new aic it went to ¿spike dextrose bag¿ screen. because it did this, the care team switched it back to the original aic ic 5269, and when the white cable was plugged back into aic ic 5269, it did the same thing. therefore, to avoid pump being off for any longer, the nurse continued through the spike/prime purge menu and impella restarted on auto as if it were a new impella cp case. the device was restored to previous p-level and patient stable. the following day palliative care was discussed; the patient was not a candidate for transplant or lvad. the patient would subsequently expire. | [5 empty] | The product received date was erroneously captured as 06-oct-2025 instead of 22-sep-2025. accordingly, correction has been made to section d9. date device returned to manufacturer. related medical device reports: this automated impella controller is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other aic and the impella cp, which is associated with the demise outcome. | Medical safety reviewed the event. the patient presented in a decompensating state, was shocked 10 times for ventricular tachycardia; the patient presented with non-st-segment elevation myocardial infarction (nstemi). the patient would eventually expire on support. the palliative care was discussed; the patient was not a candidate for transplant or left assist ventricular device (lvad). the patient had additional support in the form of veno-arterial extracorporeal membrane oxygenation (va ecmo), experienced complications related to an at-home fall injury and was in a poor state. the impella was there to support the patient. the patient expired and the impella may have contributed to the patient's demise given the hemolysis and positioning issues encountered (thus may have added to an already complex situation or compounded the situation to where the patient actually expired). regarding automated impella controllers (aics) the events describes failure of pump to be detected by the aic resulting in an alarm. the original aic was reconnected with the same outcome. however, to avoid the patient being off pump longer, the nurses continued through the spike/prime purge mean and restarted the pump on auto as if it was a new case. although interruption in support, it appears the patient expired for reasons unrelated to the aics (as noted above). the aics malfunctioned and contributed to the patient's implant experience event and should be reported as a malfunction type of reportable event. related medical device reports: this automated impella controller is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other aic and the impella cp, which is associated with the demise outcome. | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed. related medical device reports: this automated impella controller is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the other aic and the impella cp, which is associated with the demise outcome. this event describes transfer failure of the automated impella controllers. there was no report or indication of significant hemodynamic decompensation. the patient was noted to have remained stable. the impella cp pump, its positioning was suboptimal, and it was noted patient developed hemolysis. given the details of the case dissociation cannot be made although the patient¿s underlying condition appears to have contributed more to the demise outcome (cardiogenic shock, hematoma/bleeding from fall injury and ejection fraction of 5%-10%).
4380 | 1220648-2025-47392 | 23264232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264232&pc=OZD | 09/11/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1881872 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Network Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The field team reported that the automated impella controller (aic) kept going offline for long periods of time during its last case. once off patient support, the staff tried to get the remote link module (rlm) log files to investigate the connection issues. when powered on, the aic would connect to impella connect. however, it would quickly lose the connection and never had enough of a connection to get wi-fi signal strength information. when putting the aic into maintenance mode, the rlm would continuously reboot. | [5 empty] | The automated impella controller was returned. upon completion of the investigation, a supplemental mdr will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary data analysis: console logs from the reported day of event are consistent with complaint and show intermittent data streaming, multiple ¿icm not ready for file receive¿ imc log messages and a few aic reboots, presumably for troubleshooting purposes. partial rlm journals recovered from the microsd card aligned with events in aic log, and indicated several reboots, occasional breakdown of usb communication from aic to rlm, and evidence of microsd card corruption. device analysis: device was tested in danvers and the issue was reproduced: microsd card corruption in real-time rlm journal, loss of usb communication between aic and rlm, and unexpected rlm reboot(s). after replacing microsd card with a known-good one, all issues were resolved. a minor defect of the backstrap cable had been discovered during inspection of the device, but there¿s not enough evidence that would indicate its relation to the card corruption.
4381 | 1220648-2025-47393 | 23264259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264259&pc=OZD | 09/11/2025 | 01/30/2026 |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1726767 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The reporter stated that the automated impella controller (aic) presented a controller error alarm, which was resolved by transferring to another aic with no reported impact on the patient. | [5 empty] | D.9 the device was returned and the date received was added. h.6 updated investigation codes due to the device returning. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4382 | 1220648-2025-47395 | 23264308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264308&pc=OZD | 08/20/2025 |  |  | 10/10/2025 |  |  | 08/20/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653719 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that implantation of an impella 5.5 was unsuccessful due to the patient's vasculature and kinking of the impella guidewire. the patient survived. this report represents the pump. another report was submitted to represent the guidewire. refer to section h.10 for the related report number. | [5 empty] | The pump was returned for analysis. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device passed all post sterile inspection checks. | A.4 is unknown. no product was returned. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of the impella. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4383 | 1220648-2025-47394 | 23264398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23264398&pc=DQX | 08/20/2025 |  |  | 10/10/2025 |  |  | 08/20/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018"PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 9262008 |  | 0052-3005 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that implantation of an impella 5.5 was unsuccessful due to the patient's vasculature and kinking of the impella guidewire. the patient survived. this report represents the guidewire. another report was submitted to represent the pump. refer to section h.10 for the related report number. | [5 empty] | A.4 is unknown. the guidewire was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary no product was returned. product damage (guide wire): guide wire got kinked during insertion and was aborted. the cause of product damage guide wire was most likely patient condition related since the guide wire got kinked during insertion due to the patient's anatomy. device history lot device lot: 9262008 device history batch subcomponent lot: n/a device history review 4 units for sterilization, 3 for particulate, 2 spec, and 1 for deformation from guidewire batch #9262008 were dispositioned to mrb. all else passed incoming inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The guidewire was returned for analysis. the cause of product damage guide wire was most likely patient condition related since the guide wire got kinked during insertion due to the patient's anatomy.
4384 | 1220648-2025-47396 | 23265258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23265258&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650503 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Myocardial Infarction | Cardiac Tamponade | Cardiac Perforation | Pericardial Effusion | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had an effusion requiring a transversus abdominis plane. there was active bleeding from the pericardial tube determined to be from a left ventricle wall rupture. the ventricle was surgically repaired. the patient was escalated to an impella 5.5 and the patient survived. | [5 empty] | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the cause of the pericardial effusion/cardiac perforation/cardiac tamponade/hypotension/myocardial infarction was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4385 | 1220648-2025-47398 | 23265352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23265352&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661811 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 102 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed bleeding from an unknown source. the patient was transfused 20 units of packed red blood cells, nine units of platelets, 12 units of fresh frozen plasma, and four units of cryoprecipitate. later, care was withdrawn and the patient expired. | [5 empty] | Investigation summary: no product was returned for investigation. major bleed: the cause of the bleeding is determined to be patient condition because the location of the bleed is unknown. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4386 | 1220648-2025-47397 | 23265590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23265590&pc=OZD | 09/07/2025 |  |  | 10/10/2025 |  |  | 09/12/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694751 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 67 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ischemia | Pulmonary Edema | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 lost pulses in the left lower leg. an embolism was suspected so the patient was taken to surgery and systemic anticoagulation was started. after the thrombectomy the patient became short of breath, hypotensive and had high pressures the patient was treated with levo. the patient also developed bleeding at the impella site, so an additional suture was placed, and flash pulmonary edema, so lasix was given. later, the patient was successfully weaned from the pump and survived. | [5 empty] | B.3 revised date of event as it was entered incorrectly on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. g.3 date of awareness was reported incorrectly on the initial report that was submitted. date should of reflected 07-sep-2025. h.6 added health effect - clinical code as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ investigation summary: major bleed: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. ischemia/ thrombosis : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension/ pulmonary edema : the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4387 | 1220648-2025-47400 | 23265668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23265668&pc=OZD | 09/06/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1864743 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Power Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller was found to be powering on and powering down on its own. the event occurred outside of patient use. | [5 empty] | The event occurred outside of patient use. therefore, section a was left blank. the automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4388 | 1220648-2025-47401 | 23265840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23265840&pc=OZD | 08/30/2025 |  |  | 10/10/2025 |  |  | 08/30/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691879 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 63 KG | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Necrosis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced thrombosis, necrosis and difficulty removing the device. after implantation of the pump, the patient developed a clot in the chest tube, and a computed tomography showed a small bowel obstruction. the patient was treated with administration of fluids and a small bowel resection. the impella 5.5 could not be removed from the access site so the pump had to be removed from the chest via surgery. a clot was observed in the explanted pump. | [5 empty] | Investigation summary: thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. necrosis: clinical details reported the patient had a dead bowel that required resection. the cause of the injury was not determined due to insufficient clinical details. removal issue: clinical details reported that the attempt to remove the pump through the graft was met with resistance, and the chest had to be reopened in order to remove the pump. upon explant, the graft reportedly had a significant amount of clot inside of it. it was also reported that anticoagulation was not started until several days after initial implant. the product was not returned for investigation. the cause of the removal issue was most likely use related, as the resistance when removing the pump was most likely due to the clotting inside the graft, which occurred because of insufficient anticoagulation throughout the case. device history review: the pump passed all post sterile inspection criteria. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
4389 | 1220648-2025-47399 | 23265848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23265848&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645513 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced suction/low flows that were not resolved by any fluid delivery or p level adjustment. the pump was explanted and replaced by an intra aocrti balloon pump. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4390 | 1220648-2025-47402 | 23266105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23266105&pc=OZD | 08/30/2025 |  |  | 10/10/2025 |  |  | 08/30/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687336 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Tachycardia | Increased Respiratory Rate | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after successful coronary artery bypass graft, an impella 5.5 was placed for recovery. on day two of impella 5.5 support, the patient spiked a temperature. a couple days later, the patient continued with a fever requiring tylenol. additionally, the patient¿s heparin-induced thrombocytopenia panel came back positive for heparin-induced thrombocytopenia. it was further reported that when the patient was extubated. the patient had some spasm and airway restriction that led to the patient becoming tachycardiac and tachypneic. a decision was made to re-intubate so that steroids can be given. the impella 5.5 was weaned and explanted as planned. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Investigation summary: the investigation has been completed. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. fever/thrombocytopenia/tachypnea: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4391 | 1220648-2025-47404 | 23266233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23266233&pc=OZD | 08/29/2025 |  |  | 10/10/2025 |  |  | 08/29/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp encountered pump flow issues. several repositioning attempts failed to maintain a proper pump flow. in every position of the pump, 0.6 liters was shown on the supporting automated impella controller (aic). both he pump and the aic were exchanged to resolve the issue. | [5 empty] | A.2, a.4 and a.5 are unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report addresses the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report number.
4392 | 1220648-2025-47406 | 23267015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23267015&pc=OZD | 09/11/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 91 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp insertion, the device became entangled in the papillary muscle of the aortic valve, which caused the pump to be unable to be fully delivered into the left ventricle. the pump was pulled back and repositioning was attempted several times without success. the physician removed the pump to rewire to access the left ventricle, however upon removal of the pump the decision was made to replace the pump and deploy a new device. the new device was successfully implanted and was subsequently successfully weaned and explanted and it was noted the patient outcome of explant with the new pump was survived. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult patient anatomy preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks. | A4: patient weight is unknown product status: the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
4393 | 1220648-2025-47407 | 23267259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23267259&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705926 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to resistance at the innominate aortic junction. manual manipulation was performed and the aortic cannula was repositioned but there was significant calcium at the innominate. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The cause of the delivery issue was determined to be patient condition as the patient had calcification of aorta preventing successful delivery of impella. the device passed all the post sterile inspection checks.
4394 | 1220648-2025-47408 | 23267458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23267458&pc=OZD | 09/12/2025 |  |  | 10/10/2025 |  |  | 09/12/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) would not recognize a purge disc during insertion of the purge cassette during set-up for a clinical case. the aic was exchanged for a new device and the issue resolved. there was no patient harm/involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. h6 codes updated to reflect product return. | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4395 | 1220648-2025-47409 | 23267525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23267525&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595020 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed bleeding at the access site. the dilator was not locked into the sheath hub. the sheath was placed in the right front atrium. bleeding was noted around the sheath near the end of the case. the impella cp was explanted and the patient survived. | [5 empty] | A.2 and a.4 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. major bleed: excessive bleeding at insertion site was noted and resolved after removal of the impella. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
4396 | 1220648-2025-47410 | 23267643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23267643&pc=OZD | 09/10/2025 | 11/26/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661963 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 27 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented to the hospital with chest pain. in the emergency room 12 leads showed anterior st-elevation myocardial infarction. the patient was taken to the catheterization lab and an impella cp was placed to support the percutaneous coronary intervention. in the intensive care unit, the patient developed a hematoma and bleeding from the impella cp access site. a new suture was applied to the impella sheath and one-unit of packed red blood cells was given. bleeding was resolved. systemic heparin was held. it was further reported that the patient expired. duration of impella cp support was seven hours. | [5 empty] | The investigation for the reported hematoma and major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma and major bleedcould not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4397 | 1220648-2025-47411 | 23267886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23267886&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026702700 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during the support on day three, the patient experienced a spiked fever over night. there was concerns for a respiratory illness. cultures were drained and urine analysis was sent. the patient will be headed to the operating room as second case of coronary artery bypass graft (cabg). temperature has not went over 100.9. following the cabg with successful revascularization, the decision was made to explant the impella cp. the patient was stable post explant. it is unknown what steps were taken to address the fever. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4398 | 1220648-2025-47412 | 23268134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23268134&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714493 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues during placement of impella 5.5. the impella was opened, primed and torque was removed. the impella was advanced over the guidewire, however, it routed down descending aorta. it was thought that the wire prolapsed out of the ventricle/ attempts to re-cross using an echocardiogram and wireless delivery were both unsuccessful. the impella was able to be advanced using x-ray guidance. however, the patient had significant septal hypertrophy obstructing proper impella placement. as a result of the patient's anatomy, the case was aborted. | [5 empty] | Corrected implant/explant date. investigation summary: no product was returned. unable to place or position (positioning issue): the cause of positioning issue was most likely patient condition related as the patient had significant septal hypertrophy and difficult anatomy preventing successful delivery of impella. device history lot: device lot: 1951080. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | A4 is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4399 | 1220648-2025-47414 | 23268348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23268348&pc=PYX | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  | 10/31/2025 |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024414203 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 85 KG | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support. the patient was escalated from impella cp to a bipella device (impella rp flex and impella 5.5). during the support, the impella rp flex was unable to flow above p6 without suction alarms. as a result, the patient was given albumin. patient care was withdrawn and the patient passed away. there are two devices reported for this event, impella cp and impella rp flex. there was no issue with the impella 5.5 . the patient's death is reported under the impella cp and a separate report was submitted for the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump suction: patient had suction alarms. after reviewing the logs, patient had multiple suction alarms. no motor current spike or placement signal trend was observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4400 | 1220648-2025-47415 | 23268424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23268424&pc=OZD | 09/10/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711247 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 85 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during the impella cp support, with concern of rising lactate, urine output, and diffuse pulmonary infiltrates, the decision was made that the patient would be escalated to an impella 5.5 as well as receiving impella rp flex support. post explant, patient care was withdrawn and the patient passed away. the report represents the impella cp. another report was submitted for the related impella rp flex. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4401 | 1220648-2025-47416 | 23268961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23268961&pc=OZD | 09/12/2025 |  |  | 10/10/2025 |  |  | 09/12/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722234 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | Hematuria | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient presented to emergency room via emergency medical services after experiencing 3-4 days of severe chest pain and back pain. the patient was noted to be hypotensive upon arrival, and the patient was taken to the catheterization laboratory. the patient was catheterized to the culprit vessel, the left anterior descending coronary artery (lad), and the lad was stented. the patient continued to be hypotensive and then decompensated post stenting and an impella cp was placed. it was noted that the patient then had pericardial effusion and needed a transversus abdominis plane for pericardiocentesis. the patient was actively bleeding from pericardial tube, and it was determined that the patient¿s left ventricle wall ruptured. it was noted that the blood appeared to be old, and the medical team believed this to be the cause of the decompensation and cardiac tamponade prior to impella use. the patient was then sent to the operating room (or) for exploration. a left ventricular rupture was confirmed in the or and the findings were consistent with severe acute myocardial infarction. the patient was cannulated for extracorporeal membrane oxygenation (ecmo), the rupture was successfully repaired, and the patient was released back to the intensive care unit with poor prognosis. two days later, the decision was made to escalate the patient¿s support to an impella 5.5 from impella cp and ecmo. the patient was successfully decannulated from ecmo, impella cp was removed, and impella 5.5 support was initiated. the patient was noted on (b)(6) 2025 to have been anuric the previous day, and the patient was started on continuous renal replacement therapy. on the third day of support, it was reported that there was leaking inside the purge reservoir housing. moisture was noted inside the clear plastic. purge flows and pressure were maintained within normal parameters. the customer support center advised to notify the physician and that ultimately, replacing the pump would be the only resolution for the leaking. on the fourth day of support, it was noted that the patient had a computed tomography scan which revealed ischemic changes to the patient's brain. the following day it was noted that the patient experienced pink-red urine output. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. this complaint involves two impella devices: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Product status: the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4402 | 1220648-2025-47418 | 23270591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23270591&pc=OZD | 09/11/2025 | 12/05/2025 |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026727973 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male | 97 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was inserted and after moving the patient to the intensive care unit, bleeding occurred from the impella access site. blood products were administered. the impella cp was removed, surgical hemostasis was performed, and the bleeding was stable. the patient was escalated to impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported major bleed has been completed. the root cause of the injury is most likely use related since after moving the patient to intensive care unit, bleeding occurred from the impella access site.
4403 | 1220648-2025-47417 | 23270597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23270597&pc=OZD | 09/10/2025 | 11/26/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717546 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 66 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was major bleeding, infection and pulmonary embolism during impella cp patient support. while on support, the patient was in interventional radiology for a lower gastrointestinal bleed. heparin drip was running per the bedside nurse. it was noted the patient had a pulmonary embolism so that is a possible reason heparin was not stopped. there was an emobilization of the bleed. a workup for amyloidosis was pending. the sputum culture was positive for pneumonia. amyloid labs came back positive. care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported major bleed, infection, and pulmonary embolism has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the major bleed, infection, and pulmonary embolismcould not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4404 | 1220648-2025-47419 | 23270710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23270710&pc=OZD | 09/12/2025 |  |  | 10/10/2025 |  |  | 09/12/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728375 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 79 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was supporting in the icu when ectopy was observed as well as signs of hemolysis with hematuria. the pump supported til explant after 17 hours. | [5 empty] | The investigation of the reported failure mode has been completed. product was not returned for review. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4405 | 1220648-2025-47421 | 23270892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23270892&pc=OZD | 09/11/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687945 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. while placenig the device on initial axillary cut down, surgeon had a difficult time with bleeding. it was noted that the patient received plavix the am of the procedure. surgeon took time trying to stop bleeding. a jp drain was placed. there was some axillary site with some oozing. patient received 4 units of packed red blood cells, but staff stated the site of bleed was from the groin (not associated with impella). patient successfully weaned as planned and is stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: staff stated the site of bleed was from the groin. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4406 | 1220648-2025-47420 | 23270895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23270895&pc=OZD | 09/11/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698043 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock underwent implantation of an impella cp device for mechanical circulatory support. the complainant noted that the patient was transported to the cardiac catheterization laboratory (ccl) for st-elevation myocardial infarction (stemi) and experienced decompensation during the percutaneous coronary intervention (pci) of the right coronary artery (rca). a decision was reached to insert the impella cp device. as the patient's condition worsened, extracorporeal membrane oxygenation (ecmo) was also initiated. the pci procedure continued and was ultimately completed. following the procedure, the patient was transferred to the intensive care unit (icu). the patient experienced multiple episodes of decompensation while in the icu. an echocardiogram (echo) was performed, revealing pericardial effusion that necessitated urgent pericardiocentesis. despite these interventions, the patient continued to deteriorate, leading the family to decide to withdraw care. subsequently, the patient passed away. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for investigation. pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4407 | 1220648-2025-47422 | 23271053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23271053&pc=OZD | 09/11/2025 |  |  | 10/10/2025 |  |  | 09/11/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616550 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was fitted with an impella cp device to provide mechanical circulatory support. while under impella cp support, the patient was transferred to a different facility. the physician opted to leave the peel-away sheath in place due to concerns regarding potential bleeding, as the patient was receiving heparin and integrillin. at that time, the physician noted that pulses were present. upon arrival at the new facility, however, the patient was found to have no pulse in the right lower extremity, which was also cold. the facility proceeded to remove the peel-away sheath at the bedside and inserted a repo sheath into the arteriotomy. during the process of repositioning the device, the cp inadvertently retracted across the valve. consequently, a decision was made to transfer the patient to the cardiac catheterization lab (ccl) to initiate peripheral veno-arterial extracorporeal membrane oxygenation (vaecmo) and replace the impella cp. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of positioning issue was most likely use issue related since removing the peel away sheath at bedside and placing repo sheath into arteriotomy, the cp pulled back across the valve while pulling device back into proper position. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
4408 | 1220648-2025-47423 | 23273583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23273583&pc=OZD | 09/10/2025 | 12/04/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722257 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 distal end, anticontamination sleeve got torn. the registered nurse placed sterile opposites over disconnected part. the patient continued 5.5 support and was bridged to heart transplant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported product damage has been completed. the device nor sufficient clinical details were returned for evaluation. therefore, the root cause of the product damage could not be determined.
4409 | 1220648-2025-47425 | 23273584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23273584&pc=OZD | 09/10/2025 | 10/10/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  | 09/30/2025 |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406187 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 88 KG |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella cp for mechanical circulatory support during high risk percutaneous coronary intervention (hrpci). upon removal of the impella cp, the physician pulled the 14 french impella sheath out of the body without wire access. manual pressure was applied and percloses tightened, but there was still significant bleeding. contralateral access was obtained, balloon inserted up and over, and brought to the operating room for surgical repair of the common femoral artery. the pump was explanted in the procedural area and discarded. the procedure was success and the patient was stable and survived. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. access site adverse event: the cause of the access site adverse event could not be determined as limited clinical details were provided. mechanical interaction with tissue: the cause of the mechanical interaction with tissue could not be determined as limited clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4410 | 1220648-2025-47424 | 23273585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23273585&pc=OZD | 09/10/2025 | 11/26/2025 |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704436 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered positioning issues during impella cp support. impella was inserted via lfa (left femoral artery) and initially appeared to be positioned well across av (aortic valve), however continuous impella flow reduced alarms present. position assessment on transthoracic echocardiography revealed impella was caught on papillary muscle so physician attempted to reposition. during repositioning, impella came out of lv (left ventricle) and into aorta. a 0.035 guidewire placed through re-access side port and used as rail to re-insert impella. despite careful re-insertion, impella still appeared caught on papillary muscle. unable to run impella higher than p6 in this position, however waveforms and flows were appropriate for p level. physician was ok with leaving pump in this position. patient was sent to unit on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issues has been completed. the root cause of positioning issue was most likely use issue related since impella caught on papillary muscle during initial placement of the pump.
4411 | 1220648-2025-47427 | 23273586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23273586&pc=OZD | 09/06/2025 |  |  | 10/10/2025 |  |  | 09/10/2025 | 10/10/2025 |  |  |  |  | 10/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026669530 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | `the eu complainant reported that the they encountered purge leak during impella cp support. the customer informed that the glucose was leaking from the automated impella controller (aic) console, and opening the door they found the glucose stagnant around the drain box, with continuous leaks. they changed the purge cassette but even with the new cassette it continued to lose glucose. console was exchanged for another one. since exchanging the purge cassette did not resolve the issue the complaint is made against the pump. | [5 empty] | The aic was returned for investigation, but the cp is still implanted in the patient. a supplemental mdr will be filed, on receiving additional information and once investigation is completed.
4412 | 1220648-2025-47343 | 23256392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256392&pc=OZD | 09/02/2025 |  |  | 10/09/2025 |  |  | 09/02/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716196 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hemolysis and access site bleeding. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleeding issue was not determined, as sufficient clinical information was not provided. the cause of the hemolysis issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. the device passed all post sterile inspection checks.
4413 | 1220648-2025-47344 | 23256393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256393&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622773 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 64 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was placed hematoma noted and pressure applied in the cath lab before transferring to intensive care unit (icu). the groin site dressing was intact, and pedal pulses were present. in icu, bleeding was noted from distal area of blue hub on repositioning sheath. one-unit packed red blood cells was infused. small trickle from blue hub minimized by angle matching. staff were made aware of this being a result of blue hub being advanced too far in insertion site. an echocardiogram was ordered; position was visualized. impella was repositioned and pulled back two centimeters. patient was weaned successfully and explanted as planned. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 added b13 code as it was omitted from the initial report that was submitted. investigation summary: the investigation was completed. no product was returned. major bleed: the cause of the bleed was most likely use issue related since bleeding occurred as a result of blue hub being advanced too far in insertion site and small trickle from blue hub minimized by angle matching. device in wrong position (positioning issue): the cause of positioning issue cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks.
4414 | 1220648-2025-47346 | 23256409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256409&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026721376 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was taken to the operation room for escalation to impella 5.5. insertion was attempted several times but unsuccessful. the impella 5.5 placement was aborted and impella cp was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: the investigation has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to patient anatomy. device history review: the device passed all post sterile inspection checks.
4415 | 1220648-2025-47347 | 23256475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256475&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706508 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the access site. the doctor placed forward tension sutures and redressed the site. the site has slight ooze determined to be from suture holes and not sheath itself. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the bleeding is determined to be use issue because the bleeding was resolved by sutures and redressing the site. the device passed all post sterile inspection checks.
4416 | 1220648-2025-47348 | 23256521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256521&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716251 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 96 KG |  | White | Injury | Yes | No | Inaccurate Flow Rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the position of an impella 5.5 could not be verified via transesophageal echocardiogram. the impella was pulled out and reinserted. there was concern for a clot or biomaterial on reinsertion given the immediate flow saturation. the impella was explanted and a new impella 5.5 was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for evaluation. the cause of the thrombosis was unable to be determined due to insufficient clinical details. the cause of the saturated pump flow was due to biomaterial ingestion based on data log analysis and clinical details. the device passed all post-sterile inspection checks.
4417 | 1220648-2025-47350 | 23256531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256531&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728404 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 64 YR |  | Female | 62 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had slow oozing at the impella access site. the dressing was changed and pressure was held. the physician placed a purse string suture and the bleeding resolved. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the bleed is most likely use related since the physician placed a purse string suture and bleeding resolved. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4418 | 1220648-2025-47351 | 23256538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256538&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690139 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted wtih an impella cp experienced low pump flow and a cannula kink. the pump was explanted and a new impella cp was implanted to support the patient. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the cannula kink was most likely patient condition related since patient had small lv and short aortic arch small lv. the cause of low pump flow was most likely patient condition related due to cannula kink per the clinical. the device passed all post sterile inspection checks. | Medical safety/clinical reviewed the event. a 71-year-old female presented to the emergency department with nausea and vomiting for four days. an ecg demonstrated st-elevation myocardial infarction (stemi), and the heart team was activated. the patient was intubated and transferred to the cardiac catheterization laboratory. past medical history included hyperlipidemia, hypertension, and type 2 diabetes mellitus. coronary angiography identified a 100% occlusion of the right coronary artery, and a decision was made to intervene on the culprit vessel. during the procedure, the patient experienced cardiac arrest requiring four rounds of cpr and two defibrillation shocks. two high-dose vasopressors were initiated prior to mechanical circulatory support. an impella cp was placed for mechanical circulatory support; however, the device was noted to be kinked and was unable to achieve flows greater than 1.5 l/min. the impella cp was subsequently removed and successfully replaced with a second impella cp. on the day of replacement pump implantation, pulses were not detectable by doppler in either lower extremity. per the treating physician, this was attributed to peripheral vasoconstriction, and no immediate intervention was noted. two days later, pulses returned in the contralateral leg, while pulses remained absent in the impella access leg; however, the extremity was noted to be warm. the following day, plan of care documentation noted frequent ectopy and thrombocytopenia. as a result, the impella performance level was reduced to p-6. impella mechanical circulatory support (mcs) was continued with no plans for escalation of care due to the patient¿s prognosis. approximately two days later, life-sustaining care was withdrawn, and the patient expired. impella cp pump 1 is a malfunction and impella cp pump 2 is a death type of reportable event. h6 investigation type/finding/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4419 | 1220648-2025-47349 | 23256541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256541&pc=OZD | 05/30/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Ischemia | Tachycardia | Pericardial Effusion | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported in the publication: "recognizing post-cardiac injury syndrome after impella 5.5 insertion in cardiogenic shock: a case-based discussion" that a patient implanted with an impella 5.5 experienced post-cardiac injury syndrome (pcis). the patient developed chest pain, tachycardia, and hypotension post-impella insertion. the patient was treated with ibuprofen and colchicine and received a heart transplant 14 days later. | [5 empty] | Patient details were not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the pericardial effusion was not determined due to insufficient clinical details. the root cause of the tachycardia/ischemia was unable to be determined due to insufficient clinical details.
4420 | 1220648-2025-47352 | 23256546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256546&pc=OZD | 09/05/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659832 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was placed, extracorporeal membrane oxygenation (ecmo) was inserted. during ecmo insertion, the impella became dislodged into the aorta. the physician attempted to reinsert the cp after rinsing it with distilled water, but the cannula kinked. the pump failed to produce adequate flow, so a new impella cp was used. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. no product was returned. product damage (cannula kink): the cause of product damage cannot be determined as insufficient clinical details were provided. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. low pump flow: no product was returned. patient had low flow. no logs are available to review. the cause of low pump flow was most likely due to cannula kink but cannot be determined if it was due to patient condition or use issue related. device history review was completed and passed all post sterile inspection checks.
4421 | 1220648-2025-47353 | 23256549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256549&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700250 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Pericardial Effusion | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was hypotension, major bleeding, positioning issues and a pericardial effusion during impella cp patient support. there was a pericardial effusion, which was resolved with pericardiocentesis. there was a significant amount of blood from the pericardial drain placed during pericardiocentesis and the patient received a blood. anticoagulation was to be held unless the drainage dropped. the decision was made to leave the impella in overnight after attempts to wean resulted in hypotension. point-of-care ultrasound was completed by a general fellow with the inflow appearing to be too shallow and right at mitral valve. the decision was made to assess the position with an echocardiogram machine instead of ultrasound machine before repositioning. | [5 empty] | D6b explant date updated. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of the bleed was not determined due to insufficient clinical details. hypotension, pericardial effusion: the cause of the injury was not determined due to insufficient clinical details. positioning issue: the cause of the positioning issue was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4422 | 1220648-2025-47354 | 23256554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256554&pc=OZD | 09/10/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Female | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted as the pump was unable to pass through the patient's vasculature. no patient harm was reported. | [5 empty] | A.4 added weight of the patient as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ investigation summary: unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had difficult patient anatomy preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4423 | 1220648-2025-47355 | 23256559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256559&pc=OZD | 09/10/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  | 08/31/2025 |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385567 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 pump and supported by ecmo lost optical signal. the patient also suffered bacteremia and was treated with antibiotics. later, the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for evaluation. the cause of the infection was unable to be determined due to insufficient clinical details. the cause of the optical signal issue was due to sensor wear within product design life of 60 days. the device passed all post-sterile inspection checks.
4424 | 1220648-2025-47356 | 23256562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256562&pc=OZD | 09/10/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719738 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 96 KG |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient arrived at the hospital with ventricular tachycardia storm. extracorporeal membrane oxygenation and impella 5.5 was placed. in the intensive care unit, the impella was repositioned overnight with echocardiogram (echo) due to runs of vt. echo showed pump too deep. once repositioned ectopy ceased. the patient continued impella support until it was explanted as planned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4425 | 1220648-2025-47358 | 23256590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256590&pc=OZD | 09/02/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025555323 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male | 88 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu (france) had a 5.5 pump support ongoing for the patient admitted in with a coronary artery bypass graft (cabg) and bentall procedure. the 5.5 was supporting but explanted after 5 days when it was explanted after a red plug leak was observed by the medical team. the field rep was notified days after the explant. no harm is reported due to issue and patient is stable. | [5 empty] | The device or logs was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. red handle leak: the cause of the red handle leak was not determined since product/data logs were not returned and insufficient clinical details were provided. purge leak- pump- the cause of the purge leak was not determined since product/data logs were not returned and insufficient clinical details were provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4426 | 1220648-2025-47359 | 23256606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256606&pc=OZD | 09/08/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636818 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 82 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was escalated from impella cp to impella 5.5 as the patient needed more support. while on impella 5.5 the patient had ectopy requiring lido and amio. the patient had to be shocked once. additionally, the patient was heparin-induced thrombocytopenia and thrombosis (hitt) positive and argatroban was started. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B7 added information that was omitted from the initial report that was submitted. h6 added b13 code as it was omitted from the initial report that was submitted. investigation summary: the complaint investigation was completed. arrhythmia/thrombosis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4427 | 1220648-2025-47360 | 23256607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256607&pc=OZD | 09/10/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716867 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp went into cardiac arrest overnight and expired on support. | [5 empty] | The root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4428 | 1220648-2025-47361 | 23256631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256631&pc=OZD | 09/10/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485774 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 52 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | Hypovolemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was ventricular tachycardia (vt), hypovolemia, and a suction event during impella 5.5 patient support. on 5.5 insertion, the patient went in sustained vt and received one direct current shock to return to normal sinus rhythm. multiple blood products were being given since the patient became hypovolemic overnight. suction events at p9 were resolved with reducing the p-level, bedside echocardiogram was performed and the team was content with the pump¿s position. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hypovolaemia: the cause of the injury was unable to be determined due to insufficient clinical details. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no logs available. the cause of the pump suction was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4429 | 1220648-2025-47362 | 23256637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256637&pc=OZD | 09/10/2025 |  |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615308 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 62 KG | Non Hispanic | White | Death | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implante with an impella cp experienced hemolyzed labs and hematuria. the pump was repositioned and the pump's p-level was turned down. later, the patient expired on support. | [5 empty] | A review of data logs confirmed the complaint condition. no product was returned for review. the cause of the hemolysis was most likely due to the pump in improper position. the cause of the device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4430 | 1220648-2025-31861 | 23256725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23256725&pc=OZD | 03/31/2025 | 10/09/2025 |  | 10/09/2025 |  |  | 03/31/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546188 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 pump support along with extracorporeal membrane oxygenation due to a history of cardiomyopathy. the pump was supporting for 28 days when there were false in aorta alarms. the pump position was reviewed with imaging and found to be appropriate within the left ventricle. the pump remained on despite the false signal. the patient was successfully weaned from the pump and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. clinical details suggest a false position in the aorta alarm on march 31. the data log did not record any major trends or abnormalities in placement signals or motor current during the reported false position in the aorta alarm. the cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4431 | 1220648-2025-31862 | 23260540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23260540&pc=OZD | 07/12/2024 | 10/09/2025 |  | 10/09/2025 |  |  | 07/12/2024 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095635 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella 5.5 support and there were 'placement signal not reliable' alarms. motor current was pulsatile and flows unchanged. staff ordered echocardiography, and placement signal monitoring was disabled. staff monitored and support continued. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that 6 and a half days after pump was started, optical placement signal became unreliable, but then self resolved. initially issue was intermittent and kept recurring, but later psnr condition was permanent and lasted util end of support (two weeks later). after pump was disconnected, shortly prior to shutdown, console logs recorded ossi_bad_lamp_reset_bench_state error, followed by failed ¿5s¿ parameter values, including light=0.53v (indication of no light). ltlog and rtlog data show that during the events, values of contrast were oscillating between -3000 and 3000 and snr was close to 0, suggesting that light was not reaching optical bench detector. there were several occurrences of placement signal not reliable alarm (#1036, due to opm signal invalid). console was sent to danvers in connection with another complaint, where the same failure mode was investigated. during testing the issue was reproduced. light was intermittently absent from the optical pump connector. while lamp was working, the optical bench was passing ¿5s¿ parameter test: cavity length: 14333 nm snr: 12200 contrast: 57.4% lamp: 100.0% gain: 1.40 mano: 753.5 mmhg light: 2.70 v measurements of voltage and current across the lamp assembly indicated that the light bulb was defective. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4432 | 1220648-2025-31863 | 23260583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23260583&pc=PYX | 06/10/2025 | 10/09/2025 |  | 10/09/2025 |  |  | 06/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025534046 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 73 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the impella rp flex was being primed and the purge cassette was not priming. the purge cassette was replaced, and the pump was primed. support continued. additionally, there were suction alarms occurring despite no drop in flow, change in motor current, or pulmonary artery waveforms. the doctor was concerned that metrics on the pump are unreliable. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: clinical details reported that suction alarms were triggering and central venous pressure (cvp) was negative when no changes were noted in pump flows, motor current, or the premature atrial contraction (pac) readings. it was reported that the patient was having a bradycardic episode at the time of the alarms, and both resolved. data logs were reviewed, and the concentration of suction alarms was confirmed. all optical signal 5s parameters were stable and within the expected range during the entire case, suggesting the optical signal was reliable. during the suction alarms, cvp did trend more negative, nearing -50 mmhg. the first suction alarms began to trigger when the pump was turned up to p-8 and they would generally resolve when turned back down to p-7. when the pump was turned up to p-8, dp placement signal (ps) increased and the minimum cvp values decreased. this is to be expected if the pump is in a suction condition or there isn't enough volume for the respective p-level. that being said, the normalized motor current ((mcmax-mcmin)/mcmean) was calculated when the pump was running at p-7 just before turning up to p-8, and at p-8 when the suction alarms began to trigger. the two values were not significantly different, suggesting that the pump may not have been in a true suction condition when the alarms triggered. the threshold for motor current (mc) modulation to trigger the intermittent suction alarm is the same at p-7 and p-8 (>80 ma), which may explain why the alarm was triggering at one and not the other. there were no clinical details provided to suggest the patient had low volume, and the relationship between the bradycardic episode and the alarms is unknown. product was not returned for investigation. since limited clinical details were provided, product wasn't returned for investigation, and data logs were inconclusive, the root cause of the optical signal issue could not be determined. low pump flow: the voice of the customer indicated that pump flows were adequate and the suction alarms were false. based on data log review, the reported false suction alarms were found to be triggered based on motor current, and pump flow calculations were on the lower end of the expected range at times that suction alarms were triggering. for this reason, the low pump flow failure mode was added and investigated. as mentioned above, suction alarms began to trigger intermittently when the pump was increased to p-8. the normalized motor current was relatively similar between p-7 when no alarms were triggering and p-8 when they were. this could be due to the fact that the two p-levels have the same threshold of motor current modulation to trigger the alarm. pump flows were generally within specification - though at the lower end - during the suction events. all this being said, cvp did trend down during the period that suction alarms were concentrated, which could support that the pump was experiencing suction. limited clinical details were provided regarding the patient's condition at the time of the complication, and the relationship between the reported bradycardic episode and the alarms is unknown. product was not returned for investigation. since limited clinical details were provided, product wasn't returned for investigation, and data logs were inconclusive, the root cause of the low pump flows could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4433 | 1220648-2025-31864 | 23260695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23260695&pc=PYX | 06/07/2025 | 10/09/2025 |  | 10/09/2025 |  |  | 06/07/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661903 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on an impella rp flex had placement signal not reliable (psnr) alarm(s) after movement. it was noted that the pump was running fine with known psnr audio off. however, the console was exchanged. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. placement signal not reliable: based on the clinical review, it was noted that during impella support, the placement signal not reliable alarm was triggered after patient movement which was unable to be resolved for the rest of the duration of support. the product was not returned, however data log analysis revealed a drifting placement signal and drifting mean flow. the central venous pressure (cvp) was stable. there were no motor current spikes outside of p-level changes. placement signal not reliable (psnr) alarms #1051 and #1052 were observed, indicating dp signal being out of range and epm sensor failure, respectively. based on the above findings, it was concluded that the root cause of the placement signal issue was most likely sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4434 | 1220648-2025-47363 | 23262838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262838&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571024 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 68 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented to a rural outlying hospital where an electrocardiogram was performed and found the patient to be in st elevation myocardial infarction. as the facility was unable to support the patient, the patient was administered tenecteplase and transferred to the complainant hospital by air for treatment. due to weather, the patient was unable to be directly transported via helicopter to the hospital, and had to be directed to another facility and then transferred by ambulance. upon landing, the patient became confused and combative. the patient arrived to the hospital hypoxic and went into ventricular fibrillation arrest. the patient was defibrillated 4 times and administered pressors. the impella cp was emergently implanted to support the patient for a percutaneous coronary intervention (pci). right heart catheterization found the patient to have a 100% occluded proximal right coronary artery despite the administration of tenecteplase. a temporary transvenous pacemaker was inserted via a right femoral vein sheath due to the patient's symptomatic bradycardia with a heart rate in the 40s, which resulted from the inferior stemi. following defibrillation to resolve the ventricular fibrillation, the patient entered atrial fibrillation rhythm. the physician decided to proceed with right heart catheterization which revealed a revealing a pulmonary artery pulsatility index (papi) of 0.6, and if the atrial fibrillation was not resolved with cardioversion, the physician would place an impella rp flex. the patient was successfully cardioverted and the papi improved to 1.1 and the physician decided to not place an rp flex. by the end of the pci, the patient was able to be weaned from pressors and the impella was able to be weaned down to power level 5. the patient was brought back to the intensive care unit rest and recover the rest of the night, and a plan was made to explant the impella the following day. additionally, the medical team was unable to obtain pedal pulses. no interventions were performed due to the lack of limb perfusion. the patient also experienced a 25 beat run of ventricular tachycardia during impella cp support, however it was noted that the impella maintained a mean arterial pressure at 60 mmhg. the patient was given a loading dose of amiodarone, and an intravenous drip of amiodarone was started. despite the amiodarone, the patient continued to have runs of vt overnight. later that morning a plan was made to explant the pump that afternoon. the patient was successfully weaned and the impella was explanted. the patient was noted to be stable after explant. | [5 empty] | A4 added information as it was omitted from the initial report that was submitted. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. no product was returned for investigation. ischemia/tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4: patient weight is unknown. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4435 | 1220648-2025-47364 | 23262839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262839&pc=OZD | 02/04/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional |  |  |  |  | 49 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 49-year-old woman presented with chest pain and palpitations, and ultimately developed cardiogenic shock secondary to a pheochromocytoma. she required impella 5.5 preceded by impella cp and venoarterial extracorporeal membrane oxygenation (va ecmo) support as a bridge to surgical resection. after adrenalectomy, she was weaned from mechanical circulatory support, with recovery of left ventricular function. because of concern for limb ischemia, impella cp was transitioned to impella 5.5 and va-ecmo was decannulated. | [5 empty] | D3 added manufacturer fax number as it was omitted from the initial report that was submitted. d4 serial number should of been left blank on the initial report that was submitted as the number is unknown. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: ischemia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the necessary materials were not returned for analysis so the serial number could not be identified to complete a phr inspection. | Section a1,4,5,6: patient information is unknown section d: lot, serial, udi, expiration date, dates of implant/explant are unknown h4: date of manufacturer is unknown. this report is based on a literature review and reports an adverse event associated with an impella cp. the incident is described in the article titled "pheochromocytoma-induced cardiomyopathy requiring mechanical circulatory support with cardiac recovery following surgical resection", published in jacc: case reports, volume 30, issue 22 on 6 august 2025. the article did not provide specific device identifiers such as lot or serial numbers.
4436 | 1220648-2025-47366 | 23262914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262914&pc=OZD | 09/09/2025 | 11/20/2025 |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689928 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had a diagnostic catheterization procedure at an outside hospital. post procedure, the patient decompensated and coded. extracorporeal membrane oxygenation (ecmo) and impella cp were placed and the patient was transferred to another hospital for continued care. on arrival to cardiovascular intensive care unit (cvicu), a transesophageal echocardiogram was performed and a kink was noted in the proximal portion of the cannula, with inability to increase flows above 0.6. the decision was made to take the patient back to the catheter lab and replace the pump. the guidewire re-access port was utilized and a new 14fr (french) sheath was placed. the left ventricle was accessed with the pigtail, and the guidewire was switched for the 0.018. the impella was backloaded and advanced into the left ventricle. the wire was removed and the impella was started at p3 with flows of 1.6. the peel away sheath was removed and forward tension sutures placed. the patient was sent to the cardiovascular intensive care unit to rest on impella. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The following sections have been updated: b4, g3, g6, h2, h11. upon further evaluation, it was determined that the reported event no longer meets medical device reporting requirements. the initial report was submitted in error.
4437 | 1220648-2025-47365 | 23262917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262917&pc=DQX | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE, 0.018", PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 8854080 |  | 0052-3046 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the patient was admitted to the hospital with shortness of breath and was taken to the cardiac catheterization laboratory. the patient was found to have left main coronary artery and right coronary artery disease. the patient¿s ejection fraction was determined with echocardiogram to be 10-20% and the patient had an end-diastolic pressure of 40 mmhg. the impella cp was prepped; however the physician was unable to cross the aortic valve due to aortic stenosis. the impella insertion was then aborted, and the patient was taken to the operating room for a coronary artery bypass graft and aortic valve replacement. no patient harm was reported due to the impella cp aborted implant. | [5 empty] | D.1 revised brand name as it was entered incorrectly on the initial report that was submitted. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had severe aortic stenosis preventing successful delivery of impella. device history review: the guidewire passed all the post sterile inspection checks. | A2, 4: patient age and weight is unknown. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4438 | 1220648-2025-47367 | 23262918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262918&pc=OZD | 09/09/2025 | 02/19/2026 |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704212 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Failure to Sense | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that a patient was placed on impella 5.5 to stabilize while working up for ventricle assisted device vs transplant. during support pink urine was reported. in addition to suction events requiring fluids and albumin. the patient also received 1 unit packed red blood cells for bleeding from ct. the patient continues on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6: added code a070908. additional information. the placement signal reading was inaccurate ( 20s/-9). patient was hemodynamically stable and position adequate under echocardiogram. reported that the device was exhibiting a ¿placement signal not reliable¿ alarm. | The investigation for the reported hematuria, major bleed, and product suction has been completed. no product was returmed for review. the root cause of the hematuria and major bleed was not determined due to insufficient clinical details. patient had suction events which resolved over night requiring fluids and albumin. after reviewing logs, multiple suction alarms were observed. no motor current spike or placement signal trend was observed. the root cause of pump suction was most likely patient condition related since the suction alarms resolved after fluids was given to the patient.
4439 | 1220648-2025-47369 | 23262920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262920&pc=OZD | 09/09/2025 | 11/26/2025 |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712219 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during a attempt to place the impella 5.5 there was an issue with the device and wire was kinking. the 5.5 was removed and a new 5.5 device was successfully placed. | [5 empty] | The investigation for the reported delivery issue and product damage has been completed. no product was returned for investigation. the root cause of delivery issue is determined to be use issue because the second pump was inserted without any issues. the root cause of the product damage is determined to be use issue because there were no issues reproduced while inserting the second pump later. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4440 | 1220648-2025-47368 | 23262921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262921&pc=DQX | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018" PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male | 90 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella implant, left ventricle (lv) perforation apical with 0.018¿ impella wire, which resulted in cardiac effusion. immediate pericardiocentesis with pigtail via xiphoid was performed and drainage of a total of 1,5 l blood and 4 units of red blood cells given. patient at any time stable with mean arterial pressure > 60mmhg, responsive and awake. patient was bridged to coronary artery bypass graft after lv perforation. device was explant in operation room. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. cardiac perforation/ pericardial effusion: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | D4: the udi, sn and lot # of the device is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4441 | 1220648-2025-47370 | 23262925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262925&pc=OZD | 09/09/2025 | 11/26/2025 |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728655 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision for impella cp was made to stabilize a patient. it was noted that the cp was opened and in sterile field. they had right groin access and found 8f cordis sheath to be nearly occlusive of vessel. left femoral artery then accessed with micro puncture which showed equally small vessel size. cat scan images confirmed porcelain aorta with severe calcium also in bilateral innominate arteries, of note patient also had a left axillary stent placed the day prior. decision to abort impella as no safe access points could be identified. intra-aortic balloon pump was then reinserted for hopeful support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the delivery issue has been completed. the root cause of the delivery issue was determined to be patient condition as the patient had porcelain aorta with severe calcium also in bilateral innominate arteries preventing successful delivery of impella.
4442 | 1220648-2025-47371 | 23262926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23262926&pc=OZD | 09/09/2025 |  |  | 10/09/2025 |  |  | 09/09/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713185 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pericardial Effusion | Reperfusion Injury | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella cp the patient had limb ischemia of right femoral artery with impella cp in place. vessels were clamped down secondary to profound shock. reperfusion cannulas placed preemptively. due to patient increased pressor dosage, escalation to impella 5.5 resulted in removal of impella cp and reperfusion cannulas. the doctor called in to case and was thinking of extracorporeal membrane oxygenation (ecmo) but also a pericardial effusion was noted so a window was performed by the doctor and hemodynamics improved. decision at that point was not place ecmo and to place impella 5.5. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | H6 added b13 code as it was omitted from the initial report that was submitted. h10 this is one of two related reports. this report represents the pump and another report was submitted to represent the introducer. the related report number was added. investigation summary: the investigation was completed and no product was returned for investigation. ischemia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. reperfusion injury/pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4443 | 1220648-2025-47373 | 23263140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263140&pc=OZD | 08/31/2025 | 09/30/2025 |  | 10/09/2025 |  |  | 08/31/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102529 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support on impella 5.5 the impella controller reported having battery indicator at 50% shortly after unplugging from power source resuming to 100% after being plugged back in for <30mins. earlier this week battery noted to be at 84% immediately upon unplugging earlier this week. | [5 empty] | D9 updated; the product has been returned. the investigation is still ongoing. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4444 | 1220648-2025-47375 | 23263148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263148&pc=OZD | 09/02/2025 |  |  | 10/09/2025 |  |  | 09/02/2025 | 10/09/2025 |  |  | 11/30/2025 |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024423582 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Fever | Pain | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump to support the patient admitted in with cardiac history, cardiomyopathy, and being put on the transplant list. the patient had the 5.5 support ongoing when there was pain at the site and fever. there was no noted intervention or premature explant. additionally, there was a leak and pump was "sticky" but no site of purge sidearm/reservoir leak was observed or troubleshot. the pump remained on until successful weaning and the patient survived. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for evaluation. pain: clinical details noted patient was experiencing some pain at impella site but was managed with medication. the cause of the injury was not determined due to insufficient clinical details. fever: the root cause of the injury was unable to be determined due to insufficient clinical details. purge leak - pump: the cause of the purge leak - pump could not be determined as no product was returned for evaluation. device history review: the device passed all the post sterile inspection checks.
4445 | 1220648-2025-47374 | 23263149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263149&pc=OZD | 08/10/2025 | 09/30/2025 |  | 10/09/2025 |  |  | 08/31/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705219 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 70 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella 5.5 on that the patient continues to rest on support, still shaded due to infection. blood cultures are still coming back positive. the patient had a broad-spectrum antibiotic for probable gastrointestinal infection. care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B3 revised date of event as it was incorrect on the initial report that was submitted. b7 added information that was omitted from the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. g3 date was reported incorrectly on the initial report that was submitted and should of reflected 10-aug-2025. h6 added code b13 as it was omitted from the initial report that was submitted. h10 this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. the related report number was added. investigation summary: the investigation was completed. infection: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4446 | 1220648-2025-47378 | 23263150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263150&pc=OZD | 09/05/2025 | 12/09/2025 |  | 10/09/2025 |  |  | 09/03/2025 | 10/09/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701682 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had hypotension, ectopy, hemolysis, minor bleeding and hematoma during impella 5.5 support. the patient got hypotensive with map in 50's a few days into pump support. as a result, fluid bolus was given. additionally, patient encountered more ectopy as seen in the ekg and so they were weaned off of levophed. the following day, patient had no episodes of hypotension, however, pfhgb (plasma free hemoglobin) levels were noted as trending down. patient also had some bleeding at insertion site after they gout out of bed which resolved on its own. surgeon did note a hematoma near tunnel and the team decided to perform a pocket washout and add a jp drain for it. hemolysis labs were also slightly up so the pump flow was turned down to level p-5. no further interventions noted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported hypotension arrythmia,hemolysis, and access site adverse event has been completed. no product was returned for review. the data logs were reviewed and lt logs showed no motor current spike prior to reported hemolysis. no persistent positioning alarms were noted and placement signal was stable and not trending. the root cause of the hypotension arrythmia,hemolysis, and access site adverse event could not be determined as no product was returned and limited clinical details were provided.
4447 | 1220648-2025-47376 | 23263151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263151&pc=OZD | 09/02/2025 |  |  | 10/09/2025 |  |  | 09/02/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701703 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on impella 5.5 for support. during support it was noted that epinephrine and norepinephrine restarted post arrhythmia. additionally, echocardiogram (echo) and repositioning performed due to ventricular tachycardia episodes. left ventricle thrombus on echo and heparin was started. the 5.5 was successfully weaned and explanted as bridge to left ventricle assisted device. the patient survived. | [5 empty] | Correction b1. the investigation of the reported failure mode has been completed. no product was returned. ventricular fibrillation, thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4448 | 1220648-2025-47377 | 23263152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263152&pc=OZD | 09/03/2025 | 10/09/2025 |  | 10/09/2025 |  |  | 09/03/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665406 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was ectopy during impella cp patient support. while on support, it was noted the patient was still having ectopy. an abiomed representative noted the patient was in ventricular tachycardia due to heart failure symptoms. once the patient was escalated to a 5.5, the ectopy resolved, and the patient was loaded with amiodarone. | [5 empty] | E.1 added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.10 this is one of two related report numbers. another report was submitted for the impella 5.5 this report represents the impella cp. the related report number was added. investigation summary: the investigation has been completed. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4449 | 1220648-2025-47379 | 23263155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263155&pc=OZD | 09/04/2025 |  |  | 10/09/2025 |  |  | 09/04/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711577 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 112 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a fever, and a suction event during impella 5.5 patient support. while on support, the patient spiked a fever. cultures were sent and antibiotics were ordered. there was suction during a bronchoscopy and albumin was administered. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 b13 code was added as it was omitted from the initial report that was submitted. investigation summary: the investigation was completed. no product was returned. fever: the cause of injury cannot be determined since insufficient clinical details were provided. pump suction: no product was returned. no data logs were available. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks.
4450 | 1220648-2025-47381 | 23263156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263156&pc=OZD | 09/05/2025 | 10/27/2025 |  | 10/09/2025 |  |  | 09/05/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704631 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 78 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Patient-Device Incompatibility | Device Contamination with Chemical or Other Material |  |  |  | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Perforation of Vessels | Thrombosis/Thrombus | [6 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella cp support to correct shunt placement and potentially bridge to surgical repair. the impella was placed per the instructions for use in the right femoral artery and turned on to p7. the patient was hypertensive on arrival to the intensive care unit, and the pump position was confirmed. the patent had high ptt, act, and bleeding noted. protamine was reversed and heparin was started. the patient had recurrent suction alarms, despite support being at p2. position of pump confirmed on echocardiogram, confirmed to have no kinks, thrombus, or biomaterial. the pump had an impella in aorta alarms, despite confirming proper placement. blood volume was given, introducer flushed, and pump was repositioned, but the pump flow issue was not resolved. a clot was suspected in the inflow. suction alarms reported were possibly due to ventricular tachycardia issues, which resolved with medication. found to have been put on double concentrated dobutamine and suspected to be the cause of the arrhythmia. hemoglobin dropped so heparin was stopped and the patient was taken for cat scan, finding bladder perforation. stroke code was also called after change in neurological status. | [5 empty] | Corrections to the initial mfr report: h6 med dev prob code a1080104 added. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4451 | 1220648-2025-47380 | 23263157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263157&pc=OZD | 09/04/2025 | 11/26/2025 |  | 10/09/2025 |  |  | 09/04/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711573 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | Unspecified Tissue Injury | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support on impella 5.5 on september 4th, 2025, the patient had to be increased on pressors, suspected sepsis. family at bedside, unable to get labs. registered nurse states labs normal, ptt sub therapeutic but they are actively titrating. gastrointestinal (gi) bleed noted, holding systemic deep tissue injury (dti) until gi consult. bicarb in purge. intact cat scan yesterday for seizure activity. on 9/8 may be restarting heparin tomorrow if gi bleed continues to subside. patient started on continuous renal replacement therapy two days ago and extracorporeal membrane oxygenation flows increased. due to gi bleed a colonoscopy yesterday was done and needed 1 unit packed red blood cells (prbc). off systemic anticoagulation. the patient was still on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported sepsis, renal failure, and major bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the sepsis, renal failure, and major bleed could not be determined.
4452 | 1220648-2025-47382 | 23263158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263158&pc=OZD | 09/05/2025 | 11/25/2025 |  | 10/09/2025 |  |  | 09/05/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026666078 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female | 60 KG |  |  | Injury | Yes | Yes | Malposition of Device | Patient-Device Incompatibility | [5 empty] | Pulmonary Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was pulmonary edema and positioning issues during impella cp patient support. while on support, it was reported that the team would continue to run extracorporeal membrane oxygenation (ecmo) at a high flow rate to clear a pulmonary edema. intensive care tapped the effusion overnight which made no difference in filling. there was a plan to change ecmo configuration to veno-arterial-venous today in cath lab. a cat scan ordered to see if the patient could accommodate a 5.5. the thought process is the pigtail of the cp was catching the displaced papillary muscle which was tethering open her mitral valve. if 5.5 could not be accommodated, then intensive care would balloon the septum. a surgeon was being consulted. the device was removed due to incompatibility with patient¿s anatomy. | [5 empty] | The investigation for the reported edema and positioning issue has been completed.the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the edema and positioning issue could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4453 | 1220648-2025-47383 | 23263161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23263161&pc=OZD | 09/10/2025 | 11/26/2025 |  | 10/09/2025 |  |  | 09/10/2025 | 10/09/2025 |  |  |  |  | 10/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673758 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 75 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient encountered a minor bleed and eventually expired after impella cp support. patient experienced oozing and bleeding at impella site. care team member was able to resolve with manual compression. additionally, patient was transferred from carle methodist to osf st. francis for tentative plan of venoarterial extracorporeal membrane oxygenation and arrested shortly after arriving to intensive care unit. patient was resuscitated for a few minutes before arresting again and was not able to be revived after 30min of cardiopulmonary resuscitation. | [5 empty] | The investigation for the reported minor bleed and hemodynamic instability has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the minor bleed and hemodynamic instability could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4454 | 1220648-2025-47298 | 23243404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23243404&pc=PYX | 09/08/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026680005 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 127 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella rp flex had a pump stop. the team attempted to restart the impella rp flex three times without success. the pump motor current was rapidly increased to 22.48 on the automated impella controller and immediately stopped. the impella rp flex was exchanged. | [5 empty] | D9 updated with device received date. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. pump stop: data log review showed a motor current (mc) spike with a pump stop (motor current high). attempts were made to restart the pump immediately after without success. there was also a slight increase in purge pressure and slight decrease in purge flow at this time. data logs show 1 hour later the pump was turned on to p-1 with minimal flows for about 10 minutes. the pump was returned and evaluated. there was significant biomaterial wrapped around the impeller and in the inlet and outlet cages. the pump stop was reproduced. the cause of the pump stop was biomaterial ingestion since returned product showed biomaterial and the data logs showed a mc spike with the pump stop. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
4455 | 1220648-2025-47299 | 23243455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23243455&pc=PYX | 09/08/2025 | 10/07/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026683405 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | IMPELLA CP SMART ASSIST SET | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received that the device was explanted due to patient demise. the cause of death was due to cardiogenic shock related to wide open aortic valve insufficiency. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The user facility reported an impella rp flex was placed along with a left sided impella cp for a patient admitted in with shock and methamphetamine use history. the impella rp flex was placed and due to bleeding/ooze there was a mattress suture adjusted/placed and manual pressure was held. the bleed did resolve. |  |  |  |  | Additional information was received and the following fields were updated: b2, b5, d6b, h1, and h6. investigation summary: the investigation has been completed. major bleed: the cause of the injury was not determined since the product was not returned and insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4456 | 1220648-2025-47301 | 23243722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23243722&pc=OZD | 09/05/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557944 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient had impella 5.5-automated impella controller (aic) support ongoing when on day 6, the aic had an optical alarm that prompted the team to replace the aic. no harm or injury was noted. the pump remains on for support. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | D6a and 6b: updated implant and explant date d9: added device return information h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary the device was returned for evaluation. data analysis: a review of the data logs shows that approximately 16.5 hours into the case, errors begin to appear in the logs. these are followed roughly two minutes later by a yellow ¿placement signal not reliable¿ (psnr) alarm, which confirms the complaint. device analysis: a pump was connected to the console and ran at various flow levels for ~2 hours. the rt logs for this testing were reviewed, and the snr was found to be ~250. the optical fiber in the front panel optical connector was examined with an inspection camera and found to be clean and free of remarkable defect. the optical 5s parameters were measured using two different test cavities. all values are within the ranges; however, the snr is marginal. root cause: the root cause of the psnr alarms is an optical bench exhibiting low snr when the console is running a pump. | Correction: d6a should have been left blank.
4457 | 1220648-2025-47300 | 23244170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244170&pc=OZD | 09/06/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633706 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed for the support of a patient with surgical plans for a ventricular septal defect. the patient was to have an impella cp and extracorporeal membrane oxygenation support but with delivery there was a failure and perforation perhaps caused by the abiomed wire, or another .035 wire . the impella cp was aborted and instead an intra-aortic balloon pump was placed. the free wall rupture/perforation was repaired. additional information was received. the 35 wire used to insert the pigtail gc was not an abiomed guidewire. the free wall rupture (fwr) occurred during the exchanged from the 35 wire to an 18 guidewire. abiomed products were neither used nor involved in this fwr. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. cardiac perforation: the cause of the injury was not determined due to insufficient clinical details. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition. device history review: the device passed all the post sterile inspection checks.
4458 | 1220648-2025-47305 | 23244229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244229&pc=OZD | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713987 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the placement signal for a patient implanted with an impella 5.5 is far off from the patient's blood pressure. additionally, the patient required additional surgery to resolve bleeding. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the optical sensor issue was not determined due to a lack of clinical details and no product being returned. the cause of the bleeding cannot be determined as insufficient clinical details were provided. the device passed all post-sterile inspection checks.
4459 | 1220648-2025-47307 | 23244239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244239&pc=PYX | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551595 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 116 KG | Hispanic | White | Death | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Fever | Hemolysis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient implanted with an impella rp flex developed a fever and hemolysis. in response, the impella was repositioned. later, the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The pump was not returned for investigation. data logs were available for the first five hours of the case. the cause of the hemolysis was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. the cause of the fever was not determined due to insufficient clinical details. the cause of the positioning issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. the pump passed all post sterile inspection checks.
4460 | 1220648-2025-47308 | 23244304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244304&pc=OZD | 09/08/2025 | 10/30/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724475 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was inserted and after removal of the peel away oozing was noted at the site. bleeding did not improve after manual pressure, forward tension sutures or angle matching. access was obtained in the left femoral artery to angio the site where a severe bleed was noted at the initial stick. the impella cp was removed and covered stents were used to obtain hemostasis. two units of blood products were administered. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed clinical details noted that there was bleeding around the repositioning that required 2 packed red blood cells, impella removal, and covered stents to stop the bleeding. physician thought stick was good, but the patient had an unusually large rfa that failed to seal around the repositioning sheath. the product was returned and evaluated. no abnormalities were found with the repositioning sheath. the cause of the access site bleeding was patient condition related since clinical details note an unusually large rfa that did not seal around sheath and the returned product had no abnormalities. the device history review was completed and passed all post-sterile inspection checks.
4461 | 1220648-2025-47304 | 23244326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244326&pc=OZD | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738748 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp suffered bleeding from the veno-arterial extracorporeal membrane oxygenation site and a gastrointestinal bleed. three units of red blood cells and one unit of platelets were transfused. systemic anticoagulation was withheld. additionally, suction alarms occurred and flows were decreased. in response, the pump was repositioned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of bleed cannot be determined since insufficient clinical details were provided. the cause of pump suction cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks.
4462 | 1220648-2025-47309 | 23244375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244375&pc=OZD | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655191 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Fracture | Difficult to Remove | Detachment of Device or Device Component |  |  |  |  | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Transesophageal echocardiogram revealed that the motor had been pulled past the anastomosis and into the graft and there was no thrombus present on the graft. | The complainant had an impella 5.5 to support a patient. the patient was under general anesthesia, and the device removal was attempted in the operating room. there was resistance during removal, and the device was successfully removed, but it broke off at the connection between the motor and the cannula. transesophageal ultrasound confirmed that the portion beyond the cannula remained within the graft. the thoracotomy was reopened, and the remaining portion was removed. the patient survived. |  |  |  |  | The investigation of the reported failure mode has been completed. no product was received for evaluation, but photo sent. difficult/unable to remove/product damage (cannula detached from motor housing) the clinical details state that the cannula detached from the motor housing in the graft upon explant. it was noted that it was thought that the device got caught on the clavicle during removal causing a lot of force to be applied. a tool was used on explant. an image of the product was provided and the cannula was confirmed to be detached from the motor housing at the outlet cage. there was a clamp mark seen on the cannula and no kinks noted in the catheter near the motor housing. the pump was rather clean as the clinical details stated the pump was washed. the delo band remained on the motor housing and did not appear to have tears. the surface of the lbb (motor housing) looks very smooth and the amount of delo on that surface cannot be confirmed. the root cause of the difficult/unable to remove cannot fully be established because it cannot be confirmed where the additional resistance was applied and it was only speculated the device was caught on the clavicle. the root cause of the product damage (cannula detached from the motor housing) is due to a material fracture but the cause cannot be further established. device history review: the device passed all the post sterile inspection checks. | B5 updated with additional information received from the clinical site. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
4463 | 1220648-2025-47311 | 23244572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244572&pc=OZD | 09/09/2025 |  |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703397 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during patient care which required turning the patient, the impella became mispositioned, the flows dropped to .2 liters and echocardiogram was done which showed the impella in the papillary muscle. the patient was taken to cardiac catheterization lab where in attempting to free the pigtail the impella came fully across the aortic valve. the impella cp was explanted and a new device was inserted. the patient was stable post explant. additional information was received. the staff disposed of the pump. the doctor did not believe the pump exchange was the cause as the patient had been very unstable and attempting to replace the pump was "heroic measures'. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. device in wrong position (positioning issue): the cause of positioning issue was most likely use issue related since the pump malpositioned after the patient was turned over during the care. pump suction: no product was returned. flows dropped to .2l after the pump came out of position. no logs are available for review. the cause of pump suction was most likely use issue related since the flows dropper after the pump came out of position when the patient was turned over during care. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
4464 | 1220648-2025-47313 | 23244777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244777&pc=OZD | 09/10/2025 |  |  | 10/08/2025 |  |  | 09/10/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694489 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 69 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support placed for an acute myocardial infarction/cardiogenic shock patient. the pump was explanted after 19 hours due to concerns of hemolysis. the explanted pump was not replaced. this complaint involves two impella devices: automated impella controller with serial number (b)(6) and impella cp pump with serial number (b)(6) this mdr specifically addresses the impella cp pump with serial number (b)(6), which is the subject of this report. a separate report was submitted for the other device. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4465 | 1220648-2025-47312 | 23244859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23244859&pc=OZD | 09/09/2025 | 10/10/2025 |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1242777 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 69 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had automated impella controller (aic)-impella cp support ongoing for a patient admitted in with acute myocardial infarction/cardiogenic shock. the aic was replaced due to a controller failure. there was purge alarm issues. a new aic was placed without any harm or injury. this complaint involves two impella devices: impella cp pump with serial number: (b)(6) and automated impella controller with serial number: (b)(6). this report specifically addresses the automated impella controller with serial number: (b)(6), which is the subject of this report. a separate report was submitted for the other device. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4466 | 1220648-2025-47314 | 23245555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23245555&pc=OZD | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026734375 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 55 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered ischemia post impella cp support. immediately after the case concluded, the impella cp and the impella sheath were explanted and the patient began to complain of their leg tingling and numbness. a doppler was used to assess blood movement in the right foot and ankle, and pulsatility was observed but produced a weaker signal than the left foot. the physician decided to have the patient kept for observation for a few hours and would determine with further diagnostic testing if further intervention was necessary. no further intervention was noted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation was completed. no product was returned for investigation. ischemia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4467 | 1220648-2025-47317 | 23247077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23247077&pc=OZD | 05/13/2023 | 01/22/2026 |  | 10/08/2025 |  |  | 05/13/2023 | 10/08/2025 |  |  | 12/31/2024 |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325112 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that they encountered an optical signal issue. once this second pump was placed, a placement signal unreliable alarm was encountered. as a result, the decision was made to replace with a new impella cp. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Investigation summary: on 2023-may-21 cp pump sn (b)(6) and pc returned and previously investigated under the incorrect pump sn and smartsolve complaint number (b)(4) . that record has been corrected and this pump sn (b)(6) checked in here to (b)(4) . optical signal issue - upon implant, psnr occurred. decision made to use a new pump. data log review showed psnr alarm occurring within 2 minutes of case start, with snr decreased, light decreased, and contrast decreased while lamp and gain remained stable. the returned pump showed abnormal %s parameters but passed laser continuity test. a kink was found in the catheter near the kink protector in the red plug. the cause of the optical signal issue was a damaged optical fiber in the catheter due to material kink to both the catheter and optical fiber. device history lot: device lot: 1692908. device history batch: subcomponent lot: n/a. device history review: pump s/n (b)(6) passed all post sterile inspection checks.
4468 | 1220648-2025-47316 | 23247095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23247095&pc=OZD | 09/10/2025 |  |  | 10/08/2025 |  |  | 09/10/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724360 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 78 KG | Non Hispanic | Asian | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Insufficient Information | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There was oozing noted during support from multiple sites. activated clotting time (act) was not known. | The complainant reported that there was a low purge pressure alarm during impella cp support. the nurse was reporting low purge pressure alarm, which was resolved prior to the call. the purge fluid bag had just been changed; however, the purge flow has been running 30 milliliters per hour. there were no leaks noted. care was withdrawn and the patient expired. |  |  |  |  | Correction b5 and h6 health effect - clinical code. the investigation of the reported failure mode has been completed. no product returned. purge pressure low - data log review confirmed purge pressure alarms throughout case with purge pressure remains on average below 370mmhg and purge flows around 30ml/hr for the duration of the case. the cause of the purge pressure low could not be determined due to lack of returned product, inconclusive data log review, and insufficient clinical details. access site adverse event: the cause of the access site adverse event was patient condition related due to the oozing noted to be coming form multiple sites. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Information on follow up report 1 regarding minor bleeding should not of been reported as minor bleeding is not a reportable event. investigation summary: the investigation information has been updated. no product was returned. purge pressure low - after three days on support, the registered nurse reported low purge pressure alarm. purge fluid was running at 30 milliliters per hour (ml/h). no leaks were noted. data log review confirmed purge pressure alarms throughout case with purge pressure remaining on average below 370mmhg and purge flows around 30ml/hr for the duration of the case. the cause of the purge pressure low could not be determined due to lack of returned product, inconclusive data log review, and insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
4469 | 1220648-2025-47318 | 23247300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23247300&pc=OZD | 07/04/2024 |  |  | 10/08/2025 |  |  | 07/04/2024 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484137 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Ischemia | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant placed the impella 5.5 to support the patient admitted for heart failure/cardiogenic shock. while on support, positioning issues, access site injury and suction alarms were noted. one day post start of the impella 5.5 support, positioning unknown alarms were encountered. under an echocardiogram, the impella 5.5 pump was back to 4.5cm from the aortic valve to address the issue. approximately six days later and some days to following it was additionally noted there was some persistent access site pain and scant serosanguinous drainage experienced. days following, the access site was tender, and the dressing was regularly changed. oozing persisted but the pain resolved. resolution of the oozing timeframe is unknown. however, now on day 11 of impella support, suction was encountered with a brief placement signal diastolic of 14. the echocardiogram revealed the impella was measuring 6.4cm and was repositioned. the impella 5.5 device was pulled back to 5.9cm (45.5cm at site). in addition to the above, notifications via abiomed¿s long-term outcome and quality indicator (loqi) impella registry were received and reported the patient experienced pericardial effusion. pericardiocentesis was performed with jp drain placement that was eventually removed. repeat echocardiograms showed stability with last one showing small effusion. furthermore, the patient experienced chest pain with discomfort and some burning, which was described as sharp, and occurred when taking a deep breath or moving. the patient was tender to palpation to both the epigastric and sternum. morphine was needed for pain control. a chest mri was done and showed granulation tissue/scarring vs low grade muscle strain at the medial margin of the left rectus abdominal muscle just inferior to the sternum just deep to the sternotomy -nonspecific findings. relatedness to the impella device used for these events was unknown/undetermined. eventually, the patient received a heart transplant, and the pump was successfully removed. | [5 empty] | Patient-specific information section a4 and a5 are unknown. the device history review was completed and the pump passed all post sterile inspection checks. the product was not returned for review. positioning issue: the cause of positioning issue was unable to be determined as limited clinical details were provided, and no product was returned for review. pump suction issue: data logs showed the pump ran without issue during support. there were suction and positioning alarms triggered during the case but resolved quickly. the cause of pump suction issue was unable to be determined as limited clinical details were provided, and no product was returned for review. pericardial effusion: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. ischemia: in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details.
4470 | 1220648-2025-47319 | 23247365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23247365&pc=OZD | 09/06/2025 |  |  | 10/08/2025 |  |  | 09/06/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551114 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing from the chest site. two units of packed red blood cells and one unit of fresh frozen plasma (ffp) were transfused. the patient also was receiving one unit of ffp now prior to rtor for hopeful washout and chest closure. increasing pressor requirements so dobutamine was started at 2.5. the patient underwent a washout, packing removal and chest closure. the patient exhibited signs of hemolysis so the pump was repositioned at the bedside. later, the patient was successfully weaned from the pump and was explanted. the graft bled back when removing the pump. the exposed graft was cut back and sealed, and tucked into the skin pocket and was sutured closed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | D.9 the device was returned and the date was added. h.6 codes were updated as the device was returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4471 | 1220648-2025-47321 | 23247844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23247844&pc=PYX | 09/02/2025 |  |  | 10/08/2025 |  |  | 09/02/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025570403 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 92 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Inflammation | Thrombocytopenia | [10 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a bipella patient encountered ventricular tachycardia (vt), developed willebrand syndrome, cholecystitis and liver fluid buildup during impella rp flex support. the patient encountered vt when they were turned and a second run of vt the following day. additionally, the care team believed that the impella caused the patient to develop von willebrand syndrome. moreover, a liver ultrasound showed fluid around the liver, blockage in the portal duct, and bilirubin levels that kept climbing. cholecystitis was also found. no other impella related alarms were noted. no anticoagulation was given. no further information was noted, and the patient finished impella rp flex support after 17 days. additionally, the bipella patient ultimately expired and the death will be associated with impella cp pump with serial number (b)(6). this complaint involves three impella devices: impella cp pump with serial number (b)(6). automated impella controller with serial number (b)(6). this report specifically addresses the impella rp flex with serial number (b)(6), which is the subject of this report. separate reports will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event. the patient experienced arrhythmic events and thrombocytopenia while on bipella support. when turning the patient ventricular tachycardia was experienced and the following morning experienced a run of ventricular tachycardia. comfort care was discussed. day nine, the patient is only on right-sided support and noted with cholecystitis and liver fluid buildup; both linked to gallstones. the following day of the support, an unexpected automated impella controller (aic) shutdown occurred; support interrupted. the team was advised against exchanging the aic due to the concern that the pump would not restart. impella rp flex support continued. seven days later, the patient expired on support despite all efforts. the patient was maxed on vent and chemical support. there is not enough evidence to exclude the impella rp flex and the aic as an associated factor to the demise outcome. however, the aic is highly unlikely associated with the outcome due to the amount of days post event and no report or indication of acute hemodynamic compromise at the time of the aic malfunction. in this case, the patient underlying condition likely contributed most to an outcome of death. note: events involving the impella cp is reportable as a serious injury type of reportable event. correction(s): after review of the case by the medical safety, it was determined the impella rp flex is a death injury type of reportable event. b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect- impact code) have been updated, corrected accordingly. d3 added manufacturer fax number as it was omitted from the initial report that was submitted. investigation summary: the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the pump passed all the post sterile inspection checks. regarding arrhythmia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. regarding thrombocytopenia/ inflammation, the cause of the injury was not determined due to insufficient clinical details. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4472 | 1220648-2025-47322 | 23248033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23248033&pc=OZD | 09/07/2025 | 10/07/2025 |  | 10/08/2025 |  |  | 09/02/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703491 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 92 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombocytopenia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a bipella patient encountered ventricular tachycardia and developed willebrand syndrome during impella cp support. the patient encountered ventricular tachycardia (vt) when they were turned and a second run of vt the following day. additionally, the care team believed that the impella caused the patient to develop von willebrand syndrome. however, after finishing bipella support, the patient decompensated and ultimately expired. the death will be associated with impella cp pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. refer to section h.10 for related report numbers for the other pump and automated impella controller associated with this report. | The investigation of the reported failure mode has been completed. no product was returned for investigation. arrhythmia: the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | Medical safety reviewed the event. the patient experienced arrhythmic events and thrombocytopenia while on bipella support. when turning the patient ventricular tachycardia was experienced and the following morning experienced a run of ventricular tachycardia. comfort care was discussed. day 9 the patient is only on right-sided support and noted with cholecystitis and liver fluid buildup; both linked to gallstones. the following day of the support, an unexpected automated impella controller (aic) shutdown occurred; support interrupted. the team was advised against exchanging the aic due to the concern that the pump would not restart. impella rp support continued. seven days later the patient expired on support despite all efforts. the patient was maxed on vent and chemical support. there is not enough evidence to exclude the impella rp flex and the aic as an associated factor to the demise outcome. however, the aic is highly unlikely associated with the outcome due to the amount of days post event and no report or indication of acute hemodynamic compromise at the time of the aic malfunction. in this case, the patient underlying condition likely contributed most to an outcome of death. note: events involving the impella cp is reportable as a serious injury type of reportable event. correction(s): after review of the case by the medical safety, it was determined the impella cp is a serious injury type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. h6: investigation type/finding/conclusion codes remain unchanged as previously reported. related medical device reports: this impella cp report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller and the impella rp flex, which is associated with the demise outcome.
4473 | 1220648-2025-47320 | 23248068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23248068&pc=OZD | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707470 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ischemia | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed pneumonia and ventricular tachycardia. the patient was treated with heparin, levophed, and vasopressin. the patient also had bleeding that was treated with transfusion of blood products and a femstop. the patient also experienced tingling of the fingers and diminished pulses, and swelling at the impella site. heparin was stopped in response to an acute drop in hemoglobin and hematocrit and presence of melena. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The root cause of the infection/tachycardia/ischemia was unable to be determined due to insufficient clinical details. the cause of the bleeding is most likely patient condition related since patient had very calcified vessel and unable to preclose and difficult to stick. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
4474 | 1220648-2025-47323 | 23251545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251545&pc=OZD | 09/08/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026663212 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 146 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was on impella 5.5 support for 14 days when patient had ventricular tachycardia storm and was intubated. the family made the decision to withdraw care and the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4475 | 1220648-2025-47326 | 23251549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251549&pc=OZD | 09/08/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710786 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 65 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Sepsis | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that there was sepsis, hemolysis and thrombocytopenia during impella cp patient support. while on support, it was noted that a sepsis workup would be continued. the sepsis workup showed a positive blood culture and a white blood cell number of 13.7. there were overnight concerns for hemolysis on labs. the doctor repositioned the impella with echocardiogram. the pump measured 4.2cm and the doctor did not wish to adjust it further at that time. the platelet count was low while on continuous renal replacement therapy. | After insertion high afterload was noted. the impella p-level was reduced and antihypertensives were administered. echo was completed on arrival at the icu and again when hemolyzed labs were noted. hemolysis was resolved. the patient was treated for sepsis. |  |  |  |  | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received and the following fields have been updated: b5, h6 correction made: d4, e4.
4476 | 1220648-2025-47325 | 23251550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251550&pc=OZD | 08/23/2024 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511471 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a left axillary artery dissection during the pump insertion through the introducer. vessel repair was completed successfully and the patient remained on impella support until successful explant. | [5 empty] | The cause of the vascular dissection was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4477 | 1220648-2025-47324 | 23251551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251551&pc=OZD | 09/08/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692932 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during insertion of impella 5.5 on (b)(6) 2025, the impella was inserted to the right axillary artery accessed via cutdown, 10mm headshield platinum graft placed. 0.035¿ wire and al1 used to attempt to cross valve. unable to get around arch, switched to guidewire. catheter and guidewire able to cross valve, exchanged for 0.018 wire. impella loaded onto wire and inserted through graft. unable to get impella down aorta, hung up on curvature of anatomy. multiple suggestions for stiffer 0.018¿ wire and to take a picture of vessel with contrast to surgeon, but surgeon did not want to try other alternatives. decision was made to abort impella 5.5 and place impella cp instead. | [5 empty] | The investigation of the reported failure mode has been completed. difficult to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4478 | 1220648-2025-47327 | 23251552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251552&pc=OZD | 09/07/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700041 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of impella 5.5 on (b)(6) 2025, the medical doctor repositioned impella earlier today. discussion of hemolysis and weaning impella today. hemolysis suspected. am labs hemolyzed, urine output decreased, color red. therapy in progress per medical doctor. impella repositioning with echocardiogram (echo). hemolysis improved overnight with volume optimization. lvedp 0 at p-8. did not tolerate dobutamine challenge with tachycardia. given likelihood of needing ongoing support for multiple days, and that metabolic demands will increase with wakefulness and potential ethanol symptoms, tentative plan for real-time oncology review for escalation to 5.5 today. the pump was explanted in operating room. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were no alarms triggered during the case. mechanical interaction with blood: the cause of mechanical interaction with blood was most likely due to patient condition. tachycardia (arrhythmia): the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of bleeding was most likely due to high activated clotting time (act), it resolved due to reduction in act later. device history review: the device passed all the post sterile inspection checks.
4479 | 1220648-2025-47328 | 23251559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251559&pc=OZD | 09/09/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585523 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Ischemia | Renal Failure | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility had a cp pump placed to support the patient admitted in with acute myocardial infarction complicated by cardiogenic shock and the patient was transferred to the tertiary from the community. the pump was supporting with limb ischemia and the ischemia prompted distal perfusion catheter placement. additionally there was positional issues, but positioning was complicated by the ventricular septal defect, and renal issues ongoing. the team gave blood, started continuous renal replacement therapy, placed extracorporeal membrane oxygenation, and support continues. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure: the cause of the injury was not determined due to insufficient clinical details. wrong positioning: the cause of the positioning issue was most likely related to the patients condition, as clinical reported that the lv was underfilling. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b explant date added. correction: e4.
4480 | 1220648-2025-47330 | 23251566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251566&pc=OZD | 09/09/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648075 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 62 KG | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had a cp placed for the support of a high-risk percutaneous coronary intervention patient. the pump had optical signal issues but, remained on for support until explant. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. optical signal issue - data logs were not recovered. the cause of the optical issue was not determined due to no product returned, no data logs recovered, and insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4481 | 1220648-2025-47331 | 23251567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251567&pc=OZD | 09/09/2025 |  |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700069 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient blood pressure is labile and also in atrial fib. quite a bit of ectopy noted. advised the a fib could reduce filling into the left ventricle and catheter may be causing some ectopy. starting a touch of levo. patient is still in a fib. additionally there was bleeding around repo sheath into groin. massive bruising around bilateral groin. blood products given. after angio evaluation and cath lab intervention on other coronaries, bleeding resolved. impella was successfully weaned as planned and patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d4 the investigation of the reported failure mode has been completed. no product was returned for investigation. arrhythmia: the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
4482 | 1220648-2025-47333 | 23251568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251568&pc=OZD | 09/09/2025 |  |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716174 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia that self-resolved. there was also intermittent suction events so the impella position was optimized under echocardiogram. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the positioning issues could not be determined due to insufficient clinical details and lack of product returned. the cause of the suction could not be determined due to insufficient clinical details or returned product/data. the device passed all post sterile inspection criteria.
4483 | 1220648-2025-47332 | 23251569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251569&pc=OZD | 09/09/2025 |  |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710448 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 96 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. during the implant the impella 5.5 was backloaded over the wire but unable to pass impella catheter through the vessel. after multiple attempts, angiogram performed, and the patient was found to have severely stenosed and calcified artery. the impella was removed from the vessel and graft. balloon angioplasty performed to artery, followed by repeat angiogram. stenosis was corrected, and balloon removed. second attempt to advance impella catheter past graft and into the artery and some difficulties noted with passing the impella catheter so physician elected to utilize viperslide to assist with pump delivery. impella catheter advanced into the artery successfully and then subsequently into the left ventricle, support was initiated. pci was then performed successfully, and the patient was sent to the unit on impella 5.5 support. additionally bleeding and hematoma were noted at the impella access site. blood products were administered. the patient was taken for washout and hematoma evacuation. the patient continued impella support and was explanted as planned and isstable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. hematoma/ major bleed: the root cause of the bleeding and hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. delivery issue: the cause of the delivery issue is determined to be patient condition because the patient condition dues to patient had sever stenosis and calcified artery causing the difficulty in insertion of impella. device history review: the device passed all the post sterile inspection checks.
4484 | 1220648-2025-47334 | 23251571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251571&pc=OZD | 09/09/2025 |  |  | 10/08/2025 |  |  | 09/09/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698070 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was unable to cross the aortic arch due to the patient¿s anatomy due to a prior surgery. implantation of the impella was aborted. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. investigation summary: no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the the pump was not able to navigate the anatomy due to patient's due to prior surgery and the delivery was aborted. device history review: the device passed all post sterile inspection checks.
4485 | 1220648-2025-47336 | 23251575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251575&pc=OZD | 09/09/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/10/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502187 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support escalated to the 5.5 pump for the patient admitted in with cardiac symptoms and high risk surgery. on the 8th day of support the team observed bleeding from all sites and a possible gastrointestinal bleed. the heparin had been held and 1 unit of red blood cells was given. the following day the pump was explanted after a successful weaning. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be patient condition because the patient is oozing from multiple sites. device history review: the device passed all the post sterile inspection checks.
4486 | 1220648-2025-47335 | 23251576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251576&pc=OZD | 09/10/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/10/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026723662 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility cp pump placed to support the critical patient who had arrested outside of the hospital with ventricular fibrillation and a lucas device was placed. the patient exhibited signs of hemolysis and was seen with elevated plasma free hemoglobin labs and urine color change. the team left the pump on for support until the patient again ventricular fibrillation arrested and expired when cpr did not resuscitate the patient. the pump had supported for 46hours. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: data log shows some fluctuation in pump flow and placement signal while running on a steady p-level. intermittent pump position in aorta/ventricular alarms were noted when the pump p-level was changed to p2. no suction or motor current spike was reported during the case run. the cause of the hemolysis issue was not determined due to insufficient clinical information and lack of returned product for investigation. device history review: the device passed all the post sterile inspection checks.
4487 | 1220648-2025-47337 | 23251579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251579&pc=OZD | 09/10/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/10/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had a cp pump placed to support the patient with an admission for cardiogenic shock complicating acute myocardial infarction. the pump was placed and after 6 days was explanted after weaning. the nurse held pressure and observed there was thrombosis. there was no noted treatment for the thrombosis/biomaterial at the pump. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4488 | 1220648-2025-47338 | 23251708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251708&pc=OZD | 08/29/2025 |  |  | 10/08/2025 |  |  | 08/29/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657747 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 88 KG |  | White | Injury | Yes | Yes | Crack | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pain | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump as care escalation from a impella cp pump. the 5.5 was supporting when there was pain and a hematoma at the pump site/arm. there was an abdominal bleed also that prompted red blood cells to be infused during the support. the entire nitinol coil detached from motor housing while explanting. the pump was explanted and the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. product damage (cannula detached from motor housing) the returned product confirmed that the cannula was detached from the motor housing at the location of the outlet. there were no deformities in the cannula as in any clamp marks or scratches seen. the delo band remained on the motor housing and there were no indications of it ripping or incorrect potting. residual glue marks are noted on the inside of the cannula and the indentation of it being bonded to the motor housing can been seen. on the surface of the lbb (motor housing) delo also remains as evidenced by the darker material on the surface. there were no kinks seen on the catheter but the amount of slack on the catheter between the motor housing an blue butterfly is rather short. the root cause of the product damage (cannula detached from the motor housing) is due to a material fracture but the cause cannot be further established. asae: it was stated that the patient had a hematoma at the access site since transfer and it did not increase in size. there was no jp drain present and they were off systemic heparin. there are no other details surrounding the hematoma and it was unknown if it resolved. the cannula was detached from the motor housing on the returned product but that occurred on explant and there were no other damages noted. due to a lack of clinical information, the root cause of the hematoma cannot be determined. bleed/pain: the patient had abdominal bleeding as well as left arm pain. the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4489 | 1220648-2025-47339 | 23251711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251711&pc=OZD | 09/08/2025 | 10/08/2025 |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717696 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 was unable to pass through the axillary artery, and the 5.5 insertion was aborted. the physician opted to placed impella cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to patient anatomy. device history review: the device passed all the post sterile inspection checks.
4490 | 1220648-2025-47340 | 23251715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23251715&pc=OZD | 09/08/2025 |  |  | 10/08/2025 |  |  | 09/08/2025 | 10/08/2025 |  |  |  |  | 10/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849988 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a controller error alarm occurred on an automatic impella controller (aic). a loaner aic was used instead. the patient subsequently expired. the patient arrived at the catheterization lab to received percutaneous transluminal coronary angioplasty (ptca) to right posterior descending artery (rpda) and drug-eluting stent to the diagonal artery. staff reported when the patient was about to be moved off of catheterization lab table the patient began having chest pain and st elevation. angiogram revealed perforation to diagonal artery and pericardial effusion was noted. the patient went into cardiac arrest. a pericardiocentesis was performed, and the decision made to place an impella cp device for mechanical circulatory support. at the start of the emergent case, the nurse turned on the automated impella controller (aic), and the yellow controller error alarm went off. the loaner aic was used instead. impella cp was inserted and return of spontaneous circulation (rosc) was achieved after impella insertion. heparin and angiomax had previously been administered; activated clotting time (act) were 272 seconds reported. multiple units of packed red blood cells (prbc) were administered. the impella peel-away sheath was removed, and repositioning sheath was advanced. the impella cp catheter secured with forward tensions sutures and angle match dressing. impella access site was clean dry and intact with no evidence of bleeding or hematoma with right distal pulse present. the patient remained while awaiting flight transport to receiving facility. patient remained intubated. patient had a central line to the left internal jugular, arterial line to right radial. volume was being administered. approximately 600 ml of output from pericardiocentesis drain. urine was noted to be clear yellow in foley bag; knee immobilizer was in place to right leg. the patient arrived via life net flight and coded as helicopter landed. resuscitation was administered for approximately 15 minutes with rosc in the emergency department. the aic transferred to the receiving hospital's aic successfully without incident. the patient was transported to unit and multiple prbcs, platelets, sodium bicarb and albumin infusing. pericardiocentesis drain was draining continuously. the patient¿s pupils were fixed and dilated at 4. discussions were made with the family. the physician consulted for percutaneous coronary intervention to the left anterior descending/diagonal arteries to try to control perforation bleed. impella site was still clean dry and intact. frequent suction alarms were noted due to bleeding issues. knee immobilizer was in place; feet were warm. no mottling was noted. the patient was severely acidotic with continued correction measures. on-site bedside assessment with the team and the patient was taken to catheterization lab. access to left femoral artery was obtained. diagnostic left heart catheterization revealed right coronary artery stent open and diagonal stent totally occluded with no bleeding from prior perforation noted. left ventricle angiogram was performed which showed no left ventricular perforation; however, left ventricle was very small with little contractility noted. end diastolic pressure was 30. the patient went into pulseless electrical activity (pea). lucus device was started. prbcs, calcium, epi and bicarb were administered. access to left femoral vein was made with swan advanced to the pulmonary artery. return of spontaneous circulation was achieved after 20 minutes of cardiopulmonary resuscitation (cpr). the patient was taken back to the unit and went into pea after arriving. lucus device was started and return of spontaneous circulation achieved after approximately 10 minutes of cpr. at this time, the family decided to assign do not resuscitate. patient went pea and impella was turned down to performance level p-2 and all drips were stopped. time of death was called. the impella and ventilator were stopped. | It was reported that a controller error alarm occurred on an automatic impella controller (aic). a loaner aic was used instead. |  |  |  |  | Medical safety review. medical safety reviewed the event. this event describes a coronary artery perforation and pericardial effusion noted during pci prior to impella placement. patient went into cardiac arrest; pericardiocentesis performed and this time a decision made to place impella cp. during the support there was frequent suction alarms noted due to bleeding issues. the patient was eventually made dnr after multiple rounds of cpr. the patient would never fully recover and expired. the patient encountered suction due to small ventricle and most likely preexisting bleed. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. conservatively, the death should be reported on the impella cp device that in use at time of expiration. regarding the aic, there is no association to outcome. the impella device was never connected to this aic. the controller error, however, is a device malfunction and should be reported as a malfunction type of reportable event. related medical device reports: this automated impella controller device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome. correction(s): b5. event description was updated to provide complete details of the reported event. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4491 | 1220648-2025-31856 | 23232476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23232476&pc=OZD | 06/30/2025 | 10/07/2025 |  | 10/07/2025 |  |  | 06/30/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686645 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that shortly after the impella 5.5 was implanted, a 'placement signal not reliable' alarm appeared. despite this, the impella continued to support the patient successfully. the pump was eventually weaned and explanted without incident, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the cut pump (only the motor side) was sent for investigation. no investigation could be performed on the returned product. the data log showed that the placement signal was lost, accompanied by a placement signal not reliable (psnr) alarm and fiber damage trends across all 5s parameters. no motor current or spikes were reported during the time of the psnr alarm, and the placement signal was not recovered until the end of support. the cause of the optical signal issue was most likely damaged optical fiber line, based on the data log review. however, the exact location of the damaged fiber line could not be confirmed without the returned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4492 | 1220648-2025-31857 | 23232671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23232671&pc=PYX | 07/01/2025 | 10/07/2025 |  | 10/07/2025 |  |  | 07/01/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659135 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 133 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella rp flex support when the pressure differential sensor became unreliable, resulting in unavailable flow readings. the device was explanted to initiate extracorporeal membrane oxygenation (ecmo), a decision not related to the sensor issue. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. no product was returned for analysis. the data logs confirm pattern associated with sensor drift with placement signal decreasing as pump flows increased with stable motor current readings and no p-level change or ingestion events. the root cause of the placement signal issue was most likely due to sensor drift of the differential pressure sensor as evidenced by pattern in the logs. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4493 | 1220648-2025-47276 | 23233018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23233018&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623838 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male | 68 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia requiring defibrillation and revascularization with one stent. ultimately, the patient expired. | [5 empty] | Arrhythmia/ cathe root cause of the arrhythmia and cardiac arrest was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4494 | 1220648-2025-47277 | 23233051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23233051&pc=OZD | 09/08/2025 |  |  | 10/07/2025 |  |  | 09/08/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684340 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 101 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a hematoma. systemic anticoagulation was paused and the team held manual pressure. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4495 | 1220648-2025-47278 | 23233066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23233066&pc=OZD | 09/08/2025 |  |  | 10/07/2025 |  |  | 09/08/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714726 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a groin bleed and dropping hemoglobin and hematocrit levels. blood products were transfused and the pump was explanted. the patient survived. | [5 empty] | The cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4496 | 1220648-2025-31858 | 23235215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235215&pc=OZD | 07/04/2025 | 10/07/2025 |  | 10/07/2025 |  |  | 07/04/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676913 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 75 YR |  | Male | 67 KG | Non Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and, during support, there were placement signal unreliable alarms. support continued. the patient eventually expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: it was reported that the placement signal not reliable (psnr) alarm triggered during case start. pump position was confirmed, and neither the pump nor console were swapped out due to the issue. data logs were reviewed and the psnr alarm was confirmed. during case start, after the pump was connected but before going in the body, there was a sudden drop of signal-to-noise ratio (snr) to 0, contrast dropped, but its amplitude increased slightly, lamp increased to 100%, gain increased, and light appears to have dropped just slightly on the lt log. the placement signal does not return for the remainder of the case. the signatures seen in data logs could be indicative of several causes. product was not returned for investigation. since data logs were inconclusive and product was not returned for investigation, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4497 | 1220648-2025-47279 | 23235306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235306&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695304 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support experienced intermittent bradycardia. a decision was made to restart epi. additionally, the patient also had hemolysis, and it was managed with continuous veno-venous hemodialysis. the patient was successfully explanted after left ventricle ejection fraction recovery. the patient survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4498 | 1220648-2025-47280 | 23235316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235316&pc=OZD | 09/08/2025 |  |  | 10/07/2025 |  |  | 09/08/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637810 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the patient. the pump was placed and support was ongoing when limb ischemia was observed. the limb required fasciotomy and blood was given. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4499 | 1220648-2025-47281 | 23235320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235320&pc=OZD | 09/08/2025 | 10/06/2025 |  | 10/07/2025 |  |  | 09/08/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738888 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 108 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a bipella support ongoing for a critical patient when the rp flex was explanted and the 5.5 pump remained on with extracorporeal membrane oxygenation being added. after just 17 hours the patient expired. the patient had been suffering from a groin site ooze and suction. the patient had been given blood products inclusive of red blood cells, fresh frozen plasma platelets, and albumin. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned. major bleed: the cause of the bleed was not determined due to insufficient clinical details. pump suction: patient had suction alarms post procedure. after reviewing the logs, multiple suction alarms were observed throughout the run time and high motor current spike and the end most likely due to explant. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4500 | 1220648-2025-47282 | 23235329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235329&pc=OZD | 09/08/2025 |  |  | 10/07/2025 |  |  | 09/08/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578495 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma that required perclose, angioseal, and manual pressure to achieve hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
4501 | 1220648-2025-31859 | 23235338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235338&pc=OZD | 07/03/2025 | 10/07/2025 |  | 10/07/2025 |  |  | 07/03/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676040 |  | 0048-0014 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during the percutaneous coronary intervention (pci) procedure using the impella cp, the motor current waveform, as well as the aortic and left ventricular placement signals, were initially visible but suddenly disappeared. the automated impella controller was subsequently exchanged, and the pci procedure was completed. the pump was later weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned for analysis. insufficient data logs were provided. the data logs returned did not cover the entire duration of the case. the logs returned were while console was on ic2825. the logs show no identifiable patterns with 5s parameters as logs only covered about a minute of support. no issue against the pump was reported. the root cause of the optical signal issue was unable to be determined as no product was returned for analysis, insufficient data logs were returned for analysis, and insufficient clinical information was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4502 | 1220648-2025-47283 | 23235359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235359&pc=PYX | 09/08/2025 |  |  | 10/07/2025 |  |  | 09/08/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661883 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 67 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp had a placement signal not reliable alarm. the alarm was disabled. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was omitted from the initial report that was submitted. e.1 added initial report phone number as it was omitted from the initial report that was submitted. investigation summary: no product returned. placement signal issue - after four days on support, peak signal-to-noise ratio (psnr) alarm occurred. data log review showed psnr alarms occurring with both alarm 1051's and 1052's. 5s parameters are stable while ps, pap, and cvp all intermittently rail. no motor current spikes outside of p-level changes. the cause of the placement signal issue could not be determined as no product was returned, inconclusive data log review, and insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
4503 | 1220648-2025-47284 | 23235403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235403&pc=OZD | 09/08/2025 |  |  | 10/07/2025 |  |  | 09/09/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700311 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump failed to deliver. the 5.5 was replacing an implanted cp when dissection of the axillary artery was identified. the dissection was repaired surgically. the pump was explanted in response to ischemia in the patient's arm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.1 selection was added that was inadvertently omitted from the initial report that was submitted. d.1 revised brand name and d.4 model number as they were incorrectly entered on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. investigation summary: no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition due to patient anatomy. initial attempt was made with impella 5.5 and was aborted since the catheter and guide wire was unable to get around the aortic arch and switched to impella cp and similar issue arise with the impella cp. the surgeon was able to manipulate the pump down the aorta into the left ventricle. hence, the cause is determined to be patient condition. ischemia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. vessel perforation: vascular surgery was consulted and discovered axillary artery dissection which was repaired. the cause of the vascular perforation was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
4504 | 1220648-2025-47285 | 23235513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235513&pc=OZD | 09/09/2025 |  |  | 10/07/2025 |  |  | 09/09/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1824431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | Power Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The us complainant had a 5.5 pump-aic support ongoing when there was a power/battery issue observed in the support of the cgs/ami patient. the pump and aic remained in use as the patient listed for the heart transplant. the aic would not keep battery charging unless holding the plug as the connection was loose. the patient was transplanted after 14 days and the console was removed for investigation. | [5 empty] | The console was returned for investigation. the investigation found that the white wire were broken and disconnected from its terminal. additionally, the outer sheath was improperly cut, in accordance with the vendor wiring instructions. the light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose. the console logs align with the reported complaint. the issue of a loose outlet was determined to be caused by defective ac power cord set. the ac power cord set was replaced and a full functional check was performed on the console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d6a and d6b should have been left blank h5 and h8 were erroneously entered on the initial manufacturer device report.
4505 | 1220648-2025-31860 | 23235582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23235582&pc=PYX | 07/06/2025 | 10/07/2025 |  | 10/07/2025 |  |  | 07/06/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675643 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was receiving bipella support with the impella cp and rp flex when the rp flex experienced optical signal loss. no intervention or pump explant was performed in response to the malfunction, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. placement signal not reliable (psnr) alarms also noted. data log confirmed issue. optical parameters were stable throughout the case. the cause of the placement signal issue was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4506 | 1220648-2025-47286 | 23240690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240690&pc=PYX | 09/07/2025 | 11/06/2025 |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026623765 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received from field that unfortunately this patient expired on support, and the devices were not retained. | The complainant reported that during support on impella rp flex that after linen change and turning the patient, went into supraventricular tachycardia (svt) 140-150s went into svt 140-150s. cardioverted x 1 successfully, resumed lidocaine. two days later patient¿s lactic increased to 14.3, plan for va ecmo with impella cp for left ventricle unloading. |  |  |  |  | Additional information received and the following fields were updated: b2, b5, h1, h6. correction to following fields were done: d2a. the investigation of the reported failure mode has been completed. no device was received for investigation. supraventricular tachyarrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4507 | 1220648-2025-47287 | 23240691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240691&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669034 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was successfully bridged to transplant. | The complainant reported that the patient on impella 5.5 support waiting on a heart transplant was positive for heparin-induced thrombocytopenia. heparin was suspended and the patient was switched to bival. additionally, aortic signal on impella was significantly different from cuff blood pressure. the patient is still on impella 5.5 support and is stable. |  |  |  |  | Additional information received: b5, and d6b updated. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4508 | 1220648-2025-47289 | 23240699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240699&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690178 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 120 KG | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. during support it was noted that the patient was experiencing hemolysis. later it was reported that during surgery there was a pump stop and the impella was removed. the patient remained stable. | Additional information received: the surgeon cross clamped over the motor, the pump stopped, and it was able to be restarted. the pump was explanted at the end of the surgical case. the impella items were disposed of and are not available for return. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | Additional information: b5. the investigation of the reported failure mode has been completed. no product was returned. hemolysis: the cause of the hemolysis was not determined since insufficient clinical details were provided, no product was returned, and data logs were insufficient. temporary pump stop: the cause of the temporary pump stop was not determined since insufficient data logs are available, insufficient clinical details were provided, and product was not returned. device history review: the device passed all the post sterile inspection checks.
4509 | 1220648-2025-47288 | 23240701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240701&pc=OZD | 09/07/2025 | 10/14/2025 |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702736 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 80 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Contamination | Restricted Flow rate |  |  |  |  | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Perforation of Vessels | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [31 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant reported a patient placed on impella cp support to correct shunt placement and potentially bridge to surgical repair. the impella was placed per the instructions for use in the right femoral artery and turned on to p7. the patient was hypertensive on arrival to the icu, and the pump position was confirmed. the patent had high partial thromboplastin time( ptt), activated clotting time (act), and bleeding noted. protamine was reversed and heparin was started. the patient had recurrent suction alarms, despite support being at p2. position of pump confirmed on echocardiogram, confirmed to have no kinks, thrombus, or biomaterial. the pump had an impella in aorta alarms, despite confirming proper placement. blood volume was given, introducer flushed, and pump was repositioned, but the pump flow issue was not resolved. a clot was suspected in the inflow. suction alarms reported were possibly due to ventricular tachycardia (vt) issues, which resolved with medication. found to have been put on double concentrated dobutamine and suspected to be the cause of the arrhythmia. hemoglobin dropped so heparin was stopped and the patient was taken for ct, finding bladder perforation. stroke code was also called after change in neurological status. the patient was escalated to impella 5.5 support. | The user facility reported a patient taken to the catheter lab for escalation from an impella cp to an impella 5.5. the physician was having difficulty passing through vasculature, and requested the staff grab and open an impella cp. after the new impella cp was handed to the physician, they were able to pass and place the impella 5.5. the customer account did not want to keep the unused impella cp as the packaging was soiled and contaminated with blood from the pump handoff. |  |  |  |  | B1 adverse event was was inadvertently omitted on the initial manufacturer device report. b5 event description was erroneously entered on the initial manufacturer device report. e4 was inadvertently omitted on the initial manufacturer device report. h1 was erroneously selected on the initial manufacturer device report. h6 health effect - impact code f26 medical device problem code a1801, health effect - clinical code e2403 were erroneously entered on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4510 | 1220648-2025-47291 | 23240725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240725&pc=OZD | 09/07/2025 | 10/07/2025 |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725283 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 121 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support on impella cp on (b)(6) 2025, the patient had gastrointestinal bleed. 4 units red blood cells (rbc). scope to identify the bleed and blood products. heparin on hold. the pump was explanted. | [5 empty] | D9 updated. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. no physical damage was noted on the pump. (hemorrhage/bleeding): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4511 | 1220648-2025-47292 | 23240729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240729&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1882177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 92 KG | Hispanic | White | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [28 empty] | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered unexpected shutdown issues with the automated impella controller (aic) for a bipella patient during impella rp flex support. it was reported that the pump went into an unexpected controller shutdown. no other issues were noted and a standby console sat in the hallway in case shutdown happened again. additionally, the bipella patient ultimately expired and the death will be associated with impella cp pump with serial number (b)(6) this complaint involves three impella devices: impella cp pump with serial number (b)(6), impella rp pump with serial number (b)(6), this mdr specifically addresses automated impella controller with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. there were no issues related to this complaint discovered when reviewing the product history of console (b)(6). product (b)(6) was returned to field service for investigation. data analysis: on (b)(6), rp pump 551808 was connected to (b)(6) and ran uninterrupted until on (b)(6) at which point the console unexpectedly rebooted. there were no logs indicative of what caused the reboot. the console rebooted and hemodynamic support was returned until the pump was removed from the console on (b)(6), a week later. device analysis: the console was tested but the failure mode was not reproduced during testing. root cause: the cause of the unexpected reboot could not be determined because there were no logs to indicate what caused the reboot and the failure did not reproduce during testing. | Medical safety reviewed the event. the patient experienced arrhythmic events and thrombocytopenia while on bipella support. when turning the patient ventricular tachycardia was experienced and the following morning experienced a run of ventricular tachycardia. comfort care was discussed. day 9 the patient is only on right-sided support and noted with cholecystitis and liver fluid buildup; both linked to gallstones. the following day of the support, an unexpected automated impella controller (aic) shutdown occurred, support interrupted. the team was advised against exchanging the aic due to the concern that the pump would not restart. impella rp support continued. seven days later the patient expired on support despite all efforts. the patient was maxed on vent and chemical support. there is not enough evidence to exclude the impella rp flex and the aic as an associated factor to the demise outcome. however, the aic is highly unlikely associated with the outcome due to the amount of days post event and no report or indication of acute hemodynamic compromise at the time of the aic malfunction. in this case, the patient underlying condition likely contributed most to an outcome of death. note: events involving the impella cp is reportable as a serious injury type of reportable event. correction(s): after review of the case by the medical safety, it was determined the aic is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. additionally, the impella device was returned (contrary to the erroneous information on the initial medwatch report). an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this aic report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp and the impella rp flex.
4512 | 1220648-2025-47293 | 23240731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240731&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694561 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a suction event, intracranial herniation and limb ischemia during impella cp patient support. it was noted that suction alarms were resolving with albumin. additionally, previously diminished distal pulses returned in the right lower extremity. the abiomed clinical representative suggested neers for left extremity perfusion monitoring given the right leg extremity cp (nitropaste on foot for apparent decreased perfusion) the patient was undergoing a neurological assessment with a pause of sedation and pain meds. the patient¿s only remaining reflex was a cough. an urgent ct was ordered. prior to the ct, the right pupil was found to be significantly larger than the left. ct later confirmed intracranial herniation and was declared brain dead. the patient was made dnr and organ donation was being pursued. the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: suction alarms are resolving with albumin. after reviewing the logs, multiple suction alarms were observed. the cause of pump suction was most likely patient condition related since the suction alarms resolved after the patient was given albumin and fluids. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4513 | 1220648-2025-47294 | 23240807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240807&pc=OZD | 09/07/2025 | 10/13/2025 |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621739 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered pump stop and patient developed fever during impella cp support. during support, patient developed a fever and as a result antibiotics were started and blood cultures were drawn. additionally, it was reported that the patient was having a central line placed and about 20 minutes later the impella console started alarming that the motor current was high and then the pump stopped. multiple attempts were made to restart the console and an attempt for console exchange was made but pump would not restart. pump was removed and retained. this complaint is for the third pump (impella cp), please see pc-001975753 for the first pump (cp) and pc-001961913 for the second pump (5.5). | [5 empty] | The investigation of the reported failure mode has been completed. the impella pump and guidewire was returned for investigation. pump stop - data log review showed impella stopped ¿ motor current (mc) high alarm 13¿s at case end. pump attempted to be restarted more than 5 times. leading up to pump stop, slight mc uptrend, purge flow increases and purge pressure decreases leading up to pump stop. no placement signal (ps) trends. the returned pump was cut but showed a large impella purge gap, indicative of bearing wear. no other physical abnormalities noted. the cause of the pump stop was bearing wear due to the large impella purge gap seen on the returned pump along with flow metric changes leading up to pump stop. the issue occurring on day 11 of support goes against design trace matrix, which shows cp c9 has an intended design life of 8 days. fever: the root cause of the fever was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4514 | 1220648-2025-47295 | 23240850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23240850&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301954 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered automated impella controller (aic) issues: pump stop and shutdown issue. it was reported that the patient was having a central line placed and about 20 minutes later the impella console started alarming that the motor current was high and then the pump stopped. multiple attempts were made to restart the console and an attempt for console exchange was made but pump would not restart. the pump was removed. this complaint is associated with the third pump (impella cp), please see pc-001975753 for the first pump (cp) and pc-001961913 for the second pump (5.5). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4515 | 1220648-2025-47296 | 23241366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23241366&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697486 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Ischemia | Pulmonary Edema | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp a hematoma to right groin was noted with manual pressure held. the nurse reported after manual pressure right groin site felt softer and controlled. urine noted to be pink tinged. echo at bedside and the doctor repositioned impella. left radial arterial line placed and central line placed to left groin. later it was reported that no longer able to assess distal pulses in bilateral lower extremities. the doctor reported impella appeared deep on echo and the doctor repositioned the impella. chest x-ray shows worsening pulmonary edema. staff still unable to doppler pulses but vascular doppler shows flow down pt to foot. foot is warm and no mottling noted. the patient went back to cardiac catheterization lab last night for temp pacer and repeat left heart catheterization. the patient remains on dobutamine drip. later the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematoma: the cause of hematoma was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria/pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4516 | 1220648-2025-47297 | 23241375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23241375&pc=OZD | 09/07/2025 |  |  | 10/07/2025 |  |  | 09/07/2025 | 10/07/2025 |  |  |  |  | 10/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707249 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Inflammation | Ischemia | Pain | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient complained of soreness around the insertion site. no bleeding or hematoma. the following day, the insertion site had mild swelling, but improved per patient. improving right hand numbness, likely secondary to intra-operative brachial plexus stretch. no evidence of hematoma. the patient remained on support until successful explant and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. pain/ inflammation : the cause of the injury was not determined due to insufficient clinical details. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4517 | 1220648-2025-31821 | 23220125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220125&pc=PYX | 06/14/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/14/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578243 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 60 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella rp pump for support in a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced several placement signal alarms. there was a differential sensor failure. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. the data log shows an upward placement signal drift for about 30 hours, with an observed opposite drift in pump flow. none of the optical signal parameters were affected during the time of the drift. the cause of the placement signal issue was most likely sensor drift due to data log analysis showing placement signal drifting down while motor current drifting up. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4518 | 1220648-2025-31823 | 23220154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220154&pc=OZD | 06/16/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/16/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1887775 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 73 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp to support the surgical patient. the pump was explanted after 25 hours due to optical signal loss. the pump was replaced and support proceeded without harm or injury, and the automated impella controller was not replaced. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of the event are consistent with the complaint. the placement signal not reliable alarm was triggered shortly after the pump was turned on. within a few minutes of the alarm, the pump was unplugged and the automated impella controller (aic) was shutdown. console (b)(6) was returned to field service for further investigation. the reported complaint could not be reproduced. during inspecting the analog output of the charge-coupled device (ccd) from the optical bench, intermittent distortions in the signal (envelope/fringe) were observed with pump and without pump. furthermore, the "5s" parameters, particularly when tested with the test pump, showed an snr value below the limits specified in the preventative maintenance procedure. however, when tested with the test cavity, the snr value were within the specified range. the root cause for the placement signal not reliable alarm was due to the defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4519 | 1220648-2025-31822 | 23220298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220298&pc=OZD | 06/14/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/14/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671841 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 81 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal not reliable alarm occurred with their impella cp. echocardiography confirmed proper positioning. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. data logs were reviewed, and the placement signal not reliable (psnr) alarm was confirmed. at this time, signal-to-noise ratio (snr) drops to zero. the pump remained in use during psnr. optical signal parameters and placement signal returned after 20 hours of case run. upon reviewing data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4520 | 1220648-2025-31824 | 23220347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220347&pc=OZD | 06/16/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/16/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660922 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella 5.5 pump for support in a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced several placement signal alarms. placement monitoring was disabled. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. the data log shows a fiber break trend on all the 5s parameters, with a loss in placement signal accompanied by an active placement signal not reliable (psnr) alarm after 30 minutes of case start. the optical signal issue persisted until the end of the case for nine days. the cause of the optical signal issue was most likely a damaged optical fiber line, based on the data log review. the location of the damaged fiber line was not confirmed without returned product investigation or sufficient clinical information provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4521 | 1220648-2025-31825 | 23220442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220442&pc=OZD | 06/16/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/18/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026670298 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cardiomyopathy patient was on impella 5.5 support when the intensive care unit staff reported that the placement signal was suddenly gone. there were no alarms on automated impella controller screen, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. neither product nor data logs were available for investigation. since neither product nor data logs were available for investigation, the cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4522 | 1220648-2025-31826 | 23220491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220491&pc=OZD | 06/20/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/20/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583407 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 122 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | AUTOMATED IMPELLA CONTROLLER | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the use of an impella cp pump for support in a patient with acute myocardial infarction and cardiogenic shock. the patient was on multiple high dose pressors. during support, the pump experienced placement signal alarms. no issues with patient support without placement signal. ultimately care was withdrawn, and the patient expired. | [5 empty] | Added h6 coding. d10 concomitant medical products and therapy dates updated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. data log review shows signal-to-noise ratio (snr) step down to 2000 just before being disconnected from first automated impella controller (aic). placement signal not reliable alarms were triggered and the placement signal was lost for the rest of the case. the pump was not returned. the cause of the optical signal issue was not determined since data log review is inconclusive without returned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4523 | 1220648-2025-31827 | 23220532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220532&pc=OZD | 06/21/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/21/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593563 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 145 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a placement signal issue with their impella 5.5. placement signal numbers and waveforms and left ventricle signal number and waveforms became distorted (abnormally high) and then waveforms and pressures disappeared, and placement signal not reliable alarm appeared. the computed tomography surgeon did come to bedside with echocardiography and pump was repositioned, but it did not resolve the problem. however, support continued with the pump. the pump was eventually successfully weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no pump was returned. data logs were returned. no impact to pump flow, motor current and 5s parameters were seen during this time. the cause of the optical signal issue was unable to be determined since no product was returned and based on inconclusive evidence per clinical and log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4524 | 1220648-2025-31828 | 23220568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220568&pc=OZD | 06/21/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/21/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601512 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 99 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that while a patient was on impella cp support, a nurse noted intermittent "placement signal low" (ps low) alarms. the arterial line readings did not correlate with the placement signal. the device position was confirmed via echocardiography. a review of the automated impella controller revealed several intermittent "placement signal not reliable" (psnr) alarms, which self-resolved. placement monitoring was subsequently disabled. there was no harm to the patient. the impella cp was eventually explanted as part of a bridge-to-transplant strategy. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for analysis. the data logs confirm placement signal not reliable alarms and placement signal low alarms and show signal-to-noise ratio (snr) averaging 1000s with stable values for lamp and gain. the placement signal at the time of the alarms was noisy suggesting possible patient condition issue. the root cause of the optical signal issue was not determined as no identifiable data log abnormalities were observed, issue resolved on its own, limited clinical information was provided and no product was returned for analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4525 | 1220648-2025-47243 | 23220576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220576&pc=PYX | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026647323 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 127 KG | Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with pulmonary embolism was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the patient presented to the hospital for evaluation of schizophrenia, suicidal ideation, acute kidney injury, sepsis, and non-st elevation myocardial infarction. during the evaluation of the patient, the patient decompensated and required cardiopulmonary resuscitation and intubation to resolve. imaging revealed a severely enlarged right ventricle with severely reduced global systolic function. the patient was taken to the cardiac catheterization laboratory for impella rp flex insertion. on the second day of impella rp flex support, the patient experienced bloody urine, however it was noted that the cause was being assessed for urinary tract trauma and there were no impella issues. the following day, the patient was still experiencing bloody urine output and was transfused with a total of four units of red blood cells. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. the patient was noted as survived at time of explant. | [5 empty] | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4526 | 1220648-2025-31829 | 23220588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220588&pc=OZD | 06/25/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/25/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542520 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female | 69 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. the pump lost its optical signal but remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump was returned for investigation. no physical damage was noted on the returned product. the pump passed the laser continuity test for the optical signal line, and all the 5s parameters were within specification during optical bench testing with signal-to-noise ratio (snr) being around 10,000. the data log shows that a "placement signal not reliable" alarm occurred from the start of pump run. no major changes were noted in the light, lamp, and gain optical parameters. other pump used on console ic8090 were checked and found to have similar low snr readings from the start of the case. the cause of the optical signal issue was not determined since it could not be reproduced on the returned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4527 | 1220648-2025-31830 | 23220676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23220676&pc=OZD | 06/26/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 06/26/2025 | 10/06/2025 |  |  | 01/31/2026 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461381 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 136 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella cp pump to support a patient with acute myocardial infarction complicated by cardiogenic shock. the patient was initially on cp + extracorporeal membrane oxygenation (ecmo) support and was later escalated to an impella 5.5 pump. after 6 days of support, an optical signal loss alarm occurred. the pump was neither removed nor replaced. subsequently, the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product returned for analysis. the data logs confirm placement signal not reliable alarms during support. the logs show an increase in gain and contrast and an abrupt drop in signal-to-noise ratio (snr) as placement signal began reading before pump initiation. the root cause of the optical signal issue was unable to be definitively determined as no clear log pattern was identified and no product was returned with limited clinical information provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4528 | 1220648-2025-46987 | 23221040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23221040&pc=OZD | 08/29/2025 |  |  | 10/06/2025 |  |  | 08/29/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661708 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was transferred from an outside hospital and coded on arrival. an impella cp was placed and the decision was made to escalate to an impella 5.5 due to increased requirements. the patient was continued on venoarterial extracorporeal membrane oxygenation (va ecmo) and impella cp during the impella 5.5 implant. the patient's vessels were very calcified, not allowing for pump advancement into the aorta. the vascular team performed an endarterectomy in the operating room, per the physician's request. the endarterectomy was performed with stenting of the axillary artery. the graft was applied to the vessel distal to the new stent placement. an axillary sheath was used and the left ventricle access was gained with a 0.035 wire and pigtail catheter. the impella 5.5 was backloaded over the 0.018 wire, and met resistance when attempting to advance the pump into the aorta. several different wires were used as rails and viper slides. another physician was contacted to help with ballooning the vessel. despite all of these interventions and many attempts, they were unsuccessful in advancing the impella 5.5 down into the left ventricle. on fluoroscopy, the physician determined the patient's vessels were very diseased and did not want to continue forcing the pump into the vessel, potentially causing further complications. the impella 5.5 implant was aborted due to patient anatomy. the plan forward, which was executed noted use of an impella cp for left ventricle unloading and va ecmo. the patient would, however, expire approximately four days later on the impella cp device. | [5 empty] | Medical safety reviewed the event. the delivery issue/unable to deliver due to the patient's vessel condition prior to therapy involving the impella 5.5 is a malfunction and should be reported despite knowing the patient's anatomical incompatibility prior to implant. the demise outcome is unrelated to the impella 5.5. the patient was continued on venoarterial extracorporeal membrane oxygenation (va ecmo) and impella cp during the impella 5.5 implant. a new impella cp was placed for left ventricular unloading was placed with va ecmo. complications of arrhythmic events and suction were experienced. the patient expired on day four of support. there is not enough information to dissociate the impella cp in at the time of the demise and therefore, will be reported as death. however, the patient's underlying condition contributed most to this outcome of death. the patient was critically ill prior to therapy and there was down time on va ecmo which caused rapid deterioration of the patient. prognosis was "poor" and the family chose to withdraw care which led to the patient expiring. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding unable to deliver or advance, the cause was determined to be patient condition as the patient had the vessels very diseased preventing successful delivery of impella. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Patient-specific information section a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp, which is associated with the demise outcome. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
4529 | 1220648-2025-47244 | 23221267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23221267&pc=OZD | 08/27/2025 |  |  | 10/06/2025 |  |  | 08/27/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612914 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 78 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was escalated to an impella 5.5 pump from an impella cp pump for mechanical circulatory support after experiencing complications. it was reported that the patient presented to the emergency room with shortness of breath symptoms. the patient was found to have a significant clot burden in his right atrium and imaging indicated there was a possible thrombus in the pulmonary artery (pa). the patient¿s right ventricle was strained, and a reduced systolic function was observed. additionally, observed left ventricle septal flattening was consistent with the patient¿s volume and pressure overload. the initial ejection fraction was determined to be approximately 40%, and a tricuspid annular plane systolic excursion of 1.2 cm. the patient became progressively short of breath and computed tomography (ct) was performed to confirm the pa thrombus. the patient was taken to the operating room for an open chest thromboembolectomy of the main pa. despite the successful procedure, the patient¿s cardiac function progressively weakened. after returning to the intensive care unit (icu), the patient began to decompensate with vasoactive intravenous drips actively being increased. the patient was then cannulated for veno-arterial extracorporeal membrane oxygenation (va ecmo). repeat imaging revealed severely reduced bi-ventricular function and worsening septal dyskinesis. the decision was then made to take the patient to the cardiovascular laboratory (cvl) to assess coronary function and to place an impella cp for left ventricular unloading in the presence of va ecmo. in the cvl, the patient¿s lateral circumflex femoral artery was accessed, and prior to impella cp insertion the physician decided to assess the coronary arteries. the physician noted there was some bridging in the left anterior descending artery, but timi 3 flow also noted distal to the lesion and the decision was made to not intervene at that time. the impella cp was implanted, and support initiated without complications. the patient was released back to the icu for rest and recovery. upon arrival to the icu on va ecmo and impella cp support, the physician reviewed the placement of the impella cp. some imaging indicated that the impella was quite shallow, however all waveforms were appropriate, with no signs of hemolysis, and no suction events. the physician therefore made the decision not to attempt to reposition the device. the patient was able to be weaned off of intravenous epinephrine drop, however when milrinone was started, the patient experienced increased ectopy, and the medication was discontinued. the following day it was reported that the va emco flows had decreased to 3l, and with the decrease of flows the patient¿s blood lactate results increased from 3.2 mmol/l to 4.0 mmol/l. the patient also became hypotensive. repeated imaging revealed the impella shallow at 2.5 cm, but the physician declined to reposition the pump. it was noted that the patient was placed on continuous renal replacement therapy the previous night. additionally, a left lower extremity ultrasound was performed, and it revealed limited flows distal to the impella insertion site. the medical team consulted a vascular surgeon to determine to escalate the patient to an impella 5.5 or to add a distal perfusion catheter to reestablish flows down the leg. the decision was made to escalate the patient to an impella 5.5. the impella 5.5 was implanted without issue, and the impella cp was explanted by a vascular surgeon. it was noted that the patient was hemodynamically stable after arriving back at the icu. it was reported that on the first day of impella 5.5 with va ecmo support, the patient was noted for potentially having north south syndrome. the impella flows were reduced, and it was noted that the patient¿s arterial blood gas results and patient color improved. imaging was performed to confirm placement of the impella 5.5 was optimal. it was noted that prior to impella 5.5 placement, the patient¿s ejection fraction was determined to be 40-45%, with noted moderate right ventricular systolic dysfunction. a head computed tomography (ct) revealed small acute or subacute middle cerebral artery infarcts. additionally, it was noted that the patient may have small bowel ischemia on ct, however the bowels were still moving. it was also noted that the pa thrombi were significantly reduced. the following day, the patient was removed from sedation, and it was noted that the patient had no gag reflexes, cough, or withdrawal from pain. the next day it was noted that repeat echocardiography from the previous day confirmed impella was well placed. it was also noted that the patient had a severely reduced left ventricular systolic function, and the ejection fraction was unable to be calculated. later the same day, the decision was made to withdraw care. the patient¿s care was withdrawn and the patient expired. this complaint involves two impella devices: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety reviewed the event. the event with the impella cp describes shallow device positioning (x2) where the physician made the decision not to reposition as there was no hemolysis or suction. subsequently distal limb perfusion concerns were noted with the impella cp, so the decision made to escalate to impella 5.5 axillary. thereafter escalation to the impella 5.5, the patient became hemodynamically stable. the event occurring with the impella cp is a serious injury and should be reported. the following day after start of impella 5.5 mechanical circulatory support (mcs), the ct scan revealed acute and subacute mca infarct; physicians unconcerned; patient expired two days later. there is not enough information dissociate the impella 5.5. however, the infarct is most likely unrelated to the impella 5.5 due to the fact that preecmo (extracorporeal membrane oxygenation) insertion there was identified clot in right side of the heart; presumed pulmonary embolism. the patient's underlying condition likely contributed most to an outcome of death. the patient had severely reduced left ventricular systolic function, and care was withdrawn which led to the patient expiring. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the stroke/ischemia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Although the patient's presenting condition may be more likely to have impacted the event the impella 5.5 pump cannot be excluded as a possible contributing factor in the patient's demise. product status: the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4530 | 1220648-2025-47245 | 23221612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23221612&pc=PYX | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  | 01/31/2026 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517109 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella rp flex had very low flow while supporting a patient so tissue plasminogen activator was added to the purge. upon explant of the impella rp flex, biomaterial was observed. | [5 empty] | The device was returned for evaluation. the data log review confirmed the complaint. the cause of the placement signal issue/low pump flow was most likely sensor issues caused by sensor drift. the device passed all post sterile inspection criteria. | Product complaint # (b)(4). the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4531 | 1220648-2025-31831 | 23222140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222140&pc=OZD | 05/28/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 05/28/2024 | 10/06/2025 |  |  | 02/28/2026 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467754 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump exhibited an optical signal issue. after 14 days of use, on the day of explant, the pump¿s trending screen indicated a drop in cardiac output. there was no associated hemodynamic change. the pump was explanted, and the patient experienced no reported harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for investigation. no physical damage was observed on the returned product. the pump passed the laser continuity test, and all the 5s parameters were found within the specification range. the pump was run without any reported issues in a bucket of water. the data log shows the placement signal trend and pump flow trend during the reported event of discrepancy in cardiac output on the last day of support. there was no issue noted with the optical sensor parameter during the field run. additionally, there were no positioning alarm, suction alarm, or motor current spikes reported in the data log. the cause of the optical signal issue was most likely related to the patient's condition, as the trending placement signal and motor current could affect the cardiac output calculation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4532 | 1220648-2025-31832 | 23222401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222401&pc=OZD | 01/31/2023 | 10/06/2025 |  | 10/06/2025 |  |  | 01/31/2023 | 10/06/2025 |  |  | 10/31/2024 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268129 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the measured left ventricle signal waveform was outside of the expected range. this indicates an optical signal issue. there was no noted harm to the patient as a result of the failure. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. the lt log shows suction alarms and impella position unknown alarms around the time that the optical signal issue was reported. all the 5s parameters are within specification during the case. the cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4533 | 1220648-2025-31833 | 23222469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222469&pc=OZD | 02/09/2023 | 10/24/2025 |  | 10/06/2025 |  |  | 02/23/2023 | 10/06/2025 |  |  | 03/31/2024 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023210194 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported optical placement signal issues that occurred during impella support. the alarm was subsequently disabled on the console and patient support continued. there was no reported harm to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the data logs were reviewed and show that after insertion motor current pulsatility is extremely low, and placement signal shows aortic waveforms. after approximately 1 minute pulsatility returns, flow increases and the alarm resolves. this is most likely the pump being in the incorrect position, making this a positioning issue. the root cause of the positioning issue was determined to be use related. all pertinent information available to abiomed has been submitted at this time. | G3: date corrected. | D6a: corrected the implant date. d6b: corrected the explant date.
4534 | 1220648-2025-31834 | 23222527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222527&pc=OZD | 02/15/2023 | 10/06/2025 |  | 10/06/2025 |  |  | 02/16/2023 | 10/06/2025 |  |  | 07/31/2024 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257902 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. the pump reported disappearance of left ventricle optical placement signal during support. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data was reviewed and the sudden shift in optical signal metrics align with a sensor break failure. the returned pump was connected to an console and the placement signal not reliable alarm reproduced. the root cause of the optical signal issue was a damaged sensor. there was likely an interaction with another device that led to the sudden sensor damage. all pertinent information available to abiomed has been submitted at this time.
4535 | 1220648-2025-31835 | 23222663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222663&pc=OZD | 05/15/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 05/15/2024 | 10/06/2025 |  |  | 02/28/2026 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473967 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable error appeared during impella 5.5 support. the audio was disabled and support with the implanted pump was maintained. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella 5.5 pump was not returned for investigation. data logs were reviewed, and multiple events of optical signal failure was observed. on 05/10/24 snr and contrast were decreasing while gain was increasing. no placement signal issue alarms were found. later all 5s optical signals automatically resolved after 90 hrs without any troubleshooting indicating likely air gap failure. on 05/15/24 placement signal not reliable (psnr) was observed with snr decreasing, contrast oscillating and gain, light decreasing. the cause of the optical signal failure was most likely an air gap from between the automated impella controller (aic) and the patient cable plug per log review. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4536 | 1220648-2025-31836 | 23222715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222715&pc=OZD | 06/11/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 06/11/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482756 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the device exhibited an optical placement signal not reliable alarm. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for evaluation, and the root cause is unable to be determined. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4537 | 1220648-2025-31837 | 23222794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23222794&pc=OZD | 02/11/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 03/11/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1631552 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during therapy, the console experienced loss of signal, but no change in motor current was noted. the audio was disabled. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint. console was tested in danvers but the issue was not reproduced. the root cause of the console issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4538 | 1220648-2025-31839 | 23223126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223126&pc=OZD | 10/11/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 10/11/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1151533 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Death | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella cp support. during support, a 'placement signal unreliable' alarm was present. automated impella controller (aic) was not replaced. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the placement signal not reliable alarm from the complaint was identified in the data logs. console was not returned. the cause of the placement signal not reliable (psnr) was not determined since console was not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4539 | 1220648-2025-31840 | 23223188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223188&pc=OZD | 02/14/2023 | 10/06/2025 |  | 10/06/2025 |  |  | 02/15/2023 | 10/06/2025 |  |  | 04/30/2023 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022079874 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Tachycardia | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal not reliable alarm during impella 5.5 support. upon observing the reading, the signal was noted to be off axis. the alarm was disabled, and the physician was notified of the alarm. it was documented that improvements were seen following the occurrence on lower p levels and that the flows remained adequate. multiple ventricular tachycardia codes occurred overnight, and it was reported that the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was most likely biomaterial since temporarily dropping p-level resolved the placement signal waveform. however, the root cause was unable to be determined. all pertinent information available to abiomed has been submitted at this time.
4540 | 1220648-2025-31841 | 23223364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223364&pc=OZD | 12/26/2023 | 10/06/2025 |  | 10/06/2025 |  |  | 12/26/2023 | 10/06/2025 |  |  | 10/31/2025 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415436 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced persistent placement signal issues, and the alarm was disabled by staff. despite this, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no pump returned. data logs shown alarms with many suction alarms. the root cause of the optical signal issue was not determined since no product was returned for investigation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4541 | 1220648-2025-31842 | 23223440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223440&pc=OZD | 08/10/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 08/23/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512696 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support, the impella placement signal readings were low compared to the patient's noninvasive cuff pressure and arterial line. with repeat alarms and false pressure readings, the placement signal and suction alarms were turned off. positioning was verified via echocardiography, and support continued. it was reported that care was withdrawn and the patient died. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. data logs were returned for review. lt logs showing placement signal (ps) steadily declining for +50 mmhg to greater than -50mmhg. several suction alarms, impella position unknown alarms and ps low alarms were found. a0-ps signal decoupling with suction indicated and optical 5s particularly signal-to-noise ratio, contrast are widespread and variable during whole case. no major issues with gain, lamp & light. rt log on (b)(6) showing several suction alarms with aorta-ps decoupling. pulmonary arterial pressure (pap) and central venous pressure (cvp) are lost whenever suction alarms are active. rt log on (b)(6) morning with ps low/impella position unknown alarms. aorta-ps decoupling and pap, cvp lost for long time with suction/impella position unknown. imc logs are not available for the complaint timeline with second console. the root cause of the optical signal issue was not determined since product was not returned and due to inconclusive evidence per log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4542 | 1220648-2025-31843 | 23223620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223620&pc=OZD | 06/03/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 07/09/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669852 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the pump experienced a placement signal was lost after inflation of a shockwave balloon. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. during shockwave inflation, placement signals were lost and 'placement signal not reliable' alarm appeared. distance of pump from shockwave was not specified. no data logs were returned. the root cause if the optical signal issue was most likely due to a damaged sensor due to shockwave interaction as evidenced by the clinical detail confirming shockwave use. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4543 | 1220648-2025-31844 | 23223654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223654&pc=OZD | 06/22/2023 | 10/06/2025 |  | 10/06/2025 |  |  | 06/22/2023 | 10/06/2025 |  |  | 04/30/2025 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349539 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp support ongoing when the automated impella controller showed a false in aorta alarm. the placement signal alarm was disabled, and support continued. the pump was weaned and explanted, but the patient eventually expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed: data review: data showed 16 alarms 128 (impella position in aorta) were triggered within 36 hours & 39 minutes since the pump started then disappeared. there was a spike in motor current in the data logs indicating biomaterial interaction. all 5s parameters were in specification (see check list). product analysis: product was not returned for review. conclusion: the cause of the false position in aorta alarm was most likely biomaterial since proper position was confirmed and a spike in motor current was confirmed in the data logs. all pertinent information available to abiomed has been submitted at this time.
4544 | 1220648-2025-47247 | 23223677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223677&pc=OZD | 09/06/2025 | 11/05/2025 |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 97 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp had hemolyzed labs and dark red urine. while transferring the patient the impella became dislodged into aorta. the impella was explanted and the patient lost distal pulses, requiring a fasciotory and repair of femoral artery. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the thrombosis and ischemia was unable to be determined due to insufficient clinical details. the cause of hemolysis was most likely related to the patients condition, as pressure is seen trending down at the time of hemolysis, and clinical reported a high afterload the night before it was reported. the cause of the positioning issue was most likely user related, as the pump came out of position while the patient was being moved. the device passed all post sterile inspection criteria.
4545 | 1220648-2025-31845 | 23223801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223801&pc=OZD | 10/02/2023 | 10/06/2025 |  | 10/06/2025 |  |  | 10/02/2023 | 10/06/2025 |  |  | 05/31/2025 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356563 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the impella cp had continuous placement signal low alarms during support. the impella aorta placement signal does not correlate with the arterial line. a bedside echocardiogram confirmed placement to be approximately 4.0 cm in the left ventricle. the alarm was disabled and support continued. the patient eventually expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data logs show multiple occurrences of low placement signal alarms throughout the entire case run while the pump was set to a higher p-level. all 5s optical parameters were within specification limit during time of use. the cause of the optical signal issue was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
4546 | 1220648-2025-47249 | 23223809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23223809&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/08/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713536 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing for the patient admitted in with cardiogenic shock and cardiomyopathy. on the 4th day of the support there was high purge pressure and purge system blocked alarms. the team added thrombolytics to the purge and the pump remained without any issues. the patient was explanted successfully and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | G1 added manufacturing site name and address as it was omitted from the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ investigation summary: the investigation has been completed. no product or logs were returned for evaluation. high purge pressure: clinical details noted that the purge system blocked alarms were noted on day two of support. tissue plasminogen activator was administered and purge pressures and purge flows were able to resolve issue. the cause of the high purge pressure could not be determined as no product or logs were returned for evaluation. device history review: the pump passed all post-sterile inspection checks.
4547 | 1220648-2025-31846 | 23225481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23225481&pc=OZD | 04/19/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 07/25/2025 | 10/06/2025 |  |  | 10/31/2025 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407936 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp support ongoing for a patient. the pump support was ongoing with placement signal loss. there was no intervention for the optical signal. the family made the choice to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. upon review of the data logs, a definitive automated impella controller failure pattern was suspected. no problem was found with the pump. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4548 | 1220648-2025-31847 | 23225922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23225922&pc=OZD | 05/02/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 05/02/2024 | 10/06/2025 |  |  | 11/30/2025 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423348 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an impella cp implant and there were placement signal issues. the placement signal monitoring was disabled due to false impella in aorta alarms. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. according to the clinical, on the second day of impella cp support placement monitoring was disabled due to multiple false impella in aorta alarms. no product was returned for a visual inspection. data log analysis confirmed the presence of multiple false impella in aorta alarms followed by placement monitoring disabled alarms. no motor current spikes were observed and all 5s parameters appeared to be stable. the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4549 | 1220648-2025-31848 | 23226420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23226420&pc=PYX | 01/16/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 01/16/2024 | 10/06/2025 |  |  | 11/30/2025 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024434401 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp flex support and a placement signal not reliable alarm appeared. the physician attempted to pull back and reposition the impella, however the placement signal was still absent. the staff monitored the motor current, impella flows, and the patient's hemodynamics. patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. according to the clinical, hours into impella rp support the pa placement signal disappeared and a placement signal not reliable alarm was recorded. an attempt was made to reposition the pump with no success. as flow was still visible the device was not removed and the patient later expired on support. no product was returned for a visual inspection. data log analysis confirmed the following trends: lower signal not reliable, lower light, higher lamp, lower contrast with an increase in amplitude, and higher gain. the most likely cause of the optical signal issue was a damaged fiber line due to a kink that occurred during support. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4550 | 1220648-2025-31850 | 23227018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23227018&pc=OZD | 03/24/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 03/24/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608483 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm was disabled during impella cp support. the performance of the pump remained unaffected, and support continued with the implanted device. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were reviewed and lt log at end of case showed placement signal not reliable (psnr) alarm was triggered. 5s optical parameters showed oscillating contrast, with a drop in snr, light, and gain, and lamp maxed out at 100%. at time of psnr alarm and oscillating contrast, placement signal went out of range. clinical details noted that a psnr alarm was triggered towards end of case and placement signal went out. the cause of this problem cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4551 | 1220648-2025-31849 | 23227180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23227180&pc=OZD | 02/15/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 02/15/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584595 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 93 KG | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella in aorta alarm appeared during impella 5.5 support. the device exhibited optical placement issues. pump positioning was verified to be accurate, and support continued uninterrupted. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs confirm impella position in aorta alarms, and there were no motor current spikes prior. the 5s parameters were within specification and there were no data log abnormalities. the root cause of the optical signal issue was most likely patient condition causing false alarms. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4552 | 1220648-2025-31851 | 23227689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23227689&pc=OZD | 04/19/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 04/29/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1682677 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp support ongoing with an automated impella controller for a patient admitted with multiple cardiac and systemic comorbidities. the pump support was ongoing with placement signal loss. there was no intervention performed for the optical signal. the family made choice to withdraw care, and the patient eventually expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the automated impella controller issue was a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4553 | 1220648-2025-47251 | 23229536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229536&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710230 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed a fever and was unable to follow commands. the patient received a transplant and was weaned from the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 revised selection as the selection entered on the initial report was entered incorrectly. b.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. investigation summary: fever: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4554 | 1220648-2025-47252 | 23229540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229540&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712131 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 58 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced abdominal bleeding. the patient was treated with transfusions of volume and blood products, and systemic heparin was withheld. the patient¿s family decided to withdraw care and he patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks.
4555 | 1220648-2025-47253 | 23229543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229543&pc=PYX | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025610741 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp developed a fever and was unable to follow commands. the patient received a transplant and the pump was explanted successfully. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the clinical symptom was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4556 | 1220648-2025-47255 | 23229546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229546&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026651631 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of impella 5.5 the patient developed ventricular tachycardia (vt), started cardiopulmonary resuscitation (cpr), 6 defibrillations, the impella was reduced to p2, after stabilizing the patient, p level back to p7. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia :the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4557 | 1220648-2025-47254 | 23229548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229548&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686619 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 89 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced hemolysis and hypotensive episodes. in response, the pump was repositioned. the patient was treated with pressors in response to hypotension. oozing was observed and suction alarms occurred. the pump position was confirmed to be optimal. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned. the cause of the hypotension was not determined due to insufficient clinical details. the cause of the minor bleed was not determined due to insufficient clinical details. the cause of hemolysis cannot be determined since insufficient clinical details were provided. the cause of positioning issue cannot be determined since insufficient clinical details were provided. the cause of pump suction cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks.
4558 | 1220648-2025-47256 | 23229550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229550&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712333 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp the patient had some oozing both groins and patient having frank. manual pressure held and decision by faculty to hold heparin for now. spoke with fellow and rn. aortic diastolic on arterial line correlating with automated impella controller aortic diastolic pressure. echocardiogram showed stable position of impella at 3cm. the audio for placement signal alarm was silenced. the patient remained on support for an additional day with successful weaning and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematuria: the cause of the injury was not determined due to insufficient clinical details. asae: the cause of the access site adverse event could not be definitively determined as limited clinical details were provided. optical signal issue: data logs were reviewed and a persistent "placement signal low" alarm was noted. no other alarms related to placement signal were noted. no motor current spikes or trends in placement signal observed at time of alarms. the cause of the optical signal issue could not be determined as limited clinical details were provided and no product was returned for evaluation. device history review: the device passed all the post sterile inspection checks.
4559 | 1220648-2025-47257 | 23229553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229553&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716865 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp had suction alarms while supporting a patient. the impella p-level was reduced. additionally, the patient developed gastrointestinal bleed. the patient received two units of blood. suction alarms resolved since receiving blood. heparin was also stopped. the patient was brought in the operating room for planned escalation to impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product received for investigation. major bleed: the cause of the bleed was not determined due to insufficient clinical details. pump suction: data log review confirms suction alarms triggering in the case. placement signal is seen slightly trending down from case start, briefly dropping to 50 mmhg. it was reported that suction resolved after volume was given, and placement signal (ps) is seen starting to go back up after the reported date. the cause of the suction was most likely related to the patients condition, as ps is seen trending down, and clinical reported suction resolved after blood volume was given. device history review: the device passed all the post sterile inspection checks.
4560 | 1220648-2025-47258 | 23229560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229560&pc=OZD | 09/06/2025 | 02/13/2026 |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026699496 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella 5.5 support. after encountering high purge pressure, purge pressures improved to 600s-800's. additionally, thrombus was suspected but not confirmed. alteplase was used as purge solution. the patient remained on support until successful explant and survived. | The complainant reported additional information upon request: "1. from what i recall, i believe the medical team decision to utilize alteplase in the purge resolved the concern. 2. no. 3. unsure - i have copied my area manager and other acc for this account if they have additional information.". |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 updated. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, thrombosis/thrombus, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. purge pressure high: the pump was not returned for investigation. data log shows a very slow, gradual rise in purge pressure from the start of the case. the purge pressure reached above 100 mmhg and was seen to resolve after administration of alteplase as the purge solution, with a gradual resolving trend. the rise in purge pressure did not result in any purge-related alarms, and no motor current spikes were observed during the case run. the cause of the purge pressure rise issue was most likely biomaterial buildup, based on data log review and clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4561 | 1220648-2025-47259 | 23229562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229562&pc=OZD | 09/05/2025 |  |  | 10/06/2025 |  |  | 09/05/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697870 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 34 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Pneumothorax | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation requiring defibrillation and resuscitation. the patient had possible tension pneumothorax requiring decompression, and had a fem-fem bypass. | [5 empty] | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. investigation summary: cardiac arrest/ pneumothorax : the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4562 | 1220648-2025-47261 | 23229571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229571&pc=OZD | 09/03/2025 |  |  | 10/06/2025 |  |  | 09/08/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568874 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump being placed as a bailout for the patient in poor condition in the percutaneous coronary intervention and revascularization. the pump was placed but the pump was unable to fully deliver at the reposheath due to due to tissue of adipose patient. the medical team decided than to close 1st approach with help of vascular surgeon simultaneously and to insert another impella cpsa 553825a femoral contralateral for continued hemodynamic support. once this was successfully done, patient was transferred to the intensive care unit (icu) for further stabilization, where icu medical doctors together with previous team of interventionalists decided to withdrew care due to hemorrhagic shock and multiorgan failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E1 added initial reporter postal code as it was omitted from the initial report that was submitted. g1 added manufacturing site name and address as it was omitted from the initial report that was submitted. h6 added code b13 as it was omitted from the initial report that was submitted. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ medical safety reviewed the event. the event describes after a planned procedure failed, the patient went unstable during 3-vessel revascularization, the impella cp implanted with plans to leave in post procedure. the physician was unable to insert the repositioning sheath after removal of the peel-away sheet due to adipose tissue, and placement of another impella cp was implanted contralaterally for continued hemodynamic support. same day on unit, the patient would expire as the physician's withdrew care due to hemorrhagic shock and multiorgan failure. regarding the impella cp pump 1 (b)(6) there is not enough information to dissociate from the demise outcome. however, the patient's underlying condition contributed most to the death outcome. per the physician, the patient expired from hemorrhagic shock and multiorgan failure unrelated to the impella devices. care was stopped by the physician which led to the patient's demise outcome. investigation summary: the investigation has been completed. unable to deliver or advance: the pump was not returned for investigation. data log review was not required for this case run. according to the clinical details, the repositioning sheath could not be inserted after removing the peelaway sheath due to the patient¿s adipose tissue. the doctor also suggested a possible defect in the repositioning sheath, which could not be confirmed without the returned product. therefore, the cause of the delivery issue was not determined without returned product investigation and sufficient clinical details. device history review: the pump passed all post sterile inspection checks. h6 type of investigation, investigation findings, conclusion codes and component code were updated accordingly based on the completed investigation.
4563 | 1220648-2025-47260 | 23229572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229572&pc=OZD | 09/08/2025 |  |  | 10/06/2025 |  |  | 09/08/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1798913 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical team at the hospital saw a console with water damage before use while the console was in the storage area. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4564 | 1220648-2025-47262 | 23229593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229593&pc=OZD | 09/08/2025 | 10/22/2025 |  | 10/06/2025 |  |  | 09/08/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710546 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had a impella cp pump placed and supporting the patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was supporting for 40 minutes when there was purge pressure was rising and the team replaced the pc as there was a 'kink' noted. this troubleshooting did not resolve the issue and so the team replaced the pump. there was no patient harm and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. the pump was returned for investigation. after 40 minutes of support, high pressure purge (pp) noted. troubleshooting included exchanging purge cassette with no resolution. pump exchanged for a new one and no further issues arose. no data logs recovered. purge pressure high - the cause of the purge pressure high was biomaterial blocking the purge gap. without data logs, it could not be determined if this was an ingestion event or a gradual build up. product damage (purge tubing kink) - the cause of the purge tubing kink could not be determined as the returned product showed no signs of kinking, data logs were not recovered, and insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
4565 | 1220648-2025-31852 | 23229903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229903&pc=OZD | 12/06/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 12/06/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573289 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed a 5.5 pump to support a patient as escalation from the cp. the pump was placed but lost optical signal. the lost of signal did not prompt explant. the pump remained on for another day til the patient had care withdrawn and expired. (death reported on sister record for the automated impella controller aic). | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed on 12/6. at this time, snr drops to zero, lamp maxes out at 100%, light decreases, gain decreases, and contrast oscillates between +/-3,200. the signal unreliability lasts for roughly 20 minutes, and then all optical signal metrics return to values within range for the remainder of the case. product was not returned for investigation. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4566 | 1220648-2025-31853 | 23229986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23229986&pc=PYX | 12/21/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 12/21/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025480399 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there were placement signal issues and flow readings were not possible. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. initial log review showed the placement signal drifting downward before going out of range. the pump flow went to zero and the cvp went out of range when the placement signal went out of range. motor current and cvp were stable otherwise. all optical parameters were stable. the product was returned and went under a dp investigation since the log review indicated an issue with the dp, not optical, sensor. there were no abnormalities with electrical testing. there was some blood surrounding sensor as well as scratches around the surrounding metal border. the pump was run in the pressure flow loop test, where the sensor drift reproduced within the first hour of running. the lab reproduction looked similar to the field logs with placement signal drifting downward before going out of range. the root cause of the placement signal issue was sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4567 | 1220648-2025-31854 | 23230065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230065&pc=OZD | 10/20/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 12/02/2024 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1648111 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 support for a patient. the 5.5 was placed, and after four days, the optical signal was seen to be malfunctioning. the loss of placement signal did not prompt any intervention or explant. in the pump (5.5) investigation, the team made the choice to investigate the automated impella controller, which was never replaced during the support. the pump was eventually weaned and explanted, and the patient survived the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ¿ the case associated with the reported complaint was started on 15 oct 2024 with the pump (520586). ¿ after 2 days and 22 hours since the pump was started, the console displayed ¿placement signal not reliable (opm invalid signal, optical pump connected]) ¿ alarm,¿ event # 1036. ¿ event #1036 was resolved and re-occurred several times until 21 oct 2024. this confirms the reported failure mode. ¿ during event #1036, the rt log confirmed the placement signal (ps) spiked to 3000 mmhg, signal-to-noise ratio (snr) dropped, and the contrast oscillated between -3272 and +3269. ¿ the pump was not moved to a different console, and the support was continued with ic (b)(6) up to 23 oct 2024 without a placement signal issue. a similar placement signal issue was observed in the case performed before the complaint date (pump sn # (b)(6) from 15 jul 2024 to 19 jul 2024. device analysis: this console was tested after arriving in field service, and obtained a noisy envelope with a dent [without connecting either the optical test pump or test cavity with aic] and a noisy fringe pattern with a dent [with the optical test cavity connected to the aic]. this confirms the defective optical bench. the snr was low (1960) when tested with the optical test cavity length 14340.1 nm. found the front panel optical connector contaminated, unable to remove the debris upon cleaning. the root cause of the placement signal not reliable alarm was the defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4568 | 1220648-2025-31855 | 23230101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230101&pc=OZD | 06/30/2024 | 10/06/2025 |  | 10/06/2025 |  |  | 06/30/2024 | 10/06/2025 |  |  | 01/31/2026 |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505737 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support, an error message ¿placement signal not reliable¿ occurred, and the values/curve from the placement signal are missing. the complainant also reported that the pump was firmly plugged in, but the white pump cable near the red impella plug was severely knotted and externally damaged. the pump was successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the priming issue and optical signal issue has been completed. the cause of the issue was not determined since neither product nor data logs were returned. all pertinent information available to abiomed has been submitted at this time.
4569 | 1220648-2025-47263 | 23230489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230489&pc=OZD | 09/06/2025 |  |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698659 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). during the support, the following day, it was noted the purge cassette was leaking. a replacement cassette was not initially available but eventually one was located and successfully changed without harm to the patient. support was uninterrupted and continued. two days later an update noted the patient was awake, situation stabilized over the weekend. no catecholamines were needed, lactate values were in range, and the impella cp pump ran without issues. however, the next day, the patient expired on support; care was withdrawn. further details surrounding the demise outcome were unnoted. | [5 empty] | The complaint involving the purge cassette is not part of this reportable event. however, rather, it is part of the overall event story, and the aware date is based on this conservatively. as there is limited information regarding the patient's underlying condition and circumstances surrounding the demise outcome other than care was withdrawn, and no alternative cause for the event was identified as a likely cause, the mdr is being conservatively submitted. patient-specific information a4: weight and a5: ethnicity is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 added b13 code as it was omitted from the initial report that was submitted. medical safety reviewed the event and noted the following: a patient presented in cardiogenic shock and was placed on impella cp for hemodynamic support. on the following day, a purge cassette leak was observed. the purge cassette was replaced without interruption to patient support. four days later, care was withdrawn, and the patient expired. there is no evidence to suggest that the purge cassette leak or the impella cp device contributed to the patient¿s death. the device functioned as intended following purge cassette replacement, and support was uninterrupted. device status. the impella cp device has still not been received from the customer. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product returned. post procedure adverse event (hemodynamic instability): the cause of the injury was not determined due to insufficient clinical details. device history lot device lot: 1943248. device history batch subcomponent lot: n/a. device history review pump s/n (b)(6) passed all post sterile inspection checks.
4570 | 1220648-2025-47264 | 23230494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230494&pc=OZD | 09/06/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716882 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Backflow | Pumping Stopped | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered a pump stop and placement signal issue. the team initially had set up the pump but when it was plugged in, it immediately went to impella stopped / retrograde flow. no aorta waveforms were noted on the automated impella controller (aic) via ic. the ccl team exchanged aic's and attempted again with the same outcome. when the registered nurse (rn) called, i was about to tell them to exchange the catheter, when it abruptly showed an aorta placement signal and alarmed, the impella was stopped and retrograde flow appeared. i advised the rn to start the pump at p4 to assess placement signal and left ventricle waveforms. normal waveforms and motor current were seen. they were then advised to go to auto and the pump initiated auto and performed as expected. the team was confused on the pump and catheter's functionality. | [5 empty] | D9, h3 updated. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data log review shows with impella stopped ¿ reverse flow alarms with purge metrics working, low snr, low contrast, high gain, no lamp/light issues, and all pump metrics zero. this shows the pump was being read but never turned on. the pump was then unplugged and plugged back in and the new log showed the pump running without issue, 5s parameters nominal. retrograde flow - their was no defect found with the retrograde flow alarms. data logs show them triggering correctly while the pump was experiencing air gap issues and not running. optical sensor issue - the cause of the optical sensor issue was an air gap confirmed by the data logs that resolved after plugging the pump back in. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown.
4571 | 1220648-2025-47265 | 23230500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230500&pc=OZD | 09/07/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584955 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during support on impella 5.5 on (b)(6) 2024, the patient started heparin per intensive care unit cts. was previously on eliquis for pulmonary embolism. lij thrombus seen when attempting central line. bicarbonate was added to purge fluid. plan to replace a-line, attempt adding pulmonary artery catheter. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. peri-procedural adverse event(thrombosis/thrombus): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4572 | 1220648-2025-47268 | 23230501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230501&pc=OZD | 09/06/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698772 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented with worsening shortness of breath and lower leg edema. decision for impella cp was made. during cp support it was noted that there was a position in aorta alarm with flat motor current. the team attempted repositioning decision made to explant pump. the cp was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4573 | 1220648-2025-47269 | 23230503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230503&pc=OZD | 09/06/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/06/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026725083 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 57 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella 5.5 experienced oozing at site, jackson-pratt drain was placed. in addition to bleeding at right axillary site, blood products administered. there were also runs of ventricular tachycardia and arrhythmia issues were addressed with amiodarone. the 5.5 was successfully weaned and explanted. | Additional information was received that indicated the patient was needing an off-pump coronary artery bypass (opcab). |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be use issue due to patient movement because the bleeding weas noted after transfer to icu. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4574 | 1220648-2025-47271 | 23230504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230504&pc=OZD | 09/07/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/07/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 99 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient with extracorporeal membrane oxygenation (ecmo) and impella cp support had transesophageal echocardiogram (tee) done. tee images were difficult to see but it appeared that the impella was in the papillary muscle. with great effort the intensivist tried repositioning but was unable to free the impella. the cp was left running at p1 level because a higher level, suction alarms would return. the physician was aware that the impella catheter would require to be run at a higher level but will rest the heart with ecmo for now. the patient was stable. | The device was successfully weaned. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrected information provided in b5 and d4. the investigation into the reported issue has been completed. no product was returned for evaluation. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: no logs are available for the events. the cause of pump suction cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks.
4575 | 1220648-2025-47272 | 23230508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230508&pc=OZD | 09/07/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/07/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713023 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed bleeding in the groin. a stitch was added to the groin and bleeding was resolved | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. major bleed : the cause of the injury is most likely use related since an additional stitch was added to the groin and bleeding was resolved. device history review: the device passed all the post sterile inspection checks.
4576 | 1220648-2025-47274 | 23230515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230515&pc=OZD | 09/07/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/07/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720733 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed bleeding at the impella access site. lidocaine with epinephrine injected was administered around the site with good effect. the site was redressed. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleeding; the cause of the bleeding is determined to be anticoagulation management because the activated clotting time (acts) were higher than the therapeutic range (act= 238) in icu and the oozing was managed by medication injected around site with good effect and redressing the site. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4577 | 1220648-2025-47275 | 23230516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230516&pc=OZD | 09/07/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/07/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738280 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support had bleeding and hematoma at the impella access site. the physician decided to insert a swan but there was still significant bleeding. the repo sheath hub wasn¿t well seated in the radiofrequency ablation. the repo sheath was pushed forward a little and sutures were added. the bleeding and hematoma were resolved. blood products were administered. additionally, the patient had ventricular tachycardia which led to cardiopulmonary resuscitation and cardioversion. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed/ hematoma: the cause of the bleeding and hematoma are determined to be use issue because e bleeding and hematoma resolved by pushing t repo sheath hub forward a little more and adding sutures. tachycardia/ cardiac arrest: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4578 | 1220648-2025-47270 | 23230523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23230523&pc=OZD | 09/07/2025 | 10/06/2025 |  | 10/06/2025 |  |  | 09/07/2025 | 10/06/2025 |  |  |  |  | 10/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026720444 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was a delivery issue during impella 5.5 patient implant. the impella 5.5 implant was aborted due to patient vascular anatomy. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy along with resistance in axillary artery preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4579 | 1220648-2025-47205 | 23218746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218746&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  | 12/31/2025 |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441812 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 54 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient with cardiomyopathy had impella 5.5 placed for recovery. on day two of support, arrhythmia with suction event occurred. the patient would get arrhythmia if the impella would go above the p4 level. the team was going to check how to control rhythm, but the patient expired that same day. | [5 empty] | Corrections: b1, e4. the investigation of the reported failure mode has been completed. no product was returned. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: after the review of logs, multiple suctions alarms were observed along with impella in aorta and ventricle alarms. no motor current spikes were observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4580 | 1220648-2025-47206 | 23218747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218747&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563138 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 59 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the physician was able to advance the impella 5.5 past the anastomosis but unable to advance 5.5 into the aorta. the physician attempted twice but inserting the 5.5 was unsuccessful. the 5.5 was removed and impella cp was placed. | [5 empty] | E.4 added selection as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anastomosis preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4581 | 1220648-2025-47207 | 23218748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218748&pc=OZD | 09/05/2025 | 01/27/2026 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026662884 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 104 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed while supporting a patient had placement signal low alarm. a-line was not correlating with placement signal. the audio was disable. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: optical signal issue: clinical reported placement signal low alarms and a-line not correlating with placement signal. the data logs confirm placement signal low alarms. motor current is stable throughout support and 5s parameters are stable. the first rt log shows placement signal calibrated at 20mmhg instead of 0mmhg. this error in calibration is likely a contributor to placement signal values being lower than expected. the cause of the issue was not established as no product was returned and limited information regarding pump set up and support initiation was provided device history lot. device lot number: 1914794. device history batch. subcomponent lot: n/a. device history review: the pump (b)(6) passed all post sterile inspection checks
4582 | 1220648-2025-47208 | 23218749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218749&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677189 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 178 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was bleeding, arrhythmia, tachycardia, and positioning issues during impella cp patient support. while on support, the patient¿s groin was oozing and required fem-stop placement to assist in prevention of oozing/bleeding. no blood products were administered or required. additionally, the patient had multiple position alarms. a bedside echocardiogram x2 revealed a deep pump (> 5cm). the doctor pulled the device back 2-3cm. the waveforms improved but then patient developed frequent premature ventricular contractions. the pump had to be repositioned/clocked once more to remedy the ectopy. waveforms and rhythm appeared to be stable. the patient was tachycardia at 126 beats per minute but he was cardioverted for a narrow-complex tachycardia with a rate over 180 beats per minute. the family elected the patient for comfort care and the patient expired. | [5 empty] | E.1 added zip code extension as it was omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.6 code e0506 should not have been reported on the initial report. investigation summary: the investigation has been completed. no product was returned. arrhythmia, tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4583 | 1220648-2025-47209 | 23218750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218750&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026735009 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during the percutaneous coronary intervention (pci) with impella cp support, the impella accidentally pulled into the aorta. the cp was unable to be re-advanced. a decision was made to complete pci without impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was accidentally pulled out of position into the aorta by the user. device history review: the device passed all the post sterile inspection checks.
4584 | 1220648-2025-47210 | 23218752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218752&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | IMPELLA CP SMART ASSIST |  |  | ABIOMED, INC |  |  | US |  | 2026700874 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient was admitted with ischemic cardiomyopathy and peripheral vascular disease. the pump was implanted; however, patient developed limb ischemia within four hours of implant. subsequently, pump was explanted due to the complication. patient condition stabilized after explant. | [5 empty] | Correction d4: the investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
4585 | 1220648-2025-47211 | 23218753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218753&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725520 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. impella cp was prepped and advanced with cessation of advancement occurring at the takeoff of the left vertebral artery. multiple attempts were made to cross the aorta but were unsuccessful. the impella cp insertion was aborted due to the patient¿s anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A.4 added as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks.
4586 | 1220648-2025-47212 | 23218754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218754&pc=OZD | 09/04/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690401 |  | 1000482 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Device Damaged by Another Device | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the complainant experienced a delivery failure while a 5.5 pump was being positioned over the abiomed 0.018 guidewire, due to an interaction between the guidewire and the pump that prevented the guidewire from loading properly. subsequently, a new pump and wire were installed, resulting in no further malfunctions. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9: the impella device was received for investigation. once investigation results are provided a supplemental report will be submitted. correction d1, d4. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. unable to deliver or advance/device interaction or damage by another device: clinical details note that the 0.018 wire was unable to be threaded through the 5.5. using a new wire did not resolve the issue. the pump was replaced and they had no issues. the pump was returned and evaluated. there were no abnormalities inside the cannula or cages indicating a blockage. there was dried dextrose observed on the impeller. a test wire was threaded through the 5.5 multiple times without issue. the cause of the issue was not determined since no defect was found on the returned pump and the issue could not be reproduced. device history review: the device passed all the post sterile inspection
4587 | 1220648-2025-47214 | 23218755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218755&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1185159 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male | 81 KG | Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the battery of an automated impella controller would not charge and so the aic was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4588 | 1220648-2025-47215 | 23218756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218756&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-0040 |  |  | ABIOMED, INC. |  |  | US |  | 1468353 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | Pumping Stopped | Application Program Problem | Unintended Application Program Shut Down | Priming Problem |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in high risk surgery/coronary artery bypass graft was implanted with an impella 5.5 device for mechanical circulatory support. the priming of the pump was unsuccessful because the automated impella controller (aic) experienced several alarms, including controller error, open purge system, impella stoppage, retrograde flow, and air in the purge. consequently, the aic was replaced to address this problem. the newly installed aic facilitated the operation and utilization of the 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction h8. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: the data logs align with the reported complaint. device analysis: the console was returned for further investigation; however, the reported issue could not be reproduced during testing. root cause: the root cause for the controller error (opm communication stopped) [optical pump connected] ¿ alarm issued, event #1043 and the inability of the automate impella controller (aic) to shut down was due to an aic software issue. the purge system open (#407) and air in purge (#404) alarms occurred because the purge cassette was removed while the aic still detected the pump as connected, primarily due to the same aic software issue. the impella stopped ¿ retrograde flow (#18) and impella stopped ¿ mechanical/electrical issue (#115) alarms were also triggered because the aic incorrectly detected the pump as connected, again linked to the aic software issue. device history review: the device passed all post sterile inspection checks. | Correction: d2, d6a, and h6 medical device problem code 1414 was erroneously entered on the initial manufacturer device report. | The product was returned therefore d9 was updated.
4589 | 1220648-2025-47216 | 23218757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218757&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732908 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 62 KG |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had placed a cp pump to support the patient admitted in with medical and cardiac history/comorbidities. the pump was placed but kinked attempting to cross the valve and the cp was replaced, as the kinked pump also had low flows. | [5 empty] | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. a small kink was observed in the returned product. low pump flow: after reviewing the logs, multiple suction alarms and impella in aorta alarms were observed. the cause of low pump flow was most likely due to cannula being kinked in relevance to patient condition (aortic stenosis). product damage (cannula kink): the cause of the product damage cannula kink was most likely patient condition related since it was due to aortic stenosis. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was received for investigation, d9 updated. once investigation results are provided a supplemental report will be submitted
4590 | 1220648-2025-47213 | 23218889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218889&pc=OZD | 09/05/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713759 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the placement of the impella cp was aborted because of the patient's anatomical structure. the patient presented with a calcified aorta. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to calcification in aorta. cp was unable to be advanced past aorta and insertion was aborted. device history review: the device passed all the post sterile inspection checks.
4591 | 1220648-2025-47217 | 23218890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218890&pc=OZD | 08/25/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689352 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 124 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with the impella 5.5 device for mechanical circulatory support (mcs) and experienced complications. the patient would ultimately expire.the patient was initially admitted with advanced heart failure and was subsequently transferred to the unit status post intra-aortic balloon pump for cardiogenic shock. the patient was brought to operating room for impella 5.5 placement and would continue with orthotopic heart transplantation (oht) work-up. the impella 5.5 was successfully placed, and the patient returned to the unit to recover. after approximately two and a half months of support, a sudden controller error alarm occurred, and an automated impella controller (aic)-to-aic transfer was completed. the patient continued on support. approximately four days later, intermittent placement signal not reliable (psnr) alarm occurred (with the impella 5.5 and second aic). the intermittent psnr issue self-resolved. additionally noted this day was the purge flow rate issue occurred and was closely monitored. the patient continued to wait for oht but would still have intermittent psnr. instructions were given to disable the alarm. two occurrences again (10 days and 15 days later), purge flow issues were encountered. tissue plasminogen activator (tpa) was discussed. however, it is unknown if tpa was added to the purge system. support continued however four days after the latest purge flow rate issue. however, the patient expired on support while waiting for heart and kidney transplant, despite being elevated to status one. | The complainant reported that there was placement signal issue during impella 5.5/automated impella controller patient support. the patient expired on support while waiting for heart & kidney transplant, despite being elevated to status 1. |  |  |  |  | Product complaint # (b)(4). | Correction: b5. event description and b7. medical history/preexisting condition were corrected to provide all relevant details regarding the complaint. g3 on the initial report that was reported should of reflected 25-aug-2025 as the date. h6 code b13 was omitted from the initial report that was submitted and was added. medical safety review: medical safety reviewed the event. the event describes an automated impella controller (aic) failure resulting in aic-to- back up aic transfer. thereafter on the second aic and with the impella 5.5 device, intermittent placement signal not reliable (psnr) was encountered. additionally involving the impella 5.5 device, there was high purge pressure and tissue plasminogen activator (tpa) was possibly added to the purge system. the patient would eventually expire on support while waiting for heart and kidney transplant, despite being elevated to status one. the first and second aic's had a device malfunction and did not contribute to the demise outcome. although the first aic failure resulted in temporary loss of support due to aic-to-aic exchange, this was a planned switch that occurred approximately 21 days prior to the patient's demise. regarding the intermittent psnr, positioning was confirmed and support continued. based on the information at hand, the patient expired for reason unrelated to the aics despite the product malfunction. for the impella 5.5 device, high pressure at first self-resolved but noted to have continued on three occasions. tissue plasminogen activator (tpa) was possibly added to the purge. however, there was no report or indication of inadequate support. mechanical circulatory support (mcs) support continued. in this case, the patient's underlying condition likely contributed most to an outcome of death. the patient required both heart and kidney transplant which did not materialized although at status one. conservatively, the death will be reported against the impella 5.5 device as this device was in use at time of expiration. the impella 5.5 cannot be ruled out but highly unlikely due patient condition. despite the purge concerns the pump remained supporting the patient. investigation summary: investigation has been completed. no product returned for evaluation. however, event logs were received and analyzed. high purge pressure: data logs were reviewed and long-term log at time of reported low purge flows did not show any "purge pressure high" alarms. purge flows were within normal range. clinical details noted that low purge rates were observed approximately two months into support. the cause of the high purge pressure could not be determined as limited clinical details were provided. no high purge pressure was observed in the data logs, and no product was returned for evaluation. optical: clinical details noted intermittent psnr alarms occurring throughout support. data logs were reviewed and long-term log of psnr alarms showed intermittent placement signal spiking to 3000 with variable/low signal-to-nosie ratio and contrast, and all other optical parameters were stable. no motor current spikes were observed. previous long-term log from the case using a different console showed the same pattern of optical parameters. the cause of the optical signal issue could not be determined as limited clinical details were provided and no product was returned for evaluation. device history review: the pump passed all post-sterile inspection checks. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device is one of three devices associated with the event. refer to section h10 for the related report numbers. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4592 | 1220648-2025-47218 | 23218891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218891&pc=OZD | 08/25/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114572 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella 5.5/automated impella controller patient support there were intermittent placement signal not reliable alarms with erratic placement signal waveforms. the patient expired which has been reported on the pump (serial number (b)(6)). | The complainant reported the patient was implanted with the impella 5.5 device for mechanical circulatory support (mcs) and experienced complications. the patient would ultimately expire. the patient was initially admitted with advanced heart failure and was subsequently transferred to the unit status post intra-aortic balloon pump for cardiogenic shock. the patient was brought to operating room for impella 5.5 placement and would continue with orthotopic heart transplantation (oht) work-up. the impella 5.5 successfully placed, and the patient returned to the unit to recover. after approximately 2.5 months of support, a sudden controller error alarm occurred, and an aic-to-aic transfer was completed. the patient continued on support. approximately four days later, intermittent placement signal not reliable (psnr) alarm occurred (with impella 5.5 and automated impella controller-aic #2). the intermittent psnr issue self-resolved. additionally noted this day was the purge flow rate issue occurred and was closely monitored. the patient continued to wait for oht but would still have intermittent psnr. instructions were given to disable alarm. two occurrences again (10 days and 15 days later), purge flow issues were encountered. tissue plasminogen activator (tpa) was discussed. however, it is unknown if tpa was added to the purge system. support continued however four days after the latest purge flow rate issue. however, the patient expired on support while waiting for heart & kidney transplant, despite being elevated to status 1. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b5. event description and b7. medical history/preexisting condition were corrected to provide all relevant details regarding the complaint. medical safety review. medical safety reviewed the event. the event describes an automated impella controller (aic) failure resulting in aic-to- back up aic transfer. thereafter on the second aic and with the impella 5.5 device, intermittent placement signal not reliable (psnr) was encountered. additionally involving the impella 5.5 device, there was high purge pressure and tissue plasminogen activator (tpa) was possibly added to the purge system. the patient would eventually expire on support while waiting for heart and kidney transplant, despite being elevated to status 1. the first and second aics had a device malfunction and did not contribute to the demise outcome. although the first aic failure resulted in temporary loss of support due to aic-to-aic exchange, this was a planned switch that occurred approximately 21 days prior to the patient's demise. regarding the intermittent psnr, positioning was confirmed and support continued. based on the information at hand, the patient expired for reason unrelated to the aics despite the product malfunction. for the impella 5.5 device, high pressure at first self-resolved but noted to have continued on three occasions. tissue plasminogen activator (tpa) was possibly added to the purge. however, there was no report or indication of inadequate support. mechanical circulatory support (mcs) support continued. in this case, the patient's underlying condition likely contributed most to an outcome of death. the patient required both heart and kidney transplant which did not materialized although at status 1. conservatively, the death will be reported against the impella 5.5 device as this device was in use at time of expiration. the impella 5.5 cannot be ruled out but highly unlikely due patient condition. despite the purge concerns the pump remained supporting the patient. device status: additional information noted the aic was received, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this impella aic device is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the first aic device and the impella 5.5, which is associated with the demise outcome. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. data analysis: review of data logs confirmed the reported failure mode. device analysis: the console was returned for further investigation. the reported complaint could not be reproduced, even after the console was operated with the pump and purge cassette for a couple of hours. additionally, a visual inspection of the internal circuitry showed no abnormalities. further inspection did not revealed any distortion in the optical bench analog signal. the optical bench 5s parameters were within specification. root cause: the root cause of the erratic placement signal waveforms and placement-related alarms could not be determined, as the issue could not be reproduced. device history review: the device passed all post sterile inspection checks.
4593 | 1220648-2025-47220 | 23218892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218892&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732917 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. on (b)(6) 2025, during support with the impella cp, the physician initially inserted a first impella but had to abort the procedure due to insufficient pump flows. the doctor then chose to insert a second impella cp pump, which had not been rinsed in sterile water. the valve was crossed, and the auto was activated with a flow rate of 3.4 l/min. following radial access for percutaneous coronary intervention (pci), the impella was slightly retracted and kinked at the atrioventricular (av) valve as a result of aortic stenosis. the physician was unable to readvance and unkink the device. rather than removing it and running it in sterile water, the doctor opted to attempt the insertion of a third impella cp pump. | [5 empty] | The investigation of the reported failure mode has been completed. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had aortic stenosis preventing successful delivery of impella. product damage (cannula kink): the cause of the product damage (cannula kink) was determined to be patient condition as the patient had aortic stenosis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4594 | 1220648-2025-47221 | 23218893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218893&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672832 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support on impella cp on (b)(6) 2025, the patient had suspected gastrointestinal bleed from dark tarry stool. bicarbonate purge. care was withdrawn of support and the patient expired. | [5 empty] | B.7 added information that was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | Given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
4595 | 1220648-2025-47222 | 23218894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218894&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685582 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had cp placed after the patient condition had deteriorated after non-st-segment elevation myocardial infarction. the patient had a complex history inclusive of coronary artery bypass graft kidney transplant, and cardiomyopathy. after pump placed the patient had ventricular tachycardia and was shocked x3 and patient was sent to the intensive care unit on support (and with extracorporeal membrane oxygenation) where a gi bleed was observed. the bleed was seen with drop in h&h labs and red blood cells unit x1 was given. the patient expired after being placed on palliative care on day 6 of the support. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. ventricular tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: it was reported that patient had unknown source of bleeding. the cause of the bleeding is determined to be patient condition because the location of the bleeding is unknown. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4596 | 1220648-2025-47225 | 23218897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218897&pc=PYX | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025605882 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 122 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient presented for ablation. after ablation was performed the patient coded multiple times. it was noted that ejection fraction had decreased to 35% and the right ventricle was not moving well. decision made to bring the patient to the cath lab for emergent left heart catherization and impella placement. the complainant reported that the bipella patient encountered access site bleeding during rp flex support. when impella rp flex was placed there was difficulty stopping the site from bleeding after the repositioning sheath was inserted. there was already a hematoma on the right groin from the pci (percutaneous coronary intervention) access site that added to this difficulty. manual pressure was held and patient was pronounced due to other causes before manual pressure could be stopped. patient continued to decompensated however and rhythm was lost. patient coded multiple more times before the care team elected to call the code. patient expired. | [5 empty] | A.4 added weight as it was omitted from the initial report that was submitted. b.2 revised as previous selection on the initial report was incorrect. d.3 added manufacturer fax number as it was omitted from the initial report that was submitted. d.4 added model as it was omitted from the initial report that was submitted. d.7a revised as it was incorrect on the initial report that was submitted. e.1 added zip code extension as it was omitted from the initial report that was submitted. g.4 added pma/ 510(k) as it was omitted from the initial report that was submitted. h.6 added b13 code as it was omitted from the initial report that was submitted. h.10 this is one of three related reports. this report represents the impella rp flex and other reports were submitted to represent the impella cp and automated impella controller. refer to section h10 for the related report numbers. investigation summary: the investigation has been completed. major bleed/hematoma: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | There are two associated devices for the reported event. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4597 | 1220648-2025-47223 | 23218898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218898&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026722133 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed a impella 5.5 and guidewire for the support of the patient admitted in with acute myocardial infarction with cardiogenic shock. the impella 5.5 was placed as escalation from the cp. the pump was placing when the wire was advanced and had a bend. the medical doctor was concerned that the impeller was damaged due to strain he felt advancing and wanted to have a new 5.5 pump. | [5 empty] | The investigation of the reported failure mode has been completed. the product was not returned for investigation. device interaction or damage by another device: the cause of the device interaction was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Upon review of the provided information, the noted issue was associated with the guide wire. there is currently no allegation of a malfunction or serious injury associated with the pump. a regulatory report is no longer necessary.
4598 | 1220648-2025-47226 | 23218899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218899&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTMATED IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-0040 |  |  | ABIOMED, INC. |  |  | US |  | 1805843 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant indicated that they experienced an issue with the placement signal of the automated impella controller (aic). an echocardiogram revealed a pericardial effusion, prompting the physician to perform a pericardiocentesis, during which 500 ml of fluid was extracted. the right ventricle (rv) was observed to be functioning significantly better, and the pulsatility on the aic showed marked improvement. the impella cp was successfully placed in the left femoral artery (lfa) following standard procedures. the peel-away sheath was removed, and a repositioning unit was inserted, with an angiogram conducted through the side port that demonstrated good distal flow. however, as the patient was being prepared for transport, they experienced another decompensation, resulting in dissociated waveforms on the aic. a discussion was held with the physicians regarding support for the right ventricle (rv). it was decided to place the rp flex through the right femoral vein (rfv). patient expired on support. patient death has been captured under the pump complaint. | [5 empty] | Additional information: the device was received for investigation d9 updated. once the investigation is complete a supplemental report will be submitted. | The device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4599 | 1220648-2025-47227 | 23218900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218900&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709016 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 122 KG | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support phase, the patient experienced bleeding, hematoma, and pericardial effusion. a hematoma was observed in the right femoral artery (rfa), necessitating closure of the site before the cp insertion. the impella cp was successfully placed in the left femoral artery (lfa) following standard procedures. the impella pump could not be disassociated, which contributed to the bleeding that had commenced before the impella cp insertion and continued throughout the patient's support. an echocardiogram revealed a pericardial effusion, prompting the physician to perform a pericardiocentesis, during which 500 ml was removed. the right ventricle (rv) was noted to be functioning significantly better, with marked improvement in pulsatility observed on the automated impella controller (aic). the impella was placed in the lfa without complications, adhering to standard practices. the peel-away sheath was removed, and a repositioning unit was inserted, followed by an angiogram through the side port that demonstrated good distal flow. as preparations for transport were underway, the patient experienced decompensation once more, and the waveforms on the aic became dissociated. a discussion was held with the physicians regarding support for the rv (right ventricle). the decision was made to place rp flex through the right femoral vein (rfv). unfortunately, the patient subsequently expired while on rp flex. | [5 empty] | A4 added as it was omitted from the initial report that was submitted. b7 added information as it was omitted from the initial report that was submitted. h10 added related report number as it was omitted from the initial report that was submitted. this is one of two related reports. this report represents the pump while the related report represents the automated impella controller. the investigation has been completed. no product was returned for investigation. pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. hematoma: a hematoma was noted at the non-insertion site at the right femoral artery and the site was closed prior to impella cp insertion. the cause of the hematoma was determined to be patient condition because the hematoma was reported at the non-impella site before impella insertion. major bleed: bleeding was started prior to impella cp insertion as it persisted throughout patient support. the cause of the bleeding is determined to be patient condition because the patient was bleeding from the beginning even before the insertion of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
4600 | 1220648-2025-47230 | 23218901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218901&pc=OZD | 09/06/2025 |  |  | 10/05/2025 |  |  | 09/06/2025 | 10/05/2025 |  |  | 11/30/2025 |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2025438214 |  | IMPELLA CP |  | No |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced bleeding from the impella access site. perclose sutures were tightened, and oozing stopped. also, bival was held for approximately three hours. additionally, it was reported that left femoral artery distal pulses were undetectable. the cp was escalated to impella 5.5. | [5 empty] | Correction: g4, h5 the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be use issue because the bleeding was stopped by tightening the perclose sutures. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4601 | 1220648-2025-47228 | 23218902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218902&pc=OZD | 09/05/2025 |  |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719804 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient presented following coronary angioplasty requiring emergent surgical repair. the patient was centrally cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo) with plan for ecmo reconfiguration, right ventricle assistive device and impella placement with dr. (b)(6). the complainant reported a patient was implanted with an impella device for mechanical circulatory support. the physician placed the impella 5.5 using a competitor's wire, which resulted in a left ventricle perforation. the left ventricle tissue was also noted to be "very friable". the perforation was repaired and chest left open and ecmo reconfiguration was aborted. the team did not administer heparin for heparin-induced thrombocytopenia concerns. subsequently, the patient left the operating room, however, it was noted that care was withdrawn after 32 hours of support. | [5 empty] | Correction g4. the investigation of the reported failure mode has been completed. no product was returned for investigation. thrombocytopenia: the cause was not determined due to insufficient clinical details. cardiac perforation: during the delivery of impella the left ventricle (lv) was perforated with the use of a non abiomed wire for delivering the pump. the cause of cardiac perforation is determined to be use issue because of the use of the non abiomed product which is not recommended as per instructions for use. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4602 | 1220648-2025-47229 | 23218903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218903&pc=OZD | 09/06/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 09/06/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713983 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support experienced bleeding from the right axillary site where the 5.5 was implanted. the patient was taken to cardiovascular operating room for axillary site revision. surgeon found axillary artery and anastomosed graft all intact. bleeding from surrounding tissue successfully stopped. no blood products were administered in relation to this event. the patient continued 5.5 support and was explanted as planned. | [5 empty] | The investigation of the reported failure mode has been completed. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
4603 | 1220648-2025-47231 | 23218904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218904&pc=OZD | 09/06/2025 |  |  | 10/05/2025 |  |  | 09/06/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026702711 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient became unresponsive at home and cardiopulmonary resuscitation was performed by a neighbor. the patient was intubated in the emergency department and left heart catheterization was performed showing heart failure and stenosis in left anterior descending (lad). impella was placed pre-percutaneous coronary intervention and stent placed in lad. the patient experienced ventricular tachycardia and it was reported that the patient was shocked 12 times. additional informatin noted that the patient's family decided to withdraw care. the patient expired within 3 hours of admission to unit. | [5 empty] | Correction d4, g4 the investigation of the reported failure mode has been completed. tachycardia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4604 | 1220648-2025-47232 | 23218905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218905&pc=OZD | 09/06/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 09/06/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677096 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was oozing, hematoma, suction and arrhythmia during impella 5.5 patient support. while on support, oozing and hematoma at the impella site was reported. a new dressing applied. it was noted that the patient continued in atrial fibrillation. one unit of packed red blood cells and albumin were planned. the device was down to p8 due to suction episodes. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | B.7 revised as it was previously entered incorrectly on the initial report that was submitted. e.1 added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e.4 selection was added as it was omitted from the initial report that was submitted. h.6 code e0506 should not have been reported on the initial report that was submitted. the investigation has been completed. no product was returned. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the cause of the hematoma cannot be determined since insufficient clinical details were provided. pump suction: no product was returned. patient had overnight suction alarms. after the review of logs, multiple suctions were observed. no motor current spike or trend in placement signal was present. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4605 | 1220648-2025-47233 | 23218906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218906&pc=OZD | 09/04/2025 |  |  | 10/05/2025 |  |  | 09/06/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658738 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | White | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that they encountered persistent pump suction alarms and a placement signal not reliable alarm during impella 5.5 support. persistent suction alarms were encountered and the nurse (rn) stated that they have been diuresing the patient and they believed the suction could be related to volume. no central venous pressure reading was available. patient's cuff pressure (140/77) was significantly higher than the placement signal (87/27). it was advised that the sensor had likely drifted as the pump had been in for more than 40 days. discussion was held with the team on how to disable the audio for suction. additionally, placement signal not reliable (psnr) alarm was also encountered and it was recommended to disable the audio of the alarm as it did not affect pump performance. the patient outcome was reported to be stable. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. false alarm: data logs at time of persistent suction alarms showed no motor current spike or trend in placement signal to indicate ingestion. pump flows are stable at that time and optical parameters are within range. the cause of the false alarm could not be definitively determined as limited clinical details were provided and no product was returned for evaluation. optical signal issue: data logs were reviewed and at time of reported issue showed a slight drift down in placement signal over approximately 12 hours before placement signal rails to 3000 and out of range for remainder of case. the cause of the optical signal issue was due to sensor wear within pump design life of 60 days. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4606 | 1220648-2025-47234 | 23218907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218907&pc=OZD | 09/15/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 09/06/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728627 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a bipella patient encountered ischemia, pump suction, stroke, torsade's and access site bleeding during impella cp support. a fem-fem bypass was performed while patient received both impella cp and impella rp flex support to prevent ischemia. additionally, after removal of the impella rp flex, patient encountered suction alarms for which 2 units of blood were administered. patient also encountered stroke and torsade's, for which patient was shocked several times and medication was given. the cat scan image of patient's brain showed patient had cerebrovascular accident (cva) that was both ischemic and hemorrhagic. moreover, patient suffered bleeding from all sites and additional blood products were given. however, care was ultimately withdrawn and patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned. arrhythmia, ischemia, stroke: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: bleeding: the cause of major bleed was most likely patient condition related since patient was bleeding from multiple sites. pump suction: no logs are available for the event. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
4607 | 1220648-2025-47237 | 23218908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218908&pc=OZD | 08/09/2023 |  |  | 10/05/2025 |  |  | 08/09/2023 | 10/05/2025 |  |  | 11/30/2024 |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2024311240 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device was explanted as it was noted to be pointing towards the mitral valve. subsequently, the medical team published upon the reported event. there was an additional allegation of pump malpositioning and hematuria. these issues were noted by the authors to be "malrotation". | [5 empty] | The product was not returned for investigation. positioning issues: clinical reported after implanting the pump, it was observed to be close to the mitral valve. the product was not returned for investigation. the cause of the positioning issues was not determined due to insufficient clinical details and no product returned. hematuria: the product was not returned for investigation. the cause of the injury was not determined due to insufficient clinical details. the pump passed all post sterile inspection criteria.
4608 | 1220648-2025-47236 | 23218910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218910&pc=OZD | 08/21/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 08/21/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703315 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was impella inserted and started successfully. initial position was deep in papillary muscles, pulling negative, low flow, and suction. decision made to explant the impella and reinsert. impella lodged deep in papillary muscles. team proceeded to treat left anterior descending artery with 1st stent placement. the patient continued to decompensate, requiring increase in pressor support despite the impella flowing 2.7. patient went pulseless electrical activity (pea) and coded when transferring off table. patient coded, pea, return of spontaneous circulation achieved after 15 minutes. family made patient dnr. patient transferred to intensive care unit in critical condition. the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been complete. the product was not returned for investigation. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. positioning issues: the root cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: data log review confirms suction occurring throughout the case. placement signal is seen trending down, then dropping to zero. this aligns with the clinical report, as the pump was explanted and reinsertion as an attempt to troubleshoot the positioning issues. reinsertion of the pump did not solve the suction alarms. suction alarms occurred frequently and only resolved when the p-level dropped below p-4. placement signal spiked a few times, but continued to trend down to below 50 mmhg while also decreasing in pulsatility as the case progressed. the cause of the suction alarms was most likely related to the patients condition, as clinical reported the patient was decompensating, and data log review supports this with the loss in pressure, seen by the downward trending placement signal and loss in pulsatility over time. low pump flow: data log review confirms low flow levels throughout the case. in the first hour of the impella's runtime, it was achieving a maximum of 2.0 l/min at p-6, which falls below spec. the flow would occasionally spike up to higher flows, but this would be accompanied with suction conditions. placement signal was seen trending down before and after reinserting the pump, spiking occasionally, but continuing to decrease to below 50 mmhg while losing pulsatility over time. the case of the low pump flow was most likely related to the patients condition, as clinical reported the patient was decompensating, and data log review supports this with the loss in pressure, seen by the downward trending placement signal and loss in pulsatility over time. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4609 | 1220648-2025-47219 | 23218911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218911&pc=OZD | 09/05/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 09/05/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1647627 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after impella cp connecting there has been yellow screen with "optical sensor failure" recognized. the team has processed according to the automated impella controller (aic) guide - disconnected and connect the pump again, than the calibration has been performed by the aic according to the display. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 device receive date added. the investigation of the reported failure mode has been completed. the console and data logs were received for investigation. data analysis: console logs from the day of the event are consistent with the reported issue. device analysis: console was returned to field service for further investigation. the reported complaint could not be reproduced when using a known-good pump and purge cassette. further investigation of the optical bench revealed a distortion in the analog output and a small irregularity in the fringe pattern; however, the snr values were within specification when the test cavity was connected. root cause: the cause of the optical sensor failure was likely caused by a malfunction of the optical bench or its components. product history review summary: one additional complaint was found in the product history of console ic9997, not related to the reported issue. additionally, there was another complaint with the same subcomponent lot number that was related to the reported issue.
4610 | 1220648-2025-47238 | 23218912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218912&pc=OZD | 08/22/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 08/22/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713124 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | Hemodynamic instability | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 on (b)(6) 2025 the patient returned to operating room (or) for bleeding from primary impella pocket primary axillary pocket with jackson-pratt drain (jp) in place. patient was hypertensive and was initially extubated in operating room. upon arriving at intensive care unit patient with bleeding and decision made to return to operating room for exploration. surgeon opened pocket discovered branch artery bleeding and 1 clip applied to arterial bleed and transfusion of multiple blood products. hemostasis achieved. j/p drain in place. the patient developed gastrointestinal (gi) bleed, so heparin stopped at that moment. 2 units packed red blood cells with hgb 7.6. impella repositioned last night with final position at 44cm. the patient also had pulseless electrical activity (pea) arrest this morning, 6min of advanced cardiovascular life support (acls), family decision to withdrawal care and the patient expired. | [5 empty] | Corrections e4, h6. the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the root cause of the artery bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. cardiac arrest: the root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | Patient information unknown. the impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
4611 | 1220648-2025-47240 | 23218921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218921&pc=OZD | 08/25/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 08/25/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701029 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 132 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 on (b)(6) 2025, that the patient had hematoma at 5.5 site. anticoagulation remains off to prevent worsening hematoma at impella site and icd site. hematoma is currently stable. cts is aware, will continue to monitor. | [5 empty] | The investigation of the reported failure mode has been completed. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. thrombocytopenia: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
4612 | 1220648-2025-47239 | 23218923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218923&pc=OZD | 07/22/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 08/22/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665854 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 on (b)(6) 2025, the patient (pt) was on bipella support, and had ectopy unrelated to impella. the patient had renal dysfunction and was on dialysis. this was attributed to continuous renal replacement therapy and electrolyte imbalance according to the physician. on (b)(6) 2025 however, dialysis clotted off, blood was not returned and became hypotensive. epinephrine started yesterday during all of this. the pump was explanted on (b)(6) 2025. | [5 empty] | Correction e4 the investigation of the reported failure mode has been completed. no product was returned for investigation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4613 | 1220648-2025-47241 | 23218927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218927&pc=OZD | 08/19/2025 | 10/05/2025 |  | 10/05/2025 |  |  | 10/02/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814768 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | As no patient consequence was associated to this event, the regulatory decision has been populated based on previous risk review(s) for the respective product experience codes and/or analysis codes associated to the complaint. | [5 empty] | Patient information unknown the impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the investigation for the power cord has been completed. the reported issue of the "plug ripped from wall and wires hanging out" was determined to be caused by defective ac power cord set.
4614 | 1220648-2025-47242 | 23218928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23218928&pc=OZD | 07/31/2025 |  |  | 10/05/2025 |  |  | 07/31/2025 | 10/05/2025 |  |  |  |  | 10/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165236 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 121 KG | Non Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced optical signal loss and suction events. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4615 | 1220648-2025-47171 | 23217950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217950&pc=OZD | 08/04/2025 |  |  | 10/04/2025 |  |  | 08/04/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025533979 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex encountered placement signal issues and low pump flow. after two weeks of support care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4616 | 1220648-2025-47172 | 23217952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217952&pc=OZD | 09/04/2025 |  |  | 10/04/2025 |  |  | 09/04/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704417 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support pulses diminished to bilateral lower legs. impella 14 fr. peel away 13 cm sheath in place with distal perfusion cannula in place. flow returned to right leg with removal of sheath and impella. diminished pulses noted to left leg unchanged. pulses diminished to bilateral lower legs. impella 14 fr. peel away 13 cm sheath in place with distal perfusion cannula in place. flow returned to right leg with removal of sheath and impella. diminished pulses noted to left leg unchanged. | [5 empty] | Correction d4. the investigation of the reported failure mode has been completed. peri-procedural adverse event(ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella cp and introducer are both reported for this issue, as there is not enough information to dissociate the impella device as an associated factor. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4617 | 1220648-2025-47174 | 23217955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217955&pc=OZD | 09/01/2025 |  |  | 10/04/2025 |  |  | 09/01/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699753 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 52 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a retroperitoneal hematoma and loss of pulses in her right leg. the decision was made to take her to the operating room for an urgent 5.5 impella placement and surgical removal of the impella cp. the patient survived. | [5 empty] | The cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4618 | 1220648-2025-47175 | 23217958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217958&pc=OZD | 08/28/2025 |  |  | 10/04/2025 |  |  | 09/01/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623375 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to the tortuosity of the patient's anatomy. an intra-aortic balloon pump was used instead. there was no patient harm reported. | [5 empty] | A2 and a4 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter address line 2 as it was inadvertently omitted from the initial report that was submitted. investigation summary: no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had severe tortuosity in the lower limbs preventing successful delivery of impella. device history review: the device passed all post sterile inspection checks.
4619 | 1220648-2025-47176 | 23217959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217959&pc=OZD | 08/31/2025 | 02/13/2026 |  | 10/04/2025 |  |  | 08/27/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485985 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 112 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported additional information upon request: "no interventions done on the patient - care was withdrawn and pump explanted. i do not have the pump- no return kit required." | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and then escalated to an impella 5.5 device for mechanical circulatory support (mcs). subsequently the patient was suspected of having sepsis and later noted again, experienced multiple frequent runs of torsades de pointes and expired. the patient was found to be in cardiogenic shock and was started on levophed and placed in bilevel positive airway pressure. catheterization was completed and diseased left main was noted. the decision was made to place an impella cp device, which was successfully completed and 17 days later the patient was transferred for higher level care. the decision was made to postpone escalation to an impella 5.5 some time after the percutaneous coronary intervention (pci). the patient was taken back to the catheterization lab for high-risk pci which was completed. two days later the optical sensor stopped giving a signal and the purge flow became blocked. a decision was made to explant the impella cp in the operating room with vascular cutdown and have cardiothoracic surgery on standby for an impella 5.5 placement if the patient required more support. it was noted the impella cp peel away sheath remaining in place and appeared to be clotted off. during explant and after, the patient did not require more support and vitals were stable. the decision was made not to place the impella 5.5 device, and the patient was extubated and sent to the intensive care unit. however, later this day, the patient decompensated requiring reintubation and high dose pressors. the impella 5.5 was emergently placed, and the patient was taken back to the unit on impella 5.5 mcs. the following day, it was noted the patient was possibly septic and had an increase in pressor dose. lactate climbed overnight to 3.2 but was now trending down at 2.2. the impella 5.5 device was running at p-9 and maintained at p-9 until pressor support could be weaned. the patient remained sedated due to agitation, but sedation was reduced as tolerated. discussions were made regarding strategies around revascularization of the left anterior descending artery if weaning is unsuccessful. impella support continued. four days later, however, it was noted that the patient had multiple frequent runs of torsades de pointes. pressors have continued to wean down. the patient was awake and able to talk this day. the following day, pressors continued to be weaned, and the patient was still having runs of torsades/ventricular tachycardia, although less frequent. the impella 5.5 device was reduced to support level p-8 and continued to wean as noted the following day. of note, there was no swan-ganz catheter in place, and the central line was not set to transduce. all labs remained normal, and the team was working toward left ventricular assist device work up. however, the following day, the patient was reintubated as every time the impella 5.5 device was weaned down the patient went into ventricular tachycardia. pressors were started after reintubation. the patient was kept sedated. the plan was to remain as is and try to extubate the following day if tolerated. over the next two days it was noted the impella ran at high levels with an attempt to wean pressors as tolerated. an attempt to clear the lungs was made to extubate. however, the patient¿s status remained unchanged. urine output was adequate, and it was noted the patient was administered one unit of packed red blood cells. now day 12 of support the patient was noted looking septic. this day the patient expired. due to the continuous decline, the family decided to withdraw all care. |  |  |  |  | B5 updated. | Additional information was subsequently received and noted the following: the cause of death as not due to the pump. sepsis was not confirmed but rather a small infection that was treated. the cause, assumed, was the impella cp being placed with questionable sterility. and lastly, the arrhythmia presented when support was reduced. the cause of this was not stated and the reason for transfusion was not state. however, there was no continuous bleed. medical safety review. medical safety reviewed the event. the first event describes access site thrombus requiring intervention. this is a serious injury attributed to the introducer. the event involving the impella cp was purge issues; the purge pressure rose, and the patient had sat upright and damaged the purge sensor. the optical signal was unreliable. the pump was explanted. during impella cp explant and after, the patient did not require more support and vitals were stable. the decision was made not to place the impella 5.5 device, and the patient was extubated and sent to the intensive care unit. however, later this day, the patient decompensated requiring reintubation and high dose pressors. the impella 5.5 was emergently placed, and the patient was taken back to the unit on impella 5.5 mechanical circulatory support. the following day, it was noted the patient was possibly septic and days thereafter the patient would experience runs of torsades de pointes. performance level was reduced and every time the device was weaned down the patient experienced the ventricular tachycardia. the cause was not provided, unnoted. additional information noted the sepsis was not confirmed. however, a small infection was treated and the cause, assumed, was the original impella cp being placed with questionable sterility. the patient condition remained unchanged and eventually the patient expired; care was withdrawn. regarding the impella cp device, there was a device malfunction, but it did not contribute to the death outcome. the cp was explanted and patient continued support with the impella 5.5 device. additionally, there is no information that indicates or suggests that the patient expired from the small infection. the event involving the impella cp is a serious injury type of reportable event. the impella 5.5 was in at the time of the patient's demise. conservatively the death should be reported as the impella 5.5 device was in use at the time of expiration. there is no information that indicates or suggests that the patient expired from arrhythmic events. the arrhythmia presented with support was reduced. the patient's underlying condition contributed most to the death outcome. due to the continous decline in the patient's condition from the cardiac event, the family decided to withdraw all care and this led to the patient expiring. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the tachycardia and sepsis, in order to make a cause determination, all of the clinical details would have to be provided. the root cause was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella 5.5 device report is now one of three devices associated with the event story (a complaint was created for the introducer). refer to the related manufacturer report numbers as noted in section h10 for the medical device report on the impella cp and introducer. | Patient-specific information a5: ethnicity and a6: race are unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella 5.5 device report is one of two devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp.
4620 | 1220648-2025-47177 | 23217960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217960&pc=OZD | 08/24/2025 |  |  | 10/04/2025 |  |  | 08/24/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after a pump stop with an impella 5.5, a new impella 5.5 was implanted successfully. during support with the second impella 5.5 a suction alarm occurred. the pump was repositioned and running well at p7. throughout support they continued to have suction issues that either self resolved or required additional volume in addition to pump repositioning. additionally, the patient had an episode of ventricular tachycardia requiring defibrillation. the patient's urine was pinkish. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of the pump suction cannot be determined since insufficient clinical details were provided as the pump was repositioned and suction alarms were only temporarily resolved. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4621 | 1220648-2025-47179 | 23217961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217961&pc=OZD | 08/23/2025 |  |  | 10/04/2025 |  |  | 08/23/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638818 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp required a thrombectomy to restore flow to the legs. the patient also had hemolysis, so echo was utilized to pull the pump back into position. ischemia of the right lower leg required a fasciotomy but the procedure was unsuccessful and the leg required amputation. later, the pump was explanted and the patient survived. | [5 empty] | Product complaint # (b)(4). investigation summary: clinical symptoms (ischemia): in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. mechanical interaction with blood (hemolysis): the pump was not returned for investigation. according to the clinical details, an adjustment was made to the pump position during reported hemolysis on (b)(6) at 12:11 pm. urine clearing was observed after the position adjustment. data log shows a few suction alarms during hemolysis and sudden flow improvement at 10:00 am with hike in motor current. there was no position-related alarm or motor current spike noted during the hemolysis event. the cause of the hemolysis issue was most likely improper device positioning, based on the clinical details. device in wrong position: the cause of the positioning issue was not determined due to insufficient clinical information. device history lot: device batch: 1898374. device history batch: subcomponent lot: na. device history review: the pump sn(b)(6) passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4622 | 1220648-2025-47178 | 23217962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217962&pc=OZD | 08/23/2025 |  |  | 10/04/2025 |  |  | 08/23/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615331 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma on the right femoral artery which was managed successfully with manual pressure. additionally, suction alarms occurred and they were addressed by infusing two liters of lactated ringer's. the pump's position was checked with an echocardiogram and the impella was pulled back into a stable position. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of hematoma was most likely use issue related since hematoma occurred after the patient arrived to unit. the cause of major bleed was not determined due to insufficient clinical details. the cause of positioning issue cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4623 | 1220648-2025-47180 | 23217963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217963&pc=OZD | 08/22/2025 |  |  | 10/04/2025 |  |  | 08/22/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - |  |  | US |  | 2026722910 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis and oozing at the groin site. the pump position was adjusted to resolve the hemolysis, heparin was withheld, and a femstop was placed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the hematoma was not determined due to insufficient clinical details. the cause of hemolysis was most likely due to positioning issue but the cause for positioning issue is unknown. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the cause of pump suction was most likely patient condition since suction resolved after giving fluid boluses. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4624 | 1220648-2025-47181 | 23217965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217965&pc=OZD | 08/21/2025 |  |  | 10/04/2025 |  |  | 08/21/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658525 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 119 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day three of impella 5.5 support the patient experienced ventricular fibrillation and ventricular tachycardia. the patient was shocked four times. a temporary ventricular (temp v) pacing wires were placed. two days later, the temp pacer was turned off as the patient was in normal sinus rhythm. later it was reported that gastrointestinal was consulted for possible gastrointestinal bleed and anticoagulant was stopped along with one unit of blood products were administered. on day 11 of support, the patient accidentally pulled impella out of place. the team tried multiple times to cross impella back to the left ventricle without success. the impella was removed. the patient was stable and survived. | [5 empty] | H6: added code b13 and b11 to type of investigation. updated c13, and d15 based on the investigation results of the tachycardia/ ventricular fibrillation and major bleed. investigation summary no product was returned. major bleed: the cause of the bleed was unable to be determined due to insufficient clinical details. tachycardia/ ventricular fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: pt accidently pulled impella out of place overnight. team made multiple attempts to float it back across the lv without success. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient. device history lot device lot: 1919469. device history review device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4625 | 1220648-2025-47185 | 23217968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217968&pc=OZD | 08/16/2025 |  |  | 10/04/2025 |  |  | 08/16/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707152 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 105 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | Thrombocytopenia | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had intermittent ventricular tachycardia and an axillary hematoma that remained stable. the patient was treated with bivalirudin due to persistent thrombocytopenia. the patient was successfully weaned from the pump and survived. | [5 empty] | B.2 revised as selection entered on the initial report that was submitted was incorrect. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the cause of the hematoma was not determined since product was not returned and insufficient clinical details provided. the device history review was completed and the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4626 | 1220648-2025-47186 | 23217969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217969&pc=OZD | 08/13/2025 |  |  | 10/04/2025 |  |  | 08/13/2025 | 10/04/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692878 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemolysis | Sepsis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced intermittent placement signal low and suction alarms. additionally, the patient was septic and had hemolysis. later, the pump was successfully weaned from the pump and survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. optical sensor issue: clinical reported intermittent placement signal low alarms. the data logs confirm alarm events and show a gradual trend down in placement signal through support till explant. further rt log analysis showed intermittent dips in placement signal towards 20mmhg verifying correct trigger of alarms. motor current values were stable and unaffected. there were no position related alarms. no relevant patient condition or abnormality noted. no volume issues noted. the 5s parameters were within specification with contrast slightly variable. the cause of the issue was not established as no product was returned and limited clinical information was provided. mechanical interaction with blood: clinical reported hemolysis. the data logs do not show any motor current spikes outside p-level changes. no positioning alarms in the data logs and no reports of improper pump positioning. logs showed gradual decrease in placement signal through support. the cause of the issue was not established as no product was returned and limited clinical information was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4627 | 1220648-2025-47187 | 23217970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217970&pc=OZD | 09/02/2025 |  |  | 10/04/2025 |  |  | 09/02/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667310 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 90 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding from the access site. systemic heparin was withheld. the patient also experienced ventricular tachycardia requiring cardioversion and administration of albumin. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned for investigation. major bleed: access site bleeding around the sheath. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. tachycardia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1917633. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. as per qn (b)(4), this qn has been opened to investigate the legibility, completeness, and scanability of 2000193 batches. lot 1920527 was originally in the scope of this qn; however, it is not material 2000193. lot 1920527 is material 2000154; therefore, it has been added the to corresponding qn for this containment under the correct material number (qn (b)(4)) and removed from this qn. all investigation and rework activities remain applicable, as all qns are for the same label issue and require the same correction. the multiple qns were made only to separate the affected lots by material number. additional qns for this labeling issue include (b)(4). rwk qty correction: original intended rework was for qty 324 pumps; however, additional, unrelated ncs were identified on pumps (b)(6). mrb further reviewed the following lots and determined qty 3 pumps did not pass the label rework re-inspection per 0000-8400rt. no further investigation is required, as the pumps failed due to handling during the rwk process. pump with sn: (b)(6) from lot 1919492, pump with sn: (b)(6) from lot 1921852, pump with sn: (b)(6) from lot 1919480, the above pumps have been labeled as scrap and not for human use. the remaining 6 pumps in these lots passed reinspection, per the attached data sheet and will be accepted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4628 | 1220648-2025-47188 | 23217971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217971&pc=OZD | 07/19/2025 |  |  | 10/04/2025 |  |  | 07/19/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604006 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 125 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted to support a patient during coronary artery bypass graft surgery. while on support, the patient experienced atrial fibrillation with rapid ventricular response. amiodarone was administered. when the patient was later moved to a chair, they experienced atrial fibrillation with rapid ventricular response again. ventricular tachycardia was also reported. when the doctor tried to remove the impella through the graft and met resistance, the decision was made to open the patient's chest and explant that way. the patient survived. | [5 empty] | The investigation into the reported issues has been completed. no device was received for evaluation. the root cause of the arrythmia was unable to be determined due to insufficient clinical details. the root cause of the removal issue was unable to be determined as insufficient clinical details were provided. the device passed all post sterile inspection checks. | The complainant reported the patient experienced atrial fibrillation and ventricular tachycardia. the root cause of the injury was unable to be determined due to insufficient clinical details. the complainant reported difficulty removing the device. the root cause of the removal issue was unable to be determined as insufficient clinical details were provided.
4629 | 1220648-2025-47189 | 23217973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217973&pc=OZD | 07/15/2025 |  |  | 10/04/2025 |  |  | 07/15/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655582 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced multiple issues. the patient experienced bleeding that was resolved by repositioning the pump and transfusing blood products. the patient had a fever and an infection was suspected. the patient was treated with antibiotics. the pump generated a suction alarm after the patient had rolled over in bed. an echocardiogram showed the pump was in an optimal position. the patient remains on impella support. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. pump suction: clinical reported brief suction event when patient rolling in bed, resolved with repositioning supine. this was reported on overnight on day 29. rt log analysis of day 29 confirm suction alarm #14 diastolic suction. a review of the dp_max, dp_offset, and ps at mc_min show that alarm was triggering correctly. aligning with clinical information, suction resolved not too long after. no volume related issues were mentioned since patient noted to be hypovolemic on day 2 of support. there was no confirmation of pump positioning with clinical imaging at alarm occurrence nor were there pump positioning alarms. there were no issues with 5s parameters. there were other suction events later on during support with the next one occurring about 6 days later. the cause of the pump suction could not be definitively determined as limited clinical information was provided on exact timing of patient position in relation to alarms and no product was returned. optical signal issue: clinical reported placement signal low alarms which seemed to occur when patient was in certain positions. echo showed optimal pump position. reports were on day 56 onwards. logs confirm placement signal low alarms while pump was running at p5. rt log analysis shows periodic drops in placement signal and confirm correct trigger with placement signal dropping below 30mmhg. there were no issues with 5s parameters. the cause of the optical signal issue was not established as limited clinical information was provided on exact time of patient position in relation to alarms and no product was returned. the root cause of the fever/infection and major bleed were unable to be determined due to insufficient clinical details.
4630 | 1220648-2025-47191 | 23217975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217975&pc=OZD | 07/05/2025 |  |  | 10/04/2025 |  |  | 07/06/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671824 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 128 KG | Non Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | Hematuria | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient with st-elevated myocardial infarction was implanted with an impella 5.5. the patient developed hematuria that was resolved with unknown intervention. additionally, the pump generated a, "signal not reliable" alarm with absent aortic and left ventricle placement signals waveforms. later, care was withdrawn and the patient expired. | [5 empty] | The user reported absent aortic and left ventricle placement signal waveforms. the returned pump passed the laser continuity test. the data logs confirm 'placement signal not reliable' alarms and show an optical sensor drift within indicated design life at 51 days. the placement signal trended down while pump flows trended up before placement signal went out of range. the cause of the issue was determined to be due to material wear of the optical sensor. the user also reported hematuria. the cause of the injury was not determined due to insufficient clinical details.
4631 | 1220648-2025-47192 | 23217976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217976&pc=OZD | 06/13/2025 |  |  | 10/04/2025 |  |  | 06/13/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623698 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A japan complainant reported that a acute myocardial infarction/cardiogenic shock patient was on impella 5.5. it was reported that after closure, the left subclavian artery containing the original 5.5 was opened and the graft was trimmed as much as possible. the tip of the 5.5 and the end of the artificial blood vessel were submitted as specimens. (the cause is currently unknown.) however, fungus was identified, and antibiotics were initiated. the 5.5 was explanted as bridge to left ventricular assist device. | [5 empty] | Correction: d3 and h11 (investigation results) was inadvertently omitted on the initial manufacturer device report. infection: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the pump was not returned for investigation. the root cause of the infection was unable to be determined due to insufficient clinical details.
4632 | 1220648-2025-47193 | 23217977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217977&pc=OZD | 09/15/2024 |  |  | 10/04/2025 |  |  | 09/25/2024 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531917 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Anemia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Pain | Tachycardia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had bleeding, drainage, and pain at the access site. the patient was treated with oral oxycodone and transfused one unit of blood. the patient also experienced self-limiting bouts of ventricular tachycardia and was defibrillated. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The root cause of the tachycardia was unable to be determined due to insufficient clinical details the cause of the pain and anemia was not determined due to insufficient clinical details. the cause of the bleeding issue was not determined due to the lack of returned product and insufficient clinical information. the cause of the high purge pressure issue was not determined as the product was not returned for investigation, and sufficient clinical information was not provided. the cause of the hemolysis issue was not determined as the product was not returned for investigation, and sufficient clinical information was not provided. the pump passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4633 | 1220648-2025-47194 | 23217978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217978&pc=OZD | 09/05/2025 |  |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660843 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | Malposition of Device | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed hemolysis and encountered pump detection issue with the automated impella controllers (aic). the patient would eventually expire. the patient required c-pap for over 24hrs but continued to decompensate overnight, was intubated and was sent to catheterization lab for an emergent catheterization. the impella cp was implanted for protected percutaneous coronary intervention, which was completed. however, during the pci, the physician was unable to cross the proximal left anterior descending occlusion. the pci was paused for shock call. the decision was made to cannulate for venoarterial extracorporeal membrane oxygenation (va ecmo). penumbra was used and the occlusion was ballooned, and one drug-eluting stent was implanted. after the pci, the peel-away sheath was removed, touhy-borst locked at 91cm and grin was dressed, dry and intact. the patient was sent to the unit on ecpella (ecmo and impella) support. the following day it was noted that the impella was repositioned with point-of-care ultrasound for an alarm in left ventricle. hemolysis complication was noted and consequently, the impella catheter was pulled back 1cm. of note, placement signal alarm was disabled for frequent placement signal low alarms occurring. however, the impella position still appeared to be malpositioned, caught in the mitral apparatus and sitting deep at 6 cm. patient care note unrelated to the impella, noted the patient continued to require blood products related to the at home fall injury. heparin was placed on hold and possible surgery was noted for the left flank hematoma/bleed. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. ecpella support continued and the next day ecmo was weaned to 3800rpm/3.7l/min. p level increased to 4, flow 2.5l. drips were weaned downed. surgical consultation did not require intervention for hematoma/bleed. the patient received one unit of red blood cells over the night. of further note, the placement signal alarm was enabled again. ecpella support continued. two days later echocardiogram showed ejection fraction improvement 5% to 10-20%. the impella was still angled towards the mitral valve but less than initial implant. inlet at 4.7 cm from aortic valve. plasma free hemoglobin at 10mg/dl and alanine aminotransferase down trending was noted. low dose heparin drip was started. intros adjustment was noted. also noted was that the patient experienced another vt storm overnight with torsades and received one shock from implantable cardioverter defibrillator. support continued with target for heart recovery versus transfer for potential left ventricular assist device (lvad). ecpella support continued and three days later [(b)(6) 2025] ventricular tachycardia and supraventricular tachycardia was noted through the day and over the night. now transplant versus lvad work up started. the following day, ganglion block was placed to help control ectopy. this day the aic had to be switched out and when the nurse attempted to switch out the aic from the outside hospital console, pump not detected was encountered. when white cable was plugged in to aic ic 9571, instead of restarting the pump on current setting on the new aic it went to ¿spike dextrose bag¿ screen. because it did this, the care team switched it back to the original aic ic 5269, and when the white cable was plugged back into aic ic 5269, it did the same thing. therefore, to avoid pump being off for any longer, the nurse continued through the spike/prime purge menu and impella restarted on auto as if it were a new impella cp case. the device was restored to previous p-level and patient stable. the following day palliative care was discussed; the patient was not a candidate for transplant or lvad. the patient would subsequently expire. | [5 empty] | E.1 added zip code extension as it was omitted from the initial report that was submitted. medical safety reviewed the event. the patient presented in a decompensating state and was shocked 10 times for ventricular tachycardia. the patient presented with non-st-segment elevation myocardial infarction (nstemi). the patient would eventually expire on support. the palliative care was discussed. the patient was not a candidate for transplant or left ventricular assist device (lvad). the patient had additional support in the form of veno-arterial extracorporeal membrane oxygenation (va ecmo), and experienced complications related to an at-home fall injury and was in a poor state. the impella was there to support the patient. the patient expired and the impella may have contributed to the patient's demise given the hemolysis and positioning issues encountered (thus may have added to an already complex situation or compounded the situation to where the patient actually expired). regarding automated impella controllers (aics) the events describe failure of the pump to be detected by the aic resulting in an alarm. the original aic was reconnected with the same outcome. however, to avoid the patient being off the pump longer, the nurses continued through the spike/prime purge mean and restarted the pump on auto as if it was a new case. although interruption in support, it appears the patient expired for reasons unrelated to the aics (as noted above). the aics malfunctioned and contributed to the patient's implant experience event and should be reported as a malfunction type of reportable event. related medical device reports: this impella cp device (which is associated with the demise outcome) is one of three devices associated with the event. refer to the related manufacturer report numbers as noted in section h10 for the medical device reports for the automated impella controllers. investigation: the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the following investigation outcomes was noted for the adverse event(s) and device malfunction(s): hemorrhage/bleeding: the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood/hemolysis: as per the clinical details, hemolysis was reported during concerns about pump position. data log shows fluctuations in placement signal, with drops and increases in amplitude for some time during the hemolysis event. several pump position in aorta alarms were reported, with placement signal disable activity. no position-related issue was noted after the morning of 09-sep, and the device was utilized for another 10 days. the cause of the hemolysis issue was most likely improper positioning, based on data log review and provided clinical details. device in wrong position: the cause of the positioning issue was not determined since the product was not returned for investigation, and sufficient clinical information was not provided. pump not detected: as per the clinical details, the pump automatically reinitiated to the original case start steps after switching to the transfer console. the doctor switched back to the previous console, and the same issue with case start was repeated. the doctor followed all the steps, and the pump was restarted. long term log confirms that the case start was initiated on aic 5269, and the case start was reinitiated while switching to aic 9571 on the 9th day of support. the console was switched back to aic 5269, carrying the same reported case start steps as uninitiated pump, and the pump was able to start after following the case start sequence. the log was reviewed on aic 5269 and confirmed that the first case start was initiated after insertion of the purge cassette. pump plug connection was recorded after the "step 3: connect impella catheter" notification. there was an eeprom read error and an "impella defective (error reading eeprom)" alarm was issued. the pump was unplugged and replugged, with no initializing data recorded. there was a case start cancel notification; however, the priming steps were successfully performed and the pump was started. after 8 days into the case run, the pump was switched to console aic 9571 and the pump read as uninitialized, resulting in a repeated case start. the same issue occurred when the pump was switched back to console aic 5269, and the pump was restarted after the doctor followed the case start priming steps again. the cause of the pump not detected was related to eeprom corruption. h6 investigation type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. correction: complaint file coding for device problem codes was updated to add a better described code to go along with the other codes that were previously submitted for the reported event. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp device (which is associated with the demise outcome) is one of three devices associated with the event. refer to the related manufacturer report numbers as noted in section h10 for the medical device reports for the automated impella controllers. this event describes transfer failure of the automated impella controllers. there was no report or indication of significant hemodynamic decompensation. the patient was noted to have remained stable. the impella cp pump, its positioning was suboptimal, and it was noted patient developed hemolysis. given the details of the case dissociation cannot be made although the patient¿s underlying condition appears to have contributed more to the demise outcome (cardiogenic shock, hematoma/bleeding from fall injury and ejection fraction of 5%-10%). instructions for use for the related event are as follows: impella cp with smartassist system: section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section use of echocardiography for positioning of the impella catheter - ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section impella catheter too far into the left ventricle - ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4634 | 1220648-2025-47195 | 23217979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217979&pc=OZD | 09/05/2025 |  |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724364 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Tachycardia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented to the emergency department with complaints of chest pains that had been occurring for two days. the patient was diagnosed as having non- st elevation myocardial infarction but then became in st elevation myocardial infarction. the patient was then taken to the cardiovascular laboratory for treatment. the patient was noted to have chronic total obstruction of the right coronary artery and proximal left anterior descending coronary artery (lad) was determined ot be the culprit lesion. a dissection was noted post ballooning of the artery, and due to difficult patient anatomy, the wire access was lost. the patient the became hemodynamically unstable and the decision was made to place an impella cp. the percutaneous coronary intervention (pci) continued through single access at the impella access site. the lad was successfully rewired, and the stent was successfully deployed. following the pci, the patient was noted to have a hematoma at the impella access site. manual pressure was held, and the patient was monitored and prepped for transfer to another hospital for further treatment. the following day, the patient was transferred to another hospital and shortly after arrival was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo). after cannulation, dobutamine was discontinued and the patient¿s tachycardia resolved. additionally, it was noted that distal perfusion catheters were placed bilaterally for limb perfusion. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | D.3 added manufacturer fax number as it was omitted from the initial report that was submitted. e.1 added initial reporter address line 1 as it was omitted from the initial report that was submitted. h.6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. hematoma: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. tachycardia/ischemia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | A4: patient weight is unknown. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4635 | 1220648-2025-47196 | 23217980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217980&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673609 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 112 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed red/pink urine. the day before, on transthoracic echocardiography the cp measured deep. the md did not reposition as the patient was not having an issue. the impella p-level was adjusted, from p-8 to p-6 resolving the urine color. | [5 empty] | Correction e4. the pump was not returned for investigation. data log was not provided and could not be retrieved from the remote server. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was not determined, as the product was not returned and sufficient clinical information was not provided. device in wrong position: the cause of the improper positioning issue was not determined since the sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4636 | 1220648-2025-47198 | 23217981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217981&pc=OZD | 09/05/2025 |  |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665985 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 85 KG |  |  | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress | Patient Device Interaction Problem |  |  |  | Hemolysis | Hematuria | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient arrived to the hospital emergency department via emergency medical services in anterior st elevation myocardial infarction. the patient was brought to the cardiovascular laboratory, and an angiogram revealed a 100% occluded left anterior descending coronary artery (lad). the medical team was unable to full open the lad, and with thrombectomy passes the patient began to decompensate. the decision was then made to place an impella cp. the following day it was reported that the patient¿s laboratory samples were hemolyzed, and the patient was experiencing blood in their urine. additionally, the placement signal was decoupled and there were suction alarms for the impella. a plan was made to perform echocardiography to verify placement of the impella, and the medical team repositioned the impella. however, the impella twisted towards the inferior posterior wall causing suction. the physician pulled back the pump 1 cm but was unable to torque the pump around into position. the p level was lowered to p6 which resolved alarms. it was noted that the pulsatility of the apex and septum was poor and that the lad remained clotted. the attending physician determined that the patient was not a candidate for further intervention. the family decided to withdraw care of the patient. the patient¿s care was withdrawn and the patient expired. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: as per the clinical details, hemolysis and suction were reported six hours after implant on (b)(6). the doctor made some adjustments to the pump, and suction alarms resolved after the p-level was reduced to p6. total pump run time was around 13 hours. data log shows the pump was initiated at p9, with a trending placement signal and some drops in flow and motor current. after about three hours, several suction alarms were reported while running at higher p-levels, and the pump was then turned to p6. later, some improvement in flow trend was seen, followed by several "pump in aorta" alarms. the pump was explanted after another six hours of the case run. there was no conclusive evidence of pump performance issue or confirmed patient condition without relevant clinical details. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. pump suction: the cause of the pump suction issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. device in wrong position: the cause of the positioning issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. product damage (cannula kink): the cause of the cannula kink issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. clinical symptoms (hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump (b)(6) passed all post sterile inspection checks. | Although the patient's presenting condition may be more likely have impacted the event the impella pump cannot be excluded as a possible contributing factor in the patient's demise. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4637 | 1220648-2025-47197 | 23217982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217982&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673560 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information: hemolysis was managed by the primary team by reducing p-level to lowest level tolerated by patient, but the primary team did not elect to reposition. | The user facility reported that the patient on impella cp support had increased lactate dehydrogenase and indirect bilirubin. there was concern for hemolysis. the cp position was appropriate on transthoracic echocardiography (tte) the day before. the plan was to reassess position and reposition under tte later that day. additionally, the impella placement signal was dampened versus a-line tracing noted. |  |  |  |  | Additional information received and the following fields were updated. b5, h6 correction: e4. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4638 | 1220648-2025-47199 | 23217983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217983&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712051 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 62 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was successfully supported through percutaneous coronary intervention with impella cp. the patient lost distal pulses in right lower extremity. angiogram confirmed loss of flow. left femoral to right femoral bypass performed with 6french retrograde sheath in the left femoral artery and 6f right femoral antegrade sheath in the sfa. additionally, suction alarms were intermittent and were treated with albumin. the team was also concerned for sepsis. ultimately, the family made the decision for comfort care and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4, h6. the investigation of the reported failure mode has been completed. no product was returned. ischemia, sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no product was returned. patient had suction alarms and the position of pump was almost lost and was advanced back into optimal position. after reviewing the logs, multiple suction alarms and impella in aorta alarm was observed. the cause of pump suction was due to improper positioning but the cause for position issue is unknown if it is use related or patient condition related. device history review: the device passed all the post sterile inspection checks.
4639 | 1220648-2025-47201 | 23217984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217984&pc=OZD | 09/05/2025 | 10/14/2025 |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689317 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Pt came down for 5.5 planned escalation. pt prepped and draped in standard fashion. graft sewn on and anastomosis achieved without difficulty. lv access achieved with wire. cath removed. impella backloaded after priming. md inserted impella but met large resistance getting the impella across the bend of the inomenant. the surgeon continued to push and pull technique as well as fingers in the pocket manipulation. ultimately, the catheter just distal to the motor housing became bent and lost its integrity. md decided to remove impella and opt for a new 5.5. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. the pump and introducer were returned for investigation. delivery issue: the product was returned for investigation, and a kink was observed on the catheter distal to the motor housing. additionally, a cannula kink was observed near the motor housing, and a kink in the guidewire was observed as well. the cause of the delivery issue was most likely related to the patients condition, as clinical reported calcified vasculature as the cause of resistance. product damage (catheter kink): clinical reported that during insertion of the pump, resistance was met due to calcified vasculature, and the catheter became kinked distal to the motor housing. the product was returned for investigation, and a kink was observed on the catheter distal to the motor housing. the cause of the catheter kink was most likely related to the patients condition, as clinical reported calcified vasculature as the cause of resistance. device history review: the device passed all the post sterile inspection checks.
4640 | 1220648-2025-47200 | 23217985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217985&pc=OZD | 09/05/2025 |  |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719761 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 68 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient presented to the hospital on (b)(6)2025 with complaints of chest pain, nausea, vomiting. it was noted that the patient¿s troponin levels were elevated that resulted while in the emergency room. the patient was then admitted and taken to the cardiac catheterization laboratory for left heart catheterization. severe mitral valve disease was noted with a preserved ejection fraction of 50%. the patient was placed on a heparin drip, but the chest pain did not resolve. the patient was then scheduled for a 2 vessel bypass on (b)(6)2025. after coming off cardiopulmonary bypass (cpb) following the coronary bypass procedure, the patient decompensated quickly. the patient was quickly placed back onto cpb and the coronary arteries were reassessed. there were no significant issues noted with the grafts of the bypass, and the patient was again attempted to be weaned from cpb. the patient again decompensated, and the left ventricular ejection fraction dropped significantly after 10-15 minutes of being off cpb. a plan was made to cannulate the patient for extracorporeal membrane oxygenation and impella cp for left ventricular unloading, however the patient¿s femoral arteries were found to be severely diseased. the physician decided to place an impella 5.5 directly through the aorta as the patient¿s chest was still open. the impella 5.5 was inserted with no issues and support was initiated. the patient was then released back to the intensive care unit for rest and recovery. the following day, it was reported that the patient was experiencing bloody urine, and the patient was placed on dialysis. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary; the investigation has been completed. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | Product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4641 | 1220648-2025-47202 | 23217986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217986&pc=OZD | 09/05/2025 |  |  | 10/04/2025 |  |  | 09/05/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724537 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support. following the implantation, the patient encountered "significant" bleeding at the access site and was initially unable to achieve hemostasis through manual pressure alone. a femstop compression device was subsequently utilized alongside extended manual pressure, which resulted in a reduction of the bleeding. due to the urgent need for pump placement, no closure device was employed prior to the implantation. the patient was transferred to the unit on support and unfortunately passed away approximately five days later. this case presents a scenario of "significant" bleeding with an initial failure to attain hemostasis in a complex situation. although it was reported that prolonged manual pressure and the femstop device were applied, leading to a decrease in bleeding, the term "significant" remains ambiguous, as the specifics regarding the volume of blood loss and whether blood products were necessary are not disclosed. in this instance, the patient's underlying conditions predominantly contributed to the fatal outcome, and care was ultimately withdrawn, which seemingly resulted in the demise. nevertheless, based on the available information regarding the impella device, which was in use during the complications experienced, there is insufficient data to separate the device's involvement from the outcome. | [5 empty] | B.1 product problem was selected incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. minor bleed: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
4642 | 1220648-2025-47203 | 23217993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217993&pc=OZD | 09/04/2025 |  |  | 10/04/2025 |  |  | 09/04/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026666712 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | Tachycardia | Thrombocytopenia | [7 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient presented to the emergency department with chest pain on exertion. the patient underwent nuclear stress test which was found to be positive. patient was taken to the cath lab and it was revealed that the patient had 80% occlusion mid left anterior descending (lad). during percutaneous coronary intervention, the lad dissected. the patient went asystole and underwent cardiopulmonary resuscitation for 2 minutes. return on spontaneous circulation was achieved the patient was intubated and lad was re-wired. 3 stents were placed to lad, but was unable to fully cover dissection. the patient continued to decompensate requiring high dose vasopressors. impella cp discussed but due to poor femoral vessel size was not initially considered. subsequently, the decision was made to proceed with impella cp support. while on support, there was a large amounts of chest tube output and the patient received 12 blood products in the operating room. after icu transfer, the patient continued to receive blood products. the patient went into atrial fibrillation with rapid ventricular response. the patient converted back to normal sinus rhythm and back into atrial fibrillation. the patient¿s heart rate was in the 140s. an amiodarone bolus and drip was started. nephrology was consulted to place a dialysis catheter and start continuous renal replacement therapy. the impella was repositioned by the surgical team to 3.6cm below the aortic valve. the patient¿s white blood cell count was elevated; the staff was concerned that the patient was septic. the heparin drip was held as the patient had a low platelet count. their platelets were down to 33,000 and two units were ordered. later, the doctor attempted cardioversion three times but it was unsuccessful and the patient remained in atrial fibrillation with rapid ventricular response. care was withdrawn and the patient expired. | [5 empty] | Correction d4, g4. the investigation of the reported failure mode has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. tachycardia/ arrhythmia/ sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. renal failure/ thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety/clinical reviewed the event. the patient was a 73-year-old with a history of peripheral artery disease, coronary artery disease, hypertension, hyperlipidemia, and active tobacco use who presented with exertional chest pain. a nuclear stress test was positive, and cardiac catheterization revealed an 80% occlusion of the left anterior descending (lad) artery. during percutaneous coronary intervention (pci), a lad dissection occurred, and the patient went asystole requiring approximately two minutes of cpr, after which return of spontaneous circulation was achieved. the patient was intubated, and three stents were placed in the lad; however, the dissection could not be fully covered. the patient remained hemodynamically unstable with ongoing cardiogenic shock requiring high-dose vasopressors and was taken emergently for coronary artery bypass grafting (cabg) with planned impella support (indication: pccs/lcos). an unsuccessful attempt was made to implant an impella 5.5 via right axillary artery. after multiple attempts, the impella 5.5 could not be advanced past the innominate artery, and the procedure was aborted. an impella cp was subsequently implanted via axillary cutdown and was successfully placed. impost impella cp implant same day, the patient experienced a major bleeding event with significant chest tube output, requiring transfusion of approximately 12 units of blood products intraoperatively. over the subsequent days, the patient developed atrial fibrillation with rapid ventricular response, acute kidney injury requiring nephrology consultation and planned continuous renal replacement therapy (crrt), white blood cells elevated with concern for sepsis, and thrombocytopenia necessitating interruption of heparin therapy. the patient remained critically ill, intubated, and on combined mechanical circulatory support including impella (p4) and veno-venous extracorporeal membrane oxygenation (vv ecmo). it was noted despite ongoing support; there was no meaningful clinical improvement. a decision was made to withdraw life-sustaining support leading to the patient's subsequent demise. h6. investigation type/findings/conclusion remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
4643 | 1220648-2025-47204 | 23217994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23217994&pc=OZD | 09/04/2025 |  |  | 10/04/2025 |  |  | 09/04/2025 | 10/04/2025 |  |  |  |  | 10/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707474 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | Unspecified Kidney or Urinary Problem | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that there was a pump in ventricle alarm that was not resolved with repositioning. the pump was pulled back but appeared to be flipped, deflecting against the wall of the left ventricle. the pump remained implanted as the patient continued to be well supported, despite the orientation of the pump. patient had frank blood in urine with clots. the p-level was lowered to mitigate hemolysis. possible uti and cultures pending. | [5 empty] | The investigation of the reported failure modes has been completed. device in wrong place: the data logs show that there are positioning alarms seen on 9/4 with spikes in the ps and noisy 5s but that improves on 9/4. there are still multiple positioning alarms and spikes in the ps that continue throughout 9/11. due to the lack of clinical details provided, the root cause of how the pump came out of position cannot be determined. hemolysis: the data logs show that there were positioning alarms during this time period with spikes in ps, suction, and noisy 5s as the pump was not getting the best signal. the motor current was stable during this time period with no spikes or drastic changes in pump flow at the respected p-levels. the pump run remained stable throughout the rest of the case. due to a lack of product returned, the root cause of the hemolysis cannot definitely be confirmed. hematuria/urinary tract infection the root cause of the patient injuries cannot be determined due to a lack of clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4644 | 1220648-2025-31808 | 23208053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208053&pc=OZD | 01/06/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 07/28/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563996 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella 5.5 support. during this time, 'placement signal not reliable' alarms occurred. the patient had multiple medical issues and was deemed ineligible and too ill for heart, kidney, or liver transplant. the patient ultimately expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and complaint was confirmed. upon review of the data logs, it is apparent that the console is the potential cause of the issue. no problem was found with the pump during the case. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4645 | 1220648-2025-47131 | 23208088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208088&pc=OZD | 09/05/2025 | 10/02/2025 |  | 10/03/2025 |  |  | 09/05/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616592 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was placed to support high-risk percutaneous coronary intervention (hrpci). after this procedure, the impella cp was weaned and removed. it was further reported that the patient had evidence of a stroke following hrpci. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. stroke: the root cause of the stroke was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
4646 | 1220648-2025-47132 | 23208174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208174&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 09/05/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713941 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was supporting a patient for eight days. upon removal, a clot was pulled from the pigtail. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
4647 | 1220648-2025-31809 | 23208192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208192&pc=PYX | 03/27/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 03/27/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025590613 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, a 'placement signal not reliable' alarm occurred. after six days of rp flex support, care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned for analysis. the data logs were evaluated and alarm placement signal not reliable was observed. the root cause of the placement signal issue could not be definitively determined as no product was returned for analysis to confirm dp sensor failure or cause of failure and limited clinical information was provided, a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4648 | 1220648-2025-47133 | 23208203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208203&pc=OZD | 09/05/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 09/05/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693549 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 100 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was inserted to support a percutaneous coronary intervention procedure. on day two of impella placement, a transesophageal echocardiogram was done which revealed a large left ventricle thrombus. the patient was on a cangrelor drip and was transitioned to brilinta post operatively. also, all access-line sites were oozing. the physician re-sutured the impella site groin. on day three of impella support, a code stroke was called. the impella was removed due to the left ventricle thrombus and was switched to an intra-aortic balloon pump. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: it was noted that the patient was bleeding from all access sites. the cause of the bleeding is determined to be patient condition because of the bleeding from multiple sites. stroke/ thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4649 | 1220648-2025-47134 | 23208311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208311&pc=OZD | 09/03/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710567 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 76 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced persistent oozing that was managed with dressing changes. the patient also had a fever and was treated with antibiotics. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleed was not determined since product was not returned and insufficient clinical details were provided. the root cause of the infection was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks.
4650 | 1220648-2025-47135 | 23208388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208388&pc=OZD | 09/04/2025 | 10/02/2025 |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107786 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during use in the intensive care unit, the automated impella controller (aic) had an active placement signal alarm. the audible on the aic was not the normal sound of the yellow alarm. the impella cp with smartassist and the aic appeared to be functioning as expected. there was no harm reported. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4651 | 1220648-2025-47136 | 23208449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208449&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 104 KG |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with known tortuosity and lesion to the iliac artery, complicating the impella cp access/implant. the physician used a 25 centimeter 14 french peel away sheath, but the cannula would not make it past the lesion and exit the sheath. the physician exchanged for the 18 french gore dryseal sheath, but aortic stenosis prevented the crossing of the device into the ventricle. the complication was successfully managed with a balloon angiography of the aortic valve and the buddy wire was used to cross the cannula into the ventricle. the pump was placed with fluoroscopy guidance with appropriate waveforms and flows. the percutaneous coronary intervention was completed and the impella cp was successfully weaned. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was returned. difficult to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had known tortuosity and lesion to the iliac artery preventing successful delivery of impella. difficult/unable to insert through introducer: the cause of the issue was most likely patient condition since patient had known tortuosity and lesion to the iliac artery. physician opted for the 25cm 14fr peel away sheath. the impella cannula would not make it past the lesion and exit the sheath. patient anatomy or condition prior to therapy: the cause of the issue was most likely patient condition since patient had known tortuosity and lesion to the iliac artery. physician opted for the 25cm 14fr peel away sheath. the impella cannula would not make it past the lesion and exit the sheath. device history review: the device passed all the post sterile inspection checks.
4652 | 1220648-2025-47137 | 23208647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208647&pc=OZD | 09/03/2025 |  |  | 10/03/2025 |  |  | 09/03/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684311 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Pain | Sepsis | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced drainage and ventricular tachycardia. the patient was transfused three units of fresh frozen plasma, the pump was repositioned, and the patient was treated with milrinone and amio. the patient developed pneumonia and was treated with antibiotics. suction alarms occurred with patient movement. the patient remains on impella support. | [5 empty] | D6b, date of explant, has been added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the bleed was unable to be determined due to insufficient clinical details. the root cause of the arrythmia and sepsis was unable to be determined due to insufficient clinical details. the root cause of the pain was unable to be determined due to insufficient clinical details. the root cause of the positioning issue was unable to be determined due to insufficient clinical details. the cause of pump suction cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks.
4653 | 1220648-2025-47138 | 23208813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208813&pc=OZD | 09/03/2025 | 10/31/2025 |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711505 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 suffered a stroke. there was facial drooping and left sided weakness. this was treated by heparin. the patient survived the event. | [5 empty] | No product was returned for investigation. the root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4654 | 1220648-2025-31810 | 23208977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208977&pc=PYX | 05/06/2025 | 10/24/2025 |  | 10/03/2025 |  |  | 05/06/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026647341 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, a 'placement signal not reliable' alarm occurred. after fifteen days of rp flex support, care was eventually withdrawn, and the patient expired. | [5 empty] | D6b "if explanted, give date" corrected. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was not returned. the fm optical signal issue was changed to placement signal issue as the data investigation revealed it is the dp sensor, not the optical sensor contributing to the reported issue. based on the data log analysis, the root cause of the placement signal issue is most likely due to dp sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4655 | 1220648-2025-47139 | 23208987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23208987&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719851 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 111 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on support with an impella 5.5. upon exchange of the impella cp for an impella 5.5, placement signal alarm (psl) active alarm occurred while utilizing a double barrel technique. the psl alarm was ongoing since placement. after pulling the pulmonary artery catheter and central line, the patient went into atrial fibrillation. multiple suction alarms occurred overnight and into the morning. the physician was called to the beside and the impella 5.5 was repositioned to approximately 4.4. eventually, care was withdrawn and the patient expired. | [5 empty] | Correction d6a, d6b. the investigation of the reported failure mode has been completed. optical sensor issue: ongoing placement signal low alarms reported. no relevant patient condition or abnormality noted. no volume issues noted. the logs confirm placement signal low alarms and show initial placement signal value at the start of support was 0. the 5s parameters were stable through support. the cause of the optical signal issue was not established as no product was returned and limited clinical information was provided. pump suction: the data logs confirm suction alarms and correct trigger. clinical reported suction and repositioning attempt which did not resolve suction alarms. no patient volume issues were noted nor was ecmo used. the cause of the suction alarms was not established as no product was returned and limited clinical information was provided. paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4656 | 1220648-2025-47140 | 23209015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209015&pc=OZD | 09/03/2025 |  |  | 10/03/2025 |  |  | 09/03/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690306 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 64 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a catheter kink and pump stop. the catheter at the connection between red plug and the catheter became severely kinked, resulting in multiple pump stoppages. the pump was explanted and exchanged for a new impella 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned. the cause of the product damage (catheter kink) was user related. the cause of the pump stop was not determined due to no product returned and insufficient clinical details. the device passed all post sterile inspection checks.
4657 | 1220648-2025-47142 | 23209179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209179&pc=OZD | 09/03/2025 |  |  | 10/03/2025 |  |  | 09/03/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685078 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 64 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp suffered a brain bleed. a full body ct was performed in response to low hemoglobin and multiple brain bleeds were identified, in addition to air in the abdomen indicative of a perforation. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the vessel perforation and bleeding is most likely patient condition related since multiple areas of bleeding in patient's brain were shown and air was noted in the abdomen indicating high suspicion of perforation. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4658 | 1220648-2025-47143 | 23209315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209315&pc=OZD | 09/03/2025 | 02/18/2026 |  | 10/03/2025 |  |  | 09/03/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885741 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella 5.5 the supporting automatic impella controller (aic) generated a yellow alarm indicating a low battery after unplugging it from an outlet. the error self-resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 updated; the product has been returned. the investigation is still ongoing.
4659 | 1220648-2025-47145 | 23209461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209461&pc=OZD | 09/05/2025 |  |  | 10/03/2025 |  |  | 09/05/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655334 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female | 88 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that echocardiography revealed a 5cm bleed at the groin insertion site. no further information was provided on interventions performed for the bleeding. additionally, it was noted 4 days later, the patient expired on impella support. | [5 empty] | Given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. product status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H.11. added instructions for use as they were omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. major bleed: in the catheterization lab, echocardiogram measured five centimeters bleeding at the groin insertion site. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
4660 | 1220648-2025-47144 | 23209636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209636&pc=OZD | 07/08/2025 | 01/22/2026 |  | 10/03/2025 |  |  | 07/08/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636180 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 122 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced arrhythmia. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary peri-procedural adverse event/arrhythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device lot number: 1904213 device history batch subcomponent lot: n/a device history review the pump (b)(6) passed all post sterile inspection checks
4661 | 1220648-2025-47146 | 23209707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209707&pc=OZD | 08/30/2025 |  |  | 10/03/2025 |  |  | 08/30/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657638 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 108 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Asystole | Hemodynamic instability | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient presented to the hospital in stemi on (b)(6) 2025. the patient¿s proximal left anterior descending artery was revascularized with one drug-eluting stent. during the procedure, it was noted that a mid-sized diagonal branch had shut down after the stent was deployed, however the patient was stable and remained stable and was therefore released to the intensive care unit (icu). the following day, the patient had 8 out of 10 chest pain and abnormal electrocardiogram changes. the patient was placed on a nitroglycerin drip and the medical team requested the patient to be transferred to another hospital for escalation of care. the transfer was denied and overnight the patient decompensated. the patient was placed on an intravenous levophed drip to keep the patient¿s mean atrial pressure above 60 mm hg. the patient was then taken to the cardiac catheterization laboratory the following morning and an impella cp was placed to hemodynamically stabilize the patient. after impella cp placement, the patient was urgently taken to another hospital for further treatment. upon arrival, the patient was noted to have arrythmias resulting in suction alarms and nonperfusing rhythms. the impella cp position was confirmed to be in the correct position, however it was noted that the patient¿s laboratory results, and hemodynamics indicated a need for escalation of therapy. the following day, the patient was brought to the operating room (or) for an escalation to an impella 5.5 due to worsening end-organ dysfunction, a larger body surface area, and a cardiac index of 1.2 l/min/m². additionally, a temporary pacemaker was placed overnight due to ventricular standstill and asystolic events. during the removal of the cp and closure of the femoral access site, it was reported that the perclose device was unsuccessfully deployed. the decision was then made to do a cut down and surgical closure of the impella cp site. during the implant of the impella 5.5, the initial site of the right axillary was unsuccessful at the junction of the innominate, and imaging revealed a subtotal occlusion preventing passage of the device. the decision was then made to access the left axillary artery, and the impella 5.5 was inserted without difficulty and support was initiated with flows noted to be at 4 l/min. during the time of the impella cp removal and after impella 5.5 support was initiated, it was reported that the patient¿s activated clotting time (act) was above 700 seconds after a dose higher than expected was administered to the patient. the patient experienced oozing and bleeding from the impella cp site during device removal. the patient required blood products to be transfused and a protamine injection. suction alarms were present on the impella 5.5, and the p level dropped to p-4 with flows lower than expected at 1.9 l/min. imaging was performed and revealed a thrombus in the left atrium (la), which was noted to have been present at the start of the case. after the patent was volume resuscitated and volume status was stabilized, the pump was unable to be increased past p-4. the impella 5.5 position was confirmed to be correct, and it was also noted that the right ventricle function was normal. additionally, a thrombus was found with imaging at the inlet cage of the impella 5.5, and the la thrombus was no longer visualized. during this time, it was reported that there were periods of decreased left ventricle filling, and the patient experienced hypotension. heparin was re-dosed, and the pump was slowly able to be increased to p-6 with variable flows lower than expected at 2.6 l/min. the patient was released to the icu for rest and recovery, with a plan to monitor motor currents and to attempt to increase p level as able. the following day, the decision was made by the family and the medical team to withdraw care of the patient. the patient¿s care was withdrawn, and the patient expired. this complaint involves two impella devices: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Based on the information available, the impella 5.5 cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Medical safety reviewed the event. regarding the initially implanted impella cp, the event describes an arrhythmia event and major bleed resulting in intervention. this major bleed is due to known perclose failure after impella cp removal. this is a serious injury. the patient was escalated to an impella 5.5 experienced complications that included thrombosis that was noted to be in the left atrium potentially migrated to the lv., low pump flow requiring intervention, which may have added to an already complex situation. the patient expired the following day. in this case there is not enough evidence to exclude the impella 5.5 device as an associated factor. however, the heart team and family chose to withdraw care which led to the patient expiring. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, data logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical symptoms: arrhythmia, asystole, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. regarding clinical symptoms hemorrhage/bleeding, and hemodynamic instability, the cause of the was not determined due to insufficient clinical details. pertaining to the mechanical interaction with tissue, a surgical cut-down was performed for impella cp removal after an unsuccessful perclose attempt due to vessel calcifications. the cause of the mechanical interaction with tissue was most likely related to the patient's calcified vessel condition. pump suction, the pump was not returned for investigation. data log shows several suction alarms with a trending placement signal throughout the case run. there was no major trend reported in motor current or pump flow. no confirmed positioning issue was noted in the data log. the cause of the pump suction issue was not determined, as sufficient clinical information was not provided and the product was not returned for investigation. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
4662 | 1220648-2025-47147 | 23209912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209912&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703431 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 81 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella 5.5 demonstrated a ¿placement signal unreliable¿ alarm. nurse called to inform them that they will be muting the alarm and disabling the placement signal. support continued and there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. optical sensor issue: data log review showed a ~50 mmhg drop in placement signal (ps) over about 30 minutes followed by a more gradual decrease in ps until it was out of range. signal not reliable (snr) and contrast trended downward with the ps. the rest of the 5s parameters were fairly stable. a psnr alarm triggered when the ps went out of range. the optical signal issue was reported 42 days into the case, which is under the intended design life of 60 days per design traceability matrix. the root cause of the optical signal issue was most likely sensor silicone wear since the ps downward trend matches the known data log pattern for sensor wear. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4663 | 1220648-2025-47148 | 23209913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23209913&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714583 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 124 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed an impella cp to support the patient being worked up for an aortic valve surgery. the pump was placed and the physician was made aware of the suspicion for clot entrainment in the impella. the pump dropped low. the physician agreed given the patient's previous clotting instances and factor v leiden disorder. the decision was made by the physician to remove the current impella and not placed with another as the risk of another clot was extremely high. additional information was received. it was determined that the patient's factor v disorder was to be the etiology of the impella clot. no clot was observed on the impella after explant. it is believed that the impella clot was possibly embolized. however, the could not be determined conclusively. the patient had a clot that formed in his left coronary artery from the guide catheter prior to impella placement (which was the reason the patient had cardiac arrest). all arterial access sited were closed and the cardiac catheterization lab aborted as a precaution due to the patient's factor v hypercoagulation. systemic anticoagulation was started after the incident. it was determined that the patient was too high of risk for another clot to place another impella. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | H1 type of reportable event was erroneously entered on the initial manufacturer device report. | The root cause of the thrombosis/thrombus could not be determined due to insufficient clinical details. the pump was not returned for investigation. data log shows that flow dropped 10 minutes into the case run, accompanied by an active suction alarm. pump motor current dropped without any reported spike. no abnormalities were observed on the purge parameters. the cause of the low pump flow issue was not determined, as the device was not returned for investigation and sufficient clinical information. the device passed all post sterile inspection checks.
4664 | 1220648-2025-47150 | 23210029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210029&pc=OZD | 08/29/2025 |  |  | 10/03/2025 |  |  | 08/29/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728331 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support (mcs). consult for coronary artery bypass graft (cabg) surgery was completed, and the patient was deemed not a candidate. the patient was given the option of protected percutaneous coronary intervention, which was scheduled and completed the same day of the consultation. consult for coronary artery bypass graft (cabg) surgery was completed and the patient was deemed not a candidate. the patient was given the option of protected percutaneous coronary intervention, which was scheduled and completed the same day of the consultation. the patient was brought to catheterization lab for the elective protective pci. it was noted the patient had a known right coronary artery lesion and was staged for treatment. the left common femoral artery was accessed and an angiogram showed small but adequate vessel size; the decision made to proceed. the site was wired, preclosed and upsized to 14fr oscor sheath. impella prepped and inserted in standard fashion. however, during priming, the impella was recognizing the catheter as prior generation with gray power cable. the prime was aborted and restarted which was successful. upon support initiation it was noted ¿impella flow low¿ alarm, which self-resolved. however, flows were lower than expected for duration of case. all attempts to troubleshoot to no avail. a single access was used as the pci site; guide and wires were inserted. during this time some bleeding was observed from outer hub of peel away sheath. recommendations were given, and the physician proceeded with case as is. one unit of packed red blood cells was administered coronary angiogram show stenosis of the left main, proximal to mid left anterior descending, proximal circumflex and obtuse marginal 2 arteries which were treated with drug-eluting stenting and/or 1.5 rota burr. post intravascular ultrasound and angiogram showed excellent results. guide catheter and guidewires were removed. the impella cp was weaned off and removed. the access site was closed with existing perclose. the site was dressed and clean, dry, and intact. the patient was sent to the unit for recovery in stable condition. | [5 empty] | Patient-specific information a4: weight is unknown, device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This is one of three related reports. this report represents the pump and other reports will be submitted for the automated impella controller and introducer. h6 type of investigation, investigation findings, investigation conclusions and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. low pump flow: clinical details reported a "impella flow low" alarm after starting the pump followed by sustained lower than expected flows. no further clinical details were provided. data logs were reviewed and the alarm was confirmed. the pump was set to auto and the alarm was active for 10 seconds while the mean flow calculation was 2.34 liters per minute (l/min). the alarm triggers for impella cp when the pump is running in auto and the mean flow calculation is less than 2.5 l/min for 10 seconds, so the triggering criteria was met. additionally, the "impella flow reduced" alarm was active for the remainder of the case, further supporting that the pump was unable to achieve high enough flows for auto. during the case, the placement signal was generally stable and there were no other alarms. the product was not returned for investigation. the cause of the low pump flows could not be determined since insufficient clinical details were provided, data logs were inconclusive and the product was not returned for investigation. major bleed: clinical details reported bleeding from the orange portion of the peel away sheath hub after doing the single-access technique. the medical doctor continued with the procedure. the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the pump passed all post-sterile inspection checks.
4665 | 1220648-2025-47116 | 23210092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210092&pc=OZD | 09/03/2025 | 10/02/2025 |  | 10/03/2025 |  |  | 09/03/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719605 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male | 58 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The eu complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was observed that despite the intervention/stenting of the femoral artery, the pump was unable to deliver. the anatomical structure and the stenting obstructed the placement of the pump, leading to pump explant. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult complexity in femoral artery anatomy along with preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4666 | 1220648-2025-47152 | 23210139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210139&pc=OZD | 09/04/2025 | 10/02/2025 |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502501 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | Malposition of Device | Patient Device Interaction Problem | Poor Visibility |  |  | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient implanted with an impella cp. after implant, the pump showed shallow on transesophageal echocardiogram (tee) and facing towards the posterior wall. attempted to reposition but were unsuccessful, as the pump remained shallow and facing the mitral valve. the purge was changed to bicarbonate. the second day of support, the pump was repositioned after flows dropped. repositioning successfully improved flows. on the final day of support, the patient had acute agitation and coughing event with near sudden frank hematuria. lactate was sent and hemolyzed, flows dropped, and the device was shallow on tee, with the inflow just below the level of the aortic valve annulus. attempted to reposition but the pump continued to be tethered to the mitral valve apparatus. throughout support, the pump had poor positioning, consistently moving towards the mitral valve. the physician elected to proceed with explant in the setting of the improved left ventricle function. the impella cp was removed in setting of hemolysis and poor positioning consistently towards the mitral valve. the physician elected to convert to intra-aortic balloon pump for continued support. | [5 empty] | Correction: e4, h6 medical device problem code a090814 was erroneously entered on the initial manufacturer device report the investigation of the reported failure mode has been completed. no product was returned. device in wrong position and difficult to place or position - the cause of the device in wrong position and difficult to place was not determined due to no product returned and insufficient clinical details. mechanical interaction with blood - data logs were not recovered. the cause of the mechanical interaction with blood was not determined due to no product returned, no data logs recovered, and insufficient clinical details. low or blocked pump flow -the cause of the low flows was not determined due to no product returned, no data logs recovered, and insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4667 | 1220648-2025-47149 | 23210224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210224&pc=OZD | 09/03/2025 |  |  | 10/03/2025 |  |  | 09/03/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683691 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 66 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Hypovolemia | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 went into ventricular tachycardia, and the patient was resuscitated. a clot was surgically removed from the patient's abdominal cavity and the patient received multiple transfusions. the pump generated occasional suction alarms and one position alarm so 250 ml of albumin were given. as the patient was weaned from peripheral venoarterial extracorporeal membrane oxygenation (ecmo) the patient became hypotensive. there was noted swelling with a hematoma at the junction of the axillary site 5.5 and arterial ecmo cannula site. the swelling was managed with a pressure bag. bleeding was observed at the arterial cannula site. the patient survived. | [5 empty] | No product was returned for investigation. the cause of the bleed, hematoma, arrythmia, hypovolemia, and thrombosis was unable to be determined due to insufficient clinical details. the cause of positioning issue cannot be determined since insufficient clinical details were provided. the cause of pump suction cannot be determined since insufficient clinical details were provided. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4668 | 1220648-2025-47095 | 23210262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210262&pc=OZD | 08/11/2025 |  |  | 10/03/2025 |  |  | 08/11/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713062 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 103 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Arrhythmia | Fever | Low Blood Pressure/ Hypotension | Unspecified Infection | Pulmonary Edema | Hematuria | Thrombocytopenia | Thrombosis/Thrombus | [5 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support was being escalated to impella 5.5. the physician decided to do a coronary artery bypass graft (cabg) and escalation to direct impella 5.5 placement. once cabg done the impella 5.5 was advanced into graft with transesophageal echocardiogram (tee) guidance, but the physician was unable to advanced passed aortic valve. the impella 5.5 was then removed and flushed. the wire and catheter were advanced into the left ventricle (lv) and the impella was then advanced over the wire into lv with tee guidance without difficulty. the impella was started. the patient was transferred to intensive care unit. the patient went into atrial fibrillation with rapid ventricular rate (rvr) and vasopressin drop was started and the patient was converted to normal sinus rhythm. on day four of impella support, the patient began running a fever. antibiotics were started and cultures were sent. the patient was placed onto continuous renal replacement therapy (crrt). there was an episode of pulseless electrical activity requiring one round of cardiopulmonary resuscitation and atrial fibrillation with rvr requiring cardioversion. the pt experienced hypotension. concentrated, tea-colored urine was noted as the patient was given electrolytes and red blood cells for abnormal labs. the patient also was heparin-induced thrombocytopenia and thrombosis (hitt) positive. argatroban drop was administered. the patient also experienced pulmonary edema and the patient continued on impella support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation has been completed. no product was returned for investigation. delivery issue: the impella 5.5 s2 was attempted for insertion surgically via direct approach through the right artery. the cause of the delivery issue was not determined because the insertion was successful without any issues after flushing and reinserting the pump into the left ventricle. infection/ thrombosis/ arrhythmia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia/ fever/ hypotension/ hematuria/ pulmonary edema: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
4669 | 1220648-2025-47154 | 23210302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210302&pc=OZD | 09/03/2025 | 10/02/2025 |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649349 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Decrease in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a cp placed to support the patient admitted in with an acute myocardial infarction. the pump was placed but a purge leak was observed. the team replaced the purge cassette first, but cassette replacement did not resolve the issue. the team later observed there was a crack and the crack was taped and pump remains on. the patient age and lack of escalation had caused the team to leave the pump on. the patient remains on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The products, pump and purge cassette, were received and the investigation was completed. data log was not returned/could not be retrieved from the remote server. device history record review was completed, and pump passed all post-sterile inspection checks. purge cassette: no issue was found related to the purge cassette. damage was identified on the pump sidearm, with a lengthwise crack down the yellow luer, where the purge leak was reproduced. purge leak ¿ pump: the root cause of the purge leak was determined to be due to a damaged sidearm yellow luer. product damage (purge sidearm crack): the cause of the product damage was a crack in the yellow luer but the cause of the crack was unable to be determined. the supplier investigation of a similar complaint, indicated possible high tension in combination with disinfectants and certain drug ingredients like for example oil, alcohol, lipids but was unable to confirm." purge pressure low: the cause of the purge pressure low alarm was due to server purge leak from the damaged sidearm yellow luer. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Product complaint # (b)(4).
4670 | 1220648-2025-47155 | 23210306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210306&pc=OZD | 09/03/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709883 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male | 120 KG |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed thrombosis in the left atrium and left ventricle. there were suction and low flows. the patient expired on support and the attending physician noted there is no relation between the pump and the death. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device and data logs were not returned for review. the cause of low flow/suction was unable to be determined as limited clinical details were provided and no product & data were returned for review. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
4671 | 1220648-2025-47156 | 23210322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210322&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704157 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was placed but there was persistent blood leaking from the red plug that was taped and managed with bond. the leak was observed after the team had observed the need for a distal perfusion due to limb ischemia. the family and team made decision to withdraw care and patient expired after 23 hours rather than escalated to the 5.5 pump. | [5 empty] | Correction: b1, d4. the investigation of the reported failure mode has been completed. no product returned. red handle leak - the cause of the red handle leak/bleeding was not determined due to no product returned and insufficient clinical details provided. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4672 | 1220648-2025-47158 | 23210327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210327&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705075 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was escalated from a intra-aortic balloon pump to the impella cp due to worsening cardiogenic shock. it was noted that it was difficult to optimize flows due to small left ventricle and decreased volume status. the patients hemodynamics improved after correcting acid base balance and providing 2+ liters of fluid which then allowed impella flows to be increased to p-7. additionally the patient had severe gastrointestinal bleeding that required 2 packed red blood cells, 2 fresh frozen plasma and 2 cryo. however, unable to ever start heparin, patient in disseminated intravascular coagulation. hemodynamics and pressor requirements improved. team was suspicious of ischemic bowel. the cp was successfully weaned and explanted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4673 | 1220648-2025-47161 | 23210468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210468&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026731861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 64 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | False Alarm | Restricted Flow rate | Malposition of Device |  |  |  |  | Arrhythmia | Bradycardia | Ischemia | Vascular Dissection | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had bradycardic events leading to decreased flows, requiring placement of a temporary pacer. the pump began to get suction once out of automatic mode. after the peel away sheath was removed, the pump was left slightly shallow. an angiogram was taken of the femoral arteries, showing no distal perfusion to the right leg. a distal perfusion catheter was placed to restore flow. the patient was still having low pressures with intermittent suction on the cp. support was decreased to break the suction. flows continued to decline with suction and cardiopulmonary resuscitation (cpr) was initiated. the medical team could not stabilize the patient to even get them to the icu to start continuous renal replacement therapy, and determined placement of an impella rp was necessary. on bipella, placement monitoring was turned off due to inappropriate position in aorta alarms. placement of the pump was confirmed via echocardiogram (echo) and chest x-ray. the patient had increasing needs for medications. another echo was taken, showing the pump was slightly shallow, but the patient is small with a small left ventricle. suction events occurred, requiring 1 unit of red blood cells and medication. unable to pull fluid on crrt, had a sudden drop in platelets, and down trending mixed venous saturations. the pump was noted across the valve on echo. on the final day of support, the patient's pressures dropped after impella rp removal. went into atrial fibrillation. team waited for the patient to stabilize prior to removing the impella cp, and pressures increased dramatically. the right femoral artery was repaired via endarterectomy by vascular surgery to restore flow from the repositioning sheath and distal perfusion catheter. the pump was successfully weaned, site closed, and the patient survived. the patient was implanted with both right (rp) and left side (cp) support, so both the pumps are reported for all the events as its not clear on which device attributed to the events. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. low pump flow/suction: data logs were reviewed and the suction alarms were confirmed. data log review does not support it being patient related based on the stable placement signal (ps) while flows trended down. there were no motor current (mc) spikes to support biomaterial, but flows did become saturated at the end of the case. for these reasons, the cause of the pump suction/low flows couldn't be determined. false alarm/optical sensor issue: data logs were reviewed and a concentration of position in aorta alarms was confirmed ~8 hours after case start. all optical signal metrics were stable and within specification during the entire case. in reviewing one of the potential false alarms, it was confirmed that the triggering criteria was met for the alarm: mc modulation was less than 75 ma, ps pulsatility was greater than 20 mmhg, and there were no indications of the pump being in a suction condition based on dp values. product was not returned for investigation. since clinical details report that the pump was in proper position when the alarms were triggering, data log review confirmed the alarm triggering criteria was met, and product wasn't returned, the cause of the false alarms/optical sensor issue could not be determined. device in wrong position: clinical details report that the pump was intentionally positioned shallow due to the patient's size. based on this information, the cause of the device in wrong position was determined to be patient (anatomy) related. bradycardia/arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. vascular dissection: the cause of the injury was not determined due to insufficient clinical details. patient anatomy/condition prior to therapy: based on the clinical details provided, the cause of the patient anatomy/condition prior to therapy was determined to be patient related. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
4674 | 1220648-2025-47157 | 23210482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210482&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655268 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after placement of the impella cp support was initiated, the automated impella controller had multiple alarms. the physician stated the pump was in position on fluoroscopy but the pump had ¿ no flow¿ and had multiple alarms. the patient's status was critical/unstable. the decision was to implant a new impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was returned. low pump flow: patient had multiple alarms per clinical. after review logs, no suction alarms or placement signal alarms were observed. couple of impella in ventricle alarms were present. the cause of low pump flow cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4675 | 1220648-2025-47153 | 23210658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210658&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725521 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during a impella cp insertion there was access complications. it was noted that pre impella-with preclose occurred due to heavy calcium burden with noted heavy bleeding after deployment. manual pressure was held and the 14fr short sheath was inserted with hemostasis noted post insertion. the cp was successfully weaned and explanted. | [5 empty] | The investigation of the reported failure mode has been completed. major bleed : the cause of the injury is most likely patient condition related since heavy calcium burden with noted heavy bleeding after deployment. when deploying the preclose the physician felt that something wasn¿t right and the site began to bleed. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4676 | 1220648-2025-47162 | 23210984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23210984&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  | 01/31/2026 |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509624 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was bleeding, a suction event and a placement signal low alarm during impella cp patient support. while on support, there was a placement signal low alarm. the physician assistant ordered a fluid bolus for suction. additionally, the patient was bleeding from multiple sites. the nose and mouth were packed and continued to ooze. the impella groin developed a large hematoma which required manual pressure. the hematoma was managed with manual pressure and the long term repositioning sheath was advanced inside with better forward suturing. there was an angle matched dressing. there was packing applied to the mouth and rhinorocket to nose. multiple blood products were administered. in total, it was seven units of packed red blood cells, four units of cryoprecipitate and frozen fresh plasma. | [5 empty] | The investigation of the reported failure mode has been completed. optical signal issue - on the day of implant, placement signal (ps) low alarms occurred. pa ordered fluid bolus for suction. echo not taken at this time and positioning not confirmed. data log review confirmed suction and placement signal low alarms occurred without resolve. no motor current (mc) spike outside of p-level changes. ps trends down over the course of the case. the cause of the optical signal issue could not be determined due to no product returned, inconclusive data log review, and insufficient clinical details. pump suction - data log review confirmed suction alarms occurred without resolve. no mc spike outside of p-level changes. the cause of the suction alarms could not be determined due to no product returned, inconclusive data log review, and insufficient clinical details. access site adverse event - the cause of the access site adverse event could not be determined due to insufficient clinical details. patient codes: major bleed, hematoma, epistaxis - the cause of the major bleed, hematoma, and epistaxis could not be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4677 | 1220648-2025-47163 | 23211453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23211453&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1706584 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp/automated impella controller there was a placement signal low alarm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4678 | 1220648-2025-47164 | 23212492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23212492&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700054 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 80 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp the patient continues with bradycardia. ep consulted and considered possible pacer. there was rhythm changes and had to be taken for a temporary venous pacemaker. plan is to place a ppm/aicd today. the patient continued on support until successful explant and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. bradycardia: the root cause of the bradycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4679 | 1220648-2025-31811 | 23212600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23212600&pc=OZD | 04/08/2023 | 10/03/2025 |  | 10/03/2025 |  |  | 07/22/2025 | 10/03/2025 |  |  | 11/30/2024 |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293874 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. the complainant reported the impella was placed independently, and, upon arrival to the catheterization lab, placement signal was impaired and appeared to not function correctly. an impella replacement was recommended but declined by the physician. the motor current was showing good waveform, and placement was confirmed with x-ray. the patient expired after being on support for 25.27 hours. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation.the investigation has been completed. no product was returned. as per data log review, the placement signal not reliable alarm was found to be present after motor current spike. the trend of signal-to-noise ratio (snr) and contrast were found to be decreasing immediately after the event of motor current spike. the cause of the optical signal issue was most likely biomaterial ingestion with motor current spike observed affecting the snr and contrast parameters per data log review.a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer.all pertinent information available to abiomed has been submitted at this time.
4680 | 1220648-2025-31812 | 23213189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23213189&pc=OZD | 02/03/2024 | 10/03/2025 |  | 10/03/2025 |  |  | 07/25/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support and while on support there, were placement signal and left ventricle numbers unreliable alarms. the automated impella controller showed an erratic placement signal with unrealistic clinical numbers. support continued. no known patient harm or injury. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. device analysis: this console was tested after arriving in fs. unable to reproduce the reported failure mode. still found the signal not reliable was marginally passing. the front panel optical connector was contaminated. upon replacing the front panel optical connector with known good, the passing 5s parameter was obtained. the fringe pattern looks good, confirming no issue with the optical bench. the root cause for the placement signal was defective pump connector. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4681 | 1220648-2025-31813 | 23213222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23213222&pc=OZD | 08/29/2024 | 10/03/2025 |  | 10/03/2025 |  |  | 08/29/2024 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524611 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal to the pump was lost during shockwave therapy. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was investigated and the 5s parameters showed signal not reliable out of specification. clinical details mention shockwave was used. there was dried dextrose on the sensor. the pump was run and the placement signal not reliable alarm reproduced. the data logs confirm placement signal not reliable alarm and show 5s parameter pattern that aligns with a damaged sensor. the root cause of the optical signal issue was determined to be due to a damaged sensor. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4682 | 1220648-2025-31814 | 23213375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23213375&pc=PYX | 07/22/2024 | 10/03/2025 |  | 10/03/2025 |  |  | 07/22/2024 | 10/03/2025 |  |  | 01/31/2025 |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024311900 |  | 0046-0035 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp support and the placement signal was lost during attempted swan placement. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. data logs were reviewed and 5s parameters change at time of "placement signal not reliable" alarm #(b)(4) are consistent with a broken sensor face with raw optical signal dropping to zero. raw electrical placement signal remained stable at time of alarm. additionally, a small spike in motor current with a speed deviation and deviation in pump flow right before placement signal not reliable (psnr) alarm occurred. a larger motor current with concurrent speed deviation and deviation in pump flow occurred after psnr alarm. clinical details noted that at time of psnr alarm, physician was performing a wire insertion for swan placement. it was noted that placement signal was not visible on screen after alarm. the root cause of the optical signal issue was most likely due to a sensor face break. a trend chart was pulled for all impella rpsa pumps with failure mode of "optical signal issues" and a root cause of "damaged sensor". a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4683 | 1220648-2025-47166 | 23213404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23213404&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714696 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemoptysis | High Blood Pressure/ Hypertension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was in atrial flutter (af). pressors weaning as tolerated and patient remain on epinephrine and heparin. the patent was planned for explant, but remained on care for longer due to hypertension. hemoptysis noted and patient will be brought for ct. | [5 empty] | B3 date of event was revised as it was entered incorrectly on the initial report that was submitted. d6b added explant date. e1 revised initial reporter address line 1, zip code and zip code extension as they were entered incorrectly on the initial report that was submitted. h11 information was added as it was omitted from the initial report that was submitted. this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h10 for the related report number. investigation summary: the investigation has been completed. arrhythmia : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hemoptysis/hypertension : the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4684 | 1220648-2025-47167 | 23213552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23213552&pc=PYX | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026717903 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the patient was transferred from community to tertiary center on extracorporeal membrane oxygenation ( ecmo) after ecmo was weaned off patient had signs of right ventricle failure so decision was made to place rp flex. the pump was placed and the team explanted it after 26 hours due to placement signal loss. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: g1. the investigation of the reported failure mode has been completed. placement signal issue: the data logs confirm placement signal not reliable and show negative values for cvp before going out of range intermittently. this could be related to a patient condition related issue regarding patient volume. the raw optical sensor readings were also railing out of range. the pump flows were lower than expected for set p-level but the 5s parameters were stable and within specification. alarms stop when pump is turned down to p4. limited clinical information was provided on patient volume or support. the cause of the issue could not be determined as no product was returned and limited clinical information was provided. device history review: the device passed all the post sterile inspection checks.
4685 | 1220648-2025-31815 | 23213999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23213999&pc=PYX | 12/11/2024 | 10/03/2025 |  | 10/03/2025 |  |  | 12/11/2024 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274736 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 32 YR |  | Female |  |  |  | Death | No | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [18 empty] | IMPELLA RP FLEX | [18 empty] | R |  |  |  | Initial submission |  |  |  |  | 12/11/2024 |  |  |  |  | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported that placement signal not reliable alarm occurred when the patient arrived at the unit. motor current waveforms appeared to be stable at baseline. the team monitored hemodynamics and impella rp flows. the patient died on support after the patient was made a dnr and comfort care. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console was returned for investigation, and available log data is insufficient to determine exact failure mode and root cause. all pertinent information available to abiomed has been submitted at this time.
4686 | 1220648-2025-31817 | 23214057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23214057&pc=OZD | 05/31/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 05/31/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626847 |  | 1000080 |  | No |  |  | No |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. the complainant reported that the placement signal was reading lower than the a-line. the nurse was walked through disabling the alarm and support continued uninterrupted. the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows that the placement signal was 100/70 at the start of the case and gradually decreased to 60/20. a motor current spike was reported during the case without affecting the pump performance. several suction alarms were noted, with some fluctuations in motor current while running on a steady p-level. none of the 5s optical signal parameters were affected during the case. the cause of the optical signal issue was not determined since the product was not returned and sufficient clinical information was not provided regarding the patient condition during the case run. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4687 | 1220648-2025-31816 | 23214108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23214108&pc=PYX | 09/05/2024 | 10/03/2025 |  | 10/03/2025 |  |  | 09/10/2024 | 10/03/2025 |  |  | 02/28/2026 |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529162 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant had an rp flex delivered for support of a patient with respiratory failure and viral infection. the pump lost its placement signal. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were reviewed, and the placement signal not reliable (psnr) alarm was confirmed on 9/10. it was alarm 1051 that triggered, indicating the reported complication was a dp sensor problem rather than an optical sensor problem (for that reason, the investigated failure mode was different than the reported). this was further supported by the fact that throughout the case, the motor current was stable, the raw optical was stable, but the raw electric sensor was trending down, the placement signal was trending upwards cyclically, and the pump flow was mirroring the placement signal (ps). the central venous pressure (cvp) was also trending down towards the end of the case, but cvp isn't entirely calculated by the optical sensor, there is some reliance on the dp sensor, so this makes sense. product was returned and it passed electrical resistance testing. it was noted that part of the repositioning sheath was on the cannula. there were no damages noted to the dp sensor; however, the pump was unable to be run. based on the data logs, the root cause of the placement signal issue was determined to most likely be dp sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4688 | 1220648-2025-31818 | 23214272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23214272&pc=OZD | 03/08/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 03/25/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1162737 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal not reliable alarm occurred. impella flows remained the same and positioned was verified by transesophageal echocardiogram. the impella was in good position. the decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console logs from the reported day of event are consistent with complaint and show that after 3.5 days of support, console presented with placement signal not reliable alarm. the placement signal not reliable (snr) condition self-resolved after one minute but recurred 2 minutes later and remained unresolved for the duration of support. ltlog and rtlog data indicated that during this time optical snr dropped to 0 and contrast had unrealistic values oscillating between -3000 and +3000. this is consistent with loss of light received by optical bench ccd detector. the console was tested, and the issue was reproduced: initially there was no light coming out of optical pump connector, and optical bench lamp was found to be not producing light. after tapping on the lamp assembly, light came back. during console testing it¿s been determined that even with light temporarily restored, the optical system was out-of-spec due to low contrast (unrelated to this complaint), for which optical pump connector was partially responsible. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4689 | 1220648-2025-47159 | 23215968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23215968&pc=OZD | 09/03/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588706 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support they encountered a purge pressure low alarm. immediately after pump support began, a low purge pressure alarm occurred. even after replacing the purge housing, the alert did not improve, so support continued. the device has been in use in the icu for over a week, but the alarm has not improved. the patient is still on support of device. | The device has been in use in the icu for over a week, but the alarm did not improved. the patient expired. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information provided in: b1, b2, b5, d6b, d9, h1 and h6. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. purge pressure low: logs confirms the purge pressure low alarm throughout the case run. the product was returned and evaluated. no abnormalities were found upon visual inspection other than biomaterial in the purge gap. when the pump was run with a test purge cassette, the biomaterial from the purge gap cleared out, and the low purge pressure and high purge flow were reproduced, with purge pressure at approximately 125 mmhg and purge flow at around 30 ml/hr while running in a bucket of water. no leaks were observed. the issue was not determined on the returned product; however, this failure was related to the device but could not be traced to the root cause. device history review: the device passed all post sterile inspection checks.
4690 | 1220648-2025-47168 | 23216008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23216008&pc=PYX | 09/07/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026690547 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 64 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. the impella cp was having suction. this patient came in severely dehydrated and in severe metabolic acidosis. the continuous renal replacement therapy (crrt) was requiring more volume. red blood cells were given to function properly. it was not thought that the impella was causing this. the suction events were directly correlated to when they were pulling more fluid off with crrt. this issue was just temporary and did not continue throughout the entire run of bipella. the seizure like movements were likely due to lactic acidosis but was not confirmed either way and was a single event. as far as the positional alarms, the impella rp flex was positioned "optimally" and did not have any alarms associated as far as the associate clinical consultant was aware of. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the patient had a rp flex and cp support ongoing for a bipella patient. the patient was supported on the rp flex for 3 days and then the pump was explanted. during the support there were suction and alarming but, the pumps remained on for support. after being on minimal pressor, patient dropped pressures after rp removal, waited an additional 30 mins for patient to stablized, cp was removed and pressures increased dramatically with 8 of levo still on board. patient survived the procedure. |  |  |  |  | B3: revised date as it was entered incorrectly on the initial report that was submitted. b5: additional information was added. d1: revised brand name as it was entered incorrectly on the initial report that was submitted. d4: added model number as it was omitted from the initial report that was submitted. e1: added zip code extension as it was omitted from the initial report that was submitted. g3: the date on the initial report was reported incorrectly and should of reflected the date of 07-sep-2025. investigation summary: no product was returned for investigation. arrhythmia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review; the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the related impella cp is reported under mrn # 1220648-2025-47161.
4691 | 1220648-2025-47170 | 23216021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23216021&pc=OZD | 09/04/2025 |  |  | 10/03/2025 |  |  | 09/04/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654283 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient suffered a tamponade during the percutaneous coronary intervention (pci) and was given a pericardiocentesis and perforation was repaired. when peeling away the sheath after the pci the patient went into pulseless electrical activity (pea) and was given cardiopulmonary resuscitation the pump remained on but while in the icu there was a hematoma observed, the site was closed by the manta but, the bleed needed contralateral leg access and stenting along with 6 units of red blood cells. there was a presumed damage to the vessel from the sheath and access site. impella was explanted. | [5 empty] | B.3 revised date of event as it was incorrect on the initial report that was submitted. d.4 added model as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. hematoma/major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. arrhythmia : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4692 | 1220648-2025-31819 | 23216076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23216076&pc=OZD | 02/02/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 02/02/2025 | 10/03/2025 |  |  |  |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025565269 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. the day after implant the pump alarmed and the team disabled the placement signal. decision made to withdraw care and patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows that the signal not reliable (snr) was below 1500 before being inserted into the patient's body. later, the snr dropped close to 0 after the pump was placed inside the patient's body. the snr trend and placement signal correlated throughout the case run. impella was observed in the aorta/ventricular alarms were noted during the case run without any observed trend in motor current. as per the clinical data, the ps monitor was disabled after verifying the good position of the pump. the cause of the optical signal issue was not determined since product was not returned. refer (b)(4) for aic related issue. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4693 | 1220648-2025-31820 | 23216093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23216093&pc=PYX | 01/31/2025 | 10/03/2025 |  | 10/03/2025 |  |  | 01/31/2025 | 10/03/2025 |  |  | 08/31/2025 |  | 10/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025453554 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Eu complainant reported the patient was on impella rp support and during support, there were ¿placement signal not reliable¿ alarms. therapy was terminated due to missing heart recovery. due to poor prognosis the therapy was not further escalated. care was withdrawn and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. on the same day as implantation, placement signal not reliable alarm occurred, unable to be resolved. patient remained on support. 7 days later, therapy was terminated due missing heart recovery and poor prognosis. patient expired. the pump was not returned. data logs were not recovered. the cause of the placement signal issue was not determined since product and data logs were not returned and due to lack of clinical details. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4694 | 1220648-2025-31783 | 23198255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23198255&pc=OZD | 04/04/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 04/04/2024 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314437 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 pump with an associated automated impella controller lost placement/optical signal after two days of support. the pump remained on for support despite the loss. the patient eventually expired, but the patient's demise is captured in the records for the impella 5.5. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. console logs were consistent with what was reported in the complaint. instances of placement signal not reliable alarm observed in the imc logs. console was tested after arriving in field service. the reported placement signal not reliable issue was unable to be reproduced. the root cause of this issue was likely a defective optical bench, defective fiber optic cable connecting the optical bench to the pump receptacle, and/or a defective ribbon cable connecting the 4-in-1 to the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4695 | 1220648-2025-47093 | 23198346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23198346&pc=OZD | 08/09/2025 | 11/04/2025 |  | 10/02/2025 |  |  | 08/09/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587280 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 78 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult or Delayed Positioning | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that suction alarms, gastrointestinal (gi) bleeding, and positioning issues were encountered during impella 5.5 support. after five days of support, heparin was stopped overnight due to concerns for gi bleeding. a few days later, suction alarms occurred when attempting to increase the flow. the patient was transfused one unit of packed red blood cells. additionally, there was a problem repositioning the catheter through the blue suture hub. the impella catheter was unable to be advanced through suture hub, however, the repositioning button was still functional and capable of being depressed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | The pump was returned for evaluation. the cause of the bleed was not determined, as it occurred at a non-impella site, and no abnormalities related to the bleed were observed in physical product investigation. the cause of the positioning issue was a catheter kink, which was observed in physical product investigation. the cause of the pump suction was most likely patient related, as the patient required ecmo support and data logs indicate poor patient condition shown by the downward trend in placement signal and low pulsatility. the product was returned for investigation, and there was a catheter kink where the catheter lock had been positioned over. the cause of the catheter kink was not determined due to insufficient clinical details. the pump passed all post sterile inspection criteria.
4696 | 1220648-2025-31784 | 23198349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23198349&pc=OZD | 04/07/2025 | 11/19/2025 |  | 10/02/2025 |  |  | 05/22/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076066 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 117 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal not reliable alarm occurred on their automated impella controller with absent placement signals. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ic1043 was running 5.5 pump (b)(6) from 5/5 to 5/19 and on 5/10 the contrast oscillated between +/- 3276.8 as placement signal railed to 3000 mmhg (ic1043 ltlog 489409 osc contrast). both optical signals resolved shortly afterwards. there were other examples of oscillating contrast in the field (see attached). device analysis: the console was eventually returned for investigation. the oscillating contrast was not reproduced while running a test pump, the console performed as expected. root cause: the root cause of the intermittent unreliable placement signal and oscillating contrast could not be determined as the failure was not reproduced during testing. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the automated impella controller was running with a 5.5 pump from 5/5 to 5/19, and on 5/10 the contrast oscillated between +/- 3276.8 as placement signal railed to 3000 mmhg. both optical signals resolved shortly afterwards. the cause of the placement signal not reliable was not determined because the console was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4697 | 1220648-2025-31785 | 23198509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23198509&pc=OZD | 04/09/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 04/09/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED |  |  | US |  | 1484655 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that impella aic had placement signal unreliable alarm. no further information is available. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show multiple placement signal not reliable events triggered by invalid placement signal values, often coinciding with impella position unknown alarms. after testing the returned console it produced a defective optical bench is considered defective. the root cause of the aic issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4698 | 1220648-2025-47003 | 23198555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23198555&pc=OZD | 08/29/2025 |  |  | 10/02/2025 |  |  | 08/29/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711574 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 92 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on an impella 5.5 the patient had no suction but flows were subpar and albumin was administered for volume. later, no anticoagulation due to low platelets. it came down to p8 due to some suction alarms. anticoagulation continued to be held. the patient was bleeding from old chest tube sites and needed transfusions. the patient continued having suction on p6 so it was moved to p5. the patient was weaned from the pump successfully and survived. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. major bleed: patient had bleed from old chest tube. the cause of major bleed cannot be determined since insufficient clinical details were provided. pump suction: no product was returned. patient had suction alarms at p6 and was then moved to p5. after reviewing logs, multiple suction alarms were observed. the cause of pump suction cannot be determined due to insufficient clinical details. thrombocytopenia: the cause of clinical injury cannot be determined since insufficient clinical details were provided. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. h6 health effect - clinical code was updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4699 | 1220648-2025-47097 | 23198820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23198820&pc=OZD | 09/05/2025 |  |  | 10/02/2025 |  |  | 09/05/2025 | 10/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713894 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Decrease in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella 5.5 support they encountered low purge pressure and a pump stop. the patient had a sustained ¿purge pressure low¿ alarm and leaking was noted at purge disc/filter within retainer. the medical team planned to seal the retainer and increase the solution to d20. signs of pump failure were also noted. the company conducted some troubleshooting for the pump stop issue, however, as the purge pressure low alarm continued, the decision was made to explant the device. this is one of two reports. this report represents the pump. another report was submitted to represent the purge cassette. refer to section h.10 for the related report number. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | D9: added date of device return. h3: device was returned for evaluation. h6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. the clinical details describe how there were purge pressure low alarms, with tight connections and d20 added to try to resolve the leaking but that did not work, so the sidearm was sealed to prevent leaking onto the floor. an internal call with the field representatives confirmed this narrative as there was also no pulling on the patient cable or any purge connections, as it occurred very rapidly. the data logs show that the purge pressure and purge flow were stable throughout most of support until on 5-sep the purge flow rose from ~10ml/h to 30ml/h over the span of 20 minutes while the purge pressure trended down to the 300's, triggering purge pressure low alarms. this was not fixed by changing the cassette, tightening connections or adding d20. the pump was kept in for another 14 hours until it was explanted as there was no resolution of the leaking. during this time motor current remained stable. the returned product has the purge cassette tubing still connected, and when that was attempted to be unscrewed, it broke as there was a sealant around it. there was also a sealant around the sidearm retainer which had to be dremeled through to see the filter. there was an excessive amount of fluid pooled in the casing. the bottom filter stem had broke off from the filter membrane area and was cracked with cloudy regions on the fracture surface, indicating deformation prior to the fracture. it was not broken where the tubing in the patient cable is assembled to the sidearm. the root cause of the purge leak - pump and purge pressure low alarms is due to a damaged sidearm filter. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4700 | 1220648-2025-31786 | 23200331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23200331&pc=OZD | 05/22/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 05/22/2023 | 10/02/2025 |  |  | 04/30/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330850 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported abnormal placement signal readings during support, alleged against the automated impella controller. the failure investigation team found the actual issue relates to the impella 5.5 pump. the alarm was muted and support proceeded with no adverse impact, however, care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue is most likely due to the patient¿s condition. all pertinent information available to abiomed has been submitted at this time.
4701 | 1220648-2025-47098 | 23200363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23200363&pc=OZD | 08/15/2025 |  |  | 10/02/2025 |  |  | 08/15/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668075 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 60 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Hematuria | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted to support a patient after a mitral valve repair. while on support, multiple blood products were transfused in response to concerns of bleeding. additionally, continuous renal replacement was started. a computed tomography (ct) showed a stable small right parietal lobe hematoma and decreased volume subdural hemorrhage from a ct completed the previous week. impella support continues. | [5 empty] | The cause of the renal failure, hematoma, major bleeding, and hematuria was not determined due to insufficient clinical details. the root cause of the cerebrovascular accident was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. the event logs were returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4702 | 1220648-2025-47100 | 23200711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23200711&pc=OZD | 07/10/2025 |  |  | 10/02/2025 |  |  | 09/02/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695585 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Cardiac Tamponade | Cardiogenic Shock | Pericardial Effusion | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant received via abiomed long term study database that there was two adverse event. adverse event number one was reported as cardiac tamponade with relationship listed as possibly related "after the impella was removed during the impella supported pci procedure, the subject became progressively hypotensive with unclear etiology. bedside tte revealed a moderate sized pericardial effusion. a pericardial drain was placed with immediate aspiration of 600cc of bright red blood which resulted in immediate improvement to baseline. repeat tte revealed no residual effusion, and the drain was sutured in place. drain remained in place until (b)(6) 2025 when it was pulled." adverse event number two was reported as : "approximately 5 minutes after a pericardial drain had been placed, the subject because hypotensive for a second time. 100cc of blood was aspirated from the pericardial space with no resolution in hypotension. repeat tte demonstrated no residual effusion. retroperitoneal bleeding and delayed/missed extravasation from the coronaries were ruled out as causes of hypotension. iv volume was pushed and the subject was considered ¿relatively stable¿ on small dose phenylephrine. subject was transferred to the icu for further management of cardiogenic shock. in icu switched from phenylephrine to levophed gtt for bp support. levophed stopped on 12jul2025 at 16:53." the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be patient condition because of the bleeding in multiple locations with retroperitoneal bleeding along with access site bleeding. retroperitoneal hemorrhage/ cardiac tamponade/ hypotension/ pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
4703 | 1220648-2025-47101 | 23201168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23201168&pc=OZD | 09/04/2025 |  |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  | 02/28/2026 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478551 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 102 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing from the access site. the patient received a blood transfusion. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
4704 | 1220648-2025-47102 | 23201231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23201231&pc=OZD | 09/04/2025 |  |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716788 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis. repositioning the pump was unsuccessful so the pump was explanted and saturated flows were observed. the team placed a new impella cp and support proceeded. | [5 empty] | The impella cp was returned for evaluation. the returned pump was inspected and there were no related abnormalities. there was dried dextrose and dried blood which cleared after soaking. the pump was run in the lab and saturated flows did not reproduce. the cause of the saturated flow was biomaterial ingestion as evidenced by log pattern and returned pump not reproducing saturated flow. the cause of the hemolysis was determined to be biomaterial ingestion. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Product complaint # (b)(4). the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4705 | 1220648-2025-47099 | 23201362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23201362&pc=OZD | 09/03/2025 |  |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713136 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. the pump was disposed and was unable to save the pump for further investigation. the information obtained was a medical doctor conversation prior to wean/explant of the impella 5.5. the hematoma that was mentioned by another clinical representative was present and subsided after removal of the impella cp, when the patient was escalated to impella 5.5. a computed tomography confirmed no presence of a retroperitoneal bleed which is what the team was concerned about during case support due to continued blood transfusions. there was no culprit bleed found. no hematoma or problems were from the impella 5.5 per medical team. | It was reported that a patient was escalated from a impella cp to an impella 5.5. the escalated was successful. during impella 5.5 support, there was suspected bleeding from old groin sites, one of them being the old impella cp site. a computed tomography was done and no retroperitoneal (rp) bleed was found on either side. multiple transfusions were been given. continuous renal replacement therapy and a ventilator was required due to continued transfusions. vascular was consulted, however, no surgical intervention was needed. anticoagulation has been stopped due to slight bleeding noted on the computed tomography. there was also noted on field service notes of continued anemia requiring multiple blood transfusions. in addition to doctor's comments noted of suspected rp bleed, since transition to impella 5.5. the impella 5.5 was successfully weaned and explanted. |  |  |  |  | B.5 additional information was received. h.6 code f01 and e0506 was removed due to new information that was received. the investigation has been completed. no product was returned for investigation. anemia: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4706 | 1220648-2025-31787 | 23201785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23201785&pc=OZD | 08/28/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 08/28/2023 | 10/02/2025 |  |  | 05/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352562 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Convulsion/Seizure | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the device exhibited an optical signal issue. continuous false impella in aorta alarm, with pulsatile motor current. additionally, the patient experienced a seizure during treatment. no further information is available. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. product was investigated and no damage or abnormalities were found. the pump passed optical continuity testing and the 5s parameters were within normal ranges. no biomaterial found on the impeller. the logs confirmed the impella in aorta alarms throughout the case with no aortic placement signal. the alarms were occasionally accompanied by suction alarms or p-level at p-2. the root cause of the optical placement signal issue was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
4707 | 1220648-2025-31788 | 23203072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203072&pc=OZD | 07/15/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 07/15/2023 | 10/02/2025 |  |  | 05/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345271 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signaling issue occurred to a patient on impella 5.5 support. imaging showed that impella was in the proper position. the device was advanced slightly to ensure that the device was in proper place. there was no harm to the patient as a result of this incident, however, care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was most likely patient condition related with several impella position unknown alarms and false aorta alarm based on data log and clinical review. all pertinent information available to abiomed has been submitted at this time.
4708 | 1220648-2025-31789 | 23203086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203086&pc=OZD | 01/04/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 01/04/2024 | 10/02/2025 |  |  | 09/30/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418760 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and there were placement signal issues. the patient had left and right ventricle tears that were not impella related. the patient was crashed onto venoarterial extracorporeal membrane oxygenation, and the placement signal was lost. staff noticed a placement signal not reliable alarm that could not be fixed despite troubleshooting. transesophageal echocardiography confirmed impella positioning. the patient eventually expired after the family decided to withdraw care. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the most likely root cause of the optical signal issue was biomaterial. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4709 | 1220648-2025-31790 | 23203094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203094&pc=OZD | 05/22/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 05/22/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288834 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, a placement signal not reliable alarm occurred. the console was not switched. additionally, the patient experienced a gastrointestinal (gi) bleed. the physician decided to perform a gi resection and ostomy to treat the gi bleed, heparin was withheld. additionally, the patient experienced a hematoma and arrythmia during treatment, resolution unknown. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs confirms that all signals received from the optical bench were confirmed. console was returned for investigation. the root cause of the intermittent unreliable placement signal and oscillating contrast could not be determined as console was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4710 | 1220648-2025-47094 | 23203120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203120&pc=OZD | 08/09/2025 |  |  | 10/02/2025 |  |  | 08/09/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695335 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 70 KG | Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Unspecified Tissue Injury | [10 empty] | Death | Hospitalization | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 encountered high purge pressures, infection, epistaxis and oral bleeding, skin wounds and gastrointestinal (gi) bleeding. the patient's bed was changed to a ¿sand bed¿ to prevent further skin breakdown. the high purge pressure was resolved by administering tissue plasminogen activator. the patient's blood cultures tested positive for clostridioides difficile so antibiotics were administered. one unit of blood was transfused in response to the gi bleed. later, the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B.3 revised date as it was previously entered incorrectly on the initial report that was submitted. the investigation has been completed. no product was returned for evaluation. infection: clinical details noted that the patient was positive for clostridioides difficile while on impella support. in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleeding/epistaxis/soft tissue injury: clinical details noted that a patient had epistaxis, oral and gastrointestinal bleeding, and skin breakdown during support. the cause of the injury was not determined due to insufficient clinical details. purge pressure high: data logs were reviewed and long term logs showed a gradual rise in purge pressure over approximately one hour on 24-aug-2025 before "high purge pressure" alarms were triggered. purge pressure remained high for approximately two hours before purge pressures resolved. clinical details noted that tissue plasminogen activator was administered and purge pressures were resolved by the next day. the cause of the high purge pressure was due to biomaterial buildup based on data log analysis and clinical details. device history review: the pump passed all post-sterile inspection checks.
4711 | 1220648-2025-47103 | 23203213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203213&pc=OZD | 09/03/2025 |  |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702543 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 90 KG | Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed hemolysis. the patient was treated with fluid and diuretics. after a percutaneous intervention was performed a thrombosis was observed, and a thrombectomy was performed. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Investigation summary: thrombosis: the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood. there is red and pink urine throughout support in which volume was given for the red urine on the 3rd but then they were diruresising the patient on the 4th. the data logs showed that the motor current and placement was stable throughout support but there was a slight decrease in the placement signal from about 120mmhg to 80mmhg on the 3rd. flows are stable and there are no positioning issues seen. due to lack of data logs or sufficient clinical information, the root cause of the hemolysis cannot be determined. device history review: the impella cp passed all post sterile inspection checks.
4712 | 1220648-2025-47105 | 23203276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203276&pc=OZD | 09/02/2025 |  |  | 10/02/2025 |  |  | 09/02/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692915 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Difficult to Advance | [6 empty] | Ischemia | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was difficult. after the pump was successfully implanted the patient exhibited dusky fingers and a drop in blood pressure. the patient was transfused one unit of packed red blood cells, was treated medically, and the pump was replaced with an impella rp flex. | [5 empty] | The cause of the delivery issue is most likely use related since the anatomy was difficult to access the aortic valve with just tee. after several attempts, fluoroscopy was added and ventricular access was achieved. the pump was implanted per instructions for use. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the cause of the hemodynamic instability was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4713 | 1220648-2025-47106 | 23203288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203288&pc=OZD | 09/02/2025 |  |  | 10/02/2025 |  |  | 09/02/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664083 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp the patient was actively bleeding. manual pressure and aggressive angle matching in use. assessed by the physician and purse string suture placed and redressed, site hemostatic with aggressive angle matching. a safeguard was placed by the nurse. labs were drawn, inr 5.1, lactate 2.3 and heparin was stopped. there was also oozing from pac site also. all lfts extremely elevated, bnp 2300. co 4, ci 1.9, svo2 59%. increased milrinone to 0.5. plans to diurese. voided brownish colored urine. transthoracic echocardiography ordered and ump in good position, measuring at 4.9cm, pulled back by the physician to 3.7cm and locked at 86cm. plan to rest on cp overnight and reassess. team will reach out if assistance is needed. urine remains brownish and dark amber. dropped to p7 and p5. the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d6b the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: the cause of major bleed was most likely patient movement related since the bleeding occurred after transfer mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. hematuria/hemostasis: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4714 | 1220648-2025-47107 | 23203298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203298&pc=OZD | 09/02/2025 |  |  | 10/02/2025 |  |  | 09/02/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135098 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced battery issues on the supporting automated impella controller (aic). the aic battery was 50% when not plugged in. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4715 | 1220648-2025-31791 | 23203417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203417&pc=OZD | 01/21/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 01/21/2024 | 10/02/2025 |  |  | 09/30/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401294 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that impella 5.5 position was unknown and impella position in aorta alarms occurred. the alarms are not conducive to the echocardiography showing impella in proper placement and contraction of the left ventricle. the patient eventually expired while on impella support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issues and impella positioning alarms cannot be determined with no data logs or pump returned for evaluation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4716 | 1220648-2025-31792 | 23203535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203535&pc=OZD | 09/14/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 09/14/2024 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502177 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, it was reported that aorta placement signal (ps) is not correlating with art line pressures triggering placement signal low alarm (psl). nurse requested that alarms be disabled. instructions were provided to disable psl audio. additionally, the patient experienced frequent arrhythmias, the resolution is unknown. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the product was not returned. the logs were returned for only the last 4 days of the case. the issue was reported on the 4th day of support out of 22 total days. there were no "placement signal low" alarms present in the last 4 days of logs. the fm "placement signal issue" was updated to "optical signal issue" because the impella 5.5 does not have a dp sensor. the root cause of the optical signal issue was not determined since the product was not returned and insufficient data was provided. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4717 | 1220648-2025-31793 | 23203584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203584&pc=OZD | 03/14/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 04/25/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693068 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [5 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there was an optical signal issue on the automated impella controller. the complainant had a concern regarding a potential optical sensor issue due to the placement signal waveform being incongruent with the patient¿s arterial line pressures. additionally, the patient experienced an rp bleed, the customer stated that the rp bleed resolved. also, the customer stated that the patient experienced hemolysis, the resolution was not provided. no information was provided on the resolution. support continued. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. partial console logs obtained from the cloud are mostly consistent with complaint and show multiple placement signal not reliable (psnr) events (opm signal invalid). no conclusion can be made based on this incomplete data. console had not been returned for this investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4718 | 1220648-2025-31794 | 23203590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203590&pc=OZD | 07/12/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 07/12/2024 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503009 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and there were placement signal not reliable alarms. motor current was pulsatile and flows unchanged. staff ordered echocardiography, and placement signal monitoring was disabled. the staff monitored and support continued. the patient eventually expired after care was withdrawn. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no problem was found with the pump, and it is apparent that the automated impella controller was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4719 | 1220648-2025-31795 | 23203925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203925&pc=OZD | 02/27/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 02/27/2023 | 10/02/2025 |  |  | 07/31/2024 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023249956 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ventricular Fibrillation | [9 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a purge cassette leak during support that was resolved after removing/replacing the cassette. a placement signal not reliable alarm occurred days later, however, support proceeded uninterrupted. additionally, the customer reported that the patient experienced bleeding from the mouth and nose that required multiple units of blood. the nasal and nose passages required cauterization. the patient also had a fever/infection and was prescribed antibiotics. the patient also experienced rounds of ventricular fibrillation; the physician tried to reposition the device and this did not resolve the issue. care was withdrawn and the patient died. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the purge cassette issue and the optical signal issue has been completed. the cause of the priming issue was an rfid programming issue. the cause of the optical signal issue was sensor silicone wear. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4720 | 1220648-2025-31796 | 23203968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23203968&pc=OZD | 08/11/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 08/11/2024 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087214 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. while on support, the automated impella controller (aic) displayed placement signal not reliable alarm occurred. the audible alarm was turned off. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show that after 13 days of operation, console presented with placement signal not reliable alarm. console was tested in danvers, but the issue was not reproduced, and console¿s optical system operated within specification. the pattern of optical signals recorded in aic logs is consistent with intermittent loss of light, due to malfunction of either lamp or optical bench electronics. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4721 | 1220648-2025-47108 | 23205855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23205855&pc=OZD | 09/02/2025 | 11/06/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689094 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during support on impella cp on september 2nd, 2025, the physician was at the step pulling out the peel away sheath and right after inserting the repositioning sheath into the groin access. a massive bleeding in the groin occurred very quickly. the patient got into hemorrhagic shock and received directly blood donations. afterwards patient immediately received a cat scan(ct). after ct, the surgical team were immediately aware and involved. the impella was removed by surgical closure. the surgical physician saw no harm of the femoral artery. after treating the bleeding and groin closure, patient is stable and moved to a normal ward. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was returned therefore d9 was updated and h6 type of investigation and investigation findings. | Correction: e4. the investigation of the reported failure mode has been completed. the pump and cassette were returned for investigation. major bleed/access site adverse event: the product was returned for investigation, and no kinks or abnormalities were observed. the cause of the access site bleed was use related as activated clotting time (act)s were above 180 at the time of bleeding. device history review: the device passed all the post sterile inspection checks.
4722 | 1220648-2025-47109 | 23205939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23205939&pc=OZD | 09/03/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719339 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was having an non-st-segment elevation myocardial infarction percutaneous coronary intervention in the cath lab. left anterior descending artery became 100% occluded and patient coded. while chest compressions were being performed decision was made to place impella cp. while performing chest compressions, 14fr sheath was inserted. impella cp inserted. peel away sheath removed. repositioning sheath was inserted and secured with sutures. point of care arterial blood gas test showed hemoglobin had dropped to 5. four units of blood products were rapidly infused. impella flows couldn¿t go above 2 without suction. patient taken to cat scan to rule out hemothorax. there was a large hematoma around the repositioning sheath. the family made the patient a dnr, and the patient expired. duration of impella cp support was for two hours. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. the product was not returned for investigation. hematoma: the cause of the hematoma was most likely user related, as the acts were high which potentially caused the bleeding. low flow/suction: data log review showed down trending placement signal, that briefly resolved before losing magnitude and pulsatility throughout the case. there were a couple motor spikes with motor speed deviating down. motor current pulsatility decreased throughout the case due the loss in placement signal (ps) pulsatility. the cause of the low flow and suction was most likely related to the patients condition. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4723 | 1220648-2025-47112 | 23205941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23205941&pc=OZD | 09/03/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | +1(978)739-0590 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male | 75 KG |  | White | Injury | Yes | Yes | Failure to Read Input Signal | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had an aic replaced due to purge/cassette case start issues. the aic was replaced and support proceeded on the cp. | [5 empty] | Correction: e4 was inadvertently omitted on the initial manufacturer device report h5 and h8 was erroneously entered on the initial manufacturer device report. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4724 | 1220648-2025-47110 | 23205942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23205942&pc=OZD | 09/03/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026738313 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 64 KG |  |  | Death | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Ischemia | Vascular Dissection | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support on impella cp on (b)(6) 2025, the per close failed and tore the artery causing bleeding. peel away sheath was left in with direct primary care and rep sheath at the hemostatic valve and hemostasis was achieved. due to the artery dissection, peel away sheath was left in with a donor sheath from the impella peel away short sheath. vascular will be consulted in the am. pulses found in distal pt, but not at. since receiving volume and dobutamine, suction alarms present in the cath lab have subsided. plan is to let the patient rest overnight and reevaluate next steps in the am. access site is cardiac distress inventory. care was withdrawn and the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. the pump was not returned for investigation. data log review was not required for this case run. access site adverse event (hemorrhage/bleeding): the cause of the access site bleeding issue was most likely unintentional use-related, based on the given clinical details. mechanical interaction with tissue (vascular dissection): the cause of the vascular damage issue was most likely unintentional use-related, based on the given clinical details. clinical symptoms (ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4725 | 1220648-2025-47113 | 23205952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23205952&pc=OZD | 09/02/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1542953 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during support on impella cp on (b)(6) 2025, the field services informed the local team that biomed created this complaint requiring attention: ¿suction alarms/placement signal low, would not resolve. mean arterial pressure was in the 60s, while the aic was reading in the 20s.¿ | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown. | D9: updated the devices return information.
4726 | 1220648-2025-47114 | 23206364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206364&pc=OZD | 09/02/2025 | 11/18/2025 |  | 10/02/2025 |  |  | 09/02/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1736233 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 76 KG | Non Hispanic |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a console with a ac disconnection notification during the 5.5 pump support. no harm or injury note. the team plugged in console and ac power alarm remained. | [5 empty] | The console logs align with the reported complaint. the console was returned for further investigation without the ac power cord. the aic was inspected with a known-good power cord; no ac-power-related issues were observed. root cause: the reported issue of the "ac disconnected notification, console plugged in and cont to get ac disconnected notification" was most likely caused by a defective ac power cord set. however, the exact cause could not be determined because the ac cord was not returned for investigation, and the issue could not be reproduced. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. d6a and d6b implant and explant dates were added as they were omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
4727 | 1220648-2025-47115 | 23206365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206365&pc=OZD | 09/02/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 09/02/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677042 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 100 KG | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump supporting the patient with cardiomyopathy and awaiting heart transplant. the patient has been on support to date for 83 days. during the support the pump has had persistent suction resolved with patient movement/standing and infusion of fluids. additionally, there has been hemoptysis that prompted discussions of heparin removal, which did not take place. the pump also had ps issues that triggered false in aorta alarms as well as sleeve damage. none of these malfunction has prompted pump explant. the patient still awaits transplant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. product damage (anticontamination sleeve split/torn) - the cause on the product damage could not be determined as no product or images of the damage were returned. optical signal issue -the cause of the optical signal issue was patient condition related due to the patients noted odd anatomy causing the pumps positioning to not be adequate. the patient also had volume issues throughout case. pump suction - the cause of the pump suction was patient condition related due to the patients noted odd anatomy causing the pumps positioning to not be adequate. the patient also had volume issues throughout case. bleeding: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4728 | 1220648-2025-47117 | 23206379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206379&pc=OZD | 09/03/2025 | 10/14/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701799 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 50 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella cp support developed hematoma and bleeding at the access site after moving, rolling the patient, and taking knee immobilizer off. to manage bleeding, the suture pad was stitched flat. the patient had coronary artery bypass graft done and the physician decided that the patient didn¿t need the impella after surgery, so impella was surgically removed. patient is stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. access site adverse event/major bleed/hematoma: returned product was investigated, and there were no abnormalities noted on the repositioning sheath. it was noted, however, that the sterile sleeve had torn off of the touhy-borst valve. this was not reported, and deemed to be unrelated to the access site bleed. the reported contributing factors of the access site bleed were all related to the use of the product: high activated clotting time (acts), improper positioning of the repositioning unit, and patient movement. for this reason, the cause of the access site bleed was determined to be a use related issue. device history review: the device passed all the post sterile inspection checks.
4729 | 1220648-2025-47118 | 23206380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206380&pc=OZD | 09/03/2025 | 12/23/2025 |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026731965 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a cp pump placed for the support of a patient admitted in with ventricular fibrillation arrest, acute myocardial infarction with cardiogenic shock and coronary microvascular disease. the pump was placed and was supporting despite the purge cassette failing and replaced due to de-airing. the patient also developed a hematoma and hemolysis. the pump was explanted and escalated to the 5.5 pump after 2 days. | [5 empty] | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: no product returned. mechanical interaction with blood - after pump implant, within 8 hours of pump implant, patient noted to have red urine with a lactate dehydrogenase (ldh) of 1524. pump positioning noted to be correct and no known troubleshooting efforts were noted. within 10 hours, patient urine noted to be yellow. ldh not collected for remainder of pump usage. data logs show suction alarms occurring around the time hemolysis was noted. no suction alarms occur after the p-level was reduced below p-7 and no suction alarms occur after hemolysis noted to be resolved. the cause of the mechanical interaction with blood was most likely patient condition related due volume issues as suction alarms occurred at p-7 around the time hemolysis was noted. the pump passed all post sterile inspection checks. | Upon review of the complaint record, by the engineering team, it was found that the hematoma occurred during a sheath transfer prior to pump insertion. therefore, the hematoma was removed from the coding of the pump ip. a new ip was created for the introducer to capture the hematoma at the access site. the insertion kit is addressed under manufacturer report number (mwr-(b)(4). | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4730 | 1220648-2025-47119 | 23206382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206382&pc=OZD | 09/03/2025 |  |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685014 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Death | Hospitalization | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a non-st-segment elevation myocardial infarction patient was found to have severe left main disease coded when engaging vessel, decision made to place impella cp. the impella was placed per protocol without issues. during support it was noted bleeding from groin site requiring multiple dressing changes, no blood products were needed. however, interventional cardiology called to bedside, waiting on plan. in addition to plasma free hgb was also elevated with pink tinged urine. p level reduced to 6 and nurse noted clearing of urine. ultimately the patient continued to decline, coded and expired on support. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: bleeding from groin site not improved and groin site required multiple dressing changes. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4731 | 1220648-2025-47120 | 23206386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206386&pc=OZD | 09/03/2025 |  |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668110 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 106 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with st elevation and symptoms of shock decision to implant impella 5.5. the 5.5 was successfully placed. during support it was reported that shortly after the patient was extubated the patient experienced a stroke. the patient remains on support. | [5 empty] | B7 added information that was omitted from the initial report that was submitted. d6b explant date was added. h11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ investigation summary: stroke : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4732 | 1220648-2025-47121 | 23206389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206389&pc=OZD | 09/03/2025 |  |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677352 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation and ventricular tachycardia. the patient was defibrillated and treated with epinephrine, amiodarone, and lidocaine. additionally, the pump generated a suction alarm. in response, the pump was repositioned and the patient was treated with lactated ringer's and albumin iv. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned. the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of pump suction was most likely patient condition related since giving fluids resolved the suction issues per clinical. the device passed all post sterile inspection checks.
4733 | 1220648-2025-47123 | 23206392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206392&pc=OZD | 09/03/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711440 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 109 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient experienced ventricular fibrillation and major bleeding during impella 5.5 support. there was a continued gastrointestinal bleed and all anticoagulation was off. there were repeat interventions by the gastrointestinal doctor and units of packed red blood cells were required. additionally, the patient had rectal bleeding. clips were implemented, which reduced the bleeding. two units of blood products were administered. additionally, the patient continued to have episodes of ventricular fibrillation. care was withdrawn and the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: persistent gi bleed despite repeat interventions. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. ventricular fibrillation: the root cause of the ventricular fibrillation was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4734 | 1220648-2025-47122 | 23206393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206393&pc=OZD | 09/03/2025 |  |  | 10/02/2025 |  |  | 09/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1344266 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during the cath lab impella training automated impella controller couldn¿t be started and just a blue screen appeared. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4735 | 1220648-2025-47124 | 23206394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206394&pc=OZD | 09/04/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712266 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support developed a gastrointestinal bleed. systemic anticoagulants were held. two days later, gastrointestinal (gi) scope was done with a clip placed, and bleeding was resolved. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned. major bleed/ retroperitoneal hemorrhage: the cause of retroperitoneal hemorrhage cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4736 | 1220648-2025-47125 | 23206412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206412&pc=OZD | 09/04/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713142 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with placement position unknown alarm due to low pulse pressure overnight. the patient was hypotensive, but quickly resolved after fluids and stopping bumex. received one unit of blood for low hematocrit. possible blood in stool. a colonoscopy and biopsy were completed and there was no additional bleeding. patient survived. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. hypotension/ hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. major bleed: it was reported that possible blood in the stool. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4737 | 1220648-2025-47126 | 23206414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206414&pc=OZD | 09/04/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026654264 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 79 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient with a history of aortic valve stenosis was not able to have impella 5.5 placed. 5.5 insertion was unsuccessful, and case was aborted. the impella cp was then inserted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4738 | 1220648-2025-47129 | 23206416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206416&pc=OZD | 09/04/2025 |  |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665269 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma at the access site. manual pressure was held, and a fem-stop was added for support. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the hematoma is most likely use related since hematoma noted at cp access site upon arrival to icu. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4739 | 1220648-2025-47128 | 23206417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206417&pc=PYX | 09/04/2025 |  |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025529808 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 67 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during impella rp flex support no flow calculations or placement signals were displayed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The pump was returned for investigation. data logs were not available for investigation. the returned pump exhibited no visual abnormalities. the sensor electrical wires were tested for resistance and were within specification. the pump was connected to a known good console, and all optical sensor parameters were within specifications. the optical sensor passed laser continuity testing. finally, the pump was run in a bucket of water, and the readings were all stable and within expected ranges. flows were being calculated normally. since data logs were not available for investigation and returned product did not reproduce any abnormalities, the cause of the reported placement signal issue could not be determined. the pump passed all post-sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4740 | 1220648-2025-47127 | 23206418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206418&pc=OZD | 09/04/2025 | 10/01/2025 |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026725118 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was placed to support the percutaneous coronary intervention procedure. no distal pulses on impella limb were found prior to leaving cardiac catheterization laboratory. the physician took an angio of groin and distal flow was noted past impella sheath. the physician wanted to wait for labs to be resulted to decide on removal or reperfusion sheath. in intensive care unit, a blanket warmer was placed on the patient and about an hour later distal pulse was noted with doppler. no intervention was needed. additionally, lactate dehydrogenase had increase which led to hemolysis concerns. no other signs of hemolysis and impella was turned down to p3. impella was shallow on echocardiogram but there were no talks on repositioning after the team was aware. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. ischemia: due to a lack of clinical information provided, the root cause cannot be determined. device in wrong position: the data logs do not show any positioning alarms on the 5th, and there was stable placement signal (ps), pump flows and motor current. due to a lack of clinical information provided, the root cause of how the pump came out of position cannot be determined. hemolysis: due to a lack of product returned, the root cause of the hemolysis cannot be determined. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4741 | 1220648-2025-47130 | 23206419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206419&pc=OZD | 09/04/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 09/04/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703711 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 5.5 pump placed to support the patient but due to interaction of the pump and the subclavian/brachial artery and hypotension that occurred, the pump was removed after surgery/cabg. the hypotension was from the delivery through the tortuosity of the native anatomy. once explanted they hypotension resolved. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hypotension: the cause of the injury was not determined due to insufficient clinical details. difficult to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had severe curvature of the root vessel along with severe tortuosity of vessels and resistance at subclavian artery to brachiocephalic artery and the accordion-like course of the blood vessels. device history review: the device passed all the post sterile inspection checks. | Product complaint # (b)(4). b5 additional information received.
4742 | 1220648-2025-31797 | 23206552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206552&pc=OZD | 10/18/2024 | 10/02/2025 |  | 10/02/2025 |  |  | 10/18/2024 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539058 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump implanted to support a patient admitted in acute myocardial infarction and cardiogenic shock. the patient was also supported by extracorporeal membrane oxygenation (ecmo). while on support, there were placement signal issues. the pump continued to support while the family was deciding to withdrawal care or not. the patient remained on support for just under forty-three days and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation has been completed. no product was returned. the first set of logs is from 10/05 to 10/18, and the second set is from 10/18 to 10/27. during these timings the 5s parameters are all stable. there are placement signal low, suction and positioning alarms. there is a trend in the placement signal. the most likely root cause of the optical signal issue is due to the patient¿s condition. all pertinent information available to abiomed has been submitted at this time.
4743 | 1220648-2025-31798 | 23206560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206560&pc=PYX | 03/03/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 03/03/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089163 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the use of an impella rp pump for bipella support in a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced several placement signal alarms, which were discovered after removal. care was withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the automated impella controller was returned. the failure mode was confirmed. the optical signal issues were due to a consistently low snr bench that intermittently affected placement signal. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4744 | 1220648-2025-31799 | 23206571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206571&pc=OZD | 05/12/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 05/15/2023 | 10/02/2025 |  |  | 09/30/2023 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146473 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an ongoing placement signal not reliable alarm during impella cp support. proper positioning was verified via echocardiography, and the alarm was subsequently silenced. monitoring methods were discussed with the staff and patient support continued with the implanted pump. the patient¿s family eventually withdrew care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was not determined since product was not returned and based on inconclusive evidence of abnormalities from data log review. all pertinent information available to abiomed has been submitted at this time.
4745 | 1220648-2025-31800 | 23206590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206590&pc=OZD | 08/04/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 08/04/2023 | 10/02/2025 |  |  | 01/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306270 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. the complainant reported that at time of initiation, a ¿placement signal not reliable¿ alarm sounded. the pump was reported to be flowing adequately for p9 with motor current pulsatile. a decision was made to withdraw support, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed, and the 5s parameters are affected with several placement signal low alarms. impella position unknown alarms were also observed. the cause of the optical signal issue was most likely patient condition related based on the data log review. all pertinent information available to abiomed has been submitted at this time.
4746 | 1220648-2025-31801 | 23206595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206595&pc=OZD | 01/25/2025 | 10/02/2025 |  | 10/02/2025 |  |  | 01/25/2025 | 10/02/2025 |  |  |  |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531182 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. while on support, there was a placement signal low alarm. the pump was repositioned, and the diastolic placement signal became lower than the original diastolic pressure. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. since there were no abnormalities with the console and the pump was not returned for investigation, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4747 | 1220648-2025-31802 | 23206617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206617&pc=OZD | 09/11/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 09/11/2023 | 10/02/2025 |  |  | 05/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354129 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a shockwave case supported by the impella cp. the optical placement signal was lost, but the team still was able to complete the case and wean/explant. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. according to the clinical, shockwave was used on the proximal lad and after the fourth run, loss of optical sensor was noted. no attempts were made to resolve the issue. the physician believed the issue was probably related to the patient's small body stature and the proximity of her proximal lad to the optical sensor. product evaluation found no issues with the pump. the pump passed the laser test, but the optical bench test showed signal-to-noise ratio (snr) dropped close to zero and contrast stepped down. this indicated that the sensor had been damaged. data log analysis confirmed that the placement signal (ps) spiked to 3000 mmhg around 30 minutes into the case which triggered ps not reliable alarms. this event corresponded with the time shockwave was used and remained until the end of the case. additionally, the logs showed snr & contrast were within specification when shockwave was used and remained until the end of the case. the most likely cause of optical signal issue was a damaged sensor due to interaction with shockwave. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4748 | 1220648-2025-31803 | 23206620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206620&pc=OZD | 06/06/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 06/06/2023 | 10/02/2025 |  |  | 04/30/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333460 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient required to have an impella implanted as he was being weaned off the cardiopulmonary bypass (cpb). as the patient was being weaned off cpb, a suction problem was noted in impella which required the doctor to reposition the device and the flow increased. there was no harm to the patient as a result of this incident. additionally, upon initiation of support, placement signal (ps) noted to be >30 mmhg difference in value of ps and arterial line. patient was weaned off cpb onto impella support and difference remained. corrected by performing adjustment to ps. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no complaint product was returned for investigation. data logs reviewed: lt log showed placement signal issue when the case initiated, it was confirmed the site adjusted the placement signal (ps) at the beginning of the case. many impella position unknown and suction alarms occurred. no evidence of biomaterial in the beginning of the case. the cause of the optical signal issue most likely was patient condition related. no other abnormalities in the signals showed. all pertinent information available to abiomed has been submitted at this time.
4749 | 1220648-2025-31804 | 23206623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206623&pc=OZD | 11/20/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 11/21/2023 | 10/02/2025 |  |  | 08/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390723 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. the complainant reported the impella had aorta (ao) sensor drifts and suction alarms with no known cause. the doctor adjusted the ao waveforms. they were unable to flow above p-7. echocardiography showed the impella position and there was no cause for concern. patient remained on support. no known patient harm or injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were investigated but all the 5s parameters were in specification. no other optical signal related alarms were observed. to investigate the complaint further, product investigation would be necessary. therefore, the root cause of the optical signal issue was not determined since product was not returned and data logs were inconclusive. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4750 | 1220648-2025-31805 | 23206639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206639&pc=OZD | 11/10/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 11/10/2023 | 10/02/2025 |  |  | 08/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385529 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an impella 5.5 pump became unreliable during patient support. the pump was verified to be in good position, and the pump flow was maintained. the staff were instructed to disregard the placement values and to monitor the motor current values for any fluctuations. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were investigated and there was an instant rise in purge pressure along with motor current rising. the p-level was also observed to be automatically rising as an auto-response. signal-to-noise ratio (snr) and contrast dropped while gain, light, and lamp were constant. these symptoms are similar to that of a biomaterial ingestion. therefore, the root cause of the optical signal issue was most likely biomaterial. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4751 | 1220648-2025-31806 | 23206660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206660&pc=OZD | 12/10/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 09/18/2025 | 10/02/2025 |  |  | 10/31/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406301 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, a placement signal not reliable alarm occurred. the audio was disabled. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned. data logs were returned. upon data log analysis, it was found placement signal not reliable. there were no motor current spikes outside of p-level changes. the contrast was variable for the entire case runtime. the root cause of the optical signal issue was unable to be determined due to lack of sufficient clinical details, unreturned product, and inconclusive evidence from data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4752 | 1220648-2025-31807 | 23206672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23206672&pc=OZD | 06/28/2023 | 10/02/2025 |  | 10/02/2025 |  |  | 06/28/2023 | 10/02/2025 |  |  | 04/30/2025 |  | 10/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349781 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was using a cp pump to support a patient with acute myocardial infarction and cardiogenic shock. after shock wave, lost placement signal and received placement signal unreliable. motor current pulsatile and good flows. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump was returned for investigation. no physical abnormalities were observed. data log shows the placement signal not reliable alarm along and all 5s parameter was down which indicate hard failure of the sensor. placement signal was lost immediately after the shockwave device used. the cause of the optical signal issue was a damaged sensor due to the shockwave device interaction. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4753 | 1220648-2025-31757 | 23187734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187734&pc=OZD | 01/26/2025 | 10/01/2025 |  | 10/01/2025 |  |  | 01/26/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with heart failure and cardiogenic shock. the complainant reported that the impella cp had placement signal not reliable alarm at implant. no further information was provided. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows that the signal-to-noise ratio (snr) started at approximately 400 at the beginning of the case, gradually decreasing to 0, with a placement signal spike reaching 3000. no abnormalities were detected in the other 5s optical signal parameters. it is noteworthy that the pump previously used on the same console exhibited a similar trend of low snr, suggesting that console-related abnormalities may be causing the optical signal issues. the cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4754 | 1220648-2025-31758 | 23189457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23189457&pc=OZD | 02/21/2025 | 10/01/2025 |  | 10/01/2025 |  |  | 02/21/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597824 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was using a cp pump to support a patient undergoing a high-risk percutaneous coronary intervention. the pump experienced a 'placement signal not reliable' alarm appeared on the console. the positioning and motor current were verified, and support proceeded without interruption. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. product was not returned for investigation. data logs showed drops/increases in signal not reliable (snr) to correlate directly with changes in p-level, however, it never dropped below specification to trigger a psnr alarm. since product was not returned for investigation and data logs do not show clear cause for the ps not matching the ao line, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4755 | 1220648-2025-47060 | 23190138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23190138&pc=OZD | 08/30/2025 |  |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690368 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support and had a pump stop. the impella cp and extracorporeal membrane oxygenation therapies were ongoing for the postcardiotomy cardiogenic shock after the ventricular septal defect surgery. the pump was supporting the patient when suddenly there was an unexpected pump stop, displaying p-0. attempts to restart the pump at the same assist level as before the shutdown and by cycling the power were unsuccessful. therefore, the impella cp was explanted and replaced with a survived outcome. | [5 empty] | Patient-specific information a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The product with event logs were received, and the investigation was completed. device history record review was completed, and pump passed all post-sterile inspection checks. pump stop: clinical details note that on the sixth day of cp operation, the pump suddenly stopped working in the intensive care unit, displaying p-0. pump was successfully replaced. data log review showed stable pump metrics until a motor current spike and mch pump stop. the pump was turned back on successfully for about 3 minutes before stopping with a motor current spike again. there were multiple attempts to restart the pump, but they were unsuccessful. the product was returned and evaluated. there were biomaterial and foreign material observed wrapped around the base of the impeller. the foreign material was blue and cylindrical, and it was threaded up into the cannula. the foreign material was not a fiber. the cause of the pump stop was biomaterial/foreign material ingestion since biomaterial and foreign material was found wrapped around the base of the impeller and data logs showed multiple motor current spikes. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Additional information noted the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, sections d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly.
4756 | 1220648-2025-47061 | 23190377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23190377&pc=OZD | 09/01/2025 |  |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | Yes | Use of Device Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump supporting a patient with acute myocardial infarction/cardiogenic shock when there was pump repositioning and hemolysis signs observed. labs were hemolyzed and the urine color had changed. the patient suddenly became hypotensive with decoupling impella waveforms. levophed and vasopressin were started as it was noted that the patient had just been taken of levophed prior to the hypotensive event. an echocardiogram revealed that the impella was deep under the pap muscle. the tuohy borst valve was loose. therefore, the medical doctor pulled back, slack was taken out and it was repositioned back to 3.52 centimeters. the impella was then at 91 centimeters. several blood draws came back hemolyzed and drops were off and the impella was at p-7. the pump supported for 22 hours and was weaned and explanted. the patient survived. | [5 empty] | The cause of the hemolysis was use related as evidenced by the clinical confirmation of device in wrong position being deep and under pap muscle due to the tuohy borst valve not be tightened. the cause of the hypotension was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | A.4, a.5, and a.6 are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4757 | 1220648-2025-47062 | 23190567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23190567&pc=OZD | 09/02/2025 |  |  | 10/01/2025 |  |  | 09/02/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692012 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while a patient was on impella cp support, the optical sensor failed and no aorta-placement signal or left ventricle-placement signal was generated. an echocardiogram confirmed that the impella was in a good position. additionally, the patient's urine was dark colored and labs were hemolyzed. the impella was turned down from p-9 to p-6 with improvement in pulsatility. there were concerns of a gastrointestinal bleed because of melena stool so heparin was withheld and a protonix iv drip was started. also the patient had arrhythmias. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this is one of two reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number. | Investigation summary: optical sensor issue: clinical details report intermittent peak signal-to-noise ratio (psnr) alarms on (b)(6) 2025. motor current remained stable. data logs were reviewed, and the psnr alarm was confirmed. the psnr alarm was active the majority of (b)(6) 2025. from case start, signal-to-noise ratio (snr) was never greater than 3000, and before being inserted in body, snr was approximately 1100. contrast and lamp were variable, but light/gain remained stable. this indicates that there was no lack of light, but there was significant noise in the signal. the product was not returned. since the product was not returned for investigation, the cause of the optical sensor issue could not be determined. however, upon review of data logs, it is apparent that the console could be the cause of the issue. hemolysis: clinical details report that the patient had hemolyzed labs overnight between (b)(6) 2025 and (b)(6) 2025. the pump p-level was dropped to troubleshoot. the following morning, it was reported that the patient's labs were no longer hemolyzed and echocardiogram showed the pump to be in good position and good right ventricle function. around 22:30 on (b)(6) 2025, the pump was repositioned slightly. data logs were reviewed and there were no motor current spikes or improper position alarms. there was only one suction alarm during the case and it triggered early the morning of (b)(6) 2025. while running at p-9, flows rapidly trended down out of specification. the pump was dropped to p-6, and flows were initially low, but gradually improved the morning of (b)(6) 2025, getting back into specification. since the placement signal was unreliable at this time, it could not be investigated for trends. the product was not returned for investigation. based on data log review, there was a clear drop in pump flows that recovered during the reported time of hemolysis which could be related. however, insufficient clinical details were provided during the time of hemolysis symptoms to determine the cause for the trending flows. major bleed: clinical details report suspicion of a gastrointestinal bleed overnight on (b)(6) 2025. heparin was paused for a period of time. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. arrhythmia: it was reported that the patient experienced premature ventricular contractions beats and arrhythmias overnight on (b)(6) 2025. they resolved later that morning. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history review: the pump passed all post-sterile inspection checks.
4758 | 1220648-2025-31760 | 23190712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23190712&pc=PYX | 05/20/2025 | 10/01/2025 |  | 10/01/2025 |  |  | 05/20/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026635494 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had impella rp flex ongoing for a patient with an extensive cardiac history / comorbidities. the patient had been found down and unresponsive in the community. cardiopulmonary resuscitation was given, and the patient was intubated for the percutaneous coronary intervention. the rp flex lost optical signal but remained on for the support until the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and lt logs showed multiple instances of "placement signal not reliable" (psnr) alarms occurring with central venous pressure (cvp) going out of spec at time of alarms. rt log at time of first instance of psnr alarm showed optical parameters being momentarily affected with a decrease in signal-to-noise ratio and contrast, increase in amplitude of gain, and cvp going out of range, that all return to normal limits after approximately three minutes. motor current remains mostly stable but with some variability prior to reported issue. cvp is negative throughout rt log. prior to psnr alarms, pump experiences similar changes in pump metrics with motor current being more variable with suction alarms being triggered. clinical details noted that psnr alarms occurred after insertion and cvp waveform was not visible during support. the root case of the optical signal issue could not be definitively determined as no product was returned for evaluation and limited clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4759 | 1220648-2025-47064 | 23190904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23190904&pc=OZD | 08/23/2025 |  |  | 10/01/2025 |  |  | 09/26/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed an impella 5.5 with automated impella controller for support of a patient admitted in and awaiting/bridging to advanced therapies, like a heart transplant. while on support, there were optical signal loss/unreliable alarms. the patient remained on support until successful weaning and there was no patient harm. | [5 empty] | The automated impella controller (aic) was not received from the customer. therefore, the investigation of the device was not possible. this is one of two reports. this report is replacing 1220648-2025-46294 due to an incorrect serial number that is unable to be corrected within our system. a supplemental report will be sent on 1220648-2025-46294 stating we will now be reporting under this report number. any new information received regarding this aic will be reported under this report number. report number 1220648-2025-46295 was submitted to represent the pump. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: data analysis: console logs from the reported day of the event are consistent with the complaint. device analysis: console (b)(4) was returned to field service for further investigation. the reported issues could not be reproduced. the aic was run with a known good pump and purge cassette without any issue observed. additionally, measured "5s" parameter using both optical test cavities were within the specification and the analog output from the ccd array of the optical bench did not show any distortions. root cause: the root cause for placement signal not reliable was due to the intermittent failure of the optical bench lamp assembly. the root of the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. | D.9 the device was returned and date was added. g.3 date should of been 30-sep-2025 on the initial report that was submitted. h.6 codes were updated to reflect the device being returned. h.11 additional manufacturer narrative was updated due to the device being returned. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
4760 | 1220648-2025-31759 | 23191029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23191029&pc=PYX | 05/21/2025 | 10/27/2025 |  | 10/01/2025 |  |  | 05/21/2025 | 10/01/2025 |  |  | 06/30/2025 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024371174 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 83 KG |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp pump was prepped triggered unreliable optical sensor alarms. the automated impella controller was swapped to troubleshoot the issue, but the failure remained. the staff opted to proceed with the implant despite the noted optical sensor issue. care was eventually withdrawn, and the patient expired. | [5 empty] | B2 outcomes attributed to adverse event - death date added h6 evaluation codes were corrected to: 4441 unspecified vascular problem. 1802 death. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs for the case show alarm #1055 placement signal not reliable (opm invalid bench signal) before the pump was started. clinical mentioned that the console was swapped with no resolve. the logs show central venous pressure out of range and low signal-to-noise ratio and light values with max gain and lamp values. the pattern observed could be a possible fiber break but it is unable to be determined as no product was returned. the root cause of the optical signal issue was unable to be determined as no product was returned for analysis to confirm potential issue and insufficient clinical information was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4761 | 1220648-2025-47063 | 23191163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23191163&pc=OZD | 09/02/2025 |  |  | 10/01/2025 |  |  | 09/02/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that optical sensor with placement signal not reliable comes and goes on the automated impella controller. motor current remains pulsatile with appropriate flows for p-level. however, the procedure was success with the same console. | [5 empty] | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two reports. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report number.
4762 | 1220648-2025-47066 | 23191455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23191455&pc=OZD | 09/01/2025 |  |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688154 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia with positioning. there were no plans to position via echocardiogram. the next day, the pump was explanted and the patient received a transplant. | [5 empty] | B1 selection added that was omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): no product was returned. the patient experienced ventricle tachycardia with positioning. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4763 | 1220648-2025-47065 | 23191484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23191484&pc=OZD | 09/02/2025 |  |  | 10/01/2025 |  |  | 09/02/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719192 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 106 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had concerns of a gastrointestinal bleed that had begun prior to impella implantation and liver failure. heparin was withheld. ultimately, care was withdrawn and the patient expired. | [5 empty] | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the cause of the organ failure was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4764 | 1220648-2025-47067 | 23191613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23191613&pc=OZD | 10/30/2023 |  |  | 10/01/2025 |  |  | 09/02/2025 | 10/01/2025 |  |  | 06/30/2025 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379304 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Unspecified Infection | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp experienced hypotension and worsening leukocytosis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned for investigation. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. infection: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the infection was unable to be determined due to insufficient clinical details. device history lot: device lot: 1733776. device history batch: subcomponent lot: n/a device history review: device with sn: 458346 passed all post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4765 | 1220648-2025-31761 | 23191920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23191920&pc=OZD | 01/28/2025 | 10/10/2025 |  | 10/01/2025 |  |  | 01/28/2025 | 10/01/2025 |  |  | 08/31/2024 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025579080 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was using a 5.5 pump to support a patient. the pump experienced a 'placement signal not reliable' alarm appeared. despite this issue, the motor current and flow rate were both within normal limits and support continued uninterrupted. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump and data log were not provided. as per the clinical data, the placement signal was reading negative with an impella position unknown alarm. pump was kept in use for another day. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D4: corrected the udi. h4: corrected the manufacturer date.
4766 | 1220648-2025-47069 | 23192337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23192337&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714009 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient reported to the operating room for durable ventricular assist device (vad). the physician successfully removed impella through graft and closed using vascular stapler. | The complainant had placed an impella 5.5 pump to support a patient admitted in with extensive cardiac history. during impella support, the patient spiked a fever. test results showed positive blood cultures and antibiotics were given. the patient had the original impella 5.5 replaced with new impella 5.5 to manage the infection. |  |  |  |  | B5 updated based on the additional information received. the investigation of the reported failure mode has been completed. no product was returned for investigation. infection: the root cause of the infection was unable to be determined due to insufficient clinical details. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4767 | 1220648-2025-47070 | 23192870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23192870&pc=OZD | 06/20/2025 |  |  | 10/01/2025 |  |  | 06/20/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648758 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp became combative during sedation weaning and kicked off the purge line that connected to the red plug. the original plan was to upgrade this patient but this made the upgrade emergent. the patient was brought straight to the operating room for impella 5.5 implant. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation has been completed. product damage (red plug detachment): clinical details reported the patient kicked off their purge line that connected to the red plug. no product was returned for analysis. the cause of the issue was determined to be due to an unintended use issue as evidenced by the clinical summary of the patient causing damage while in a combative state. device history review: the pump passed all post sterile inspection checks.
4768 | 1220648-2025-47071 | 23193014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193014&pc=OZD | 06/23/2025 | 01/23/2026 |  | 10/01/2025 |  |  | 06/23/2025 | 10/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655543 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Fever | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient implanted with an impella 5.5 developed a fever and had positive blood cultures, so was started on antibiotics. impella support continues. | [5 empty] | D6b, explant date, has been added. | A.4 and a.6 are unknown. the impella device was not returned, and the investigation has been completed. the device history review was completed and the pump passed all post sterile inspection checks. bacterial infection/ fever: the cause of the injury was not determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section d9 has been updated.
4769 | 1220648-2025-31762 | 23193043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193043&pc=OZD | 05/02/2025 | 10/01/2025 |  | 10/01/2025 |  |  | 05/02/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567481 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 67 KG | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support a persistent adjust left ventricle (lv) signal alarm began to appear. the staff walked through how to disable the lv signal. there was no interruption to support. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows the "adjust lv signal offset" notification at the case start, which reappeared after 24 hours of the case run. the "adjust lv signal offset" alarm was resolved on (b)(6) 2025. the imc log was reviewed, and it was noted that the lv offset was applied after the first "adjust lv signal" alarm was issued. a second notification appeared after 24 hours with no action taken during that time. it was confirmed that the "adjust lv signal offset" was resolved after the lv offset was applied on (b)(6) 2025. the instructions for use suggest that the "adjust the lv placement signal" notification appears first when a suggested adjustment is calculated. if a suggested adjustment is calculated, then a second notification will appear after 24 hours of pump use. the cause of the optical signal issue was use issue, as the lv offset was not initiated after the second notification, and the notification persisted until the successfully applied lv offset. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4770 | 1220648-2025-47072 | 23193096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193096&pc=OZD | 01/29/2025 |  |  | 10/01/2025 |  |  | 01/29/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597161 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring defibrillation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not returned, and the investigation has been completed. the device history review was completed and the pump passed all post sterile inspection checks. tachycardia: the root cause of the reported tachycardia was unable to be determined due to insufficient clinical details.
4771 | 1220648-2025-47073 | 23193240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193240&pc=OZD | 01/07/2025 |  |  | 10/01/2025 |  |  | 01/07/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569088 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered ventricular tachycardia and ventricular fibrillation episodes during impella 5.5 support. the patient was treated with amiodarone and lidocaine. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the tachycardia and ventricular fibrillation was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks.
4772 | 1220648-2025-47074 | 23193380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193380&pc=OZD | 07/06/2025 |  |  | 10/01/2025 |  |  | 07/06/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687180 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced thrombocytopenia and tested positive for heparin induced thrombocytopenia. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation into the reported issue has been completed. there was no device received for evaluation. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4773 | 1220648-2025-47075 | 23193493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193493&pc=OZD | 08/29/2025 | 01/22/2026 |  | 10/01/2025 |  |  | 08/29/2025 | 10/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703377 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Pulmonary Edema | Tachycardia | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced tachycardia and pulmonary edema. the p level of the pump was increased and the patient was treated with medication. the patient also had a low-grade fever, and was being treated with micafungin. impella support continues. | [5 empty] | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the fever and pulmonary edema was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary tachycardia : in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. fever / pulmonary edema : the cause of the injury was not determined due to insufficient clinical details. pump stop: clinical reported patient endured sudden pump stop while moving in bed 59 days into pump run. there were no lt logs between 6-oct and 19-oct. the data logs confirm #119 pump stop alarm and show failed restart attempts. logs do not show any motor current spikes outside p-level changes. there were no placement signal trends or positioning alarms before pump stop. though clinical mentioned pump stop occurred during patient movement, the cause of the issue was not established as no data log failure patterns were identified and no product was returned. device history lot device lot : 1958968 device history batch sub-component lot: n/a device history review the pump s/n: (b)(6) passed all the post sterile inspection checks.
4774 | 1220648-2025-31763 | 23193519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193519&pc=OZD | 11/16/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 11/16/2023 | 10/01/2025 |  |  | 09/30/2025 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024402974 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that while on impella support there was a placement signal not reliable alarm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. as per clinical when transferring patient to the hospital second console alarmed placement signal not reliable alarm immediately after plug connect. data logs were reviewed and no placement signal or 5s parameter issue was observed on the first console. after automated impella controller (aic) transfer placement signal unreliable alarm was found due to large air gap with signal not reliable to zero, contrast, light decreasing while lamp and gain was increasing. this issue resolved after 6 hours in the second console. the cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug after transfer of aic per clinical and log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4775 | 1220648-2025-47076 | 23193552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193552&pc=OZD | 09/01/2025 |  |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586738 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device prior to impella 5.5 escalation and impella rp flex implantation for bipella mechanical circulatory support (mcs). approximately one day after start of the bipella support, fluid boluses were given for suction alarm. the next day (overnight) the nurse called the emergency call center for assistance with impella 5.5 intermittent suction alarm at p-8 which resolved at p-7. pulmonary artery and central venous pressures were not available. the nurse reported the physician verified the position and stated right heart shows improvement. the inferior vena cava was not dilated via point-of-care ultrasound. the physician ordered fluids and a unit of blood product. 1l difference between the rp flex and 5.5 flows was noted. additionally, later this day, the impella 5.5 access site was noted bleeding. the patient received one unit of packed red blood cells, the dressing was change and a safeguard compression device was placed over site. the following day bipella support ended as the impella rp flex was explanted. the patient continued on impella 5.5 mcs. | [5 empty] | Additional information was received and noted the impella 5.5 device was explanted after a total of approximately six days of mechanical circulatory support. the device was successfully weaned with a survived patient outcome. based on the additional information d6b was updated accordingly. e1 added zip code extension as it was omitted from the initial report that was submitted. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the major bleed, the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | Latest update notes the device was explanted. however, the explant date with outcome remains unnoted. as of the time of this mdr writing, the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for impella 5.5 with smartassist system: section general patient care considerations - ¿assess access site for bleeding and hematoma.¿ section potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
4776 | 1220648-2025-47014 | 23193553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193553&pc=OZD | 08/21/2025 |  |  | 10/01/2025 |  |  | 08/21/2025 | 10/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713995 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Difficult or Delayed Positioning | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. post extubating placement signal alarm was noted. the impella 5.5 was pulled back one centimeter (cm). alarms resolved with repositioning. the team elected to leave the impella at 6.2 cm due to poor visualization on echocardiogram and resolution of placement signal and suction alarms. additionally, the patient experienced oozing from all sites. heparin-induced thrombocytopenia (hit) panel was sent and systemic anticoagulation was put on hold. the patient was administered two units of packed cells and platelets. the patient continued on support until the device was successfully weaned. | Additional information was received. difficult to place or position: there was no difficulty placing the device and it was optimally placed in the operating room. device repositioning: transthoracic echocardiogram windows were suboptimal post surgical, however we were able to visualize the bend on the impella 5.5 in the left ventricle pod#2 and repositioned successfully by pulling back one centimeter to optimize. no further positioning issues and no ill effect on the patient. major bleeding/hemostasis/blood transfusion/thrombocytopenia: double valve replacement on a critically ill patient with long bypass run and with multiple risk factors, endocarditis and multiple co-morbidities including renal disease. discussed high risk of mortality in the operating room with the family prior to the start of the case. the patient had no additional options other than this high risk procedure with anticipated intraprocedural and post procedural risk. anticipated to require multiple blood products. bleeding and transfusions unrelated to device. the patient was stabilized and ultimately did great and left the hospital. |  |  |  |  | The investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed/thrombocytpenia: the cause of the injury was not determined due to insufficient clinical details. difficult to place or position: the logs confirm placement signal low alarms and show no significant change or abnormalities with the 5s parameters. the logs show an applied placement signal offset of about 7mmhg after an adjust left ventricle (lv) notification was triggered. the log analysis shows that the placement signal low alarms were triggered correctly and logs align with clinical details of pump being deep. the cause of the issue was unable to be determined as no product was returned and limited clinical information was provided on how pump's position became suboptimal. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B5 additional information was received. e1 added zip code extension as it was omitted from the initial report that was submitted.
4777 | 1220648-2025-47077 | 23193628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23193628&pc=OZD | 09/01/2025 |  |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728584 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after the impella cp was implanted there was oozing from the impella access site. a femostop was applied and the bleeding resolved. on day three of impella support the patient had upper gastrointestinal bleed so heparin was held. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A.4 added weight of the patient as it was inadvertently omitted from the initial report that was submitted. b.7 added information that was inadvertently omitted from the initial report that was submitted. d.7a revised as it was previously entered incorrectly on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. investigation summary: major bleed: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4778 | 1220648-2025-47039 | 23194241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23194241&pc=OZD | 08/26/2025 | 09/30/2025 |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711529 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump support placed for the patient admitted in with acute myocardial infarction with cardiogenic shock and placed on support, but explanted after 24 hours. patient had hemolysis signs and had been treated with infusion of haptoglobin. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all the post sterile inspection checks.
4779 | 1220648-2025-47078 | 23194314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23194314&pc=OZD | 09/02/2025 | 01/23/2026 |  | 10/01/2025 |  |  | 09/01/2025 | 10/01/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723582 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | The complainant had a impella cp pump placed to support the critical patient. it was reported that when patient was undraped at the impella site on left femoral artery found to have a large hematoma and was still bleeding around repo sheath. repo sheath was sutured from the back holes and pulled up front so attempted to resuture pump where it wasn¿t pulling the pump up as much. got hemostasis when patient or pump was not being touched but still had a slight ooze with improvement to hematoma. while still on the table patient continued to have hypotension episodes despite being maxed on levophed and had to receive several pushes of epi. pulse pressures are very narrow and suction alarms with hypotension occurred. the patient was given vaso and albumin. it was also reported that the cp was removed and replaced with the 5.5 due to high stick and need for covered stents. the patient survived. | Clinical rationale: an impella cp device was inserted into the left femoral artery in a 78-year-old male patient presenting in acute myocardial infarction (ami) and cardiogenic shock (cgs), on multiple inotropes and vasopressors, in scai stage d shock, prior to initiation of support. the impella was inserted for ventricular support for a protected percutaneous coronary intervention (pci). during the procedure, the access site was noted to have a large hematoma and was bleeding around the repositioning sheath. the repo sheath was sutured from the back and pulled up front. hemostasis was achieved by resuturing the pump. there was still a slight ooze, but improvement to the hematoma. the patient experienced hypotensive episodes despite being maxed out on levophed, and several pushes of epinephrine. pulses pressures were narrow with suction alarms. the patient was sent to the intensive care unit (icu), where vasopressors and albumin were started for heart failure. the following day, the impella cp was weaned with an escalation of therapy to an impella 5.5. the post-procedure outcome is survived at explant. high-risk interventions require intense blood thinners, increasing the risk of active bleeding. bleeding is a known risk due to the impella's anticoagulation and purge requirements. the hypotensive episodes were most likely due to the clinical condition of a critically ill patient in cardiogenic shock, dependent on multiple inotropes and vasopressors, along with the complications of stage d shock. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product or logs were returned for evaluation. hypotension, major bleed: clinical details noted that patient had oozing from axillary jp drain that required 4 units of blood products. patient was having continuous hypotensive episodes despite medications. the cause of the injuries was not determined due to insufficient clinical details. access site adverse event/hematoma: clinical details noted that patient had a hematoma at the impella access site and was bleeding around the repo sheath. repo sheath suturing was adjusted and hemostasis was able to be achieved with slight oozing and improvement to hematoma. the cause of the access site adverse event was a user related issue as bleeding resolved with re-suturing at access site. pump suction: clinical details noted that patient was experiencing suction alarm with hypotension episodes while on support. no additional details were provided. the cause of the pump suction could not be determined as no product or logs were returned for evaluation and limited clinical details were provided. device history lot. device lot: 1967150. device history batch. subcomponent lot: n/a. device history review pump sn (b)(6) passed all post-sterile inspection checks. | B5 was updated and h6 health effect - impact code was updated. | Correction: b1, b3, b7, h6 additional failure modes added. | Component codes were updated.
4780 | 1220648-2025-31764 | 23194315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23194315&pc=OZD | 12/22/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 12/22/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536187 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, placement signal was lost. the motor current and flows remained unchanged. audio was disabled. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows that the placement signal spiked to 3000 with a placement signal not reliable (psnr) alarm at 52 hours of pump run. no trends on the 5s optical signal parameters justified the known pattern of sensor/fiber damage. no abnormalities were reported on motor current or positioning issue during case run. the placement signal failure mode was changed to an optical signal issue, as the cp pump only has an optical sensor, which is responsible for calculating placement monitoring. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4781 | 1220648-2025-31765 | 23195003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23195003&pc=OZD | 02/19/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 02/19/2023 | 10/01/2025 |  |  | 10/31/2024 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023272383 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a optical signal loss while on the 5.5 support. the team confirmed the clinical numbers and positioning. the choice was made to monitor and not replace/explant due to discussion of palliative/care/comfort conversations. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for investigation. the lt log shows the placement signal trend consistent with sensor wear. an abrupt drop of the placement signal of about 50mmhg occurs. then the placement signal remains low until it fails completely on day 39 of support and triggers a placement signal unreliable alarm. the cause of the placement signal issue was sensor silicone wear. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4782 | 1220648-2025-31766 | 23195484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23195484&pc=OZD | 05/23/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 05/23/2023 | 10/01/2025 |  |  | 03/31/2025 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024319075 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant's biomed department reported upon an automated impella controller (aic) with low p levels/suction. the team would like the aic evaluated for the possible malfunction. there was no noted harm or injury. ifu states to use a backup console for these failures. the 5.5 pump remained on for support despite the optical signal alarm. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the field lt log was reviewed and motor current was steady and pulsatile for the majority of the run. a week into support, suction began to occur. two weeks into support, drops in p-level were confirmed. 2 days later, the pump was disconnected while running and automated impella controller (aic) was swapped. the suction/p-level drops resolved and successful support continued for 17 more days without issue. the cause of the issue was biomaterial since the issue resolved with the same pump after aic swap. no trends in placement signal to indicate patient condition related. refer to es2020-082 for additional biomaterial pattern details, a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4783 | 1220648-2025-31767 | 23195733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23195733&pc=OZD | 06/21/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 06/21/2023 | 10/01/2025 |  |  | 12/31/2024 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024307330 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a 5.5 pump support ongoing for care escalated from the cp. the placement signal was lost on the 5.5 but, the pump remains on for support despite signal loss. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data review: data logs showed pump was run on 2 consoles: ic1435 & ic4475. both consoles' logs showed placement signal spiked to 3000 mmhg, snr & contrast were out of specification for entire case. product analysis: product was not returned for review. the root cause of the optical signal issue was unable to be determined as limited clinical details were provided and no product was returned for review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4784 | 1220648-2025-31768 | 23196284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196284&pc=OZD | 02/05/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 02/05/2024 | 10/01/2025 |  |  | 11/30/2025 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422021 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported using an impella 5.5 pump to support a patient undergoing a high-risk percutaneous coronary intervention. after three days of support, the placement signal was lost. the loss of optical signal did not prompt treatment and the pump remained on to support the patient. a decision was made to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was not returned for investigation. data logs were reviewed and 5s parameters were out of specification. at the time of incident, placement signal went up to 3000, snr went to 0, lamp increased to 100, light was decreased, gain was stable and contrast was intermittently barcoding. no motor current spikes were observed. these changes in parameters are similar to that of a damaged fiber line. therefore, the root cause of the optical signal issue was most likely a damaged optical fiber line. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4785 | 1220648-2025-31769 | 23196377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196377&pc=OZD | 06/25/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 06/25/2024 | 10/01/2025 |  |  | 07/31/2025 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024371827 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. while on support, there were placement signal not reliable alarms. support continued. there was no reported harm or injury to the patient. the patient was successfully weaned from the pump and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. on the same day as impella insertion, the placement signal not reliable alarm was triggered. the pump was not explanted due to the complication. neither product nor data logs were returned. the root cause of the optical signal issue could not be determined because not enough information was provided and no product nor data were returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4786 | 1220648-2025-31770 | 23196564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196564&pc=PYX | 05/18/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 05/18/2024 | 10/01/2025 |  |  | 01/31/2026 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025461547 |  | 1000323 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported using an impella cp pump along with an impella rp as bipella support for a patient undergoing a high-risk percutaneous coronary intervention. while on support, the rp lost the optical signal. despite this, the pump remained on to support the patient. due to the poor prognosis for the patient, the family withdrew care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. there was a downward trend of placement signal followed by a loss of placement signal about 4 hours after implant. no product was returned. the data logs confirm a placement signal not reliable (dp signal out of range) alarm and a loss of placement signal during support. the 5s parameters were stable. the pattern observed in the lt log was characteristic of sensor drift. the root cause of the placement signal issue was due to a dp sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4787 | 1220648-2025-47079 | 23196577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196577&pc=OZD | 09/02/2025 |  |  | 10/01/2025 |  |  | 09/02/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723041 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male | 74 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during the removal of cp it was noted that the patient had bleeding complications and had to do surgical cutdown . the physician attributed bleeding to bivalirudin infusion not being stopped prior to going to the operating room. the patient received two units of red blood cells to correct blood loss post operatively. site stabilized with cutdown and no further issues. patient was escalated to 5.5 and patient is stable post explant. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: bleeding from surgical cutdown site and removal of impella cp. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4788 | 1220648-2025-47080 | 23196584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196584&pc=OZD | 09/02/2025 |  |  | 10/01/2025 |  |  | 09/02/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026717006 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp the cath lab staff were unable to obtain distal pulses on foot on impella side. extremity was warm and had capillary refill. unaffected extremity without any intervention had a pulse. after patient arrived in intensive care unit a popliteal pulse was found in the impella leg. the surgeon consulted was not concerned and believed the leg was perfused but had already planned to exchange femoral cp for impella 5.5 via axillary artery and patient was taken to the operating room later that day for the procedure. it was reported that the patient was brought in for impella 5.5 placement and surgical removal of impella cp. patient prepped and right groin cut down exposing ight common femoral artery. after 5.5 implanted the cp was pulled to descending and turned to p-0 then explanted with distal control. surgical repair of vessel was done and closure of skin site. no additional information provided and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. vessel perforation: the cause of the vascular perforation was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4789 | 1220648-2025-47081 | 23196585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196585&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) was plugged into red outlet and once aic was unplugged the battery immediately went from 100% to 50%. once the aic was plugged back in after some time the aic was able to be charged fully to 100% again. customer was concerned about the safety of the aic if patient remained on current aic. the aic was replaced and there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. data analysis: overall, the battery data embedded in the log is consistent with the reported issue. device analysis: the reported complaint could not be reproduced, even after the console was power cycled and operated with the pump and purge cassette for a couple of hours. communication between the power battery manager (pbm) and the batteries was monitored, revealing no issues. additionally, a visual inspection of the internal circuitry showed no abnormalities. root cause: the root cause of the "battery level low" alarm was most likely an erroneous trigger of the invalid data sample in the smbus communication between the pbm and the batteries. however, the specific component responsible was not determined. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4790 | 1220648-2025-47082 | 23196594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196594&pc=OZD | 08/30/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674590 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was implanted the patient developed groin bleeding. on the second day of support, the impella cp was removed due to the bleeding. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4791 | 1220648-2025-47083 | 23196598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196598&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026733035 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female | 65 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the rn was unable to doppler pulse for the patient on impella cp support. vascular duplex study ordered to formally assess flow down impella extremity, unable to doppler pulse. the next day, the nurse was able to doppler pulse. the patient remained on support until successful explant and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4792 | 1220648-2025-47084 | 23196599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196599&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690433 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The device was explanted and there were no access site complications reported. | The complainant reported that impella cp was inserted during pre-percutaneous coronary intervention (pci) for acute myocardial infarction. pci was performed in auto mode. the flow rate was at p9 level. this remained the same for over two hours, but the doctor noticed hematuria and lowered the support level to p3, after which the hematuria subsided. per the physician, the high support level in auto mode, combined with the pump position facing the posterior wall, led to the hematuria, which the physician believed was due to impella-induced hemolysis. the patient condition is stable. |  |  |  |  | Additional information received b5 and d6b updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4793 | 1220648-2025-47085 | 23196601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196601&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657599 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was unsuccessful due to the patient's anatomy. the impella insertion was aborted and the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the delivery issue is determined to be patient condition because physician could not cross the valve with wire and catheter due to patient anatomy. the device passed all post sterile inspection checks.
4794 | 1220648-2025-47086 | 23196659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196659&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657758 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed as care escalation from the cp pump. the 5.5 was supporting when there was a bleed from the jp drain at the pump site, 1 unit of red blood cells were given. it was also reported that the patient had a fever and blood cultures were drawn. the patient remained on support until successful explant and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: no product was returned for investigation. fever: the cause of the clinical symptom was not determined due to insufficient clinical details. major bleed: bleeding from graft anastomosis was reported. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
4795 | 1220648-2025-47087 | 23196660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196660&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710526 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 77 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing at the access site and a central line-associated bloodstream infection. the pump was explanted upon receipt of a heart transplant, and the patient survived. | [5 empty] | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the root cause of the infection was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4796 | 1220648-2025-47088 | 23196661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196661&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698160 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 82 KG |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp could not be completed due to anatomic limitations beyond the iliofemoral and into the aorta. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to patient anatomy and the pump was unable to was unable to be advanced further beyond abdominal aorta despite multiple attempts and the insertion was aborted. the device passed all post sterile inspection checks.
4797 | 1220648-2025-47090 | 23196662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196662&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621297 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Female | 108 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp the patient had bleeding and hemolysis. the pump has been repositioned based on transthoracic echocardiography. four units of red blood cells were given to the patient. the pump was removed due to the failure modes and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. major bleed: the cause of major bleed was unable to be determined as limited clinical details were provided and no product was returned for review. device history review: the device passed all the post sterile inspection checks.
4798 | 1220648-2025-47089 | 23196664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196664&pc=OZD | 09/02/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712351 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted wtih an impella cp experienced bleeding and a hematoma in the right groin. a femostop was used to obtain hemostasis and multiple blood products were transfused. the pump was surgically removed via cut down, the hematoma was evacuated, and the vessel was repaired. a new impella cp was implanted to continue supporting the patient. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. investigation summary: major bleed/ hematoma: the cause of the injury is most likely use related since significant bleeding and hematoma to right groin was noted upon arrival to intensive care unit from catheterization lab. device history review: the pump passed all the post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4799 | 1220648-2025-31771 | 23196670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196670&pc=PYX | 06/26/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 06/26/2023 | 10/01/2025 |  |  | 11/30/2024 |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289408 |  | 0046-0035 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there were placement signal issues. the team attempted to re-zero without resolution. the patient was unable to tolerate dialysis due to end stage renal disease. the family decided to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. rp pump was not returned for investigation. data logs shows patient condition decreasing trend in optical signal based on 5s parameter value of snr with placement signal not reliable alarm. meanwhile, patient was on bipella support. upon further review of the cp pump logs, it was found that the left heart pump cp was under several suction conditions during the case run. the cause of the optical signal issue was most likely patient condition related based on log analysis and was related to management of bipella since the left sided device was in suction during the optical signal issues. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4800 | 1220648-2025-47091 | 23196677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196677&pc=OZD | 09/03/2025 |  |  | 10/01/2025 |  |  | 09/03/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732875 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thromboembolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed for the support of the patient. during support a small embolism was found in the femoral and iliac arteries. the corresponding physician believed this was due to the patients activated clotting times being too low. it was noted that thrombus/biomaterial was observed in the patient or on the device. vascular surgery was consulted and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. thromboembolism: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4801 | 1220648-2025-47092 | 23196698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196698&pc=OZD | 09/04/2025 |  |  | 10/01/2025 |  |  | 09/04/2025 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704483 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was provided. the patient developed a graft infection post impella removal from the axillary artery. the patient was brought back to the operating room to get the graft removed. | The complainant reported that the patient experienced major bleeding during impella 5.5 support. it was noted that there was an unknown source of bleeding. one unit of blood was administered. later the patient was successfully explanted and survived. |  |  |  |  | Additional information provided b5, h6 updated. the investigation completed. no product was returned. major bleed: the cause of major bleed cannot be determined since insufficient clinical details were provided. infection: the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4802 | 1220648-2025-31772 | 23196905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196905&pc=OZD | 09/18/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 10/19/2023 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1082391 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal not reliable alarm appeared on impella connect. audio had been turned off. p7- flowing 3.8 l/min. a fellow confirmed the alarm via echocardiography. the pump was removed the same day but not related to the placement signal issue, rather a bridge to a left ventricular assist device. the patient survived the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imc logs revealed that a placement signal not reliable alarm had occurred. real time (rt) logs confirmed that immediately prior to the placement signal not reliable alarm, the optical bench reported an oscillating contrast value far outside the normal range of contrast. device analysis: console (b)(6) was sent back to field service (fs) for further investigation. the issue could not be reproduced in the lab when running a known-good pump and purge system. a root cause has not been yet identified; however it is known that an oscillating contrast is the result of a hardware fault on the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4803 | 1220648-2025-31773 | 23196906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196906&pc=OZD | 08/16/2023 | 10/24/2025 |  | 10/01/2025 |  |  | 08/16/2023 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1550078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller appears to have very low signal-to-noise ratio (snr) in the field and marginally passing snr during the last maintenance in field service. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console show no optical signal related alarm being issued. device analysis: section 23 of the preventative maintenance procedure was performed and the console was not up to specification. the fringe pattern was analyzed and shown to be noisy. conclusion: the root cause for the optical signal issue was low signal-to-noise ratio resulting from a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | H4 device manufacture date updated.
4804 | 1220648-2025-31774 | 23196909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196909&pc=OZD | 09/26/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 09/26/2023 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103727 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the console alarmed that the pump placement signal was unreliable. this happened before with the same console on a different pump. with the last pump, the placement signal came back. the pump was operating as expected and there was no suspected damage to the optical sensor, it was previously working fine during this case. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis; console imc logs confirm that numerous placement signal not reliable alarms were issued. rt logs from the console confirm that the signal-to-noise ratio (snr0 was significantly low. device analysis: section 23 of the preventative maintenance procedure was performed and the console found not to be up to specification. further analysis of the analog output from the optical bench revealed it to be defective. conclusion: the root cause for placement signal not reliable alarm was low snr due to a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4805 | 1220648-2025-31775 | 23196914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196914&pc=OZD | 07/12/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 09/09/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316946 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump along with an automated impella controller support ongoing. the support was successful for 49 days when the optical signal was lost. the team had plans to explant the following day and the explant was performed. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on 7/11, while running pump (b)(6) experienced "placement signal not reliable" that intermittently resolved before quickly returning consistently over next 24 hours. as the console is shutting down, it performs a 5s on itself twice noticing that the lamp was outputting 0.5v as if the lamp was off. the next two times the console is booted up to run pumps (b)(6), there was again intermittent placement signal reliable through both cases. the rt logs for all three cases show the placement signal intermittently shooting to 3000 as optical sensor dropped to -3276.8 and contrast oscillates between +/- 3276.8. device analysis: additional testing was unnecessary as the two following cases provided sufficient evidence of reproducing oscillating contrast. root cause: the cause of the unreliable placement signal was a defective optical bench lamp. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4806 | 1220648-2025-31776 | 23196916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196916&pc=OZD | 07/20/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 08/23/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137668 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal was lost and/or triggered an alarm on multiple occasions during impella 5.5 support with an automated impella controller. the alarm was disabled with no impact on the functionality of the pump. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on (b)(6), pump 511932 is started up on (b)(6). on (b)(6), there is 1036 error and invalid low signal-to-noise ratio (snr) which caused a placement signal not reliable (psnr) alarm. the placement signal not reliable alarm resolved on (b)(6). this failure mode occurred again on (b)(6) resolving on (b)(6). the imc logs were confirmed by the rt logs as snr drops to 20 as optical sensor falls to -3276.8, and placement signal becomes 3000 during these periods as contrast oscillates between +/-3276.8. device analysis: console was eventually returned for investigation. the placement signal reliable alarms were reproduced during two separate tests. root cause: the cause of the unreliable placement signal was a defective optical bench lamp. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4807 | 1220648-2025-31778 | 23196944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196944&pc=OZD | 09/13/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 09/13/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076066 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported there were placement signal not reliable alarms occurring intermittently. the aorta and left ventricle signals were erratic at times. support continued. the investigating engineer found the issue likely stemmed from the automated impella controller. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis the imc logs show there were placement signal not reliable alarms on 9/12 and 9/13 confirming the complaint. there were additional invalid placement signals which would have been the erratic signal seen clinically. the rt logs confirm the imc logs with the placement signal spiking to 3000 mmhg for a prolonged period of time triggering the placement signal not reliable (psnr) alarm and the invalid placement signals were the rapid peaking to 3000 mmhg and then resetting to the baseline. device analysis: the failure mode was not reproduced during testing. additionally the signal-to-noise ratio (snr) was very strong at 15950 and the fringe pattern had no abnormalities. the console performed as expected throughout testing and was most likely not the cause of the failure. root cause: the cause of the placement signal issues could not be determined since failure mode was not reproduced. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4808 | 1220648-2025-31777 | 23196945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196945&pc=OZD | 08/04/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 08/04/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1109479 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 / automated impella controller (aic) support, and there were placement signal unreliable alarms present. the staff lost placement signal, but support continued. the engineer found that this issue likely stemmed from the aic, not the pump. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on 8/4 at 17:07, placement signal not reliable alarm triggers resolving at 13:27 the following day. on future boot ups, the lamp outputted no voltage failing 5s. reviewing the rt logs, the console has oscillating contrast during the placement signal not reliable alarm with placement signal jumping to 3000mmhg and optical sensor dropping to -3276.8 mmhg. this confirms what the complaint described. device analysis: the symptoms seen in the field were not reproduced but this is an intermittently reproducing symptom. root cause: the cause of the unreliable placement signal was a defective optical bench lamp. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4809 | 1220648-2025-31779 | 23196950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196950&pc=OZD | 09/18/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 09/18/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1309663 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal was alarming in aorta falsely and placement signal unreliable alarms occurred. the alarms were muted as placement signal does not affect the pump performance. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: while pump 525446 was operated by (b)(6), there were frequent signal-to-noise ratio (snr) deficiencies captured by the imc logs. this was confirmed by the lt logs for this case that showed significantly low snr throughout the case. the pump remained on this console for the entire case despite the placement errors. the console also had low snr on the following case with pump 533917. device analysis: the console was returned for investigation and the failure mode was reproduced. running a 5s test on the console, the snr was confirmed to be very low. the charge-coupled device (ccd) array's envelope was shown to not be a smooth envelope without a pump. root cause: the cause of the placement signal issues was a defective optical bench with low snr. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4810 | 1220648-2025-31780 | 23196953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196953&pc=OZD | 09/24/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 11/08/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815750 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an intermittent placement signal alarm was triggered during impella support. repositioning was attempted, but the issue remained. the alarms were disabled, and support was maintained with the implanted pump. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data review: console logs were consistent with what was reported in the complaint. multiple placement signal low alarms [event set #439] observed in the imc logs. ltlog data at the time of these placement signal low alarms show that placement signal was less than 150mmhg. device analysis: console was tested after arriving in field service. the reported placement signal low issue was unable to be reproduced while testing this console on an engineering lab bench. defects of the optical pump connector and optical bench were observed but their effect on the signal quality were minimal and would not explain the observed symptoms. conclusion: the root cause of the placement signal low alarms was unable to be determined due to the inability to reproduce. the observed defects of the optical connector and optical bench were likely unrelated to the reported placement signal low issues. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4811 | 1220648-2025-31781 | 23196963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196963&pc=OZD | 12/10/2024 | 10/01/2025 |  | 10/01/2025 |  |  | 12/13/2024 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1409579 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during patient support, an impella cp pump experienced complications with the signal following activation of the planned percutaneous coronary intervention. it was documented that the aorta waveform overshot and a placement signal is unstable alarm appeared. proper positioning was confirmed, and support was maintained with the implanted pump. there was no resulting patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ic7090 had consistent low signal-to-noise ratio (snr) alarms affecting placement signal while operating pump (b)(4). lt logs confirmed sub 2000 snr for the entirety of the case. the pump was moved over to ic9732 and the failure mode went away. device analysis: ic7090 was returned for testing and the console's low snr was confirmed. root cause: the optical signal issue was caused by a low snr optical bench which in turn intermittently lost placement signal as reported by the user. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4812 | 1220648-2025-31782 | 23196976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23196976&pc=OZD | 10/31/2023 | 10/01/2025 |  | 10/01/2025 |  |  | 10/31/2023 | 10/01/2025 |  |  |  |  | 10/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137166 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller that lost placement signal. the signal had been lost in prior to inservice/training use. the console was removed and replaced. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console imc logs confirm that the placement signal not reliable (opm invalid signal, optical pump connected]) - alarm was issued multiple times. rt logs show that both signal-to-noise ratio (snr) and contrast were low at the time of the placement signal not reliable alarm. device analysis: analysis of the oscilloscope analog output of the charge-coupled device (ccd) array showed an anomaly in the envelope, which likely affected the snr value. conclusion: the root cause of the placement signal not reliable alarm was a defective optical bench with an electrical defect on the ccd array. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4813 | 1220648-2025-47011 | 23176787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23176787&pc=OZD | 08/30/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 08/30/2025 | 09/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716252 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Hematuria | Thrombocytopenia | [33 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that an impella 5.5 pump experienced bleeding, suction event and hemolysis during patient support. the patient received fluids as suction was noted. additionally, the patient experienced ventricular tachycardia and was shocked. it was noted that the patient had significant bleeding from the nose and mouth with an unknown cause. anticoagulation was paused. the patient was positive for heparin induced thrombocytopenia and continued continuous renal replacement therapy. there was a report of pink-tinged urine. there was no foley available, so the team was unable to evaluate. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. epistaxis/thrombocytopenia: the cause of the injury was not determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the bleeding was not determined as it occurred at a non-impella site, and no physical product was returned. pump suction: data log review confirms suction occurring at the time which it was reported. at this time, placement signal is trending down. there were no other abnormalities in the data logs noted at this time, a suction did not occur for the remainder of the case, indicating that administering fluid resolved the issue. the cause of the suction was most likely related to the patients condition, as placement signal was trending down at the time of suction, and it resolved after administering fluid, indicated that the patient likely had volume issues. device history lot: device lot number: 1951097. device history batch: subcomponent lot: n/a. device history review: pump sn (b)(6) passed all post sterile inspection criteria. | Corrections: e4, g4, h6.
4814 | 1220648-2025-47012 | 23178305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23178305&pc=OZD | 08/22/2025 |  |  | 09/30/2025 |  |  | 08/22/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026720616 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 123 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp flex and impella cp (bipella) for mechanical circulatory support following mitral valve and tricuspid valve surgery. due to the patient¿s severe pulmonary hypertension and poor right ventricular function, the decision was made to place the impella rp flex status post surgery. during chest closure, prior to impella rp flex implant, the patient¿s hemodynamics worsened and the decision was made to implant the impella cp after the impella rp flex was inserted and support initiated. the right internal jugular vein was accessed for the impella rp flex, and the impella rp flex was implanted without issues. the impella rp flex position was noted to be optimal when imaged with fluoroscopy, however it was noted that the placement signal of the impella rp flex was unreasonably high with a pulmonary artery catheter pressure higher than expected and pump flow was lower than expected for the p level of the pump. the placement signal issues did not resolve with adjusting the y-axis of the pump, and the medical team was advised by company support that there was a potential issue with the optical sensor which was causing an inaccurate flow calculation. the impella cp was placed in the right femoral artery without issue, and the patient was brought back to the intensive care unit for rest and recovery. the patient then received volume resuscitation, after which the placement signal and flows returned to normal and expected levels. it was advised to the medical team by company support that the optical sensor had likely been impeded by a vessel wall and correlated with hypovolemia. the following day, the patient was transferred to another hospital for advanced care. it was noted that the patient¿s right leg was cold, and pulses were unable to be found with doppler. a plan was made to obtain computed tomography of the head, chest, and pelvis and then take the patient to the operating room for a chest washout, right femoral artery cutdown, and remove both impella pumps. if needed, the patient would then be cannulated for extracorporeal membrane oxygenation (ecmo). after the chest washout was performed, the patient¿s hemodynamics improved, and it was determined the patient did not require ecmo after impella pumps were explanted. the impella cp was emergently removed due to critical limb ischemia. ultimately, the patient required above the knee amputation of the right leg due to complications of the limb ischemia. after both the impella rp flex and the impella cp were explanted, it was decided that the patient did not need to be cannulated for ecmo as the patient was hemodynamically stable and heart function had recovered enough for the patient to be stabilized with pressors and intubation. this complaint involves two impella devices: pump 1: impella rp flex, with serial number (b)(6). pump 2: impella cp with serial number (b)(6). this report specifically addresses impella cp. a separate report was be submitted for the impella rp flex. refer to section h.10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
4815 | MW5176758 | 23178427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23178427&pc=PYX | 09/25/2025 | 09/26/2025 |  | 09/30/2025 |  |  |  | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | PYX | [5 empty] | ABIOMED, INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Physician | [6 empty] | 71 YR |  | Male | 55 KG |  | Asian | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Exsanguination | Cardiac Perforation | [11 empty] | Death | [28 empty] | Initial submission | [9 empty] | Perforation of left ventricle leading to death.
4816 | 1220648-2025-47015 | 23178882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23178882&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713868 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 69 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient implanted with an impella cp had a left femoral artery hematoma at the access site. pressure was held and the hematoma was expressed. it was soft and shrinking in size. there was also bleeding from a puncture site next to the impella which was treated with a suture, a safeguard dressing, and tegaderm. later, the impella leg lost doppler pulses on the impella leg. the vascular team was consulted. per the team, the patient has peripheral artery disease. the patient experienced continuous suction the impella seemed to be in the papillary muscle. when the pump was repositioned, the motor current would flatten. automated impella controller waveforms were showing signs of flow saturation. the decision was made to replace the impella cp with another impella cp. four days later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer as it was not explanted when the patient died. therefore, investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella cp and purge cassette were returned for evaluation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the cause of the bleeding and hematoma was most likely interaction with the ecmo cannula, which caused issues when inserting the introducer sheath. the cause of the suction was most likely an adverse event related to the patients condition, as clinical details state the patient was on ecmo support, and placement signal is seen trending down throughout the case. the cause of the saturated flow was most likely biomaterial ingestion, but the cause of the biomaterial was not determined due to insufficient clinical details. the device passed all post sterile inspection criteria. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4817 | 1220648-2025-47016 | 23179493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23179493&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694604 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 103 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on an impella 5.5 there was bleeding from the pacer insertion site and oozing from the impella and swan sites. fresh frozen plasma and packed red blood cells was given to the patient. the partial thromboplastin time was 138 without heparin or other systemic anticoagulation, suspecting underlying coagulopathy or medication interaction. ultimately, the decision was made to stop efforts and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: bleeding was noted at pacer insertion site, impella site and swan sites. the cause of the bleeding was determined to be patient condition because of the bleeding from multiple sites. hemostasis was achieved at impella site by placing femstop. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4818 | 1220648-2025-47017 | 23179703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23179703&pc=OZD | 08/31/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722900 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 119 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp, there was oozing noted around the insertion site, and positioning alarms were noted so an echocardiogram was ordered. the depth was measuring at 5 centimeters but the medical doctor stated that he will not reposition the pump because he likes the waveforms and that the patient has more fluid to give. one unit of packed red blood cells was being transfused for hemoglobin of 7 and fem stop was in place currently with good hemostasis noted upon my rounding. the plan was to explant the pump which was done. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B.7 added information that was omitted from the initial report that was submitted. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. major bleed: the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks.
4819 | 1220648-2025-47019 | 23180159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23180159&pc=OZD | 08/31/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026713150 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 67 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that attempts were made to place an impella 5.5 and after several unsuccessful attempts, the impella 5.5 placement was aborted. the impella 5.5 could not advance past the left subclavian due to the patient anatomy. an impella cp was placed instead. | [5 empty] | The investigation has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the pump would not advance past the left subclavian due to patient anatomy despite multiple attempts with different wires and catheters to advance the pump into the left ventricle. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
4820 | 1220648-2025-47018 | 23180241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23180241&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed for emergent high risk percutaneous coronary intervention. the plan was to leave the impella in overnight, but the patient became increasingly agitated and would not stop moving. a hematoma formed at the access site, and the decision was made to wean and explant. | [5 empty] | G.4 added pma/ 510(k) as it was omitted from the initial report that was submitted. investigation summary: the investigation of has been completed. hematoma: the cause of the injury is most likely use related since the patient became increasingly agitated and would not stop moving. a hematoma was formed at the access site. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4821 | 1220648-2025-47020 | 23180983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23180983&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708193 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced impella site bleeding. massive transfusion protocol (mtp) was initiated. the patient was administered five units of packed cells, one unit of cryotherapy, three units of fresh frozen plasma and one unit of platelets. the physician sutured and redressed the site. the impella was repositioned under echocardiogram. the right ventricle looked good. trial wean of extracorporeal membrane oxygenation (ecmo) at bedside with flows at 2.0 now with increasing impella support. urine output was decent and pressers were weaned down. the patient is going to the operating room for escalation to an impella 5.5 and removal of veno-arterial ecmo. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Correction g4. the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: impella site bleeding was reported. the cause of the bleeding is determined to be use issue because the bleeding issue is resolved by suturing the site and redressing. device history review: the device passed all the post sterile inspection checks.
4822 | 1220648-2025-47021 | 23181090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181090&pc=OZD | 08/05/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583657 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced a stroke. prior to start of impella mcs, the patient experienced an acute myocardial infarction and underwent coronary artery bypass graft surgery under intra-aortic balloon pump support. postoperatively, the patient was discharged to the room on extracorporeal membrane oxygenation and intra-aortic balloon pump. five days later, a left ventricular hematoma was detected with reoperation. the impella 5.5 device was inserted thereafter, but before weaning from cardiopulmonary bypass. the patient was on the impella mcs for 14 days and after removal of the impella 5.5 device, the patient developed a right cerebral infarction on the side where the impella was inserted. additional information received noted treatment for the stroke is unknown. per the physician, the cause of the stroke is unknown. the physician also noted since the cerebral infarction occurred on the side where the impella 5.5 was inserted, it is likely that the stroke was caused by the impella 5.5 device. the outcome of the patient was noted as survived but in critical condition. | [5 empty] | Patient-specific information a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the stroke, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6. investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation.
4823 | 1220648-2025-31724 | 23181210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181210&pc=OZD | 02/14/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 02/14/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600235 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Unknown | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and, during support, there were placement signal alarms. there was also a yellow placement signal alarm on the automated impella controller. support continued, and the clinical representative confirmed the pump was not explanted due to the issue. the pump was weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and evaluated. the 5s parameters were within spec. a slight kink was found in the catheter near the red pump plug. the pump was run and during a portion of the run, the catheter was kinked at the slight kink found. no significant changes in 5s parameters were seen from the kink test. the pump was also uson tested at 50 mmhg, 100 mmhg, and 150 mmhg and was able to hold pressure without issue. the patient cable was able to be twisted slightly relative to the kink protector, and the inside of the red handle showed the motor and communication cables twisted, but no abnormalities with the optical fiber. data log review showed a sudden drop in signal-to-noise ratio, which was then near zero throughout the rest of the case. contrast steps down and gain steps up at this time. lamp is steady around 91% during the case. light is also steady around 2.7 v. the root cause of the optical signal issue was not determined since the data logs did not show a pattern of a characteristic optical failure and the returned product functioned without issue and the optical parameters were within spec. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4824 | 1220648-2025-47022 | 23181245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181245&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719087 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed a severely high plasma free hemoglobin level, which was an indication of impella malposition. transesophageal echocardiogram was performed to check the position, and the impella cp was found to be pointing towards the mitral valve. a decision was made to take the patient to the catheterization lab to reposition the impella but the impella cp was unable to be repositioned. a new impella cp was placed successfully. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. hemolysis: after reviewing logs, suction alarms along impella in aorta alarms were observed. the hemolysis was most likely use issue related due to improper positioning per clinical. device in wrong position (positioning issue): the cause of positioning issue cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4825 | 1220648-2025-31725 | 23181276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181276&pc=OZD | 08/15/2024 | 10/27/2025 |  | 09/30/2025 |  |  | 03/26/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495821 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, a yellow alarm "placement signal not reliable" occurred. at the same time, the position waveform disappeared. since the motor waveform was still there, it was determined that there was no problem with the drive and pump remained in place. the patient eventually expired while on support. | [5 empty] | B2: added patient date of death. h6: corrected the health effect clinical and effect code.db6: corrected the explant date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. after one day on support, a placement signal not reliable (psnr) alarm occurred. pump remained implanted. pump was returned. data logs were not recovered. returned pump was unable to reproduce psnr alarm, showed nominal 5s parameters, passed laser continuity test, passed electrical resistance testing, and showed no visual damage. the cause of the optical signal issue was not determined since issue could not be recreated during investigation and data logs from the field were not recovered. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4826 | 1220648-2025-31726 | 23181532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181532&pc=OZD | 06/04/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 07/18/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677209 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient undergoing a high-risk percutaneous coronary intervention for circulatory support. shortly after implantation, a ¿placement signal¿ alarm occurred. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. data log review showed placement signal not reliable alarms and placement signal gradually ramps up to 3000mmhg over the course of 12 hours. the root cause of the optical signal issue was not determined since product was not returned and data logs did not match any known failure patterns. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4827 | 1220648-2025-31727 | 23181539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181539&pc=OZD | 01/27/2023 | 09/30/2025 |  | 09/30/2025 |  |  | 01/27/2023 | 09/30/2025 |  |  | 08/31/2024 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262225 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant stated that the impella rp exhibited an optical sensor issue. ultimately, the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs were reviewed and show the dp sensor starting off out of range and slowly drifting back to an acceptable value and stabilizing. the root cause of the placement signal issue was determined to be sensor drift, and the cause of the sensor drift was unable to be determined as no product was returned for review. the root cause of the introducer issue was unable to be determined as no product was returned for review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4828 | 1220648-2025-47024 | 23181574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181574&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026716175 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that tan impella 5.5 had a slow leak near the purge filter unit by the square housing. there were no changes in purge pressure or purge flow or any other parameters. the team monitored the area closely until the patient was explanted as planned. no patient harm was reported. | [5 empty] | No product returned for investigation. it was reported that there was a slow leak near the purge filter with no change in purge pressure or flow. the data logs do not show any purge pressure alarms or changes in purge pressure or purge flow. they are stable the whole case. due to a lack of product returned, the root cause of the purge leak - pump cannot be determined. the device passed all post sterile inspection checks. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4829 | 1220648-2025-47025 | 23181577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181577&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716506 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during advancement of the impella cp, resistance was encountered due to an unknown aortic stenosis. the doctor decided to remove the impella. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the pump had resistance was encountered due to an unknown aortic stenosis preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4830 | 1220648-2025-47023 | 23181585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181585&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707186 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 72 KG | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed a hematoma at the axillary site. hemoglobin was decreased to 7.1 and the patient was transfused one unit of packed red blood cells. additionally, the impella generated a "placement signal unreliable" alarm. later, the pump was explanted when the patient received a heart transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4831 | 1220648-2025-47026 | 23181605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181605&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725410 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced multiple episodes of sustained ventricular tachycardia. the patient was defibrillated and treated with dobutamine, amio, and pressors. later, care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary : no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history lot : device lot: 1963644, device history batch : subcomponent lot: n/a, device history review : device with sn: (b)(6) passed all post sterile inspection checks.
4832 | 1220648-2025-47027 | 23181683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181683&pc=OZD | 09/20/2025 |  |  | 09/30/2025 |  |  | 09/20/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1671526 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) supporting an implanted impella 5.5 generated a backup controller alarm that resolved without intervention. no patient harm was reported. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4833 | 1220648-2025-31728 | 23181740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181740&pc=PYX | 03/07/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 03/07/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540957 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella rp flex support. the flow calculation appeared to be disabled at the time of the alarm. the staff were advised on the alarm, and support continued uninterrupted with the implanted pump. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and lt log of full case showing upward trend in pump flow with downward trend in placement signal (ps). motor current and motor speed remained stable throughout case. placement signal not reliable (psnr) alarm is triggered multiple times after ps goes out of range. rt log at time of first psnr alarm shows stable optical sensor, motor current and motor speed. ps is observed to be trending down slightly then going out of range to 3000. pump flow drops to 0 right after psnr alarm is triggered. clinical details noted that when the psnr alarm was observed, flow calculations were disabled. patient was still being supported by pump as motor current and motor speed were stable during and after noted alarms. patient remained on support for an additional six days before weaning and explant. based on investigation, the failure mode of "optical signal issue" was changed to "placement signal issue" based on data log review indicating the dp sensor is the likely cause of the issue. the root cause of the placement signal issue is most likely due to sensor drift of the dp sensor based on returned data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4834 | 1220648-2025-31729 | 23181753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181753&pc=OZD | 12/05/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 12/05/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566064 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for circulatory support. during support, there were intermittent placement signal unreliable alarms. the impella was running without issue so support continued. the audio for the alarm was silenced. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were investigated and the placement signal not reliable (psnr) alarms were confirmed. particularly early in the case the signal-to-noise ratio (snr) is quite abnormal and low. as the case progressed, the snr improved, though still at relatively low values. additionally, contrast had an abnormally large variability during this period of low snr. lamp, gain, and light weren't really affected by the issue. pump flows were within specification and only seven suction alarms triggered throughout the case, indicating patient condition was not the cause. data logs for other pumps run on the same console were referenced, and there was a pump run after this one with low snr causing psnr alarms throughout the case as well. product was not returned for investigation. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4835 | 1220648-2025-31731 | 23181812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181812&pc=OZD | 10/12/2023 | 09/30/2025 |  | 09/30/2025 |  |  | 11/13/2023 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1616687 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported there were placement signal issues with support of the impella 5.5 and an automated impella controller (aic). an investigation was performed, and it was determined that the optical signal issue was due to the aic's "low signal-to-noise ratio (snr)" failure mode (due to an internal defect of the optoelectronics). the patient survived the event, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis. imc logs confirmed the reported complaint. there were multiple occurrences of placement signal not reliable alarms. the placement signal not reliable alarms coincided with multiple occurrences of the console identifying the signal-to-noise ratio (snr) measured by the optical bench being below the minimum value required. rt logs revealed that the snr measured by the optical bench had periodically dropped to zero. device analysis console ic9284 was sent back to field service for further investigation. the console was tested, and the optical parameters were measured with an optical test cavity. the optical bench was found to have an snr that was very close to being below specification. the snr was (b)(6). if the console was used with a pump that further reduced the snr, as suggested by the data logs, a placement signal not reliable event would occur, as had occurred during the reported complaint. the root cause of the complaint was identified as a faulty optical bench. the fiber optic cable between the optical pump connector and the optical bench was also identified as potentially causing further degradation of the optical signal. the output of the charge-coupled device (ccd) array of the optical bench was measured with an oscilloscope while a pump was connected, and noise was observed. the observed noise was determined to be a potential cause of the low snr. the output of the ccd array was not measured without a pump connected. the cause of the aic issue was a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4836 | 1220648-2025-31730 | 23181816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181816&pc=PYX | 09/25/2023 | 09/30/2025 |  | 09/30/2025 |  |  | 09/25/2023 | 09/30/2025 |  |  | 01/31/2025 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024325142 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the device exhibited optical signal issues. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. according to the clinical, placement signal not reliable alarms occurred. the pump was turned down to p-2 and placement signal monitoring was disabled as troubleshooting measures. a chest x-ray confirmed pump position to be unchanged. the pump was returned to p-6. it was also noted that at some point during the case the patient had been kicking his legs. the impella access site was the left femoral. the pump was not removed due to this issue. product evaluation confirmed electrical and optical parameters were within specification. dp sensor drift was also reproduced in a bucket of water in the lab. data log analysis confirmed dp sensor drift occurred before placement signal went out of spec with loss of flow metrics. however, once the flow metrics returned, they continued to drift downwards with noisy placement signal. additionally, placement signal unreliable alarms were also recorded. the cause of the placement signal issue was dp sensor drift. a trend chart was pulled for placement signal sensor drift complaints with rp flex pumps from 9/25/22-9/25/23 and confirmed this was the only occurrence, so no trend could be established. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4837 | 1220648-2025-31732 | 23181835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181835&pc=OZD | 07/07/2023 | 09/30/2025 |  | 09/30/2025 |  |  | 09/29/2023 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024311972 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The customer reported that the device exhibited optical sensor issues. the pump was not removed. care was withdrawn per the family's required and the patient died. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data analysis: on the complaint date, (b)(6) 2023, console ic8398 received the placement signal not reliable (snr) alarm, which confirms the reported failure mode. though the snr was low during the case start. device analysis: the console was tested on an engineering lab bench. reproduced the reported issue ¿placement signal not reliable (optical signal-to-noise ratio) ¿ alarm¿ after 2 minutes and 6 seconds since the pump started. obtained a fringe pattern with and without a pump connection that had a spike, which confirms the defective optical bench. also found the front panel optical connector with debris. the root cause for the placement signal not reliable alarm was due to a defective optical bench. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4838 | 1220648-2025-47028 | 23181911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23181911&pc=OZD | 08/31/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712260 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella 5.5 generated a placement signal not reliable alarm. the pump was repositioned to resolve the error. no patient harm was reported. | [5 empty] | A.4, a.5, and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4839 | 1220648-2025-31733 | 23182007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182007&pc=OZD | 05/15/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 06/28/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124467 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable error appeared during impella 5.5 support. the audio was disabled and support with the implanted pump was maintained. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show three occurrences of placement signal not reliable alarm, two of them self-resolving. lt/rt log data indicates that during the placement signal not reliable (psnr) condition, signal-to-noise ratio was low and contrast had unrealistic values oscillating between -3000 to +3000, which is consistent with momentary loss of light. pump was not returned for investigation; however it¿s been concluded that the issue was most likely related to the console. console was not sent in for this investigation but had been later investigated under for the same failure mode, which concluded that optical bench was defective. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4840 | 1220648-2025-31736 | 23182064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182064&pc=OZD | 05/22/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 05/22/2024 | 09/30/2025 |  |  | 01/31/2026 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462820 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the aorta waveform disappeared during impella cp support, and the device exhibited an optical signal issue. the functionality of the pump remained unchanged, and support continued without interruption. the patient eventually expired while on impella support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. no data logs were returned for evaluation. clinical details noted aorta waveform had disappeared from the automated impella controller but motor current was noted to be stable and didn't have any issues. the root cause of the optical signal issue cannot be determined as no product or data logs were returned for analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4841 | 1220648-2025-31734 | 23182077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182077&pc=OZD | 06/11/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 06/11/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080191 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm occurred. impella and automated impella controller were working appropriately. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable¿ alarm appearing 3.5 days after pump start, and remaining unresolved for the duration of the case (12 days). shortly before console was powered down, a controller error and 5s parameters recorded in log were failing specification (light parameter of 0.5v corresponds to no light on the detector). lt log data indicates that during that time optical contrast had erroneous values between -3000 and +3000, and the signal-to-noise ratio (snr) was close to 0. this pattern is consistent with loss of light on the optical bench detector. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4842 | 1220648-2025-31735 | 23182123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182123&pc=OZD | 10/22/2023 | 10/10/2025 |  | 09/30/2025 |  |  | 10/22/2023 | 09/30/2025 |  |  | 09/30/2025 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084302 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | IMPELLA 5.5 | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impella controller exhibited an impella position unknown alarm due to extremely low pulsatility. | The complainant reported that an impella 5.5 pump was implanted for circulatory support. the complainant reported there were placement signal issues. impella position unknown alarm was occurring due to extremely low pulsatility. support continued. the patient survived the event, and no patient harm has been reported. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on (b)(6) 2023, pump (b)(6) was started up on ic1165. within seven hours of running, the user received many impella position unknown alarms as referenced in the complaint. this continued intermittently throughout the duration of the case. the rt logs from the impella position unknown alarm show fluctuating placement signal which the user would have seen. additionally, later there were frequent 1036 errors in the logs (though not triggering an alarm for the user) starting on (b)(6) and continuing intermittently. this was caused in the rt logs by barcoding up to 3002 mmhg for the placement signal and -3276.8 mmhg for the raw optical sensor. logs for cases before and after were examined but the failure modes did not occur. device analysis: the console was returned for preventative maintenance (pm) on 4/26/24 (this complaint was not created yet) but the console passed pm including the 5s in section 23 which would influence placement signal if faulty. the console was tested but the failure mode did not reproduce and was able to run a known-good pump as expected. root cause: the cause of the impella position unknown alarms was not determined since there is nothing in the logs to explain the fluctuating placement signal. additionally, the symptoms were not reproduced during other cases and during console testing therefore it is unlikely the failure mode was caused by the console. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4843 | 1220648-2025-47029 | 23182179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182179&pc=OZD | 08/30/2025 |  |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632560 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | A complainant reported that a patient on an impella cp experienced an alarm for an increase in purge pressure and a decrease in purge flow. the pump stopped urgently. attempts to restart were unsuccessful and the patient expired. | [5 empty] | Product complaint #: (b)(4). the impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this is one of two reports. this report represents the pump and another report represents the automated impella controller. refer to section h.10 for the related report number. | Medical safety/clinical reviewed the event and noted the following: the event involves and impella cpsa that stopped working, and the patient died after that. the first impella cpsa was inserted for amicgs on august 21st. since long-term use was expected, the impella 5.5 was considered, but the blood vessel diameter was too small, so the option was abandoned (this pump was never used/inserted). the first impella cpsa pump 1 was replaced on august 27th (8 days later) given the need for long-term use and the automated impella controller (aic) was also switch for connection to the new impella cp pump 2 to reduce the time that circulatory support is terminated. there was no report or indication of device malfunction, deficiency or related adverse events for the impella cp pump and the first aic. impella cp pump 2 was inserted and approximately three days later on august 30, 2025, the pump was stopped, and the patient passed away after that. this patient was notedly already in a "difficult" condition. the purge pressure increased and the purge flow rate decreased, and then the pump stopped. several attempts were made to restart the pump, and all were unsuccessful. the patient died of circulatory failure. given the high purge pressure with the pump stopping leading to circulatory failure, there is not enough information to exclude the impella cp pump 2 as an associated factor in the patient's demise. note: automated impella controller is a malfunction type of reportable event. serious adverse event and demise outcome are associated with the pump. note: h6 investigation type/findings/conclusions remain unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10. | The investigation of the reported failure mode has been completed. the impella device was returned for evaluation. high purge pressure (pp): data log review showed a 5 minute rise in pp that remained high through the rest of the case. there was no motor current spike prior to the pp rise, but the pp rise did begin after the p-level was turned down to p-2 from p-6 and back to p-6 again. product was returned and evaluated. there was biomaterial observed in the purge gap. high purge pressure reproduced and blood ingress was observed in the purge line above the motor housing. the high purge pressure resolved after cutting the catheter above the motor housing. the cause of the high purge pressure was biomaterial buildup since biomaterial was seen in the purge gap of the returned pump. pump stop: data log review showed multiple "impella stopped - motor current high" alarms. the first alarm had a slight rise in motor current (mc) prior to the stop but no mc spike. the pump was able to be turned on again but only in short periods and ultimately stopped after some failed attempts to restart. there were mc spikes at some of the pump stops. the returned product had blood ingress past the motor housing and the pump stop was reproduced in the lab. the cause of the pump stop was blood ingress since blood ingress was observed past the motor housing and the pump stop occurred after there was high purge pressure. cardiac failure: the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4844 | 1220648-2025-47030 | 23182272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182272&pc=OZD | 08/30/2025 | 01/07/2026 |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1569441 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Heart Failure/Congestive Heart Failure | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | A complainant reported that a patient on an impella cp experienced an alarm for an increase in purge pressure and a decrease in purge flow. the pump stopped urgently. attempts to restart were unsuccessful and the patient expired. the physician requested to have the console inspected as well. refer to the related report for the death and the pump. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this is one of two reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to section h.10 for the related report number. | Medical safety/clinical reviewed the event and noted the following: the event involves and impella cpsa that stopped working, and the patient died after that. the first impella cpsa was inserted for amicgs on august 21st. since long-term use was expected, the impella 5.5 was considered, but the blood vessel diameter was too small, so the option was abandoned (this pump was never used/inserted). the first impella cpsa pump 1 was replaced on august 27th (8 days later) given the need for long-term use and the automated impella controller (aic) was also switch for connection to the new impella cp pump 2 to reduce the time that circulatory support is terminated. there was no report or indication of device malfunction, deficiency or related adverse events for the impella cp pump and the first aic. impella cp pump 2 was inserted and approximately three days later on august 30, 2025, the pump was stopped, and the patient passed away after that. this patient was notedly already in a "difficult" condition. the purge pressure increased and the purge flow rate decreased, and then the pump stopped. several attempts were made to restart the pump, and all were unsuccessful. the patient died of circulatory failure. given the high purge pressure with the pump stopping leading to circulatory failure, there is not enough information to exclude the impella cp pump 2 as an associated factor in the patient's demise. note: automated impella controller is a malfunction type of reportable event. serious adverse event and demise outcome are associated with the pump. correction: after review of the case by medical safety/clinical, it was determined the automated impella controller is a malfunction type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this is one of two devices associated with the event. refer to the related manufacturer report number(s) as noted in section h10.
4845 | 1220648-2025-47031 | 23182564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23182564&pc=OZD | 08/12/2025 |  |  | 09/30/2025 |  |  | 09/25/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168797 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Difficult to Maintain | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a loaner automatic impella controller (aic) did not pass the electrical check upon arrival. the plug ground resistance was too high, and the cord was broken. another aic loaner was provided to the customer. no patient was involved. | [5 empty] | The automated impella controller (aic) was received from the customer and an investigation regarding the returned device has been completed. the aic logs were reviewed and did not provide any information relevant to the reported failure mode. the aic power cord was inspected and the high resistance of the grounding wire was confirmed. the ac plug was disassembled and it was observed that the strain relief was overtightened to the point that it compromised the wiring inside the cord jacket.
4846 | 1220648-2025-47034 | 23183042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183042&pc=OZD | 09/01/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695467 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had impella cp placed for support, was transferred to another facility. upon arrival, impella site was bleeding. heparin drip was paused and femstop was applied. the bleeding was resolved. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. major bleed : the cause of the injury is most likely use related since impella site oozing upon arrival. device history review: the device passed all the post sterile inspection checks.
4847 | 1220648-2025-47033 | 23183084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183084&pc=OZD | 08/30/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670273 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma and an ischemic leg with compartment syndrome. surgery was performed to resolved the ischemia. later, the pump was explanted and the patient survived. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the thrombosis and ischemia was unable to be determined due to insufficient clinical details. the cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
4848 | 1220648-2025-47036 | 23183134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183134&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026719227 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hematuria | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5 the patient experienced mild ectopy that resolved without event. after the impella 5.5 was implanted, the patient had pink and red urinary output. the day following implant, the patient had atrial fibrillation overnight. the patient was transfused two units of platelets and treated with amiodarone. later, the patient was successfully weaned from the pump and survived. | Additional information was received that indicated the patient was needing an off-pump coronary artery bypass (opcab). |  |  |  |  | The cause of the kidney failure and hematuria was not determined due to insufficient clinical details. the root cause of the arrhythmia/paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received; b5 and h6 updated
4849 | 1220648-2025-47035 | 23183168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183168&pc=OZD | 09/01/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026690880 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump supporting the patient with an acute myocardial infarction with cardiogenic shock diagnosis. the patient experienced new torsade's and was defibrillated x1 shock and had cardiopulmonary resuscitation. the pump was still explanted that day as planned. the patient survived the procedure. | [5 empty] | E.1 revised initial reporter address line 1, initial reporter city, zip code and zip code extension since the initial report was submitted. e.4 added selection as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4850 | 1220648-2025-47037 | 23183546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183546&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684797 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for a patient admitted in with acute myocardial infarction complicated by cardiogenic shock. it was reported that a small hematoma was noted and manual compression was applied and the hematoma resolved. the patient was transferred to intensive care unit and again started waking up and trying to sit up flexing his hips and knees. the dressing was noted to be more saturated and again a hematoma was noted. manual compression was applied with near complete resolution of the hematoma. the patient is on integrilin as well but planning to stop if hematoma does not resolve. the patient remained on support for a few more days until successful weaning of the pump and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Investigation summary: the investigation has been completed. hematoma: the cause of the injury is most likely use related since the patient began waking up trying to sit up flexing his hips and knees. at that point, oozing from the dressing was noted and a small hematoma was noted that was not there prior. device history review: the pump passed all the post sterile inspection checks.
4851 | 1220648-2025-47038 | 23183630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183630&pc=OZD | 09/01/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698897 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 116 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported impella cp support was ongoing for a patient admitted in with st-segment elevation myocardial infarction and was air lifted to the tertiary center from community. the impella cp was supporting but, there were signs of hemolysis treated by pump repositioning. the p level was adjusted, and dialysis was done. there was additional suction noted. after 77 hours, the pump was weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. pump suction: data log review shows several suction alarms with overall downward trend in the placement signal. there was no motor current spike or major positioning issue reported. clinical details note that additional volume was given during the suction alarms. the cause of the suction issue was most likely related to the patient's condition, based on data log review and provided clinical details. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was most likely related to the patient's condition, based on data log review and provided clinical details. device history review: the device passed all the post sterile inspection checks.
4852 | 1220648-2025-31737 | 23183727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183727&pc=OZD | 02/16/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 02/16/2024 | 09/30/2025 |  |  | 03/31/2024 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023186037 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, a 'sensor not reliable' message was displayed. multiple attempts were made to resolve the issue by repositioning the pump with no success. the decision was made to leave the pump in the patient and wean them off of support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned for a visual inspection. no data logs were returned for analysis. the cause of the placement signal issue was not determined because no product or logs were returned and not enough information was given. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4853 | 1220648-2025-31738 | 23183744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183744&pc=OZD | 08/24/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 08/24/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186357 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal not reliable alarm occurred. the alarm was disabled. there was no reported harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data analysis: on 8/3, the 5.5 pump was inserted into the automated impella controller. it was then run for over 72 days before being removed by the user. during that time, there were instances of intermittent placement signal spikes, including triggering placement signal not reliable on 8/13 for two days, 8/15 for four days, and 8/24 for an hour. the lt logs for these times show the placement signal jumps to 3000 mmhg and the contrast oscillates between +/- 3276.8%. this confirms the findings from the complaint. device analysis: the failure mode was not reproduced during testing and the oscillating contrast did not reoccur. the 5s was taken and contrast was low with marginally passing signal-to-noise ratio which could have influenced the placement signal abnormalities. the ccd array's fringe pattern looked as expected. root cause: the cause of the oscillating contrast was not determined since failure mode was not reproduced. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4854 | 1220648-2025-31739 | 23183797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183797&pc=OZD | 06/10/2025 | 10/10/2025 |  | 09/30/2025 |  |  | 06/10/2025 | 09/30/2025 |  |  | 02/28/2026 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that there was a placement signal reading prior to the implant of the cp. however, after implantation, the signal reading was lower than both the femoral and radial arterial lines. the pump position was verified by echocardiography. no patient harm was reported. | [5 empty] | D4: serial number is unknown; expiration and manufacture date entered in error. supplemental report to omit this information from the initial report. | D4: device specifics unknown, therefore unique identifier (udi) # is unknown. this report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product or data logs returned for analysis. the root cause of the optical signal issue was not determined as no product or data logs were returned for analysis. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4855 | 1220648-2025-47032 | 23183817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23183817&pc=OZD | 08/31/2025 | 10/30/2025 |  | 09/30/2025 |  |  | 08/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710681 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 108 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Difficult to Insert | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Thrombosis/Thrombus | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in st elevation myocardial infarction (stemi) was implanted with an impella 5.5 pump from an impella cp pump for escalation of therapy. it was reported that the patient presented to an outside hospital in stemi on (b)(6) 2025. the patient¿s proximal left anterior descending artery was revascularized with one drug-eluting stent. during the procedure, it was noted that a mid-sized diagonal branch had shut down after the stent was deployed, however the patient was stable and remained stable and was therefore released to the intensive care unit (icu). the following day, the patient had 8 out of 10 chest pain and abnormal electrocardiogram changes. the patient was placed on a nitroglycerin drip and the medical team requested the patient to be transferred to another hospital for escalation of care. the transfer was denied and overnight the patient decompensated. the patient was placed on an intravenous levophed drip to keep the patient¿s mean atrial pressure above 60 mm hg. the patient was then taken to the cardiac catheterization laboratory the following morning and an impella cp was placed to hemodynamically stabilize the patient. after impella cp placement, the patient was urgently taken to another hospital for further treatment. upon arrival, the patient was noted to have arrythmias resulting in suction alarms and nonperfusing rhythms. the impella cp position was confirmed to be in the correct position, however it was noted that the patient¿s laboratory results, and hemodynamics indicated a need for escalation of therapy. the following day, the patient was brought to the operating room (or) for an escalation to an impella 5.5 due to worsening end-organ dysfunction, a larger body surface area, and a cardiac index of 1.2 l/min/m². additionally, a temporary pacemaker was placed overnight due to ventricular standstill and asystolic events. during the removal of the cp and closure of the femoral access site, it was reported that the perclose device was unsuccessfully deployed. the decision was then made to do a cut down and surgical closure of the impella cp site. during the implant of the impella 5.5, the initial site of the right axillary was unsuccessful at the junction of the innominate, and imaging revealed a subtotal occlusion preventing passage of the device. the decision was then made to access the left axillary artery, and the impella 5.5 was inserted without difficulty and support was initiated with flows noted to be at 4 l/min. during the time of the impella cp removal and after impella 5.5 support was initiated, it was reported that the patient¿s activated clotting time (act) was above 700 seconds after a dose higher than expected was administered to the patient. the patient experienced oozing and bleeding from the impella cp site during device removal. the patient required blood products to be transfused and a protamine injection. suction alarms were present on the impella 5.5, and the p level dropped to p-4 with flows lower than expected at 1.9 l/min. imaging was performed and revealed a thrombus in the left atrium (la), which was noted to have been present at the start of the case. after the patent was volume resuscitated and volume status was stabilized, the pump was unable to be increased past p-4. the impella 5.5 position was confirmed to be correct, and it was also noted that the right ventricle function was normal. additionally, a thrombus was found with imaging at the inlet cage of the impella 5.5, and the la thrombus was no longer visualized. during this time, it was reported that there were periods of decreased left ventricle filling, and the patient experienced hypotension. heparin was re-dosed, and the pump was slowly able to be increased to p-6 with variable flows lower than expected at 2.6 l/min. the patient was released to the icu for rest and recovery, with a plan to monitor motor currents and to attempt to increase p level as able. the following day, the decision was made by the family and the medical team to withdraw care of the patient. the patient¿s care was withdrawn, and the patient expired. this complaint involves two impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella 5.5, with serial number (b)(6). this report specifically addresses the impella 5.5 and a separate report will be submitted for the impella 5.5. | [5 empty] | B.1 added selection as it was omitted from the initial report that was submitted. medical safety reviewed the event. regarding the initially implanted impella cp, the event describes an arrhythmia event and major bleed resulting in intervention. this major bleed is due to a known perclose failure after impella cp removal. this is a serious injury. the patient was escalated to an impella 5.5 and experienced complications that included thrombosis that was noted to be in the left atrium potentially migrated to the left ventricle, low pump flow requiring intervention, which may have added to an already complex situation. the patient expired the following day. in this case there is not enough evidence to exclude the impella 5.5 device as an associated factor. however, the heart team and family chose to withdraw care which led to the patient expiring. related medical device reports: this impella 5.5 device report is one of two devices associated with the event. refer to section h10 for the related report number. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. h6 medical device problem code a150205 was incorrectly reported as a150206 in the initial mdr and is now corrected. h6 medical device problem code a24 was not included in the initial mdr and is now correctly added. h10 related report number was not included in the initial mdr and is now correctly added. investigation summary clinical symptoms(thrombosis/thrombus): in order to make a cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms(low blood pressure/ hypotension): the cause of the injury was not determined due to insufficient clinical details. access site adverse event (hemorrhage/bleeding): as per the given clinical details, activated clotting time (act) was supratherapeutic (>700) following a higher-than-expected dose of heparin. due to concerns about bleeding, protamine was administered after the implant. the cause of the access site adverse event was most likely use-related, specifically the administered dose of heparin, based on the clinical details. difficult to deliver or advance: access via the right axillary approach was unsuccessful due to subtotal occlusion at the innominate junction, so insertion was performed via the left axillary artery. the cause of the delivery issue was most likely related to the patient's condition, specifically narrowed blood vessels, based on the clinical details. pump suction and low or blocked pump flow: according to the clinical details, the patient experienced suction alarms and low pump flow, requiring the p level to be dropped as low as p4, with flows around 1.9 l/min. the patient also required blood products due to supratherapeutic act, periods of decreased l eft ventricular filling, and hypotension. despite attempts to stabilize volume status and resuscitate, the pump flow remained low. data log shows the patient was unable to tolerate higher p levels without suction during the first day of implant, with fluctuations in placement signal and pump flow noted during the case run. no evidence of positioning issues or motor current spikes was observed. the cause of the low pump flow and pump suction issue was most likely related to the patient's condition, based on data log review and provided clinical details. device history review : the pump sn (b)(6) passed all post sterile inspection checks. | Based on the information available, the impella 5.5 cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
4856 | 1220648-2025-31740 | 23184181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184181&pc=OZD | 03/22/2023 | 09/30/2025 |  | 09/30/2025 |  |  | 03/22/2023 | 09/30/2025 |  |  | 01/31/2025 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024297885 |  | 1000100 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The customer reported that the device exhibited an optical signal issue. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the root cause of the optical signal issue cannot be determined do to insufficient clinical information and no data logs or product returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4857 | 1220648-2025-31741 | 23184224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184224&pc=OZD | 05/07/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 05/07/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026651067 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump support ongoing for a patient. after three days there were placement signal issues despite proper pump position. the pump continued to run despite the optical signal issues. there was no interruption to support. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product and data log were not provided for the investigation. the cause of the optical signal issue was not determined since product was not returned sufficient clinical information was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4858 | 1220648-2025-31742 | 23184269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184269&pc=OZD | 05/05/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 05/05/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624628 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 101 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp support ongoing for a patient. they reported that the placement signal was lost, but the pump remained on for support and had not been replaced. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows several placement signal low alarms with a stable placement signal of 25-100 throughout the case run, without affecting any 5s optical parameters. no issues were noted regarding motor current or positioning during the case. the cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4859 | 1220648-2025-31743 | 23184409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184409&pc=OZD | 01/13/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 01/13/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1414254 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella cp support because the patient stopped taking their anti-rejection medication and needed additional support. while on support, there were placement signal not reliable alarms and there were no waveforms on the automated impella controller (aic). the motor current and flows were monitored and in good range. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable¿ alarm (#1036, due to invalid signal from opm), and also large number of events 1036 (¿process sample for opmin valid signal error: calculator placement signal 1036¿ translating to ¿algs suspended opm signal invalid¿, and error messages ¿ossiopsens set invalid_bad_snr_sample_mask¿ (invalid sample due to snr below minimum) throughout the case. ltlog and rtlog data showed that during most pf the case, optical bench contrast was oscillating between -3000 and +3000 values (allowed range 0-100) which is a calculation artefact due to loss of light. complaint (b)(4) (24-21543) was initially categorized as pump-related event, and investigated as such, which means console ic7155 was not requested to be returned to abiomed, and remained in the field. by the time data analysis performed by durables investigator consulting on this case confirmed that the issue might have been automtaed impella controller (aic) related, another placement signal not reliable (psnr) event occurred on 2024-01-29, involving same console and pump 482007. this was recorded as aic-related (b)(4), and console was returned to abiomed where it was investigated. from 24-22162-2: ¿the console was tested in the lab. and the issue was eventually reproduced, while analog output of optical bench charge-coupled device (ccd) was monitored. the images captured on oscilloscope showed distortion of the ¿fringe¿ pattern, which, with diminished values of signal-to-noise ratio (snr), hindered the position detection algorithm. unlike previously observed occurrences of similar issues (oscillating contrast), they could not be resolved by short mechanical vibrations applied to the bench or lamp assembly, but only by power-cycling optical bench. analysis of the data associated with one of prior complaints involving the same console (24-21543) revealed the same signatures of optical signals, therefore leading to the opening of second investigation (¿-2¿) against the console (originally, 24-21543-1, was entered against the pump). the issue addressed by 24-21543-2 had the same failure mode and root cause as this investigation.¿ the cause of the aic issue was a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4860 | 1220648-2025-31745 | 23184422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184422&pc=OZD | 03/04/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 04/17/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1736295 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump placed to support a patient. the pump was supporting along with a protek duo, when, after five days, a signal loss occurred. the placement signal not reliable alarm did not prompt the team to explant the pump. the pump remained on for support. the pump was successfully weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show large number of ¿placement signal not reliable¿ events 1036 (x2500) and several alarms #1036 (x40). ltlog data indicates low signal-to-noise ratio (snr) level throughout the case, and rtlog data shows multiple instances of unreliable raw placement signal (corelated with coinciding low values of snr) and low contrast levels. inspection of the console performed in danvers revealed significant contamination of the optical pump connector fiber. testing of the optical bench revealed passing 5s parameters and no distortion of the charge-coupled device (ccd) output, however value of measured optical output power (2.5uw) was below specification (2.7uw). it¿s been concluded that reported issues with placement signal were due to combination of contaminated optical pump connector fiber, and optical bench lamp assembly performance below specification. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4861 | 1220648-2025-31744 | 23184424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184424&pc=OZD | 05/18/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 05/18/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664245 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an implanted impella cp was intermittent during patient support. no intervention was required and support continued uninterrupted. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. only the lt log was retrieved, and it suggests that the placement signal was lost a few minutes after pump insertion, and it completely returned to normal in a few hours. signal-to-noise ratio was seen to spike to 50000, and contrast and gain were lost to 5000 and 500. the light and lamp were not affected during the loss of placement signal. the before ran pump was reviewed on same console, and it was seen that the same pattern of loss in placement signal occurred. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4862 | 1220648-2025-31746 | 23184483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184483&pc=OZD | 05/17/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 05/17/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657604 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp for the patient found down/syncopal and coding. after 25 hours, the team escalated to an impella 5.5 pump and extra corporeal membrane oxygenation. the 5.5 was supporting when the team observed signal alarms for pump position, but the pump was in position. the placement signal issue did not prompt explant or intervention. after 36 hours on the 5.5 the patient expired when care was withdrawn. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. clinical details state that after impella insertion, there were impella in aorta alarms despite the pump being correctly positioned based on echo. it was thought to be due to left ventricle (lv) activity. patient had also just been cannulated for venous-arterial extra corporeal membrane oxygenation (va ecmo) prior to 5.5 insertion. data log review showed the majority of impella in aorta alarms within 15 minutes of pump start-up. there were no abnormalities seen with the optical 5s parameters, and the impella in aorta alarms were triggering correctly based on the algorithms requirements. however, the stable motor current suggests that the pump did not move out of correct position. the motor current had very low pulsatility likely due to the lv inactivity per clinical details with the patient just being cannulated on va ecmo too. the variation in placement signal in combination with low motor current pulsatility likely contributed to the alternating impella position unknown and impella in aorta alarms. no product was returned. the root cause of the optical signal issue was most likely patient condition since the patient had lv inactivity and had just been cannulated for va ecmo, the pump was in the correct position based on echocardiography, and data log review does not suggest true positioning issues. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4863 | 1220648-2025-47040 | 23184824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184824&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705209 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 87 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump supporting the patient who had been admitted in with cgs and cardiomyopathy. the pump was supporting while awaiting transplant or left ventricular assist device (lvad).the pump had a purge sidearm leak; however, the pump was not replaced due to this. the pump stayed on for 2 more days until the lvad was placed. | [5 empty] | The investigation of the reported failure mode has been completed. purge leak - pump: data logs were reviewed and no abnormalities were noted. purge pressure remained within the expected target range and stable throughout the case. there was 1 "purge system open" alarm, however, it triggered at the end of the case after the pump had been disconnected. product was not returned for investigation. based on the clinical details provided, the leak was coming from the purge sidearm. however, without returned product and additional clinical details, the specific origin of the leak and potential cause could not be determined. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4864 | 1220648-2025-47041 | 23184841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184841&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716289 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella 5.5 had a loss of signal. the pump remained on until the patient received a heart transplant. | [5 empty] | The device was not returned for investigation. data logs were recovered. the cause of the optical signal issue was a fiber bend due to the noted possible kinking observed in the white cable with a 5s pattern matching that of known fiber bends. this is an adverse event related to an unintended use error. the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4865 | 1220648-2025-47042 | 23184852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23184852&pc=OZD | 09/02/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710170 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp fail to deliver for a patient admitted in with plan for a hrpci. the native anatomy prevented the delivery. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had narrowed anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4866 | 1220648-2025-47043 | 23185320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185320&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026709083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female | 95 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp suffered an anoxic brain injury from prolonged hypo perfusion during her initial insult. heparin was withheld. later, care was withdrawn and the patient expired. | [5 empty] | The root cause of the stroke was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4867 | 1220648-2025-47044 | 23185348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185348&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688012 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 and rp flex support ongoing for a bipella support for the patient admitted in with known cardiac history and came in for coronary artery bypass graft re-do. the impella 5.5 pump supported for a week when the patient suffered from ventricular tachycardia and cpr was administered. the family decided to stop resuscitation efforts and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 added selection as it was omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4868 | 1220648-2025-47045 | 23185394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185394&pc=OZD | 09/01/2025 | 09/29/2025 |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663375 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 92 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Cardiac Arrest | Hemolysis | Obstruction/Occlusion | Hematuria | Vascular Dissection | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was placed impella cp. prior to placement, the patient had asystole and ventricular fibrillation requiring cardiopulmonary resuscitation before rhythms showed ventricular tachycardia. the impella cp was rapidly threaded, loaded, and placed. the patient continued to be unstable. after about 30 minutes of stability, and vasopressor support, a left heart catheterization as performed. the diagonal graft went down and the left internal mammary artery (lima) had a 80-90% blockage with no apical flows. only the mid segment was perfusing. no interventions could be performed due to cardiac arrest. upon explant of the impella cp, the perclose failed and manual pressure, trauma was caused to the vessel, requiring femoral artery surgical repair. the patient received a unit of blood and the placement of a jp blub drain. | [5 empty] | Correction: e4, h6. the investigation of the reported failure mode has been completed. no product was received for evaluation. data review: data logs showed pump ran without issue during support. there were positioning and suction alarms triggered in early of the case but resolved quickly. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. cardiac arrest/vascular dissection/arterial occlusion/hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4869 | 1220648-2025-47052 | 23185754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185754&pc=OZD | 09/02/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696396 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after impella cp was placed and the patient was taken to intensive care unit, bleeding was noted. there was extensive bleeding systemically due to hyper coagulopathy with ptt >200 which included the impella access site. all anticoagulants were turned off and bleeding was resolved. it was further reported that the patient was taken for a cat scan (ct) and while in the ct, the patient had impella completely removed into the descending aorta by mistake. the team tried to put the impella back, but it was unsuccessful. the impella cp was removed and impella 5.5 was implanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned. major bleed: the cause of major bleed was most likely use issue related since the bleed was resolved by stopping the anticoagulants. device in wrong position (positioning issue): the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the patient. device history review: the device passed all the post sterile inspection checks.
4870 | 1220648-2025-47053 | 23185758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185758&pc=OZD | 09/02/2025 |  |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711879 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient implanted with an impella 5.5 had a purge leak. the pump was explanted and replaced. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The pump was returned for investigation. no damage was observed on the returned pump and the purge leak could not be reproduced. there was no evidence of fluid leak in this case. no issue was found on the returned product. the device passed all post sterile inspection checks. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4871 | 1220648-2025-47046 | 23185855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185855&pc=OZD | 09/01/2025 |  |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716189 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | Ischemia | Myocardial Infarction | Pulmonary Edema | Obstruction/Occlusion | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 lost pulses in the right foot. all stents and grafts were open. the iliac was occluded and flow was slow. a viabahn stent was placed in the right iliac to restore flow down the right leg. chest x-rays showed improvement in the pulmonary edema and the patient was transfused packed red blood cells. later, the patient was successfully weaned from the pump and survived. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the cause of the ischemia, myocardial infarction, obstruction/occlusion, pulmonary edema were unable to be determined due to insufficient clinical details. clinical details indicate that during placement of the impella 5.5 s2, the wire and catheter were difficult to advance past the anastomosis. a straight glide catheter with an angled glidewire was ultimately used, followed by a c1 and safety j-wire to cross the aortic valve. additional difficulty was noted with the inlet making the turn from the axillary artery into the descending aorta, and the wire bent in the axillary artery, making it difficult to visualize on x-ray. the cause of the delivery and advancement difficulties was most likely related to challenging patient anatomy and procedural complexity. the cause of the suction alarms was most likely related to patient condition based on data log review and provided clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4872 | 1220648-2025-47047 | 23185960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23185960&pc=OZD | 09/01/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 09/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed pre- percutaneous coronary intervention on a st-segment elevation myocardial infarction patient. it was noted that the patient went into ventricular fibrillation without pulse rhythm. advanced cardiac life support protocol without success. resuscitation stopped and the patient expired. | [5 empty] | E.1 added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e.4 selection was added as it was omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4873 | 1220648-2025-31749 | 23186151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186151&pc=OZD | 04/04/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 04/04/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and there were placement signal issues. the motor current and flows were stable. support continued. the investigating engineer found that the issue likely occurred with the automated impella controller. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show ¿placement signal not reliable¿ alarms on (b)(6) 2024. device data shows that the two occurrences were characterized by different failure modes, and different internal alarm numbers: #1048 (placement signal out of range) on (b)(6) 2024 and #1036 (placement signal not reliable) on (b)(6) 2024. rt log data confirms that the nature of two evets was different: first one had placement signal reaching ¿-200¿, unchanged signal-to-noise ratio (snr) and contrast, while the second one had no placement signal (-3276.8 or 0x8000), snr close to 0 and oscillating contrast. because the events the complaint refers to were separate manifestations of different failure modes with different root causes, each investigation addresses one of these issues: this investigation refers to event on (b)(6) 2024, and separate investigation (b)(4) addresses issue on (b)(6) 2024 (most likely use-related). the event on (b)(6) 2024 is consistent with failure of optical bench (either lamp assembly of opm electronics). later reported complaint (b)(4) (investigation (b)(4)) involved the same console, and the same issue was observed upon console boot-up the next day ((b)(6) 2024). console ic6250 was returned to danvers on (b)(6) 2024 where it was tested, and then had its optical bench replaced prior to sending back to customer. the cause of the issue was a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4874 | 1220648-2025-47048 | 23186659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186659&pc=OZD | 07/17/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695575 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received notification via the long-term outcome and quality indicator (loqi) impella registry regarding a 80 yr patient with an impella cp. it was reported that after the impella sheath was removed, subject was noted to have ischemic right lower extremity with loss of pulse. vascular surgery immediately consulted, and subject was taken emergently to the or. and found to have occlusion of the right common femoral and external iliac arteries. in the or the subject underwent a right iliofemoral endarterectomy with profundaplasty and stenting of the right common and external iliac arteries. 6 days status post surgery subject¿s right foot noted to be ¿warm and well-perfused, no sensorimotor deficit. during this time while at recovery suite, subject was noted to have left facial droop and left-sided aphasia. a stat head ct was performed and found a large acute infarct of the right pica territory cerebellum. stroke etiology ¿likely periprocedural versus cardioembolic in the context of known atrial fibrillation.¿ on discharge continued to have rle weakness and dysarthria. subject readmitted on (b)(6) 2025 and reported no residual symptoms from cva. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information in a5 is unknown. | Correction e4, g2. the investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia/stroke/paroxysmal atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4875 | 1220648-2025-47051 | 23186708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186708&pc=OZD | 09/02/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688318 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 67 KG |  | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | Cardiac Perforation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient experienced hematoma, thrombosis and left ventricle perforation during impella 5.5 pump support. while on support, there was a small hematoma at the access site. additionally, an echocardiogram was performed and the impella appeared shallow and 1-2cm from the valve. the patient¿s waveforms were normal. the impella was advanced about 4 centimeters. soon after that, the patient¿s waveforms were uncoupled and mc was flat. the decision was made to pull the impella into the aorta at bedside and see if the patient remained stable. upon pullback, the waveforms coupled and mc returned. within a minute, the patient began to decompensate and there was effusion noted on the echocardiogram. the patient continued to decompensate in tamponade and the patient was emergently brought to operating room for sternotomy. the team found the left ventricle was punctured. the left ventricle was sewn closed. the impella was left in. | [5 empty] | D6b explant date provided. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4876 | 1220648-2025-47056 | 23186725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186725&pc=OZD | 08/01/2025 |  |  | 09/30/2025 |  |  | 08/01/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679767 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 113 KG | Hispanic | White | Injury | Yes | No | Restricted Flow rate | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Pain | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge system blocked alarm and suction alarms. tissue plasminogen activator was added to the purge and albumin was administered to resolve the alarms. the patient also experienced deep vein thrombosis and pain in the impella arm. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). investigation summary: no product was returned. data logs are not available. suction: clinical details note suction events throughout the case. on (b)(6) 2025, suction alarms resolved after administering albumin. on (b)(6) 2025, the suction alarms resolved after repositioning. on (b)(6) 2025, suction events were reported when turning p-level above p4. suction alarms are disabled at p-4 and below on impella 5.5. the cause of the suction alarms was not determined since product/data logs were not returned and insufficient clinical details were provided. purge system blocked: clinical details note a "purge system blocked" alarm on (b)(6) 2025 with pp 1054 and pf <1. the cause of the blocked purge system was not determined since product/data logs were not returned and insufficient clinical details were provided. thrombus: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injuries was unable to be determined due to insufficient clinical details. pain: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history lot: device lot: 1926432. device history batch: subcomponent lot: n/a. device history review: the pump sn (b)(6) passed all post-sterile inspection checks.
4877 | 1220648-2025-47054 | 23186818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186818&pc=OZD | 09/02/2025 |  |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710807 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp experienced a "purge pressure high" alarm. due to the patient's poor status, the physician decided to proceed with lysis rather than explanation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the impella device was retuned for investigation. purge pressure high - data logs were not recovered. the returned pump was soaking in warm water and tergazyme for 10 minutes and high pressure purge (hpp) was reproduced. no biomaterial was found in the purge lumen. catheter was cut which resolved hpp. keyence imaging showed a film of biomaterial found in the purge gap. the cause of the purge pressure high was biomaterial in the purge gap. without the data logs, ingestion or build up could not be determined, device history review: the device passed all the post sterile inspection checks. | D9: device received for investigation. once investigation results are complete a supplemental report will be submitted.
4878 | 1220648-2025-47057 | 23186824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186824&pc=OZD | 09/02/2025 |  |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710787 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon implantation of an impella cp the placement signal ao was significantly lower than patient¿s artline. initial ao pressure was 103/60 and initial impella systolic blood pressure on placement signal was 44 with a map in the 30s. suction alarms was also present. the impella cp was exchanged with another impella cp to support the patient. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 revised zip code as it was previously entered incorrectly on the initial report that was submitted and added initial report phone number as it was inadvertently omitted from the initial report that was submitted. h.6 type of investigation and investigation findings codes were updated due to the device being returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4879 | 1220648-2025-47059 | 23186833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186833&pc=OZD | 09/03/2025 |  |  | 09/30/2025 |  |  | 09/03/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669764 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male | 78 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Pulmonary Edema | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella 5.5 support and during support the patient had a low-grade fever and pulmonary edema. no treatment reported. later the patient was successfully weaned from support and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. fever/ pulmonary edema: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4880 | 1220648-2025-47050 | 23186864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186864&pc=OZD | 09/02/2025 | 11/18/2025 |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660242 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 76 KG | Non Hispanic |  | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | Material Split, Cut or Torn |  |  | Abdominal Pain | Hemolysis | Low Blood Pressure/ Hypotension | Ischemia | Hypovolemia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a 5.5 pump placed to support patient with known disease/atrial fibrillation and new cardiogenic shock. the patient had the 5.5 pump placed as care escalation from an intra-aortic balloon pump. the pump was supporting when there was pain in the abdomen determined to be from the renal and splenic infarct, which may have been due to impella use. the pump remained on and sleeve age was observed on day 42. the pump remained on despite this. | [5 empty] | The investigation for the reported product damage, positioning issue and optical signal issue has been completed. data log shows the pump was initialized on (b)(6) for around 56 days. there were some drops in snr noted, without impacting any pump performance, while running on a steady p-level. the gain was slightly affected during the snr shifts. the pump was then transferred to (b)(6), where similar snr drops were noted, again without impacting pump performance. a psnr alarm was noted about 10 days into the pump run on console (b)(6). the gain and light values were also impaired during the psnr event, which resolved itself within 10 hours. the root cause of the positioning issue, optical signal issue, and product damage has been completed. | Corrections: d6b, e4, h6 . the investigation of the reported failure mode has been completed. no product was received for evaluation. clinical symptoms(infarction, ischemia): the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms(abdominal pain, low blood pressure/ hypotension, hypovolemia): the cause of the injury was not determined due to insufficient clinical details. product damage (anticontamination sleeve split/torn): the pump was not returned for investigation. data log review was not required for this case run. the cause of the product damage was not determined, as the product was not returned and sufficient clinical information was not provided. optical sensor issue: data log shows the pump was initialized for around 56 days. there were some drops in signal not reliable (snr) noted, without impacting any pump performance, while running on a steady p-level. the gain was slightly affected during the snr shifts. the pump was then transferred to a different console, where similar snr drops were noted, again without impacting pump performance. a psnr alarm was noted about 10 days into the pump run on console the gain and light values were also impaired during the psnr event, which resolved itself within 10 hours. this optical sensor failure trend did not match any known failure trend; therefore, the root cause of the optical sensor failure was not determined without returned product. device in wrong position: the cause of the positioning issue was not determined without returned product and sufficient clinical details. mechanical interaction with blood(hemolysis): data log does not indicate any abnormalities related to motor current spike, suction, or positioning issue during the time of the reported hemolysis. the cause of the hemolysis issue was not determined without returned product investigation and sufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4881 | 1220648-2025-47049 | 23186880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186880&pc=OZD | 09/02/2025 |  |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649164 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Female | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | (b)(4) on 9/2/2025 stable, only lots of short vt's. not sure if it is from impella or related to mi. the complainant reported that the patient experienced ventricular tachycardia during an impella cp pump support. during support, the patient had more than one incident of ventricular tachycardia. the team increased administration of metoprolol because of the ventricular tachycardia. the following day the patient was weaned successfully and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | G.1 added manufacturing site name and address as it was omitted from the initial report that was submitted. h.6 added b13 code as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. tachycardia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
4882 | 1220648-2025-47058 | 23186900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186900&pc=OZD | 09/03/2025 |  |  | 09/30/2025 |  |  | 09/03/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724407 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was found to have decompensated heart failure and was requiring multiple inotropes. the family wanted everything to be done for the patient¿s wellbeing, so a decision was made to place impella cp for further support. after impellla was placed in the cath lab, there was bleeding at the access site. bleeding was managed by tightening perclose and angle. it was further reported that the patient care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. the product was not returned for investigation. access site adverse event hemorrhage/bleeding): data log review was not required for this case. according to the clinical details, bleeding was managed by tightening the perclose and adjusting the angle. the cause of the access site adverse event was most likely related to unintentional use related, based on the given clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4883 | 1220648-2025-47055 | 23186902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186902&pc=OZD | 09/02/2025 |  |  | 09/30/2025 |  |  | 09/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026710543 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 65 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support experienced supraventricular tachycardia. the patient was cardioverted twice. the patient continued impella support and was explanted as planned. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4884 | 1220648-2025-31747 | 23186940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186940&pc=PYX | 07/31/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 07/28/2024 | 09/30/2025 |  |  | 01/31/2026 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517145 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp flex support ongoing for a patient. the patient was post-transplant and in right heart failure. the pump was placed but the team had optical signal loss. this occurred on the day of implant, but the pump remained on for support despite the placement signal not reliable alarm on day three. the patient survived the event, and no patient harm has been required. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. impella rp flex was returned for evaluation. pump was returned with the catheter cut in two places, no biomaterial was present on impeller. electrical and optical testing were unable to be performed as pump was cut. data logs from field were reviewed and showed placement signal go out of range to 3000 at start of case, with central venous pressure (cvp) also out of range at 300 at the same time. both cvp and placement signal return after approximately 1 hour out of range. once cvp and placement signal return, cvp remains within a normal range throughout the rest of the case. placement signal is observed to be drifting upward with a simultaneous decrease in mean flow during case. clinical details noted that placement signal and pump flows were inaccurate soon after pump initiation, while pump was in proper position and motor current was within a normal range. the root cause of the placement signal issue was most likely due to sensor drift. the failure mode reported to be an "optical signal issue", as the root cause of the issue was due to dp sensor drift was no issue with the optical sensor was observed, the failure mode investigated was changed to "placement signal issue.¿ a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4885 | 1220648-2025-31748 | 23186995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23186995&pc=OZD | 03/15/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 03/15/2024 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323483 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and there were placement signal alarms. the placement signal not reliable alarm occurred, and the account continued to monitor hemodynamics. the pump was not removed or replaced and was flowing perfectly in patient. it was found by the investigating engineer that the issue possibly occurred with the automated impella controller and not the pump. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the following errors: "algs suspended opm signal invalid," "algs suspended opm crc error," and "invalid_bad_snr_sample_mask" due to the signal-to-noise ratio (snr) falling below the minimum threshold. subsequently, several psnr and controller errors were observed. lt log data indicates that the unreliable placement signal persisted throughout most of the case, with varying intensity, which correlated with low snr. this led to placement signal not reliable (psnr alarms (#1036, due to low signal-to-noise ratio) and controller error (#1045, opm comm crc invalid) on 3/15/2024. rt log data also showed low contrast values and low snr values. further analysis of console logs revealed similar symptoms (psnr and low snr during the case) with pump 473839/7 on 2024-04-13 and pump 499991/7 on 2024-05-08. device analysis: the console ic1316 was sent back to field service for further investigation. the reported complaint could not be reproduced. however, upon inspecting the analog output of the charge-coupled device (ccd) from the optical bench, intermittent distortions in the signal (envelope/fringe) were observed. when the signal was not distorted, the measured "5s" parameters were within specifications. root cause: the root cause of the issue was identified as intermittent distortions in the optical bench signal, likely caused by a malfunction in the optical bench or its components. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4886 | 1220648-2025-31750 | 23187514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187514&pc=OZD | 02/11/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 02/11/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599863 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. while on support, a placement signal unreliable alarm appeared. the staff indicated that the signal suddenly disappeared, but no change in motor current was noted. the staff was advised that the signal was not necessary for support and the audio was disabled. support continued uninterrupted with emphasis placed on monitoring hemodynamics, flows, and motor current. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. data logs show the placement signal intermittently going out of range for the first seven days of the case. when the placement signal was out of range, the contrast was oscillating between 3000/-3000. upon review of the logs, it is apparent that the console is the potential cause of the issue. please refer to 25-41442-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4887 | 1220648-2025-31751 | 23187575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187575&pc=OZD | 01/01/2025 | 10/10/2025 |  | 09/30/2025 |  |  | 01/01/2025 | 09/30/2025 |  |  | 07/31/2025 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574071 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. while on support, a placement signal not reliable alarm occurred. the audio was disabled and the team closely monitored the motor current trends. there was no reported harm or injury to the patient. | [5 empty] | H4: corrected the manufacture date. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. log review showed the placement signal not reliable (psnr) alarm with the placement signal going out of range before returning ~32 hours later. when the ps went out of range, contrast was oscillating, gain decreased, and light decreased. all of the parameters went back to their original values once the placement signal returned. upon this data log review, it is apparent that the console is the potential cause of the issue. please refer to 25-40251-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4888 | 1220648-2025-31752 | 23187584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187584&pc=OZD | 01/02/2025 | 10/17/2025 |  | 09/30/2025 |  |  | 01/02/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587392 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump experienced intermittent optical signal issues during patient support. two days into support, the pump triggered impella position unknown alarms that self resolved. the staff opted to monitor the situation and pump support remained uninterrupted. there was no noted patient harm. | [5 empty] | D10 concomitant medical products and therapy dates updated. h4 device manufacture date updated. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs were reviewed and as reported there was a controller error on 1/2 that resolved within a minute. at this time, all optical signal metrics railed to some factor of +/- 16384. product was not returned for investigation. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4889 | 1220648-2025-31753 | 23187596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187596&pc=OZD | 01/31/2025 | 10/10/2025 |  | 09/30/2025 |  |  | 01/31/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601211 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. while on support, a placement signal not reliable alarm occured. the audio was disabled. there was no reported harm or injury to the patient. | [5 empty] | D4: corrected udi. h4: corrected manufacturing date. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows that the signal not reliable (snr) was 1000 outside of the patient's body. the snr dropped to 400 as soon as the pump was inserted into the patient's body. the pump was set to p-level 8, and a downward drift in snr was reported after the start. after 10 hours, several ps spikes were reported at 3000, with the snr being on the low end. the cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4890 | 1220648-2025-31754 | 23187615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187615&pc=PYX | 08/12/2024 | 09/30/2025 |  | 09/30/2025 |  |  | 08/12/2024 | 09/30/2025 |  |  | 11/30/2025 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024434387 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a few hours after placing an impella rp flex while the patient was recovering in the intensive care unit, the placement signal was suddenly lost with an associated placement signal unreliable alarm. all other parameters remained unchanged. the nurse closely monitored the motor current and estimated flows. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no problem was found with the pump and it is apparent the automated impella controller was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4891 | 1220648-2025-31755 | 23187662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187662&pc=OZD | 03/05/2025 | 09/30/2025 |  | 09/30/2025 |  |  | 03/05/2025 | 09/30/2025 |  |  |  |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592775 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support, and, during support, there were placement signal issues. echocardiography was performed, and the impella appeared to be in appropriate position. support continued. no known patient harm or injury occurred. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. data logs show a drop in signal-to-noise ratio (snr) with contrast barcoding and placement signal (ps) going out of range for approximately 120 hours. ps returned for the rest of the case. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4892 | 1220648-2025-31756 | 23187703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23187703&pc=OZD | 05/01/2023 | 09/30/2025 |  | 09/30/2025 |  |  | 05/02/2023 | 09/30/2025 |  |  | 10/31/2024 |  | 09/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023274811 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 pump to support a patient undergoing a high-risk percutaneous coronary intervention. while on support, a placement signal failure occurred. the family transitioned to comfort care as the patient was not making any significant improvements. he continued to go in and out of vtach after repositioning the pump for a second time. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the issue was most likely a damaged optical fiber line from excessive pull during cardiopulmonary. resuscitation. all pertinent information available to abiomed has been submitted at this time.
4893 | 1220648-2025-31700 | 23165159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23165159&pc=OZD | 10/04/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 10/04/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165515 |  | 004334 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant was using a cp pump to support a patient undergoing high-risk percutaneous coronary intervention (hrpci). the associated automated impella controller (aic) indicated that the device was inserted without detecting any placement signal. the case was aborted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. product was not returned for investigation. the data logs logs from the reported day of event are consistent with complaint and show that after pump was started. console logs were analyzed y software sme, who concluded that the issue was use-related, and did not correspond to the expected use case scenario. the reason for multiple pump disconnections/reconnections without prompts or alarms had not been identified. there was no issue found with the console. the device functioned/operated to specification. reported issue was use-related (unexpected disconnections of pump). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4894 | 1220648-2025-46988 | 23165217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23165217&pc=OZD | 08/31/2025 |  |  | 09/29/2025 |  |  | 08/31/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638941 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 108 KG |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Atrial Fibrillation | Renal Impairment | [10 empty] | Required Intervention | Life Threatening | Death |  |  |  | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that an impella cp was placed after the inability to implant an impella 5.5 due to the patient's anatomy. two days following that, the patient was having diminished urinary output and volume overload. a nephrology was consulted for continuous renal replacement therapy (crrt). following, the impella was advanced overnight from 2.8 centimeters (cm) to 4.3 cm in the left ventricle, despite not having any impella related issues. support was continued. the patient was listed as "very" critically ill and had to be cardioverted at bedside. the next day, the patient had down time on veno-arterial extracorporeal membrane oxygenation (ecmo) with a suction event, which caused rapid deterioration of the patient. the physician was not concerned with the atrial fibrillation at this time as the patient does well once cardioverted. support was continued, and two units of packed red blood cells were given overnight for ecmo flows. the following day, the patient would, however, expire. the family chose to withdraw care due to the poor prognosis. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, paroxysmal atrial fibrillation, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. for renal failure, the cause of the injury was unable to be determined due to insufficient clinical details. pump suction, no product was returned. the impella was advanced overnight from 2.8 to 4.3 centimeters in the left ventricle, despite not having any impella related issues. the patient had down time on extracorporeal membrane oxygenation (ecmo) with a suction event, which caused rapid deterioration. after reviewing logs, multiple suction alarms were observed and impella in aorta alarm at the same time. the cause of suction was patient condition related since patient had volume overload, patient was also deteriorating and ecmo support. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this impella cp report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h.10 for the report on the impella 5.5. | Medical safety reviewed the event. the delivery issue/unable to deliver due to the patient's vessel condition prior to therapy involving the impella 5.5 is a malfunction and should be reported despite knowing the patient's anatomical incompatibility prior to implant. the demise outcome is unrelated to the impella 5.5. the patient was continued on venoarterial extracorporeal membrane oxygenation (va ecmo) and impella cp during the impella 5.5 implant. a new impella cp was placed for left ventricular unloading was placed with va ecmo. complications of arrhythmic events and suction were experienced. the patient expired on day four of support. there is not enough information to dissociate the impella cp in at the time of the demise and therefore, will be reported as death. however, the patient's underlying condition contributed most to this outcome of death. the patient was critically ill prior to therapy and there was down time on va ecmo which caused rapid deterioration of the patient. prognosis was "poor" and the family chose to withdraw care which led to the patient expiring. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5. | Section h6: type of investigation inadvertently reflected b18-device discarded instead of b17-device not returned. it is unknown if the device has been discarded and unavailable for return. therefore, section h6 investigation coding (type, findings and conclusion) was corrected, repopulated accordingly. should the device or any new information be received, a supplemental mdr will be filed.
4895 | 1220648-2025-46990 | 23165381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23165381&pc=OZD | 07/17/2025 |  |  | 09/29/2025 |  |  | 07/17/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694973 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Unknown |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had atrial fibrillation with rapid ventricular response and amiodarone was started. it was further reported that while removing the impella using standard explant technique, the cannula detached from the motor housing while being pulled through the graft conduit. the entire device was able to be recovered. impella support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. due to lack of product returned, and sufficient clinical information, the root cause of the product damage (cannula detached from motor housing) cannot be determined. the device passed all post sterile inspection checks.
4896 | 1220648-2025-46989 | 23165483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23165483&pc=OZD | 07/17/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 07/17/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663213 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp encountered hemolysis. the device was repositioned to resolve the issue. later, the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, but the event logs were received. the investigation has been completed. hemolysis: a review of the device's data logs confirmed that the pump was in an improper position at the time that the patient endured hemolysis and the positioning alarms. no other issues were found on the logs. the pump was not returned for evaluation. based on the clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. positioning issue: the root cause of positioning issue was most likely use related since the touhy-borst valve was not tightened properly. device history review: the pump passed all post sterile inspection checks. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
4897 | 1220648-2025-46991 | 23165616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23165616&pc=OZD | 07/20/2025 |  |  | 09/29/2025 |  |  | 07/20/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665432 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Fluid/Blood Leak | Device Sensing Problem |  |  |  |  | Fever | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was implanted to support a patient admitted for acute decompensated non-ischemic heart failure. while on support, there were instances of suction events. some of these events self-resolved with ambulation and others were attributed to position, nipride and bumex doses. in one instance of a suction alarm, the device¿s position was double checked and pulled back one centimeter. additionally, a bleed was suspected so the patient was scheduled for colonoscopy and anticoagulants were withheld. the pump generated multiple placement signal low alarms so placement signal monitoring was disabled and audio alarms for suction were turned off. the patient had a fever and positive blood cultures, and a small amount of purge was leaking at the side arm. impella support continued. | [5 empty] | The cause of the suction events could not be determined due to inconclusive data log review and insufficient clinical details. the cause of the optical signal issue was sensor wear, a durability issue. the cause of the purge leak could not be determined due to no product returned and insufficient clinical details. the root cause of the fever, infection, and bleeding experienced was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components.¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4898 | 1220648-2025-31701 | 23165947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23165947&pc=OZD | 03/20/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 03/20/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137162 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while the patient was on support, there were ongoing placement signal alarms. the clinical staff muted the audio alarms and continued support. no patient harm or injury was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs were consistent with what was reported in the complaint. the root cause of the placement signal not reliable issue, causing oscillating contrast, was due to a malfunction of optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4899 | 1220648-2025-31702 | 23166042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23166042&pc=OZD | 05/13/2023 | 09/29/2025 |  | 09/29/2025 |  |  | 05/13/2023 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient implanted with an impella 5.5 pump for circulatory support. during support, the pump lost its placement signal. the issue remained unresolved, and the pump¿s position was subsequently monitored using flow and motor current readings, as well as echocardiographic imaging. patient support continued with the implanted pump, and there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the logs confirmed that placement signal unreliable alarm was triggered on the complaint occurrence date. the failure mode could not be reproduced given the optical bench had already been replaced at the time of this investigation. the root cause of the placement signal unreliable alarm was most likely due to a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4900 | 1220648-2025-46992 | 23166451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23166451&pc=OZD | 07/29/2025 |  |  | 09/29/2025 |  |  | 07/29/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687836 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 130 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed vancomycin-resistant enterococci (vre). later, the patient was successfully weaned from the pump and survived. additional information was received. the vre was treated inhouse at the intensive care unit. the pump was discarded. | [5 empty] | The investigation has been completed. no product was returned for investigation. infection: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
4901 | 1220648-2025-31703 | 23166467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23166467&pc=OZD | 05/22/2023 | 09/29/2025 |  | 09/29/2025 |  |  | 05/22/2023 | 09/29/2025 |  |  | 01/31/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332209 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, a ¿position in aorta¿ alarm was observed, although echocardiography confirmed that the pump was properly positioned. the alarm was eventually disabled. despite this, the pump remained operational throughout the support period until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the false position in aorta alarm was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
4902 | 1220648-2025-46993 | 23167094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23167094&pc=OZD | 08/01/2025 |  |  | 09/29/2025 |  |  | 08/01/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695292 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Device Sensing Problem | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient has been on impella 5.5 support for over two months, waiting for heart and kidney transplants. the patient was reported to be having less ectopy since milrinone was discontinued. while on support, the impella placement signal drifted and did not correlate with blood pressure. two days later, the placement signal alarm was was disabled. the patient developed a new deep vein thrombosis (dvt) in the right arm. three days later, bivalirudin dosage was increased due to the dvt. a kink was found on the impella near the red plug. foley anchors were re-attached to the skin to reduce tension on the kink. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Additional information received: d6b explant date added. | The investigation of the reported failure mode has been completed. the pump was not returned for investigation. data logs show that the placement signal spiked and then started to drift to -50 for about one hour before dropping to 3000 on (b)(6) 2025 (the 34th day of pump use), without significantly impacting the 5s optical signal parameters. the pump was running steadily at p-4 during the optical sensor issue, with no notable motor current trend. pump flow showed an opposite gradual upward trend from the initial drop, correlating with the placement signal trend. there was an active psnr alarm, and the placement signal was not recovered during the rest of the case run. the cause of the optical sensor issue was determined to be sensor wear, based on the data log review. product damage (catheter kink): the cause of the catheter kink could not be determined, as returned product investigation and sufficient clinical information were not provided. (thrombosis/thrombus/arrhythmia): the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4903 | 1220648-2025-46995 | 23167307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23167307&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714012 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 57 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was successfully bridged to transplant. | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced a few beats of ventricular tachycardia (vt). the impella was pulled back 0.5 centimeters and was optimized. there was no vt overnight or the next morning. subsequently, the patient was having increased runs of vt again. the impella was repositioned again. there were no acute concerns noted. the patient continues on support. |  |  |  |  | Additional information received b5 updated. the investigation of the reported failure mode has been completed. no product was returned. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4904 | 1220648-2025-46994 | 23167594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23167594&pc=OZD | 08/31/2025 |  |  | 09/29/2025 |  |  | 08/31/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693729 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding at the access site requiring 2 units of packed red blood cells (prbc). the patient was on multiple anti platelet therapy, heparin and lovenox with hypertension. manual pressure was used to control bleeding along with holding heparin and anti-platelet therapy to allow coagulation labs to normalize. a hemo static suture was used around the site to control hemostasis as well. the patient remained stable without bleeding after intervention. additionally, it was reported that it was not possible to increase impella flow past p2 without suction post coronary artery bypass graft. position and volume status were confirmed. possible material on inlet coronary bypass graft was identified during echocardiogram. during support, the patient's urine was noted to be changed to pink/red. motor current flat and placement signals decoupled. echocardiogram confirmed correct positioning. the impella was exchanged using side port without difficulty. the new impella resolved flow and waveforms. urine color retired to normal. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the pump was returned for investigation. no physical damage was found on the returned product. the pump passed the repo sheath leak test with pressure at 200 mmhg. the pump cannula was further inspected, and significant biomaterial was observed on the impeller. the pump also experienced low flow during testing in the water bucket. pump flow returned to specification after the biomaterial was flushed out while running in the water bucket. data log shows several suction alarms followed by flow drops. the pump was manually turned to p0, and upon restart, a motor current spike to 1500 was reported, followed by another spike while running on a steady p-level with flow saturation. mechanical interaction with blood (hemolysis): the cause of the hemolysis issue was biomaterial ingestion, based on returned product investigation and data log review. pump suction: the cause of the suction issue was biomaterial ingestion, based on returned product investigation and data log review. access site adverse event (hemorrhage/bleeding): as per the clinical details, impella site bleeding developed after the peel-away sheath was removed and the impella sheath was placed, requiring the sheath to be positioned at a high angle to maintain hemostasis. the patient, who was on multiple antiplatelet therapies, heparin, and lovenox, with hypertension, required 2 units of prbc. bleeding was managed with manual pressure, holding anticoagulation and antiplatelet therapy, and placement of a hemostatic suture. the cause of the bleeding issue was most likely related to use, based on returned product investigation and the provided clinical details. clinical symptoms code (high blood pressure/hypertension): the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
4905 | 1220648-2025-46997 | 23167669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23167669&pc=OZD | 08/30/2025 | 10/14/2025 |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667324 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 90 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the impella 5.5 device was plugged into the automated impella controller, there was a error message ¿impella defective: do not use. replace impella.¿ the decision was made to replace the pump prior to insertion. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4906 | 1220648-2025-46996 | 23167781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23167781&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699654 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Fever | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on an impella cp that the impella was advanced into the left ventricle and the patient immediately had ventricular tachycardia (vt) as soon as impella crossed the valve and the impella was pulled back into ascending and the vt did not resolve so the patient was shocked once at 200 joules along with 100 of lidocaine and 150 of amiodarone was given. the impella was advanced across the valve and the wire was removed. the patient had vt again and the doctor pulled the impella back into ascending with no resolution of vt so the patient was shocked again once at 200 joules. the doctor attempted to utilize a stiff buddy wire to assist with advancing the impella across the av but was unable to do so. the doctor removed the impella fully and then reinterested the same pump again. the impella backloaded onto the wire and was placed across the valve without issue. the wire was removed and the impella was started in auto. the following day, the patient spiked a fever so infectious workup was being pursued. no further information was provided. later, the pump was explanted successfully and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. tachycardia : the root cause of the injury was unable to be determined due to insufficient clinical details. fever : the cause of the injury was not determined due to insufficient clinical details. unable to deliver or advance : the cause of the delivery issue was unable to be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4907 | 1220648-2025-46998 | 23167863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23167863&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713152 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 58 YR |  | Male | 82 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella 5.5 support for worsening hemodynamics and increased central congestion. during support, the patent had a few episodes of ventricular tachycardia. position was verified and the patient was started on an anti arrhythmic. the impella 5.5 was successfully weaned and explanted as bridge to transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary: no product was returned for investigation. tachycardia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1946944. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks. as per qn (b)(4), pump with sn (b)(6) from batch# 1946944 found to have package has wrinkles / creases. rework of this unit has been successfully completed on rework workorder (b)(4), an "a" has been added to track sterilization. unit will be sent back to sterilizer. completed qn closure checklist has been added to the qn . this qn will now be closed.
4908 | 1220648-2025-46999 | 23168066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23168066&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636986 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on an impella cp, the impella cp was prepped and the peel away sheath was inserted. the doctor tried multiple attempts and different wires but was not able to advance past stenosed bifurcation. the doctor aborted the case and there was no reported patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had stenosed bifurcation preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
4909 | 1220648-2025-47000 | 23168137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23168137&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728361 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 88 KG | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon implant of an impella cp, there were low flows and loss of pressure. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Investigation summary: optical sensor issue: clinical details reported placement signa readings were inaccurate and there were placement signal low alarms. the logs confirmed placement signal low alarms at the start of the case. the logs showed the pump was in suction condition when placement signal low alarms occurred. the patient was also on veno-arterial extracorporeal membrane oxygenation (va ecmo) flowing at 3.8 liters per minute. no calibration issue at the start of support was noted. the cause of the issue was not established as no product was returned and limited clinical information regarding patient condition and management was provided. low or blocked flow: the logs showed low flow for set p-level flowing at about 0.8 liters per minute at p2 which should be at least 1.1 liters per minute. the patient had concurrent va ecmo during support. there were suction alarms during support. no position related issues were noted. the cause of the issue was patient condition related as evidenced by the low flow with concurrent va ecmo support. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
4910 | 1220648-2025-47002 | 23168842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23168842&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688375 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Malposition of Device | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Nerve Damage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed a hematoma at the access site, nerve discomfort, and bleeding from the ventricular tachycardia ablation site and the abdomen. the patient was transfused packed red blood cells and heparin was withheld. additionally, there were positional issues and high purge pressure alarms that prompted repositioning of the pump and tissue plasminogen activator to be added to the purge. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to h.10 for the related report number. | Investigation summary: no product was returned. data logs only provided were from 20-sep-2025 to 22-sep-2025. purge pressure high: on the 9th, there was purge pressure high alarms triggered and tissue plasminogen activator was infused, the pump position was adjusted and the systemic anticoagulation was switched from heparin to bivalirudin, and about a day later it resolved. there were only data logs provided from 20-sep-2025 to 22-sep-2025 with no issues seen. due to the lack of data logs provided during the time frame and no product returned, the root cause of the high purge pressure cannot be determined. device in wrong position: it was stated that when there were purge pressure alarms, the positioning was checked an the inlet was positioning towards the papillary and there was obstruction of the inlet window. the device was pulled back 1.5 and the inlet became free from structures. due to a lack of clinical details provided, it cannot be determined how the device came out of position. hematoma: due to a lack of clinical details provided, the root cause cannot be determined. nerve injury/major bleed: the cause of the injury could not be determined due to insufficient clinical information. device history review: the pump passed all post sterile inspection checks.
4911 | 1220648-2025-47001 | 23168942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23168942&pc=OZD | 09/04/2025 |  |  | 09/29/2025 |  |  | 09/04/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885741 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced alarms from the supporting automated impella controller (aic). while changing the purge cassette, the cassette did not prime and so an air in the purge system alarm was triggered. the decision was mad to replace the purge cassette with a new one. during the exchange, the aic triggered a controller error alarm. the alarms self resolved within ten seconds. the controller was not changed out and was operational at its pervious settings. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller. another report was submitted to represent the pump. refer to section h.10 for the related report number.
4912 | 1220648-2025-47004 | 23169521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23169521&pc=OZD | 08/30/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699031 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 77 KG | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support a nurse contacted the local team noting a drop in purge flow and an increase in impella flow. the local team reviewed the waveforms and events with csc, and it was determined that most likely there was biomaterial or clot ingested. both the intensivist and cardiologist were notified and were advised of the findings. the device was removed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of the reported failure mode has been completed. no product was received for investigation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. saturated pump flow: data logs were reviewed and real time (rt) log at time of reported issue showed suction alarms with a slight and gradual saturation in pump flows. approximately 20 minutes later, a small motor current spike followed by more saturated flows was observed. additionally, a drop in purge flow and increase in purge pressure occurred at the same time. saturated flows were able to resolve after about 3 hours when p-level was lowered to p-2 then immediately turned back up to p-5. the cause of the saturated flows were due to biomaterial ingestion based on returned data log analysis. device history lot : device lot: 1931548. device history batch : subcomponent lot: n/a. device history review: the device passed all the post sterile inspection check.
4913 | 1220648-2025-47005 | 23170720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23170720&pc=OZD | 06/25/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665943 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, the pump was inappropriately alarming ¿impella in aorta¿. this alarm has been occurring seemingly since implant, although automated impella controller metrics do not appear to suggest incorrect placement. per the team on transthoracic echocardiogram (tte), the impella has advanced too deep. the plan is for formal tte later today for repositioning. no active alarms or other impella related issues currently. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Additional information: d6b explant date. the investigation of the reported failure mode has been completed. the pump was not returned for investigation. device in wrong position: the cause of the positioning issue was not determined due to insufficient clinical information. optical sensor issue: the cause of the optical signal issue was not determined, as the product was not returned and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks.
4914 | 1220648-2025-47006 | 23171499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23171499&pc=OZD | 08/30/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714602 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 79 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a st-elevation myocardial infarction patient was taken to the cardiac catheterization laboratory for impella cp placement. the impella cp was placed however in the cardiac catheterization laboratory it was noted that they found that the patient had a perfed coronary. a covered stent was placed, however, blood continued to fill the pericardium. the patient deteriorated and was unable to stabilize. ultimately, the patient expired in the cardiac catheterization laboratory. | [5 empty] | The investigation has been completed. no product was returned for investigation. cardiac perforation: in the catheterization lab, it was found that the patient had a perforated left ventricle. the cause of the cardiac perforation was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
4915 | 1220648-2025-47008 | 23173033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23173033&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after set up and insertion, the console indicated "optical signal system error". it is noted that they were advised to remove and replace the catheter and use another automated impella controller for this case. there was no reported patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. the pump was not returned for investigation. optical sensor issue: since the pump wasn't returned for investigation and data log review was inconclusive, the cause of the optical sensor issue could not be determined. device history review: the device passed all the post sterile inspection checks. | Section a is unknown and e.1 initial reporter name is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two reports. this report represents the pump. another report was submitted for the automated impella controller. refer to section h.10 for the related report number.
4916 | 1220648-2025-31704 | 23173187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23173187&pc=PYX | 03/13/2023 | 09/29/2025 |  | 09/29/2025 |  |  | 03/13/2023 | 09/29/2025 |  |  | 10/31/2024 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023270938 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a bipella cp-rp support in which there were two reported malfunctions. the placement signal was unreliable on the cp after insertion of the rp. the pump remained on while being monitored. additionally, the team noted the first automated impella controller was replaced due to an optical sensor failure at the setup of the cp. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause for the optical signal issue is due to a broken sensor face. the data logs support this root cause due to the placement signal going out right at activation of the pump and the 5s parameters are similar to that of a broken sensor face. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4917 | 1220648-2025-47007 | 23173394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23173394&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1662880 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after set up and insertion, the console indicated "optical signal system error". it is noted that they were advised to remove and replace the catheter and use another automated impella controller for this case. there was no reported patient harm. | [5 empty] | Section a is unknown and e.1 initial reporter name is unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two reports. this report represents the automated impella controller. another report was submitted for the pump. refer to section h.10 for the related report number. | H6: updated codes based on the results of the investigation. h11: updated based on the results of the investigation. investigation summary: the device was not returned for evaluation. data analysis: console imc logs were not available. however, after reviewing both the lt and rt logs, the placement signal not reliable event was observed. root cause: the root cause for the aic indicated optical sensor system error was not determined as the console was not returned.
4918 | 1220648-2025-31705 | 23173590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23173590&pc=OZD | 05/26/2023 | 09/29/2025 |  | 09/29/2025 |  |  | 05/26/2023 | 09/29/2025 |  |  | 05/31/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330265 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced placement signal issues. pump position was verified by echocardiography. despite this, the impella continued to provide effective hemodynamic support, and pump function was not affected. the patient subsequently expired; the event was not related to the device issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned pump was connected to an aic and placement signal were stable and normal. the field data logs were reviewed and the 5s parameters were stable throughout support. the root cause of the issue was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
4919 | 1220648-2025-31706 | 23173758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23173758&pc=OZD | 06/11/2023 | 09/29/2025 |  | 09/29/2025 |  |  | 06/11/2023 | 09/29/2025 |  |  | 12/31/2024 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315987 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported ongoing positioning alarms during impella cp support. it was noted that the implanted pump was triggering both false and positive aortic alarms during use. the device exhibited an optical sensor alarm. as troubleshooting was unsuccessful in resolving the alarms, placement monitoring was disabled. support continued with the implanted pump until the patient showed signs of heart recovery, at which point the pump was explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. according to the clinical, on the second day of impella cp support, false impella in aorta alarms as well as real aortic alarms were recorded. imaging confirmed the pump was in optimal position. the decision was then made to disable placement monitoring and have the nurse monitor motor current. no product was returned for a visual inspection. data log analysis confirmed the false impella in aorta alarms. additionally, all the 5s parameters were within normal bounds. the automated impella controller console software version was 11.1 during the time of use. the most likely cause of the optical signal issue was patient condition related due to software with false 'impella position in aorta' alarm observed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4920 | 1220648-2025-31707 | 23173949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23173949&pc=OZD | 04/09/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 04/09/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2025584972 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was unreliable. despite these issues, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for evaluation. data logs were reviewed and showed placement signal reading became higher than normal range for approximately 15 minutes before suddenly resolving. no hard failures of the optical parameters were noted at time of increased placement signal. the root cause of the optical signal issues could not be definitively determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4921 | 1220648-2025-31708 | 23174017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174017&pc=OZD | 06/03/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 08/04/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598011 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced a placement signal not reliable alarm occurred, the position waveform disappeared. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. the data logs review for the pump showed placement signal trend down while flows trended up. the root cause of the optical signal issue was most likely sensor silicone wear due to the drift of the placement signal and flows seen in the data logs. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4922 | 1220648-2025-31709 | 23174070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174070&pc=OZD | 06/06/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 08/05/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1086076 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during support, the optical signal on the automated impella controller (aic) shockwave was lost. despite this, the pump remained in use for continued support. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint and show an controller error alarm. the console was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. additionally, upon inspecting, intermittent distortions in the signal (envelope/fringe) were observed, while the measured "5s" parameters. the root cause of the controller error was due to a firmware issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4923 | 1220648-2025-31710 | 23174193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174193&pc=OZD | 02/17/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 04/01/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1137162 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 support ongoing for a patient with extensive cardiac and systemic comorbidities awaiting transplant. after 39 days the pump was explanted with transplant. during the support the optical signal had been lost. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data analysis: imc logs and lt logs confirmed that the placement signal was unreliable on the date of the reported complaint. there were multiple subsequent placement signal not reliable (psnr) events, and alarms were triggered for each psnr event. lt logs revealed that the psnr events coincided with an oscillating contrast value far outside the expected range of contrast values. this oscillating contrast issue has been previously observed in a few other consoles that exhibited similar errors. a root cause has not been yet identified; however it is known that an oscillating contrast is the result of a hardware fault on the optical bench. device analysis: console was sent back to field service for further investigation. the issue could not be reproduced in the lab. this issue is a low probability event and is very difficult to reproduce. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4924 | 1220648-2025-47009 | 23174210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174210&pc=OZD | 08/30/2025 |  |  | 09/29/2025 |  |  | 08/30/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713135 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented for possible viral component to myocarditis and impella 5.5 implant. the 5.5 was successfully placed. during support, arrhythmias were noted with moving. echocardiogram showed position stable and dobutamine ruled to be the cause of arrhythmias. however the next day it was also noted that the patient was resting on ventilator due to continued arrhythmias. in addition, there was a placement signal not reliable alarm. aberrant waveforms noted on aic - specifically aorta/left ventricle. the patient continued on support until the 5.5 was successfully weaned and explanted. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: optical sensor issue: clinical details report a placement signal not reliable (psnr) alarm with abnormal placement signal (ps) waveforms after a hypertensive episode on (b)(6) 2025. data logs were reviewed and the psnr alarms were confirmed. signal not reliable (snr) was widespread (0-8000) throughout the case, but light/gain/lamp/contrast remained stable, indicating there was no lack of light to the optical sensor. in reviewing the rt log at case start, it was noted that as the ps trended up toward 100 mmhg, the ps became more and more unreliable and snr began to drop more frequently/lower. there was also an instance on 30/aug where the ps dropped below 80 mmhg and the snr improved dramatically. these are indications that the console could be responsible for the reported optical signal issues. it was reported that the patient had a hypertensive episode and arrhythmias throughout support, however the unreliable ps noted in data logs seems to be more frequent than the patient conditions described. product wasn't returned for investigation. upon reviewing data logs, it is apparent that the console is the potential cause of the issue. the cause of the optical sensor issue could not be conclusively determined as the pump was not returned for investigation. arrythmia: it was reported that the patient was having arrhythmias throughout support. in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1947664. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
4925 | 1220648-2025-31711 | 23174231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174231&pc=OZD | 02/14/2024 | 10/10/2025 |  | 09/29/2025 |  |  | 02/14/2024 | 09/29/2025 |  |  | 09/30/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409428 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced a placement signal not reliable alarm. the alarm was disabled. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | D4: corrected the udi. h4: added manufacturer date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for evaluation. a kink in the catheter was found just proximal to the motor housing the 16cm marking. the laser light continuity test was performed and it was found that there was a fiber break at the cannula kink with laser light coming from the kink and no light at the senor face. data logs were reviewed and it was found that the optical 5s parameters showed a pattern related to a fiber break. gain also decreases slightly and lamp is at 100% contrast at time of break. the root cause of the optical signal issue is due to a catheter kink that caused the optical fiber line to break. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4926 | 1220648-2025-31712 | 23174398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174398&pc=OZD | 07/23/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 07/23/2024 | 09/29/2025 |  |  | 11/30/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430967 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced a placement signal alarm. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were returned for evaluation. the root cause of the optical signal issue cannot be determined as limited clinical details were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4927 | 1220648-2025-31713 | 23174467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174467&pc=OZD | 03/20/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 03/20/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2025458206 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the use of a cp pump to support a patient undergoing high-risk percutaneous coronary intervention. during support, the pump experienced on going placement signal not reliable alarm. the audio was disabled. despite this, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs from the reported day of event are consistent with complaint and show 16 alarms ¿placement signal not reliable." the console was tested in danvers, but the issue was not reproduced. analyzed data is consistent with defective optical bench or optical bench lamp assembly. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4928 | 1220648-2025-31714 | 23174656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174656&pc=OZD | 06/04/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 06/04/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, placement signal¿not reliable alarms were observed. despite these alarms, the pump remained operational throughout the procedure until weaning and explantation. no patient harm was reported. the investigating engineer determined that the issue was likely related to the automated impella controller (aic). | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data log shows that the contrast was oscillating. this console was tested, the 5s parameter indicates that the signal not reliable. the root cause of the placement signal not reliable alarm was the malfunction of the optical bench. the root cause for the controller error and loss of placement signal is due to an optical bench communication protocol anomaly, resulting in misalignment of data communication packets. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4929 | 1220648-2025-31715 | 23174690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174690&pc=OZD | 03/26/2023 | 10/24/2025 |  | 09/29/2025 |  |  | 03/26/2023 | 09/29/2025 |  |  | 01/31/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024294319 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | H4 device manufacture date updated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned and data logs show corruption occurred with the impella system. the root cause optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D6a: corrected the implant date. d6b: corrected the explant date. | G3: corrected date.
4930 | 1220648-2025-31716 | 23174714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174714&pc=OZD | 06/08/2023 | 09/29/2025 |  | 09/29/2025 |  |  | 06/08/2023 | 09/29/2025 |  |  | 04/30/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330696 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump and data logs was not returned. per the clinical information provided, the root cause of the optical signal issue was not determined since product was not returned and insufficient clinical information was provided. all pertinent information available to abiomed has been submitted at this time.
4931 | 1220648-2025-31717 | 23174824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23174824&pc=OZD | 02/05/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 02/05/2024 | 09/29/2025 |  |  | 10/31/2025 |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415027 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited sporadic placement signal fluctuations. despite this, the flows remained unchanged and the patient remained stable. pump position was confirmed. the placement signal feature was disabled. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and "impella position in aorta" occurred. the root cause of the optical signal issues cannot be determined as the pump was not returned and the pattern of the optical parameters are not consistent with known fiber/sensor break or air gap patterns. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4932 | 1220648-2025-31718 | 23175100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23175100&pc=OZD | 10/14/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 10/14/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553896 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was unreliable, with erratic waveforms noted. the alarm audio was disabled. despite these issues, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. impella cp returned for evaluation. the root cause was changed to 'optical signal issue' which is more fitting for the reported failure mode as the complaint is against the optical sensor of the pump. no product was returned for analysis. the data logs confirm placement signal not reliable alarm. no problem was found in the pump as the root cause of the optical signal issue. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4933 | 1220648-2025-31719 | 23175371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23175371&pc=OZD | 01/01/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 02/10/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133146 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed optical signal. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are incomplete, but observed pattern of optical signal is consistent with temporary loss of light, either produced by optical bench. console was tested in danvers, but issue was not reproduced. additional inspection and testing of the console revealed persistent contamination of the optical pump connector and out-of-spec optical power caused by light ¿leaking¿ into the cladding of optical fiber due to small bend radius. the root cause of the issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4934 | 1220648-2025-31720 | 23175479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23175479&pc=OZD | 01/09/2025 | 10/10/2025 |  | 09/29/2025 |  |  | 01/09/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582085 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, the placement signal was lost. despite these issues, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. impella cp was returned for evaluation. upon visual inspection, a minor kink in the catheter next to the red handle was noted. optical parameters of returned pump was measured and no hard failures of the optical sensor was observed. pump passed laser continuity testing, no ingested biomaterial observed on returned pump. pump was run in the lab for approximately 5 minutes without issue or alarms. data logs were reviewed and showed repeated suction alarms occurring while pump was running at p-7. the root cause of the optical signal issue cannot be definitively determined as limited clinical details were provided on patient's condition at time of issue and no issues with returned pump were observed. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | D4: corrected udi. h4: added the manufacturer date.
4935 | 1220648-2025-31721 | 23175955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23175955&pc=OZD | 11/16/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 11/25/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1701427 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during impella support, the optical signal was intermittently lost but resolved spontaneously. later, the team escalated to the 5.5 pump. upon investigation, the team determined that the console should also be reviewed, even though it was not replaced. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs showed the console displayed ¿placement signal not reliable. this confirms the reported failure mode. the console was returned and console was tested. the root cause of the placement signal not reliable was the defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4936 | 1220648-2025-31722 | 23176537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23176537&pc=OZD | 02/10/2025 | 09/29/2025 |  | 09/29/2025 |  |  | 02/10/2025 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1756407 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed placement signal unreliable. the audio was disabled. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of the event are consistent with the complaint. console was sent back to field service for further investigation. the reported complaint could not be reproduced. while measuring the "5s" parameter using optical cavity found that the contrast and lamp power values were low than the specified values. the root cause for the low placement signal and suction alarms, was most likely due to the malfunction of the optical bench or its components. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4937 | 1220648-2025-31723 | 23176543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23176543&pc=PYX | 06/12/2024 | 09/29/2025 |  | 09/29/2025 |  |  | 06/12/2024 | 09/29/2025 |  |  |  |  | 09/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025483301 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm occurred on the impella rp. the waveforms and flows were still present on screen. drastic change in numbers for pulmonary artery waveforms eventually resolved. no change in p-level or motor current occurred. the patient survived the event, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs show that one hour into the case the placement signal (ps) begins to drift upwards for ~3 hours, it then stabilizes. central venous pressure (cvp) is relatively stable at this time and motor current (mc) is reflective of p-level changes. there are a few occurrences throughout the case where both optical and dp sensor are affected. 15 hours into the case there is another severe instance of dp sensor drift over a period of 2 hours. cvp and mc and p-level are stable at this time. ps drifted from 70 to 150 mmhg. it then drifts back down over one hour. there is a brief instance of placement signal not reliable (psnr) optical at the start of the case which is resolved. the 5s parameters are relatively stable and within normal ranges throughout support. there are suction alarms throughout the case but they do not occur during the drift. pump was run in a pressurized flow loop at after running for 5 days the ps had a 5 hour rise from ~20 to 60 mmhg. dp drift was reproduced in the flow loop. the pump passed electrical testing. the returned 5s parameters were within normal limits. the root cause of the placement signal issue was likely dp sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4938 | 1220648-2025-46937 | 23163384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163384&pc=OZD | 08/25/2025 |  |  | 09/28/2025 |  |  | 08/28/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026655541 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient was on long term support of impella 5.5 as bridge to durable lvad. on anticoagulation for duration of support. on removal in or during durable lvad procedure clot found on inflow and outflow of device. pump was discarding but did receive images from md. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: d4, g4. the investigation of the reported failure mode has been completed. no product was returned for investigation. thrombosis: the root cause of the thrombosis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4939 | 1220648-2025-46938 | 23163385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163385&pc=OZD | 08/25/2025 |  |  | 09/28/2025 |  |  | 08/28/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674555 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with 100% occluded mid left main and mid-left anterior descending. decision was made to place impella cp. it was noted that the cp was placed, however the staff reported that when the impella cp was started, the pump did not go into auto but rather started at p-0. it was about 20-30 seconds before the staff noticed that the patient was not being supported. staff said that it caused no harm or procedural problems, they just had to turn up the p-level. the patient was on two pressors and the patient¿s lactic had risen. the critical care team placed a swan, a-line, hemisphere, and triple-lumen catheter, but by that time the patient¿s creatinine had risen further and his lactic acid was now 14.1. physician prepared to cannulate to extracorporeal membrane oxygenation, however they believed the patient was too far gone. by that point the patient¿s lactic was 17, the patient was on 5 vasoactive drips and his svo2 was 50%. the decision was made to medically/chemically treat the patient. the patient¿s family soon decided to withdrawal care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4940 | 1220648-2025-46940 | 23163778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163778&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700990 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 support for mechanical circulatory support (mcs). ejection fraction on initial presentation was 20-25% but declined to 15% per intraoperative transesophageal echocardiogram. the patient underwent coronary artery bypass graft surgery x4 vessels but the patient failed to come off cardiopulmonary bypass. the decision was then made to place direct aortic impella 5.5 which was successfully completed. it was noted there was some bleeding in chest cavity unrelated to impella. hemoglobin drop as low as 7.8 intraoperatively and 4 units of fresh frozen plasma, 6 units of packed red blood cell and 3l of cellsaver were administered. the impella 5.5 position was confirmed in good orientation at 5.2 cm from aortic valve. the patient was sent to the unit on support. the patient arrived at the unit, and after a couple of hours, it was noted the pulmonary artery catheter was not rationale values for thermodilution output and index due to tricuspid regurgitation. the patient's pulsatility was reduced at times. the physician at bedside and ordered additional blood products for volume resuscitation. transthoracic esophageal echocardiogram showed the device at the appropriate depth and orientation. later in the afternoon the patient would experience sustained ventricular tachycardia. suction and profound hypotension followed. the patient was not defibrillated as the patient began to self-convert before code cart could be brought to room. hypotension followed, however, with mean arterial pressure as low as 20 mmhg. cardiopulmonary resuscitation was initiated briefly and the impella device was dropped to p-2 performance level. return of spontaneous circulation was noted and device returned to p-8 after assessing for appropriate waveforms. adjustments made to drips at that time: milrinone turned off, and amiodarone initiated. pulse pressure, map and other parameters normalized and device pulling close to 5 l/min of flow at this time. this same day the nurse also reported the patient was taken back to the operating room. the patient was unstable and bleeding. the nurse was unsure where the bleeding was coming and did not have much information to share regarding this event. however, it was noted a couple of clots were present and removed. the patient would continue on support. two days later it was noted that the patient continued to slightly improve each day and was hemodynamically stable. two days later the device was explanted. the patient was successfully weaned. ejection fraction was 25%. no inotropic or pressor support was required after explant, and the blood pressure and urine output were stable. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, the major bleed, the nurse is unsure where the bleeding was coming and does not have much information to share. the cause of the bleeding is determined to be patient condition because of the unknown location of the bleeding. regarding the thrombosis, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4941 | 1220648-2025-46939 | 23163779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163779&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1199223 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller console was properly upgraded for the implantation. however, after approximately 30 minutes, the soft buttons were no longer working or were no longer responding. the console was replaced without issue. there was no report or indication of patient involvement. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
4942 | 1220648-2025-46944 | 23163781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163781&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678218 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) as a bridge to transplant. after approximately 36 days on support, the patient experienced ventricular tachycardia. point-of-care ultrasound was completed and noted the impella device positioned as 5.4cm and mid-ventricular space. the patient also had low potassium which appeared to be the issue per the physician. however, this was not confirmed. the patient continued on support for an additional seven days before being successfully weaned and bridged to transplant. | [5 empty] | Patient-specific information a4: weight is unknown. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ventricular tachycardia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4943 | 1220648-2025-46941 | 23163782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163782&pc=OZD | 08/29/2025 | 10/08/2025 |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026728331 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | No | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup | [8 empty] | The us complainant had a cp pump being placed via the single access at the femoral for the patient admitted in for a hrpci. during support initiation process, impella recognized as old generation with grey power cable. the device was disconnected and prime restarted which was successful. lower than expected flows were observed in p auto without suction alarm. all attempts made to trouble shoot with after load reduction due to high mean arterial pressure (map). volume was given and impella repositioned which was all unsuccessful. case resumed as is which was successful. the patient was reported to be stable. | [5 empty] | B.7 information was added as it was inadvertently omitted from the initial report that was submitted. d.3 manufacturer name should not have contained numbers behind it in the initial report that was submitted. manufacturer fax number was added as it was inadvertently omitted from the initial report that was submitted. d.4 model name was added as it was inadvertently omitted from the initial report that was submitted. d.7a revised selection as it was entered incorrectly on the initial report that was submitted. d.9 the device was returned and date returned was added. e.1 initial reporter facility name, initial reporter address line 1, initial reporter phone and zip code extension were added as they were inadvertently omitted from the initial report that was submitted. h.6 codes were updated as the device was returned. h.10 related report number was added as it was inadvertently omitted from the initial report that was submitted. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | There are two associated devices for the reported event. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4944 | 1220648-2025-46943 | 23163785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163785&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713998 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 102 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced ventricular tachycardia. the patient subsequently expired. the patient was initially implanted with an impella cp with extra corporeal membrane oxygenation (ecmo). after approximately six days on these therapies, the decision was made to escalated to an impella 5.5 which was completed and decannulate ecmo. the next day after start of the impella 5.5 support, the patient was getting ready to be taken to the catheterization lab for intervention and went into refractory ventricular tachycardia. the critical care team was unable to break the ventricular tachycardia., and the patient expired. the family made the decision to withdrawal care at this point. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ventricular fibrillation, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4945 | 1220648-2025-46945 | 23163787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163787&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714620 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an unsuccessful placement of an impella cp device to a patient in cardiogenic shock from an acute myocardial infarction. accessed was gained and impella cp device was prepped. vascular access in the groin became limited and the physician was unable to upsize to a 14fr. sheath as the anatomy was too small. the decision was made to abort the implant. the status of the patient following the event was unnoted. however, there was no report or indication of an adverse event or outcome. | [5 empty] | The investigation of the reported failure mode has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient anatomy was too small preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4946 | 1220648-2025-46946 | 23163790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163790&pc=OZD | 06/24/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026675635 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that experienced thrombocytopenia, recurring with a "probable" relationship to the impella device used: ¿the patient developed recurring thrombocytopenia during this admission. admission platelets were too clumped to quantify. first quantifiable platelet count 68 k/ul (on (b)(6) 2025 at 0937), the patient developed reportable thrombocytopenia on (b)(6) 2025 at 2304 with a platelet count of 36 k/ul. admission nadir platelets 13 k/ul (on (b)(6) 2025). platelets were above reportable values briefly on (b)(6) 2025 with a count of 52 k/ul before dropping back down to 45 k/ul on (b)(6) 2025 at 0606. the patient again went above reportable values on (b)(6) 2025 with a count of 71 k/ul before dropping back down to 41 k/ul on (b)(6) 2025 at 1742. the patient remained below the reportable threshold for the remainder of their admission. they received two units of platelets on (b)(6) 2025, respectively. heparin-induced thrombocytopenia assay was negative.¿ additionally, the patient also experienced recurring non-sustained ventricular tachycardia with a "probable" relationship to the impella device used: ¿the patient developed recurrent episodes of non-sustained ventricular tachycardia. the patient experienced many of these episodes upon any movement. point of care echocardiography confirmed impella position. the patient's underlying rhythm was sinus with paroxysmal atrial fibrillation.¿. the cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4947 | 1220648-2025-31699 | 23163792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163792&pc=OZD | 10/18/2024 | 09/28/2025 |  | 09/28/2025 |  |  | 10/18/2024 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump lost optical signal. despite this, support was able to continue using the implanted pump. no patient harm resulted from the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the failure mode was changed from optical signal issue to positioning issue after data log review confirmed the clinical report that repositioning the pump resolved the complication. product was not returned for investigation. data logs were reviewed and the reported sensor issues on 10/17 were confirmed. several times throughout that day, signal not reliable (snr) dropped to zero and contrast also dropped while gain, lamp, and light were unaffected. based on clinical details and data logs, the root cause of the positioning issues was determined to most likely be a use issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4948 | 1220648-2025-46947 | 23163794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163794&pc=OZD | 08/08/2025 |  |  | 09/28/2025 |  |  | 08/08/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287313 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 80 KG |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 and supported by an automated impella controller (aic) experienced a loss of optical signal. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4949 | 1220648-2025-46948 | 23163796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163796&pc=OZD | 08/12/2025 |  |  | 09/28/2025 |  |  | 08/12/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1579730 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 56 YR |  | Male | 83 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a patient escalated from an impella cp to an impella 5.5 pump-aic as a bridge to recovery due to cardiomyopathy. while on support, a ¿placement signal unreliable¿ alarm was noted with patient movement. it resolved when the patient settled. at this time, the patient is still on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4950 | 1220648-2025-46949 | 23163801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163801&pc=OZD | 08/25/2025 |  |  | 09/28/2025 |  |  | 08/25/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642749 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had positive blood cultures. it is unknown how the patient was treated. later, the patient was successfully weaned from the pump and survived. | [5 empty] | B.2 revised selection as it was reported incorrectly on the initial report that was submitted. b.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings, ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the investigation has been completed. the root cause of the sepsis was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4951 | 1220648-2025-46950 | 23163803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163803&pc=OZD | 08/24/2025 |  |  | 09/28/2025 |  |  | 08/24/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523939 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 80 KG | Non Hispanic | Asian | Injury | Yes | No | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Cardiac Arrest | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp encountered positioning alarms when the patient coughed. the patient experienced ventricular tachycardia and supraventricular tachycardia requiring defibrillation. when the patient laid flat the impella migrated and had to be pulled back into position. the patient also had hematuria and hemolyzed labs. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the positioning issue was not determined due to insufficient clinical information. the cause of the hemolysis could not be determined because the product was not returned for investigation and sufficient clinical information, including the complete data log, was not provided. the pump passed all post sterile inspection checks.
4952 | 1220648-2025-46952 | 23163804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163804&pc=OZD | 08/22/2025 |  |  | 09/28/2025 |  |  | 08/22/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594977 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 81 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. another perclose device was deployed to obtain hemostasis. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. major bleed : the cause of the injury was not determined since the product was not returned and insufficient clinical details provided. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4953 | 1220648-2025-46951 | 23163805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163805&pc=OZD | 08/23/2025 |  |  | 09/28/2025 |  |  | 08/23/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 61 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted because the wire could not pass the iliac artery near bifurcation. no patient harm was reported. | [5 empty] | No product was returned for investigation. the cause of the delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to heavy iliac disease. due to aortofemoral disease too severe, interdisciplinary consensus was to abort case to avoid risk of aortic perforation. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4954 | 1220648-2025-46954 | 23163807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163807&pc=OZD | 08/13/2025 |  |  | 09/28/2025 |  |  | 08/13/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667321 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Difficult to Advance | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was implanted to support a patient admitted for cardiomyopathy. while on support, the impella 5.5 was advanced but could not pass the bend of the innominate artery. after several unsuccessful attempts with a buddy wire, a v18 was placed through the introducer and into the descending aorta which then allowed for the 5.5 to pass the bend at the innominate. once in the ascending artery, the v18 was removed and impella 5.5 was advanced into the left ventricle. a large hematoma was noted at the impella insertion site. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the delivery issue was most likely a use issue. the cause of the hematoma is most likely use related since large hematoma noted at impella insertion site upon arrival. the device passed all post sterile inspection checks.
4955 | 1220648-2025-46955 | 23163808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163808&pc=OZD | 08/16/2025 | 01/07/2026 |  | 09/28/2025 |  |  | 08/16/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626883 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | Aortic Valve Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented for successful 3 vessel coronary artery bypass grafting however quickly decompensated after chest was closed. cardiopulmonary resuscitation performed and patient reopened. decision for impella 5.5 insertion was made. the 5.5 was successfully placed. during support it was noted hat the patient experienced several bouts of non-sustained ventricular tachycardia (vt). pump was shallow however cardiology did not want to reposition. the patient with continued with bursts of vt yet self resolving. additionally, the patient had atrial fibrillation (afib), attempted cardioversion but unsuccessful, amio was on board and rate was more controlled however still atrial fibrillation. it was also noted that the patient had some aortic insufficiency, however was hemodynamically stable. the 5.5 was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary pump was not returned for investigation. usb data drive was returned with data logs. data log review was not required for this case run. peri-procedural adverse event: aortic valve insufficiency/ regurgitation: the cause of the injury was not determined due to insufficient clinical details. ventricular fibrillation, arrhythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history lot device batch: 1905620 device history batch subcomponent lot: na device history review the pump sn561304a passed all post sterile inspection checks.
4956 | 1220648-2025-46956 | 23163809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163809&pc=OZD | 08/08/2025 |  |  | 09/28/2025 |  |  | 08/08/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688991 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced sepsis and gastrointestinal bleeding. the patient was transfused one unit of packed red blood cells and treated with antibiotics for suspected septic shock. additionally, the patient had black stool for which systemic anticoagulants were put on hold. | [5 empty] | A4, a5, and a6 are unknown. no product was returned for investigation. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details and no product was returned for analysis. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
4957 | 1220648-2025-46959 | 23163811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163811&pc=OZD | 08/09/2025 |  |  | 09/28/2025 |  |  | 08/09/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628396 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 110 KG |  |  | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while implanted with an impella cp a pericardial effusion present. the p level was decreased and bicarb was added to the purge fluid. bleeding was present from the ecmo cannula. five units of fresh frozen plasma and five units of packed red blood cells were transfused. systemic heparin was withheld. later, care was withdrawn and the patient expired. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product was not returned for investigation. the cause of the pericardial effusion was not determined due to insufficient clinical details. the cause of the bleed was not determined due to insufficient clinical details. the cause of the suction alarms was most likely patient related, as clinical states the patient was on ecmo through the case, and data logs show a loss of placement signal pulsatility and downward trend. the pump passed all post sterile inspection criteria.
4958 | 1220648-2025-46958 | 23163812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163812&pc=OZD | 08/09/2025 |  |  | 09/28/2025 |  |  | 08/09/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703436 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Inaccurate Flow Rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Fever | Hemolysis | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of an impella 5.5 the patient became febrile and received empirical treatment for a urinary tract infection. after experiencing ventricular tachycardia, the patient received intravenous amio. proper positioning of the pump was confirmed after the patient experienced hemolysis. later, the pump was explanted and the patient survived. | [5 empty] | A4 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for evaluation. the root cause of the fever, tachycardia, and thrombosis was unable to be determined due to insufficient clinical details. the cause of the hemolysis could not be determined as no product was returned for evaluation and limited clinical details were provided. the cause of the pump flow saturation cannot be determined as limited clinical details were provided and no product was returned for evaluation. the device passed all post-sterile inspection checks.
4959 | 1220648-2025-46960 | 23163813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163813&pc=OZD | 08/11/2025 |  |  | 09/28/2025 |  |  | 08/11/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705923 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female | 94 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Pumping Stopped | [6 empty] | Arrhythmia | Fever | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted in a patient admitted with cardiogenic shock. while on support, the patient experienced ventricular tachycardia requiring defibrillation. the patient also had a fever and positive blood cultures. a washout of the right knee and debridement was performed, and the patient was put on antibiotics. the pump generated an ¿impella stopped¿ alarm and several unsuccessful attempts were made to restart the pump. the patient stated when they were moving from the chair to the bed, they felt the impella cord tug on their stomach and the pump stopped. due to the complications encountered, the device was removed and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4960 | 1220648-2025-46961 | 23163814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163814&pc=OZD | 08/13/2025 |  |  | 09/28/2025 |  |  | 08/13/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025593517 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a pump prolapse into the aorta. the pump's catheter was pulled back into the descending aorta and the cannula was straightened. the pump then floated into the left ventricle. a hematoma was noted in the right groin so pressure was applied and the hematoma resolved. due to low platelet count systemic heparin was withheld, and the patient was transfused red blood cells and platelets. an ultrasound at the bedside showed the impella was under the papillary and the patient had a type a aortic dissection. the impella cp was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position after wire removal. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the cause of the hematoma was not determined due to insufficient clinical details. the pump passed all post sterile inspection checks.
4961 | 1220648-2025-46741 | 23163815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163815&pc=OZD | 08/13/2025 | 09/22/2025 |  | 09/28/2025 |  |  | 08/13/2025 | 09/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708838 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | Malposition of Device | Patient Device Interaction Problem |  |  |  | Stroke/CVA | Hemolysis | Thrombosis/Thrombus | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that after impella was placed and started in auto, impella had suction alarms. albumin was administered resolving suction alarms. the patient also had a cat scan scan when arriving at intensive care unit revealing possible stroke. bival was turned off. an ischemic stroke was confirmed, and patient was neurologically intact. anticoagulation was restarted. on day six of impella support, labs were trending up, urine slightly pink tinged and labs coming back hemolyzed. impella cp inlet appeared to be pointed toward the mitral valve with pigtail in papillary muscle. the medical doctor pulled the impella back and opted to run it shallow to avoid having inlet hitting the mitral valve. on day seven of cp support, motor current (mc) with spike noted and immediately following spikes in mc, suction alarm present. mc following spike noted to be flat and no longer pulsatile. flows dropped from 1.5 to 0.8l. volume status was adequate, and echocardiogram done that showed inlet across aortic valve. impella at p-2 with veno-arterial extracorporeal membrane oxygenation (va ecmo) in place flowing around 2.5l. plan had already been for cp and va ecmo removal that day, so device removed as planned. upon removal it appeared there was questionable thrombus ingestion within the outlet of impeller. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections: e4. the investigation of the reported failure mode has been completed. the impella device was returned for investigation. data logs were not recovered. returned product had biomaterial wrapped around the impellor blade but showed no other physical abnormalities. mechanical interaction with blood: the cause of the mechanical interaction with blood was a pump positioning issue. device in wrong position: the cause of the pump being too deep requiring repositioning was not determined due to insufficient clinical details. low pump flow and pump suction: the cause of the low pump flow and pump suction was biomaterial cause by ingestion, due to biomaterial found on the returned pump and the mc spike noted in the clinical details. clinical symptoms: stroke, thrombosis: the root causes of the stroke and thrombus were unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | Product complaint # (b)(4). investigation summary: impella cp sn (b)(6) returned. after 6 days on impella support, ldh trended up with urine slightly pink tinged. echo confirmed inlet was pointed at the mitral valve with pigtail in papillary muscle. md pulled impella back and left shallow to avoid mv interaction. urine output returned to yellow following repositioning. data logs were not recovered. returned product had biomaterial wrapped around the impellor blade but showed no other physical abnormalities. -mechanical interaction with blood - the cause of the mechanical interaction with blood was a pump positioning issue. -device in wrong position - the cause of the pump being too deep requiring repositioning was not determined due to insufficient clinical details. after 7 days on cp support, mc spike occurred followed by suction alarms. flows dropped from 1.5 to 0.8l. data logs were not recovered. returned product showed biomaterial in the inlet cage and wrapped around the impellor blade. -low pump flow and pump suction - the cause of the low pump flow and pump suction was biomaterial cause by ingestion, due to biomaterial found on the returned pump and the mc spike noted in the clinical details. clinical symptoms - stroke, thrombosis - in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root causes of the stroke and thrombus were unable to be determined due to insufficient clinical details. device history lot: device lot: 1955125. device history batch: subcomponent lot: n/a device history review: pump s/n (b)(6) passed all post sterile inspection checks.
4962 | 1220648-2025-46963 | 23163817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163817&pc=OZD | 08/07/2025 |  |  | 09/28/2025 |  |  | 08/07/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675462 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Inflammation | Thrombosis/Thrombus | Swelling/ Edema | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella cp on (b)(6) 2025, the patient remains intubated and just returned from a pan ct scan to evaluate for source of bleeding. blood was noted in the oral cavity. the patient had undergone multiple bronchoscopies. the plan includes continued evaluation, colonoscopy, and gradual weaning of ecmo support as tolerated. the patient with the new extensive large multi compartment intracerebral hemorrhage involving the right front lobe and basal ganglia, with intraventricular extension into the third and fourth ventricles. ct also shows diffuse cerebral edema. the patient underwent bronchoscopy today revealing clots within the lungs. additionally, the impella cp appears deep in the left ventricle with inlet near mv apparatus; plan to reposition later today under echocardiogram guidance. patient with new extensive large multi compartment intracerebral hemorrhage involving the right front lobe and basal ganglia, with intraventricular extension into the third and fourth ventricles. computed tomography also shows diffuse cerebral edema, multifocal infarcts, and a 16 mm leftward midline shift with evidence of right subfalcine and trans tentorial herniation. given the grave prognosis, family elected to withdraw care and the patient expired. | Additional information has been provided: "the patient experienced a new large intracerebral hemorrhage with midline shift. following this clinical change, the family elected to withdraw care. the device is not available for return at this time. no return kit needed. no additional information available regarding this abiomed product for this patient." |  |  |  |  | The investigation was completed. no product was return. the root cause of access site bleeding was most likely patient condition related since the bleed was occurring from non-impella site like oral cavity along with intracerebral hemorrhage involving the right front lobe and basal. edema: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review passed all post sterile inspection checks. | Corrected: h6 health effect - clinical code. | B5 updated.
4963 | 1220648-2025-46964 | 23163818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163818&pc=OZD | 08/08/2025 |  |  | 09/28/2025 |  |  | 08/08/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1670339 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 115 KG | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the console automated impella controller (aic) had a purge sensor issue/malfunction during use in the intensive care unit. it was noted that when the nurse was changing purge fluid and tubing, the aic was not detecting the purge disc sensor. however, the procedure was completed successfully. there was no harm to the patient reported. | [5 empty] | Corrections: d6a and d6b should have been left blank. h5, h8. the investigation of the reported failure mode has been completed. data review: console logs were consistent with what was reported in the complaint. device analysis: console was tested after and the purge disc detection issue was able to be reproduced. further inspection of the console revealed that the purge disc flag was missing. conclusion: the root cause of the purge disc detection issue was the observed broken purge disc flag (missing at blue retainer). device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
4964 | 1220648-2025-46957 | 23163819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163819&pc=OZD | 08/16/2025 | 09/22/2025 |  | 09/28/2025 |  |  | 08/16/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711487 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Female | 41 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Detachment of Device or Device Component | Patient Device Interaction Problem |  |  |  |  | Fever | Hemolysis | Low Blood Pressure/ Hypotension | Renal Failure | Thrombocytopenia | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella 5.5 placed to support the patient experiencing dilated cardiomyopathy. while on support, it was noted that there was a temperature overnight and that cultured would be monitored. the patient was put-on broad-spectrum antibiotics. next, the patient experienced hypotension. epinephrine was added and the impella was increased to p-7. additionally, continuous renal replacement therapy was started and continued. several suction events and alarms occurred. in some instances, the impella was decreased to p-6. hemolysis was noted. at this time, the patient is still on support. | It was further reported that the patient was approved for heart transplant. will continue on kidney and heme management of heparin induced thrombocytopenia (hitt), pf4 and sra remain positive, plans to restart plex m-w-f schedule for transplant prep. the patient was successfully bridged to transplant. during explant of 5.5, utilized for bridge to recovery, impella motor housing detached from impella cannula. blue button lock mechanism unlocked. clot noted in vascular graft. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received: b5, d6b, and h6 updated. device received d9, h3 updated correction e4. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. pump suction: the data logs show that the suction alarms that are being triggered are #14 characteristic curve throughout the case. based on the clinical details and data logs, the root cause of the pump suction is due to the patients condition. mechanical interaction with blood (hemolysis): the data logs that there are suction alarms with slightly variable flows. there are no motor current spikes or sustained positioning alarms. the pump was returned with the cannula detached from the motor housing. no impeller damages were noted or cannula kinks. the product was unable to be hemolysis tested but there was no other defect found regarding the device. due to a lack of information regarding what resolved the hemolysis and the contributing device, the root cause of the hemolysis cannot be established. product damage (cannula detached from motor housing): the clinical details state that the cannula detached from the motor housing in the graft upon explant. it was noted that there was no resistance on explant or insertion. no additional tools were used on explant and that there wasn't any anatomical challenges but there was significant clot in the graft. the returned product confirmed that the cannula was detached from the motor housing at the location of the outlet. there were no deformities in the cannula as in any clamp marks or scratches seen. there was a scratch and damage to the pump housing but that is most likely from the hemostat after the detachment occurred. the delo band remained on the motor housing and there were no indications of it ripping or incorrect potting. residual glue marks are noted on the inside of the cannula and the indentation of it being bonded to the motor housing can been seen. on the surface of the lbb (motor housing) delo also remains as evidenced by the darker material on the surface. there was a slight kink seen in the catheter near the red pump plug. the root cause of the product damage (cannula detached from the motor housing) is due to a material fracture but the cause cannot be further established. pia-0000655 was opened to further investigate this issue. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia/hypotension/fever: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4965 | 1220648-2025-46966 | 23163820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163820&pc=OZD | 08/18/2025 |  |  | 09/28/2025 |  |  | 05/27/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026616460 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 114 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during the support patient had ventricular tachycardia storm and was shocked 6 times. the patient is awaiting orthotopic heart transplantation (oht). this patient have been on impella from (b)(6) 2025. | [5 empty] | Investigation summary ==> 5.5 pump sn (b)(6) (cut) returned on 9/12/2025. device will go to academic research team after investigation. product evaluation: returned complaint 5/5 pump visually inspected. cannula kink observed. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review ==> the pump sn (b)(6) passed all post sterile inspection checks.
4966 | 1220648-2025-46967 | 23163821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163821&pc=OZD | 08/18/2025 | 09/22/2025 |  | 09/28/2025 |  |  | 08/18/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026661709 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 113 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Patient-Device Incompatibility | [5 empty] | Atrial Fibrillation | Low Blood Pressure/ Hypotension | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella 5.5 support for mechanical circulatory support. the pump was reportedly mispositioned, and the physician declined to reposition despite urging from the manufacturer representative onsite. the patient went into ventricular tachycardia and atrial fibrillation (afib) with rapid ventricular rate (rvr), and was given amiodarone, versed, and bicarbonate before returning to nsr. the patient had additional episodes of afib/rvr and as cardioverted, becoming hypotensive and requiring increased pump support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. clinical symptoms ( tachycardia, atrial fibrillation): the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms ( low blood pressure/ hypotension): the cause of the injury was not determined due to insufficient clinical details. device in wrong position: pump was not returned for investigation. data log review was not required for this case run. the cause of the improper positioning issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. patient anatomy or condition prior to therapy: the cause of the patient anatomy or condition prior to therapy issue was not determined due to insufficient clinical information. device history review: the device passed all the post sterile inspection checks.
4967 | 1220648-2025-46965 | 23163822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163822&pc=OZD | 08/08/2025 |  |  | 09/28/2025 |  |  | 08/08/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026629132 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 99 KG |  | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella cp the patient exhibited a sudden drop in pressure and suction alarm on the impella. the impella found it to be shallow, just barely across the valve. impella was successfully repositioned under fluoro. vasopressors / inotropes increased. map eventually recovered. the nurse reported that the patient has occasional rounds of ventricular tachycardia (vt). impella flow increased to p-6 and levo gtt increased. the patient was shocked 13 times overnight. the patient decompensated with further vt and team was unable to resuscitate. the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. pump suction: clincial reported before patient moved from cath lab table, there was a sudden drop in pressure and suction alarm on the impella. impella found to be shallow, just barely across the valve. impella was successfully repositioned under fluoro. no product or data logs were returned. the cause of the issue was due to an unintended use error as clinical imaging suggest positioning issues. device in wrong position: clinical reported pump found to be shallow. no product or data logs were returned. the cause of the issue was not established as limited clinical information was provided. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and passed all post sterile inspection checks.
4968 | 1220648-2025-46968 | 23163823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163823&pc=OZD | 08/18/2025 | 09/22/2025 |  | 09/28/2025 |  |  | 08/18/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671904 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient encountered atrial fibrillation (afib), hemothorax and optical signal issue during impella 5.5 support. the patient was given amiodarone as a result of the afib. the optical signal issue was noted via a false impella in aorta alarm while patient was coughing. finally, heparin was discontinued due to no improvement in hemothorax despite the fact that bleeding stopped. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. optical signal issue: based on the clinical details and data log analysis, the root cause of the optical signal issue is most likely due to patient movement. arrythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4969 | 1220648-2025-46969 | 23163824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163824&pc=OZD | 08/18/2025 | 09/22/2025 |  | 09/28/2025 |  |  | 08/19/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702474 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female | 105 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after the patient was in intensive care unit on impella cp support, the patient had runs of ventricular tachycardia. the impella position was checked. on day two of impella support it was noted a left ventricle (lv) thrombus. cardiology assessed for aspiration of lv thrombus. the patient continued with impella support and was bridged to left ventricular assist device as planned. | [5 empty] | The investigation of the reported failure mode has been completed.no product was returned for investigation. thrombosis/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4970 | 1220648-2025-46971 | 23163826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163826&pc=OZD | 08/20/2025 | 09/26/2025 |  | 09/28/2025 |  |  | 08/20/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635464 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient encountered thrombus and placement signal issue during impella cp support. a day after impella placement, patient's echocardiogram showed a left ventricle (lv) thrombus. additionally, position alarms went off after patient had aggressive coughing. impella was pulled back using point-of-care ultrasound. position in aorta alarms were still noted when lv signal dampened but team was comfortable with position. the next day, erroneous position alarms were still noted and pump placement was verified. the nurse gave instructions on how to disable placement monitoring. however, after 3 days of support, care was withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. thrombosis: the cause of the injury was not determined due to insufficient clinical details. optical signal issue: the data logs show that about 34 hours into support there were impella position in aorta and impella position in ventricle alarms that were triggered. due to lack of clinical information on imaging confirmed to look at positioning and no product returned, the root cause cannot definitely be determined. device history review: the device passed all the post sterile inspection checks.
4971 | 1220648-2025-46972 | 23163828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163828&pc=OZD | 08/21/2025 | 02/24/2026 |  | 09/28/2025 |  |  | 08/21/2025 | 09/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705992 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Fluid/Blood Leak | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella cp that the impella measurement in the left ventricle appropriate does appear to be under the mitral valve. attempted to reposition but was unsuccessful. the bleeding was noted from red plug at the grey connection of the purge side arm. ic did have initial issue placing companion sheath and thought is that the catheter was punctured by the micro puncture needle. called by cvicu and made aware that patient had developed possible gastrointestinal (gi) bleed and was taken to computed tomography scan (ct). ct revealed gi bleed and patient taken to ir for coiling. ct also revealed no vascular injury from impella site. team was unable to find a pedal pulse in either foot, unsure at this time if there is limb ischemia. bilateral lower extremities warm with good cap refill. team is reconsulting vascular to determine if intervention is needed. clinical contacts given. patient to remain intubated overnight, team weaning vasopressors as able. patient bleed is associated with gastroepiploic artery. nurse states that they are now able to doppler distal pulses in both legs. plan today is to attempt to wean drips as patient tolerates. some concern for development of sepsis. the pump was explanted. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 updated; the product has been returned. the investigation is still ongoing.
4972 | 1220648-2025-46974 | 23163831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163831&pc=OZD | 08/21/2025 |  |  | 09/28/2025 |  |  | 08/25/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026713381 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing coronary artery bypass surgery was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was an episode of ventricular tachycardia episode for which the patient had to be intubated. patient had suction alarm and impella was repositioned to address the issue. pump was weaned and explanted as planned. patient survived the procedure. | [5 empty] | Product complaint # (b)(4). investigation summary: no product was returned. arrhythmia: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): pump was repositioned. the cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: no product was returned. patient had suction alarms. after reviewing the logs, suction alarms were observed. the cause of pump suction was most likely use related since the suction was resolved by repositioning of pump. device history review: device (sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
4973 | 1220648-2025-46975 | 23163832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163832&pc=OZD | 08/23/2025 |  |  | 09/28/2025 |  |  | 08/23/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705929 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Perforation | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was escalated from a cp to a 5.5. during support it was noted that there were suction alarms and the patient received 500 albumin for low volume in addition to fluids. it was also noted that the patient went to surgery for mitral valve and coronary artery bypass graft. which the patient had large bilateral pleural effusions and pericardial effusion. chest was opened and the patient had significant inflammation that was referred to dressler syndrome. 400ml bloody pericardial effusion was drained in addition to the bilateral pleural effusions. myocardium was very inflamed and it was decided to abort the cabg and only operate on the mitral valve. surgical mode utilized after antegrade cardioplegia. the valve was replaced. the heart was de-aired and clamp removed. there was a large amount of blood in the chest. the physician mentioned he saw the impella outside of the left ventricle and pulled back the impella 5.5. they repaired with large felt plegetted suture. they reassessed and noted that the impella was no longer across the valve. several attempts were made to re-cross the valve without success. impella 5.5 was removed from the body. it was noted that the patient did not required support post cp as the patient was doing well. | [5 empty] | The investigation of the reported failure mode has been completed. cardiac perforation/pericardial effusion: the cause of the injuries was not determined due to insufficient clinical details. positioning issues: upon advancing the pump, the patient's anatomy had sharp turn from subclavian into aortic arch. the product was not returned for investigation. the cause of the positioning issue was determined to related to the anatomy of the patient, as clinical mentioned the patient's anatomy had a sharp turn from the subclavian into the aortic arch. pump suction: the product was not returned for investigation. data log review confirmed suction alarms occurring in the case. prior to the first alarm, placement signal had a very slight tend downward. suction alarms suspended for over 12 hours once the patient received volume and albumin, and placement signal briefly increased, but suction started occurring again later on the second day. around that time, left ventricle (lv) perforation was reported, which likely contributed to volume issues, causing suction. placement signal was seen trending down again at this time. there were no spikes in motor current observed. the cause for the pump suction was most likely related to the patients condition, as clinical reported the patient had volume issues and an lv perforation which also likely contributed to decreased volume in the heart. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4974 | 1220648-2025-46979 | 23163872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163872&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710426 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented for reperfusion caths and impella cp for unloading. bilateral repo sheaths were placed with much difficulty accessing vessels but was eventually successful. frequent chugging and suction on extracorporeal membrane oxygenation (ecmo) device during process with multiple implantable cardioverter-defibrillator shocks as patient would go into ventricular fibrillation but was shocked out of it. it is of note during this procedure that air was suspected to get into ecmo circuit, likely not from repo sheaths but from central line access. upon trouble shooting the ecmo challenges, the physician proceeded to move forward with cp placement. left groin was dilated up to 14fx13cm peel away sheath. j wire and pigtail were used to try and gain access to left ventricle but could not pass over the arch. it was quickly realized that the pigtail catheter was in a false lumen of a large aortic dissection that was thought to extend from bifurcation of illiacs to about aortic arch and that inserting the impella, femoral would not be an option. dissection likely occurred during traumatic ecmo cannulation. it was later revealed there was no arch involvement. decision was made to place patient on left atrial veno-arterial extracorporeal membrane oxygenation despite local team¿s advocation for alternative access for cp insertion. cp insertion was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of delivery issue is determined to be patient condition. device history review: the device passed all the post sterile inspection checks.
4975 | 1220648-2025-46978 | 23163873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163873&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026685598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered access site bleeding. night shift nurse said groins were oozy all night and had to hold pressure. groins were stable the following morning. systemic heparin switched to heparin in purge. additionally, after transferring patient from ems gurney to icu bed, they began having suction events requiring reduction of p level to p4. echo showed impella inlet was 3.4 cm from aortic valve annulus with pigtail caught in mitral valve. physician attempted to rotate the impella without success. also, team started empiric antibiotics for presumed sepsis. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D3 added manufacturer fax number as it was omitted from the initial report that was submitted. e.1 added zip code extension as it was omitted from the initial report that was submitted. g4 added pma/510(k) as it was omitted from the initial report that was submitted. h.6 added code b13 as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned. major bleed: the patient's groin was oozing and was resolved with manual pressure. the cause of the bleed was not determined due to insufficient clinical details. sepsis: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: after transferring the patient to the intensive care unit bed, the impella inlet was 3.4 centimeters from the aortic valve annulus with pigtail caught in the mitral valve. the cause of the positioning issue was determined to be an unintentional use error as the pump was most likely pulled out of position by the user during the transfer of patient. device history review: the device passed all the post sterile inspection checks.
4976 | 1220648-2025-46976 | 23163874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163874&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026688199 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was being supported on intraaortic balloon pump (iabp) however, after ultrasound confirmed vessel size decision was made to place on impella 5.5. the iabp was successfully removed and 5.5 placed. the patient was supported for a total of 159.68 hours. it was noted that the 5.5 was explanted without on-site impella support. it was noted that explant went well and impella came out without difficulty. after they opened the cutdown pocket and removed the clotting material, she noted a gap between the graft and vessel at the distal area of anastomosis. she stated she added a few stitches and patient did well. she feels this gap was likely there the whole time and the clotting material kept it from bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: bleeding from a gap between the graft and vessel at the distal area of anastomosis. bleeding was controlled by placing few stitches. the cause of the bleeding is determined to be use issue because the complication was managed by placing stitches at the site. device history review: the device passed all the post sterile inspection checks.
4977 | 1220648-2025-46977 | 23163875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163875&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026681931 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a impella 5.5 procedure was aborted due to not being able to pass through vasculature. decision was made to place the impella cp. | [5 empty] | Correction g4. the investigation of the reported failure mode has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult patient anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4978 | 1220648-2025-46980 | 23163876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163876&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026665929 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision was made to place the impella 5.5 as a bridge to ventricle assisted device. however, it was noted that the impella 5.5 pump was primed and placed on back table. the physician made 1 attempt to advance wire but wire would not advance, portable c-arm brought in and 035 wire advanced to aortic root, multiple wire/catheter combinations attempted by cardiologist with no success, unable to cross bio-prosthetic atrioventricular. following additional attempts decision made to abort insertion and patient to remain on vasoactive support, graft ligated via surgical clips and surgical wound closed. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction g4. the investigation of the reported failure mode has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4979 | 1220648-2025-46981 | 23163877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163877&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713851 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while the nurse was getting the patient back to bed using a hoyer lift, the impella 5.5 cannula got caught underneath the bed and the wire from the red port got pulled out causing the impella to stop abruptly ¿impella stoppage¿, pump was pulled across the valve and explanted last night and an emergent femoral intra-aortic balloon pump was placed. the next morning after rounds, decision was made to re-insert an impella 5.5. there was no harm reported and the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The initial mdr 1220648-2025-46981 was submitted under an incorrect device (impella 5.5 with serial number (b)(6)) due to an administrative error. upon internal review, it was determined that the reported event involved a different device. this mdr is being replaced by 1220648-2025-46527, which has been submitted under the correct device (impella 5.5 with serial number (b)(6)) to ensure accurate reporting.
4980 | 1220648-2025-46982 | 23163878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163878&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026719258 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella was unable to pass through axillary artery. impella was removed and graft used for peripheral ecmo cannulation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction g4. the investigation of the reported failure mode has been completed. difficult to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
4981 | 1220648-2025-46983 | 23163879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163879&pc=OZD | 08/29/2025 | 02/13/2026 |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025487691 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported additional information upon request: "heparin was stopped temporarily and the oozing stopped. physician said this was normal oozing and not related to the graft and due to the nature of his cutdown and pocket. pump is not available for return. no further information is available." | The complainant reported that during support on impella 5.5, the patient had oozing at the incision sight with pressure dressing applied. the physician stopped the heparin. there were concerns of preload with cap of 9. the physician added fluids and another unit of blood. subsequently, it was noted that the patient's right lower leg was still ischemic and required amputation below the knee. the patient status was noted as too unstable for the operating room. vascular surgery was declining to operate at the time. hemodynamics has continued to decline slowly throughout the day requiring increased pressor support. the patient expired on support after do not resuscitate (dnr) status. |  |  |  |  | B5 updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d4, g4. the investigation of the reported failure mode has been completed. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4982 | 1220648-2025-46984 | 23163880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163880&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026699595 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella 5.5 the nurse observed that the patient was experiencing access site bleeding and requested advice related to the care thereof. the patient received one unit of blood but was not started on systemic anticoagulation. labs were pending. the nurse was informed that we could not advise on that matter being that the implanting physician chose to leave the peel-away sheath in place against best practice advice of the primary support on-site. there was support offered to support removing the sheath and inserting the repositioning sheath if the care team chose that in the plan of care; however, additional information noted that the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be use issue because the peel away sheath was in place against best practice advice of the primary abiomed support on-site and oozing occurred after transfer to icu. device history review: the device passed all the post sterile inspection checks.
4983 | 1220648-2025-46985 | 23163881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163881&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026714587 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support on impella cp, the patient's family planned for comfort measures only due to the patient's neuro status/stroke and patient wishes. care was withdrawn and the patient expired. | [5 empty] | Correction: g4. the investigation of the reported failure mode has been completed. no product was returned for investigation. stroke: the root cause of the stroke was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4984 | 1220648-2025-46986 | 23163975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23163975&pc=OZD | 08/29/2025 |  |  | 09/28/2025 |  |  | 08/29/2025 | 09/28/2025 |  |  |  |  | 09/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691731 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 63 KG | Non Hispanic | White | Injury | Yes | Yes | False Alarm | Malposition of Device | Device Sensing Problem |  |  |  |  | Atrial Fibrillation | Hematoma | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced a hematoma, atrial fibrillation and sensing issue with the device. the patient was admitted from the heart failure clinic with worsening symptoms. right heart catheterization was performed and the patient was scheduled from an impella 5.5 which was successfully implanted. the patient was transferred to the unit on impella mcs. once on the unit, a bedside echocardiogram was performed and showed the impella positioning shallow at 1.7. the impella 5.5 was advanced to approximately 4cm without difficulty. additionally, a hematoma was and consequently, heparin was suspended. the patient¿s hemoglobin dropped and packed red blood cells transfusion was unsuccessfully attempted as the patient had a transfusion reaction requiring epi pen injection due to hypotension and there was associated nausea. the plan was to go to the operating room in the afternoon for hematoma evacuation which appeared to have been completed. heparin remained off. urine output was adequate and amber in color. hydralazine had been given prior to the attempted blood transfusion. the patient remained status 2 for transplant, currently awaiting a donor. milrinone was reduced and patiently tolerated this well. impella mcs continued. seven days later device sensing issue was noted. false ¿impella position in aorta¿ alarm presented with no correlating event for the patient. the patient was doing fine. three days later, however, the patient experienced an arrhythmic event; patient was in atrial fibrillation and had to be off oral and intravenous anticoagulation since hematoma evacuation the prior week. impella mcs continued for an additional two days before explant with a survived outcome and bridge to transplant. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Complaint device coding was updated to better describe the reported event. therefore, h6 medical device problem code were updated/repopulated accordingly. investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the following was determined: hematoma/hypotension: the cause of the injury was not determined due to insufficient clinical details. paroxysmal atrial fibrillation: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. false alarm/optical signal issue: clinical reported false 'impella in aorta' alarms. position was confirmed via echo. the data logs confirm occurrence of 'impella position in aorta' alarm with no motor current spikes outside p-level changes. the cause of the issue was patient condition related as no log abnormalities were noted and position was confirmed via echo. device in wrong position: clinical noted pump was shallow. no information was provided on how pump became shallow. no patient abnormalities noted. the cause of the issue was not established. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4985 | 1220648-2025-46924 | 23162403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162403&pc=OZD | 08/28/2025 |  |  | 09/27/2025 |  |  | 08/28/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639698 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced hemolysis. over a two-day period, the patient had increased shortness of breath, activity intolerance, and lower extremity swelling and was readmitted for non-acute myocardial infarction cardiogenic shock (stage c). the impella cp was implanted for acute support with plans to escalate to an impella 5.5 device when feasible as bridge to recovery/consideration of left ventricular assist device. during impella mcs the following day, the patient¿s urine tea colored and concentrated. concerns for hemolysis. lactate dehydrogenase trended and elevated. transthoracic echocardiography (tte) was being discussed for placement verification. actions, if any, taken to manage the hemolysis was unnoted. however, later this same day the patient was successfully bridged to impella 5.5 as per the plan. the impella 5.5 mcs continued for approximately 12 days before being successfully bridged to the left ventricular assist device with a survived outcome. there was discussion contemplation at the time for right ventricular but not acted upon. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the clinical details state that "around 0600 urine became tea colored/concentrated. loose concerns for hemolysis. lactate dehydrogenase (ldh) trended and elevated. transthoracic echocardiogram (tte) potentially later today for placement verification. no alarms or issues with the impella at this time." the data logs do not show any motor current or placement signal spikes or trends. pump flows are stable and there are no positioning alarms. due to lack of product returned and limited clinical details, the root cause cannot be determined. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4986 | 1220648-2025-46925 | 23162405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162405&pc=OZD | 08/28/2025 |  |  | 09/27/2025 |  |  | 08/28/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707144 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and during support the pump stopped requiring immediate replacement. an impella cp device was initially placed due to increasing pressor support requirements and upward trending lactic after admission. subsequent discussions were made to upgrade to an impella 5.5. however prior, the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo) due to the continued increasing inotropic and pressors. during morning shift change the arterial va ecmo cannula was accidentally removed which resulted in a continuous left femoral artery bleed requiring emergently taking the patient to the or for cutdown and repair. the decision was made at this time to proceed with impella 5.5 upgrade since vascular surgery was assisting with left femoral artery repair and impella cp removal via cutdown. the impella 5.5 was successfully placed and the patient was sent to the unit on mcs. approximately two days later, the local team received a called emergently to the unit after it was reported patient stoop up from chair to be placed back in bed, a pop was felt and heard by the nurse at bedside and immediately following being back in bed the automated impella controller (aic) began to alarm ¿pump stopped.¿ the nurse taking care of patient reported having attempted restarting the pump multiple times without success. local team advised to exchange the aic but this did not resolve the issue. decision was made to explant impella at bedside. upon explant it was observed that thin clot was visible inside the impella cannula surrounding the impeller blades. the patient was stable after explant and after observing him for 30 minutes the decision was made to continue monitoring the patient for now. the patient was being worked up for a durable ventricular assist device. of note, there were no reports of any prior alarm on impella other than impella position unknown. the status of the patient following the event was stable. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical reported pump stop alarms after patient was moved and placed back in bed. upon explant, there was clot observed inside cannula and around impeller blades. the data logs from the case show sudden alarm #119 impella stopped alarms and show the failed restart attempts during which there were alarm #115 impella stopped (rpm deviations) alarms with motor current spikes to 3000ma indicating an electrical open issue. no product was returned for analysis. the cause of the issue was due to an electrical open issue with clinical information confirming a use-related incident during patient movement leading to pump stop. regarding the thrombosis, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
4987 | 1220648-2025-46926 | 23162414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162414&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/27/2025 |  |  | 08/29/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685593 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient admitted in with known cardiac history and new cute myocardial infarction complicated by cardiogenic shock. the pump was placed and a rp flex was added for bipella support. the pump was placed however there were signs of hemolysis and suction persistent throughout the 40hour support. the cp was then explanted but the rp flex remained on for the support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump suction: intermittent suction alarms were reported throughout the case. volume was given throughout the case. the product was not returned for investigation. data log review confirmed suction alarms occurring throughout the case. hemolysis/mechanical interaction w/ blood: data log review showed suction alarms occurring throughout the case, as well as 'impella position unknown' alarms, which are triggered by low placement signal readings. the cause of hemolysis was most likely an adverse event related to the patients condition, as the placement signal is seen trending down throughout the case, and clinical details indicate the patient was experiencing volume issues. device history review: the device passed all the post sterile inspection checks.
4988 | 1220648-2025-46927 | 23162416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162416&pc=OZD | 08/28/2025 |  |  | 09/27/2025 |  |  | 08/28/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. - 1220648 |  |  | US |  | 2026670859 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support there was a massive leakage from the purge sidearm. the pump was replaced. there was no patient injury. the procedure outcome was reported as calm and stable with great results. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4989 | 1220648-2025-46928 | 23162417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162417&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/27/2025 |  |  | 08/29/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728646 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 66 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered cardiac arrest and ventricular fibrillation during impella cp support. after the impella placement, the patient started coding and a total of nine epinephrine pushes were given along with levosimendan, vasopressors, and neo drips being started. the patient was going in and out of ventricular fibrillation and was being shocked out of it, but the patient never regained a pulse with the conversion. after 45 minutes of coding and unable to achieve return of spontaneous circulation, the care team decided to call it. care was stopped and the impella was turned off. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4990 | 1220648-2025-46907 | 23162418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162418&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/27/2025 |  |  | 08/29/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654185 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient with complicated small aortic root and calcified valve have a 5.5 pump placed to support the patient and have surgery/coronary artery bypass graft for the coronary microvascular disease. the pump placement did occur despite anatomic limitations, but during the placement caused ventricular tachycardia and the patient was cardioverted x1. when in the operating room the patient suffered a the left ventricle perforation that was treated by chest tubes for the bleeding and surgical topicals and fibrillation. the patient did expire after the 3 days of support despite all efforts. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac perforation: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
4991 | 1220648-2025-46929 | 23162421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162421&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/27/2025 |  |  | 08/29/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634443 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 119 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had cp placed for the patient admitted in with acute myocardial infarction with cardiogenic shock the pump was in poor position and flows were low. the team pulled the pump to reposition and after explant observed the pump to have clot adhered to it. the team questions a coagulation disorder, and the pump was replaced. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: the cause of the positioning issue was unable to be determined due to insufficient clinical details. low pump flow: after reviewing logs, no suction alarms were present. impella in aorta alarms were observed. the cause of low pump flow cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
4992 | 1220648-2025-46931 | 23162422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162422&pc=OZD | 09/02/2025 |  |  | 09/27/2025 |  |  | 09/02/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679389 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 81 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 developed an unspecified infection with elevated white blood cell count. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 required intervention was selected incorrectly on the initial report that was submitted. a new selection was added. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. the root cause of the infection was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks.
4993 | 1220648-2025-46933 | 23162424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162424&pc=OZD | 09/04/2025 |  |  | 09/27/2025 |  |  | 09/04/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704589 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed heparin-induced thrombocytopenia (hit). heparin was withheld. additionally, the patient had sanguineous drainage from the right pleural chest tube. multiple transfusions were given and ct was done to assess for the source of the bleed. the patient continued on support until successfully weaned from the pump as planned. | [5 empty] | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4994 | 1220648-2025-46932 | 23162425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162425&pc=OZD | 09/03/2025 |  |  | 09/27/2025 |  |  | 09/03/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100964 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Operating System Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered automated impella controller (aic) issues. it was reported that the aic failed to shut down and made a loud buzzing noise for five minutes while stuck on a loading screen. no patient harm was reported. | [5 empty] | D.9 the device was returned and the date received was added. h.6 updated investigation codes due to the device being received. h.10 related report number was added. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | A4, a5, and a6 are unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Investigation summary: data analysis: console logs from the reported event date, aligned with the complaint, indicating an unexpected controller shutdown event alarm on (b)(6) 2025. the initial case was opened on (b)(6), and on (b)(6), a shutdown was requested, with the shutdown sequence initiated. however, the sequence was not completed, and upon reboot, the automated impella controller (aic) displayed the unexpected controller alarm. no alarms related to loud buzzing sounds were observed in the logs. device analysis: the console was tested through multiple power cycles, but the issue could not be reproduced. during testing, no shutdown-related problems or irregularities within the aic were observed. no power supply interruptions were detected between the power battery manager and the 4-in-1 assembly. the aic was operated with a known good pump and purge cassette, with no issues noted, and no noise or loud buzzing sounds were observed. additionally, the speakers were also inspected, and no issues were found. root cause: the unexpected controller shutdown alarm was caused by prior inability to shut down due to incompletely executed shutdown process, after console shutdown was initiated. issue could not be reproduced during testing, and root cause was unable to be determined. product history review summary: one complaint related to the reported complaint was discovered while reviewing the product history of the console, however the issue was mostly use related.
4995 | 1220648-2025-46934 | 23162434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162434&pc=OZD | 08/13/2025 | 10/31/2025 |  | 09/27/2025 |  |  | 08/13/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885195 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 102 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while supporting a patient implanted with an impella 5.5 the automated impella controller (aic) generated a battery low alarm. the aic was plugged into multiple outlets until the alarm resolved. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report was submitted in error using the wrong device serial number and incorrect dates. a new initial report has been opened with mdr reference # 1220648-2025-46934. this report should be disregarded.
4996 | 1220648-2025-46935 | 23162437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162437&pc=OZD | 08/29/2025 |  |  | 09/27/2025 |  |  | 08/28/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026694637 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hematuria | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient presented to the emergency department in cardiac arrest with initial rhythm of ventricular fibrilation. cardiopulminary resusication (cpr) was initiated. the patient was shocked twice with return of spontaneous circulation (rosc). the patient was intubated. initial ecocardiogram revealed atrial fibrillation and the patient subsequently arrested again with ventricular fibrillation. cpr was resumed. after numerous rounds of epi, rosc obtained. the complainant reported that the patient was implanted with an impella device for mechanical circulatory support. during support the patient experienced hematuria, lost pulses and gastrointestinal (gi) bleeding. hematuria was noted in foley catheter. the team was encourage to evaluate position with echo imaging and reposition as needed. additionally, the patient was having weak and faint distal pulses, and decision was made to place distal catheter perfusion for impella insertion site. lastly, lactate was rising due to a gi bleed, possibly ischemic bowel, and the gi team was consulted. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction d4, g4. the investigation of the reported failure mode has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
4997 | 1220648-2025-46936 | 23162439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23162439&pc=OZD | 08/05/2025 |  |  | 09/27/2025 |  |  | 08/05/2025 | 09/27/2025 |  |  |  |  | 09/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026638057 |  | 000000000001000660 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male | 77 KG | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was placed in a patient with a nonischemic st elevated myocardial infarction and ventricular tachycardia with left bundle branch block. there were persistent suction events that prompted the team to treat for thrombosis. the team added tissue plasminogen activator to the purge but when flows did not improve the pump was explanted. upon explant, a thrombus was seen at the cannula and inlet. | [5 empty] | No product or data logs were returned for evaluation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the cause of the low pump flow could not be definitively determined as no product or logs were returned for evaluation. the pump passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
4998 | 1220648-2025-31688 | 23153152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23153152&pc=OZD | 06/19/2024 | 09/26/2025 |  | 09/26/2025 |  |  | 06/19/2024 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114582 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient. during support, placement signal unreliable were observed on the associated automated impella controller. motor current is normal, and echo performed, and impella is in appropriate position. there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. console logs from the reported day of event are consistent with complaint and show multiple placement-related events. data logs indicate periodically low snr values correlated with placement signal pulsatility. inspection and testing of the console revealed defect of the optical bench which resulted in periodically low snr values. the root cause was a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
4999 | 1220648-2025-31689 | 23153205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23153205&pc=OZD | 02/28/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 02/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025602494 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal unreliable were observed. despite this issue, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was returned for investigation. no physical damage was noted on the returned product. data logs were not provided for this case run, and they were also not retrievable from the remote server due to missing console information. the cause of the optical signal issue was not determined since issue was not reproduce on returned product and data log was not provided for this case run. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5000 | 1220648-2025-31690 | 23153444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23153444&pc=OZD | 06/27/2024 | 09/26/2025 |  | 09/26/2025 |  |  | 06/27/2024 | 09/26/2025 |  |  | 02/28/2026 |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502877 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient implanted with an impella cp device for mechanical circulatory support. during impella support, an error mode occurred shortly after placement in the pump. the device exhibited an optical sensor error. the patient was in poor condition. the pump was not replaced due to the patient¿s condition. the patient ultimately expired during support due to a subdural hematoma. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. a "placement signal not reliable alarm" occurred shortly after pump placement. the patient was in poor condition, so the pump was not removed. the pump positioning was checked via echocardiography. the patient expired after 80 hours of support, most likely due to a subdural hematoma. upon review of data logs, it is apparent that the console is the potential cause of the issue. a localized defect of the charge-coupled device was found. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5001 | 1220648-2025-31691 | 23153492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23153492&pc=OZD | 03/06/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 03/06/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 120605 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 115 KG |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump escalated to the 5.5 pump for a patient. the 5.5 was supporting for over two weeks when the pump's optical signal was lost. the investigating engineer determined the issue likely stemmed from the automated impella controller (aic). care was withdrawn and the patient died. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. analysis: with 5.5 pump, the automated impella controller (aic) had intermittent placement signal alarms in the logs that became persistent on 3/5 and continued through the end of the case. rt and lt logs show that after 3/5, the contrast oscillated between +/-3276.8 (should be 0-100% range) as placement signal went to 3002 mmhg triggering the alarms and optical sensor dropped to -3276.8mmhg. the pump remained on this console despite the loss of placement signal and was not transferred. after this case, there was a bad lamp on the next boot up which was investigated under another complaint. a malfunctioning lamp is a potential cause of oscillating contrast linking these two investigations. when the console was returned for investigation, the lamp was not emitting light as 0.47v were measured which is the amount of voltage across the charged-coupled device array when the lamp is outputting no light. replacing the optical bench with a known-good resolved the issue. root cause: the loss of placement signal was due to a defective lamp on the optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5002 | 1220648-2025-46892 | 23153633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23153633&pc=OZD | 07/30/2025 |  |  | 09/26/2025 |  |  | 07/30/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694634 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered delivery issues with an impella cp. the patient was brought to the catheterization lab and an angiogram was obtained. a 14 french peel away sheath was placed after sequential dilation of the vessel. the left ventricle was accessed with a 5f al1 and table j, then o18 placement wire was exchanged for the j wire. the impella was purged and zeroed and then inserted into sheath. at the level of bifurcation of the aorta and illiac arteries, the impella would not cross. three attempts were made to the pta area with 7 millimeter (mm) and 8 mm balloons, but the impella would not cross the lesion. suggestion of shockwave were made but the doctor did not want to proceed with placement of the impella. as a result, an intra-aortic balloon pump was placed for the procedure. | [5 empty] | The investigation was completed. there was no product returned. the root cause of delivery issue was determined to be patient condition because the impella would not cross the lesion despite multiple attempts and could not cross the level of bifurcation of the aorta and illiac arteries. a device history review was completed and the pump passed all post sterile inspection checks. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
5003 | 1220648-2025-46893 | 23154202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23154202&pc=OZD | 07/18/2025 |  |  | 09/26/2025 |  |  | 07/18/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694608 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | Priming Problem |  |  | Anemia | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Renal Failure | Sepsis | Respiratory Failure | Multiple Organ Dysfunction Syndrome | Mitral Valve Insufficiency/ Regurgitation |  |  |  | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was transferred from an outside hospital on veno-arterial extracorporeal membrane oxygenation (va-ecmo) and intra-aortic balloon pump (iabp). the patient was urgently escalated from the iabp to an impella 5.5 due to decompensation. an echocardiogram was performed yesterday evening, showing the impella at 5.14 centimeters (cm) below the aortic valve annulus, directed on/near the mitral valve leaflets. the pump continued to have good waveforms and flows at p6, with lactate dehydrogenase (ldh) down trending from 5000 yesterday to 3000 today. the physician elected to leave the pump in position while weaning va-ecmo, believing the impella will naturally go towards the mid left ventricle (lv). creatinine was rising and was planning to start continuous renal replacement therapy (crrt). the following day, the patient was started on crrt. ldh continued to downtrend. two days later, the patient continued on crrt with ldh down trending and stable at 409. the nurse later called in for suction and discussed reasons for suction. central venous pressure (cvp) was 12 and they are getting a formal echocardiogram. overnight suction was reported when attempting to go down to p4. the pump was showing towards the mitral valve with anterior leaflet abutting cannula. the next day, the nurse stated they were getting a cardiac output (co) of 13.1 and the impella would not allow her to enter that co. continual renal replacement was working to be discontinued. the following day the patient required emergent care and was and reintubation overnight due to hypoxia and aspiration. the patient received two bronchoscopies to clear pulmonary secretions. the aspiration event continued the next day and the patient was reintubated. daily ventilator trials were continued and ldh was down trending to the 460s. low cvp and negatively trending diastolic pressure of the left ventricle was noted the following day. no alarms were noted but discussed volume with the nurse. the next day, the overnight team increased the p-level to p7 and pressures increased and there was good flow on the impella. three days later, the patient was worked up for a left ventricular assistance device (lvad) as they were not a candidate for a heart transplant given renal failure. also concerned for multi-organ dysfunction going into lvad. the patient continued on continuous veno-venous hemofiltration. the next day the patient remained on crrt. left ventricle diastolic trending continued negatively and discussed preloads needed. two days later, the patient remained on high flow oxygen and appeared short of breath. the patient continued on fluid removal with cvvh with no reported suction alarms. they continued to monitor fluid balance. arterial blood gas was within normal limits with anemia. the patient had lower hemoglobin and platelets and will receive one unit of packed red blood cells (prbcs). the next day the patient was reintubated for decreased saturation and shortness of breath, the patient was no longer being considered for a transplant so it is more likely that the patient may receive blood products as necessary. currently, the impella was running at p8 with high rate cvvh. the patient continued to have increased pulmonary hypertension requiring nitric support. ldh was trending up. the last echocardiogram showed the pump in stable position. a computed tomography scan showed pneumonia. the customer contacted the emergency call center regarding an air in purge alarm. the support staff walked the nurse through a purge cassette change, resolving the air in purge alarm. a mitral clip was completed and the patient continued on dialysis. a few days later, the patient was aspirated and required re-intubation. there was a small bump in lactate noted and a drop in blood pressure requiring a low dose of levophed. crrt was paused at this time for nephrology to begin looking at hemodialysis. the following day, the patient failed the fifth extubation attempt and will likely to receive a tracheostomy. two days later, cvp was soft and crrt fluid/hour was reduced to 50 from 100. the next day, a small hematoma was developing at the impella insertion site. heparin was paused. the next day the parient was returned from the operating room with tracheostomy and percutaneous endoscopic gastrostomy (peg) tube placed, with wash out of the axillary access site. heparin was still on hold the next day and the nurse had been transfusing the patient so hemoglobin was stable. an infectious work up was pending. overnight pressures were very low and possibly were attributed to elevated white blood cell count and form of sepsis. infectious work up was still pending. there were concerns with bleeding and heparin held. a computed tomography scan did not show signs of excessive bleeding and heparin was restarted. a few days later, hemodialysis was started for renal support. two units of blood were given for bleeding around the hemodialysis catheter site. the nurse called the customer support line with questions on heparin use with bicarbonate purge solution. the patient had oozing from the tracheostomy and dialysis sites. the following day, bleeding continued at the tracheostomy site and systemic heparin was paused due to bleed. the patient ended up receiving a breathing treatment from the respiratory therapist and continued to have bleeding with anticoagulants held. a vent with dialysis was running. days later, the patient was resting in bed with a tracheostomy collar and received subcutaneous (subq) heparin doses but remained off intravenous anticoagulation. impella weaning attempts have been unsuccessful and the patient remains on p7 support. the next day, hemodialysis was continued, and subq heparin was stopped due to excessive bleeding from the dialysis site and they were unable to wean the impella. the following day, bleeding at dialysis site continued and the patient was restarted on subq heparin. a slow attempt to wean the impella in the next few days was done. bleeding at dialysis site resolved the following day. the team was practicing caution and continued subq heparin. the next day, low hemoglobin was noted and slow bleed at dialysis site continued. they are holding systemic and subq heparin. they planned for rehabilitation with the impella 5.5 and probable placement of durable lvad once the patient was physically capable. anticoagulation was still held due to bleeding problems. the patient had bouts of ectopic runs yesterday due to depleted electrolytes. echocardiogram was repeated and transduce cvp to continue monitoring for good impella position. the patient was receiving hemodialysis as needed. the next day, ectopy was stopped since electrolyte replacement. days later, the bedside nurse was concerned about placement signal low alarm. arterial line waveforms look better than the impella waveforms. an echocardiogram was obtained and the impella was in the mid ventricular space at an appropriate depth. there was concern for the automated impella controller (aic) waveforms as pressure on the aortic placement signal was lower than cuff pressure. the patient was being treated off the blood pressure cuff pressure. neurology concerns were resolved and the plan was to infuse prbcs due to blood loss from traumatic foley insertion. the nurse called customer support due to intermittent placement signal low alarms. the patient had been very hypotensive. the impella connect was reviewed and likely a patient issue versus a pump movement issue. the nurse had requested an echocardiogram to confirm position and left ventricle function. the patient was weaned off the ventilator and was on a tracheostomy collar now. over the weekend, the patient reported trouble breathing but was feeling much better. heparin was stopped as the patient was bleeding again. subq heparin restarted after previous concerns for tracheostomy bleed the following day. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: correct impella catheter position ¿catheter angled toward the left ventricular apex away from the heart wall and not curled up or blocking the mitral valve¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The root cause of the arrhythmia and sepsis was unable to be determined due to insufficient clinical details. the cause of the renal failure, respiratory failure, multiple organ dysfunction, anemia, respiratory failure, hypotension, bleeding, and mitral valve incompetence was not determined due to insufficient clinical details. the cause of the suction issue was not determined as the product was not returned for investigation, and sufficient clinical information was not provided. the cause of the improper positioning issue was not determined due to insufficient clinical information. the cause of the access site adverse event issue was not determined as the product was not returned, and sufficient clinical information was not provided. no issues were noted with the optical signal parameters used to calculate the placement signal. the cause of the optical signal issue was most likely related to the patient condition, based on the data log review and the provided clinical details. no problem was found with the pump regarding the air in purge system issue. the air in the purge system was found to be related to the cassette. the device passed all post sterile inspection checks. | The root cause of the arrhythmia and sepsis was unable to be determined due to insufficient clinical details. the cause of the organ failure/anemia/hypotension/bleeding/mitral valve incompetence was not determined due to insufficient clinical details. the cause of the suction issue was not determined as the product was not returned for investigation, and sufficient clinical information was not provided. the cause of the improper positioning issue was not determined due to insufficient clinical information. the cause of the hematoma was not determined as the product was not returned, and sufficient clinical information was not provided. the cause of the optical signal issue was most likely related to the patient condition, based on the data log review and the provided clinical details. no problem was found with the pump regarding the air in purge system issue. the air in the purge system was found to be related to the cassette. the device passed all post sterile inspection checks.
5004 | 1220648-2025-46894 | 23154241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23154241&pc=OZD | 08/23/2025 |  |  | 09/26/2025 |  |  | 08/23/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693671 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 75 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced oozing at the access site and pump suction alarms. the pump's p-level was lowered and heparin was withheld due to a nosebleed. care was ultimately withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Medical safety review was completed with the following additional information: in this case, patient underlying condition likely contributed most to an outcome of death (severe liver failure). the impella is contraindicated for vsr. due to the vsr is why impella was implanted which is off label. however, as the impella 5.5 device was present at the time of the event and the patient experienced complications associated with this device, the impella 5.5 device cannot be dissociated from the demise outcome. however, ultimately the patient's care was withdrawn, which led to the patient expiring. note: events involving the 5.5 occurred prior to bipella support. the impella rp flex supported the patient as intended without reported or indicated issue and is unrelated to the demise outcome. | The investigation has been completed and the product was not returned for investigation. the data log review confirms suction alarms triggering throughout the case. epistaxis: the cause of the injury was not determined due to insufficient clinical details. pump suction: clinical details reported suction alarms triggering on the sixth day, and the decision was made to lower the p-level, but not to administer fluid. the product was not returned for investigation. data log review confirms suction alarms triggering throughout the case. starting 2 days before suction was reported, placement signal began gradually trending down from about 140 mmhg to less than 100 mmhg. the p-level was decreased from p-9 to p-7. clinical details also mentioned an impella rp flex was implanted on the eighth day of the case, and around this time in the data logs, placement signal began to rise gradually. suction alarms were triggering for the entirety of the case, most frequently on the sixth day, as mentioned, and on the thirteenth day ((b)(6) 2025). at this time, placement signal spikes briefly to 300 mmhg before dropping to about 80 mmhg, triggering suction and 'impella position unknown' alarms. the cause of the suction was most likely an adverse event related to the patients condition, as the patient required additional support from the impella rp flex insertion, and data logs show a downward trend in placement signal at times of frequent suction. the device history review was completed and the pump passed all post sterile inspection criteria.
5005 | 1220648-2025-46895 | 23154430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23154430&pc=OZD | 08/27/2025 |  |  | 09/26/2025 |  |  | 08/27/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612269 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Heart Failure/Congestive Heart Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with acute myocardial infarction with cardiogenic shock was brought to the hospital and percutaneous coronary intervention was done with impella cp support. in the intensive care unit, bedside echocardiogram was performed and revealed a collapsed left ventricle (lv) around the impella device and a very dilated right ventricle (rv). the physician suspected the patient had a pulmonary embolism (pe) and was given three doses of cath flow and initiated a heparin drip. the patient was placed on the ekosonic endovascular system for the pulmonary embolism. the physician also informed the staff that the impella needed to be removed from the lv as the physician was concerned that the impella could have been worsening the rv strain. the staff turned off the device and pulled it into the descending aorta and the impella was placed in surgical mode. the impella cp remained pulled back into the descending aorta due to blood thinners and placement monitoring was disabled. the next day, the impella was removed. additional information was received. the patient has been discharged from the hospital and was in rehabilitation. there was no issue with the placement signal on the device. it was appropriately functioning. the physician independently made the decision to withdraw the pump from the lv but did not want to proceed with the explant at that time so the device remained on at p-1 to avoid constant appropriate alarms from the automated impella controller. the placement signal was disabled by the care team itself. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. pulmonary embolism/ cardiac failure: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks.
5006 | 1220648-2025-46896 | 23154467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23154467&pc=OZD | 08/29/2025 |  |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720735 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 67 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia requiring defibrillation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation has been completed. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
5007 | 1220648-2025-46897 | 23154652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23154652&pc=OZD | 08/29/2025 |  |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720716 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 49 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the axillary site. the patient was transfused with two units of packed red blood cells. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation has been completed. no product was returned for investigation. major bleed: some oozing overnight at axillary site. the cause of the bleeding was determined to be patient condition because patient was having significant ostial left anterior descending artery disease. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5008 | 1220648-2025-46899 | 23154883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23154883&pc=OZD | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725381 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that during implantation of an impella cp there was a concern of a possible perforation. the pump was pulled back and no perforation was seen. a new impella cp was then implanted. | [5 empty] | Investigation summary: the impella cp was returned from evaluation. device in wrong position: no damage was noted on the returned pump. clinical details noted that the physician had a concern for a possible perforation after the impella was inserted. the physician pulled the pump back across the atrioventricular and imaging showed no perforation. the decision was made to explant the pump rather than attempt to re-insert. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by the physician. device history review: the pump passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5009 | 1220648-2025-46900 | 23155041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23155041&pc=OZD | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670859 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male | 106 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during patient support, the impella cp leaked from the purge sidearm. the pump was exchanged for a new one to continue patient support. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. d9 has been updated to reflect product was returned for investigation. investigation summary: the cause of the purge leak pump was a puncture hole found in the blue reservoir membrane due to reasonably foreseen misuse.
5010 | 1220648-2025-46898 | 23155052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23155052&pc=OZD | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713855 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Renal Failure | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | While on support, it was noted that automated impella controller alarms were addressed. the alarms appearing were for intermittent suction, which self-resolved, and the impella in ventricle. | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring shocking and administration of amiodarone. the patient was placed on continuous renal replacement therapy. |  |  |  |  | B.1 selection was added as it was inadvertently omitted from the initial report that was submitted. b.5 information was added that was inadvertently omitted from the initial report that was submitted. h.6 added medical device problem codes due to information added. the investigation has been completed. no product was returned for investigation. pump suction - at some point during support, intermittent suction alarms occurred that were self resolving. data log review showed suction events that resolved when the p-level was lowered, indicating insufficient patient volume for the requested pump performance. the cause of the pump suction was lack of patient volume due to suction alarms resolving with lowering the p-level. device in wrong position - at some point during support, impella in ventricle alarms occurred while the patient was in sustained ventricular tachycardia. an echocardiogram confirmed the pump positioning was adequate at case beginning at 39 centimeters. on day four of support, the pump catheter positioning was noted to be 38 centimeters. the change in centimeter marker was not discussed. data log review showed the impella position unknown alarms toward the case end on (b)(6)2025. the cause of the device in wrong position could not be determined due to no product returned inconclusive data log review, and insufficient clinical details. tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. renal failure: the cause of the renal failure was not determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5011 | 1220648-2025-46901 | 23155513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23155513&pc=OZD | 08/30/2025 |  |  | 09/26/2025 |  |  | 08/30/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026728360 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during explant of an impella cp, the pump was pulled back into descending aorta. vascular surgery closed the site and a left lower extremity fasciotomy was performed. later, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the pump. another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number. | B.3 revised as incorrect date was reported on the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the investigation has been completed. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device history review as completed and the passed all the post sterile inspection checks.
5012 | 1220648-2025-46902 | 23155514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23155514&pc=PYX | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815750 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an automated impella controller (aic) generated a controller failure alarm during patient support. a new aic was used to support the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller (aic) was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the aic. a separate report was submitted for the pump. refer to section h.10 for the related report number. | The root cause of the controller failure alarm was a pud communication issue.
5013 | 1220648-2025-46904 | 23155989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23155989&pc=OZD | 08/27/2025 |  |  | 09/26/2025 |  |  | 08/27/2025 | 09/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658204 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Injury | Yes | No | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp encountered low pump flow, optical signal issue, pump suction, thrombus and gastrointestinal bleeding during impella cp with smart assist support. it was reported that there was suddenly visible change in the automated impella controller (aic) metrics: dropping of motor curve values, low impella flow at p7 (1,3) and no left ventricle (lv) signal. after that continuous suction alarm was also at p1. the pump was removed. after removal, biomaterial was visible in the pump inlet. it was also reported by the physician that the patient had gastrointestinal bleeding was not impella related. however, there is not enough information at this moment to disassociate the gastrointestinal bleeding from device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report represents the pump. another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number. | Product complaint # (b)(4). the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. impella cp sn (b)(6) returned for investigation. low pump flow/pump suction/optical sensor issue: clinical details note that there was a suddenly visible change in the aic metrics with a drop of motor curve values, low impella flow at p7 (1.3), and no lv signal. data log review showed a decrease in minimum pump flow and minimum motor current as described in the clinical. at the same time, there was a slight increase in purge pressure. there were no motor current spikes around this time. the low pump flow appears to resolve in under ten minutes but is followed by multiple suction alarms, which is when the lv signal disappeared as expected during a suction alarm. there were no abnormalities with the optical parameters during the case but there was variation in ao ps throughout the case. the product was returned and evaluated. there were no abnormalities with the optical sensor. there was a small dried piece of biomaterial/blood in the inlet and no significant biomaterial in the outlet. there were no cannula kinks and the pump ran without issue in a bucket of water. low pump flow: the cause of the low pump flow was not determined since low flow did not reproduce on returned product and insufficient clinical details were provided. pump suction: the cause of the suction was not determined since there was no conclusive evidence on returned product and insufficient clinical details were provided. optical sensor issue: the cause of the optical sensor issue was not determined since there were no abnormalities on the returned product and insufficient clinical details were provided. thrombosis: clinical details note that there was biomaterial visible in the inlet after explant. a dried piece of blood/biomaterial was observed in the inlet of the returned pump. the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: clinical details note there was gastrointestinal bleeding noted that was not impella related during the case. the cause of the injury was not determined due to insufficient clinical details.
5014 | 1220648-2025-31692 | 23156226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23156226&pc=OZD | 06/04/2024 | 09/26/2025 |  | 09/26/2025 |  |  | 06/04/2024 | 09/26/2025 |  |  | 02/28/2026 |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509555 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient. during support, placement signal unreliable were observed. the placement alarm was disabled. decision made to withdraw care, patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. there was no product returned for investigation. in reviewing the data logs, there is clear evidence that the placement signal pressure readings start to trend downward toward the end of the case. since the placement signal low alarms were confirmed to be set off as they should, the optical sensor data was all within specification throughout the case, and the patient condition was clearly deteriorating, the root cause was determined to be patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5015 | 1220648-2025-46905 | 23156249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23156249&pc=OZD | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655054 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a cerebral hemorrhage. an emergency craniotomy was performed. impella support continues. | [5 empty] | The root cause of the stroke was unable to be determined due to insufficient clinical details. the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5016 | 1220648-2025-31693 | 23156365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23156365&pc=OZD | 10/19/2023 | 09/26/2025 |  | 09/26/2025 |  |  | 10/19/2023 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1368535 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient. during support, the associated automated impella controller (aic) showed placement signal not reliable alarms. pump position was confirmed. it was confirmed the placement signal issue is an aic issue, not a pump issue. patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the product was returned for investigation. in reviewing the data logs, there is clear evidence that the raw optical sensor drop. the cause of the automated impella controller (aic) issue was most likely a defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. this report is being filed as part of a retrospective review of historical records.
5017 | 1220648-2025-46906 | 23156391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23156391&pc=OZD | 08/29/2025 |  |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026712179 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a cardiac patient was admitted in for a balloon aortic valvuloplasty (bav) and angioplasty of the multivessel. there was delivery issues with the impella cp as it was unable to cross the valve. the diameter was not appropriate and the pump was removed and another bav was performed so that the pump could be delivered. no harm from the delivery failure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. | The investigation has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the patient had aortic stenosis and bicuspid aortic valve repeated to deliver impella. the pump was unable to pass through the vasculature and the insertion was aborted. the device history review was completed and the device passed all post sterile inspection checks.
5018 | 1220648-2025-31694 | 23156423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23156423&pc=PYX | 07/25/2024 | 09/26/2025 |  | 09/26/2025 |  |  | 07/25/2024 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025503229 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp flex pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal unreliable were observed. decision was made not to change pump out as motor current had normal flows. decision made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. the pump was returned for investigation. pump passed all electrical testing. the impeller was clean of biomaterial and the dp sensor had no external damages. data logs show dp sensor drift with an upward trend in the placement signal with no active psnr alarms. no issues were noted with the optical signal parameters. pump was sent for 24 hour flow testing and sensor drift pattern was observed in the flow test data as well confirming the sensor drift issue. the root cause of the placement signal issue was a sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. this report is being filed as part of a retrospective review of historical records.
5019 | 1220648-2025-31695 | 23156490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23156490&pc=PYX | 02/07/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 02/07/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025567877 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp flex pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal unreliable were observed. decision was made not to change pump out as motor current had normal flows. decision made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. no product was returned for analysis. the data logs show that placement signal was drifting down and pump flow was drifting up before placement signal went out of range and flow calculation was disabled. the placement signal did not recover for the remaining duration of the case. the root cause of the placement signal issue was determined to be due to sensor drift as evidenced by drift observed in the logs. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5020 | 1220648-2025-46903 | 23157791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23157791&pc=OZD | 08/27/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167085 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that optical signal issues were encountered on an automated impella controller (aic) while supporting a patient on impella cp support. the pump was explanted shortly afterwards. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller (aic) was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this report represents the aic. another report was submitted to represent the pump. refer to section h.10 for the related report number.
5021 | 1220648-2025-46910 | 23158326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23158326&pc=OZD | 08/27/2025 | 10/13/2025 |  | 09/26/2025 |  |  | 08/27/2025 | 09/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026722787 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Fever | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an impella cp was placed on a patient for hemodynamic support. during support, it was noted that the patient had a fever and was being tested for covid. the fever continued into the next day where it was noted that the renal team was being consulted. no further information was provided about the fever. additionally, it was noted that when the patient sat up in bed the pump metrics jumped. motor current jumped and left ventricle placement signal was above the aorta placement signal and completely separated. pump flows read 4.0/4.0. it was determined there was pump saturation and the impella cp needed to be replaced. the impella cp was removed and a new impella cp was placed. the patient remained in stable condition. | [5 empty] | Investigation summary: the investigation has been completed. the impella cp was returned for evaluation. fever: clinical details noted the patient had a fever while on support. in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. saturated flows: the returned product showed bearing wear indicated with a large gap present at purge gap. data logs were reviewed and long term logs showed a motor current spike with speed deviation and immediate pump flow saturation. one minute later, another motor current spike and then a "impella stopped - motor current high" alarm. the pump was able to be immediately restart and flows were fully saturated for approximately one hour before p-level was turned down from p-7 to p-5 and flow slowly became less saturated. clinical details noted that when the patient sat up in bed, the pump metrics jumped. motor current was increased to 11000. left ventricle placement signal above aorta signal and completely separated. the pump flows also were reading 4.0/4.0. the cause of the saturated flows was due to bearing wear based on data logs and returned product analysis. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
5022 | 1220648-2025-46909 | 23158346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23158346&pc=OZD | 08/27/2025 |  |  | 09/26/2025 |  |  | 08/27/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674925 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that a patient was admitted to the hospital with viral cardiomyopathy. an echocardiogram revealed the patient's ejection fraction to be 5-10%. the patient was then brought to the cardiac catheterization laboratory for right heart catheterization (rhc) and left heart catheterization (lhc). rhc and lhc were completed with no issues, and no disease was noted on coronary angiography. of note, during rhc, it was observed that the patient¿s femoral vessels were small. however, the physician felt the impella could be inserted. the physician inserted the 14 french peel away sheath without issue, and the pigtail was used with a 0.035 wire to cross the aortic valve. the wire was exchanged, catheter was removed, and the impella was backloaded on to the device. the sheath was flushed and the device was inserted through the sheath with no problem. the patient was noted to have high and small aortic arch. while the wire was in the ventricle, the impella kinked and folded on itself. the wire came back across the valve and was no longer of use, as the impella was stuck in the pigtail of the device and wrapped around itself. the physician was able to manipulate the device to free the pigtail, but it was still kinked backwards. after multiple tries, the physician was able to straighten out the impella. the femoral vessel was then imaged with the impella sheath in, and it was noted there was no distal flow beyond the sheath. due to the patient¿s anatomy and ischemia observed with the sheath in, the physician decided to abort the implant of the impella cp and not attempt with a new pump. the impella sheath was removed and the two perclose devices were tightened to close the impella access site with no issues. the patient sent to the intensive care unit for further management. | [5 empty] | Product complaint # (b)(4). the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. h6 codes have been updated to reflect product was returned, and investigation is underway. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5023 | 1220648-2025-46911 | 23158468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23158468&pc=OZD | 08/29/2025 |  |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678058 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient admitted in with plans to perform a coronary artery bypass graft. the pump was explanted after just 28 hours due to limb ischemia. the patient remained stable and survived. additional information was received. the patient had a history of peripheral artery disease prior to starting the procedure. the staff needed to use a doppler in order to find distal pulses. no imaging is available. however, the hospital did perform an arterial us of the lower extremities and the doctor had said there was some flow distally so it ws not completely absent. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of the reported failure mode has been completed. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5024 | 1220648-2025-46912 | 23159195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159195&pc=OZD | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585026 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during a percutaneous cardiac intervention to the left anterior descending coronary artery (lad). it was reported that the patient was admitted to the hospital hypotensive and was found to have culprit lesion in the proximal lad. after a successful procedure, the patient was admitted to the intensive care unit (icu) for recovery. after the patient was transferred to the icu, it was noted there was oozing at the impella access site. the decision was made to remove the impella cp to achieve hemostasis. upon removal, the patient was hypotensive and required two units of blood cells to be transfused. no further information was available on patient outcome. | [5 empty] | A.4 is unknown. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. major bleeding: bleeding from the access site after transfer to icu and attempted to angle match. manual pressure but the hemostasis was not achieved. the cause of the bleeding is determined to be use issue because the oozing occurred after the patient movement to icu and hemostasis was achieved by removal of impella. hypotension: the cause of the clinical symptom was not determined due to insufficient clinical details. device lot: 1865629. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
5025 | 1220648-2025-46914 | 23159199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159199&pc=OZD | 06/24/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574854 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured recurring ventricular tachycardia with a "possible" relationship to the impella device used: ¿the patient developed recurring episodes of ventricular tachycardia. on (b)(6) 2025, the patient developed sustained ventricular tachycardia. the patient underwent direct-current cardioversion eight times. the patient lost pulsatility on three of these episodes and lost flows on the impella once. impella placement was confirmed via point of care ultrasound. throughout the admission, the patient had several more episodes of non-sustained, often self-limiting episodes of ventricular tachycardia. native rhythm was sinus.¿ it was further reported that the patient¿s family decided to withdraw care and the patient expired. | [5 empty] | A4: patient weight is unknown. based on the information available, the impella 5.5 cannot be excluded as a possible contributing factor to the patient¿s outcome. product status: the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure investigation summary: no product was returned. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: no product was returned. no clinical about low flow provided. pump was in optimal position. after reviewing logs, multiple suction alarms along with impella position unknown and in aorta alarms were observed. the cause of low pump cannot be determined since insufficient clinical details were provided. device history lot: device lot: 1847637 device history batch: subcomponent lot: n/a device history review: device sn:(b)(6) passed all post sterile inspection checks.
5026 | 1220648-2025-46913 | 23159200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159200&pc=OZD | 08/28/2025 | 10/06/2025 |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095116 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the weaning process of the impella cp there was a placement signal that appeared to have some type of electrical interference with intermittent placement displays prior to dropping signal completely at level p-2. the console displayed controller error with switch to backup controller. however, the impella was removed successfully. there was no harm to the patient reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5027 | 1220648-2025-46915 | 23159205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159205&pc=OZD | 08/25/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713386 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | Pain | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient was admitted to the hospital a week prior to impella support for initiation of milrinone. the day before impella support was initiated, the patient acutely decompensated and required cannulation for extracorporeal membrane oxygenation (ecmo) and mechanical ventilation. it was suspected that the patient was in septic shock from an infected pacemaker site infection. a transesophageal echocardiogram was performed, and it revealed a severely dilated left ventricle, spontaneous echocardiographic contrast in the left ventricle, aortic valve not opening, and severe tricuspid regurgitation and aortic insufficiency. the patient was then emergently placed on an impella 5.5 the impella 5.5 was placed, and support was initiated without complications, and the pump power level was increased to p6 with lower flows and spontaneous echocardiographic contrast in the left ventricle noted. the following day, the patient was extubated and decannulated from ecmo. the patient was noted as stable. three days following, the patient was positive for heparin induced thrombocytopenia. heparin was removed from the purge solution, and the patient was switched to agratraban. three days later, it was reported that there was a viable purge leak at the distal end of purge filter. a plan was made to monitor purge flow and pressure for any changes. two days later, there was a report there was no active purge fluid leaking from the purge filter. it was noted there were a few bubbles visible around the filter, but no further leaking. purge pressures were noted to be normal, and the purge flow was noted to be stable. the patient reported pain at the impella insertion site, and a plan was made to do a bedside washout of the impella site. it was noted that the previous day the jp drain at the site was removed. heparin was withheld for preparation of the washout. the device was subsequently weaned and explanted successfully during surgery for a durable left ventricle assist device implant. | [5 empty] | A.4, a.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B.1 added selection as it was inadvertently omitted from the initial report that was submitted. h.6 code a141101 was reported incorrectly on the initial report that was submitted. the code has been changed to a141101. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h.6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product returned. purge leak- after four days on support, visible purge leak occurred at the distal end of the purge filter. two days later, the registered nurse stated the leak was no longer visible. the cause of the purge leak was a visible leak found on the distal end of the purge filter; however, the cause of the leaking could not be determined due to no product returned and insufficient clinical details. patient symptoms - pain + thrombocytopenia: the causes of the patients access site pain and thrombocytopenia were not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks.
5028 | 1220648-2025-46916 | 23159208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159208&pc=OZD | 08/28/2025 |  |  | 09/26/2025 |  |  | 08/28/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support impella connect alarms occurred. | The complainant reported the patient presented from an outside after multiple shocks on the implantable cardioverter defibrillator. the patient was placed on an intra-aortic balloon pump (iabp) for support and was being worked up for heart transplant candidacy. while on unit the patient¿s right and left ventricles function continued to diminish. the physicians made the decision to remove the labp and place protek duo and impella 5.5 for more stability during heart transplant waiting. the patient was awaiting heart recovery. approximately seven days after start of the impella mechanical circulatory support, the impella connect was coming in and out and notifying local team that a new pump is turned on and then connecting. this created false alarms for the abiomed team particularly overnight and when not on site. the is a non-patient effect issue. support continued uninterrupted and uncompromised. continuing support, the most recent ejection fraction showed 13%. the patient had been placed on extracorporeal membrane oxygenation and noted to have remained on high ecmo and impella 5.5 support (performance level p-9) for recovery. additional information updates were received and noted the patient experienced a gastrointestinal bleed. one unit of red blood cells and one unit of platelets were administered. support continued and a plan for recovery was still noted. however, the family discussed and eventually decided to go comfort care. they ultimately decided to remove the endotracheal tube for terminal withdrawal of care. mechanical circulatory support devices were left on. the patient expired within 20 minutes of extubation. |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received regarding the impella pump explant date/end of support after approximately 20 days and the survival at explant was noted as "expired." section b5 was updated to provide the complete event details and additional information received. h6: health effect: clinical/impact and device problem/component coding remain unchanged. medical safety review was completed and noted: the event describes a false alarm/buzzer failure that had no (direct or indirect) impact on the patient's support. the automated impella controller remained in use supporting the patient. there was no interruption or compromised to the impella mcs. in this case the patient expired for reasons (underlying condition, termination of care) unrelated to the impella aic non-patient impacting event. the event will be reported as a malfunction type of reportable event. associate director of post market clinical review noted the following: agree this occurrence does not contribute to patient outcome. device status: additional information also noted the automated impella controller was returned, received, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. correspondingly, section d9: device received by manufacturer with receipt date and h6: investigation type, findings and conclusion codes were updated accordingly. related medical device reports: this automated impella controller report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
5029 | 1220648-2025-46917 | 23159209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159209&pc=OZD | 08/29/2025 |  |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694267 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump sto support a patient. the cp was placed and was supporting when there was limb ischemia. the leg was found to cold and have lacking distal pulses. the cp was explanted and replaced by a impella 5.5. the patient had many suction / volume issues that was resolved with albumin and saline. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: after reviewing logs, multiple suction alarms were observed. placement signal (ps) is not pulsatile and is trending down towards end of support. the cause of pump suction was most likely patient condition related since the suction issue was resolved by giving the patient albumin and saline. device history review: the device passed all the post sterile inspection checks.
5030 | 1220648-2025-31696 | 23159342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159342&pc=PYX | 04/12/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 04/12/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548157 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an implanted impella rp flex pump rose dramatically during patient support. it was noted that the patient had a dual function swan for pacing. the motor current remained stable and support continued uninterrupted. the pump was eventually weaned and explanted, and the patient survived the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. placement signal issue: impella rp flex returned for evaluation. upon visual inspection, no abnormalities noted on returned pump and no damage to dp sensor observed. pump was run in pressure flow loop and placement signal drifted up and out of range with pump flows drifting to 0 with a placement signal not reliable (psnr) "dp sensor out of range" alarm triggered. motor current remained stable during flow test. data logs were reviewed and lt logs show placement signal drifting up during case with a simultaneous drift down of pump flow. motor current remains stable during drift and central venous pressure (cvp) values remained stable during support. clinical details noted that placement signal show a large rise in the field and pump flows were greatly reduced but motor current had remained stable throughout support indicating pump was still supporting patient. the root cause of the placement signal issue was due to dp sensor drift based on returned product testing and data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5031 | 1220648-2025-46920 | 23159378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159378&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026716114 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 116 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered optical signal issues and the patient had atrial fibrillation (afib) during impella 5.5 support. it was reported that from time of insertion of pump into the body, both aortic and ventricular placement signals¿ systolic values were off from peripheral arterial line measurements and that impella values were lower by about 20-25. left ventricle (lv) signal was adjusted and no change was noted. prior to insertion, aortic placement signal values were between 0 and -1. flows were good and there were no other impella related concerns. additionally, the patient encountered afib overnight and was given amiodarone as a result. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. arrythmia the cause of the injury was not determined due to insufficient clinical details. optical signal issue the data logs show that after starting the pump the placement signal (ps) was stable around 60mmhg but there were spikes in it with small dips in flow, and suction alarms. impella position unknown alarms were also triggered, and the lvp can be seen lower than the ps values and diastolic suction is occurring. the raw optical sensor was steady along with all 5s parameters. there is the potential that the pump is slightly out of position, but since the product was not returned and there is insufficient clinical information provided, the true cause of the optical signal issue cannot be determined. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5032 | 1220648-2025-46921 | 23159379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159379&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026729557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male | 85 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Unspecified Infection | Tachycardia | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp pump placed to support the patient admitted in for high-risk percutaneous coronary intervention. while on support, the patient had long runs of ventricular tachycardia, which required intubation. it was noted the patient was having more ekg ectopy. there was an instance of non-sustained ventricular tachycardia where the patient required shocks. the patient had a fever and the cardiology team assumed there was an infectious process going on. care was withdrawn and the patient expired. | Additional information was provided by an abiomed representative noted that the infection was confirmed via urosepsis and treated with broad spectrum antibiotics. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 was updated with additional information provided. h6 health effect impact code updated per additional information. the investigation was completed and no product was returned. the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the injury was not determined due to insufficient clinical details. the device history review was completed and passed all the post sterile inspection checks.
5033 | 1220648-2025-46922 | 23159383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159383&pc=OZD | 08/29/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689529 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Asystole | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered limb ischemia and asystole during impella cp support. during the placement of impella cp, an access-port was utilized to assess distal flow and noted to have no flow, thus, a distal right-ulnar line was placed by the doctor and a reperfusion line was utilized via a left brachial art-line to mitigate distal limb ischemia. it was noted that this vascular surgery resolved limb ischemia. the patient was put on comfort care and went asystole. the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. asystole: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5034 | 1220648-2025-46923 | 23159385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159385&pc=OZD | 08/26/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689368 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that during support of impella 5.5 on (b)(6) 2025, the patient was stable and doing well. started with antibiotics. chest x-ray showed large hematoma. will continue to support over the next few days. on (b)(6) 2025 the plan was to go to the operating room (or) tomorrow to wash out and remove impella. spoke with physician assistant and critical care doc about lower flows from impella due to high after load. the patient was taken to the or for weaning and removal of impella with the physician. left ventricle showed recovery and no issues arose in the removal. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the product was not returned for investigation. hematoma: the cause of the injury was not determined due to insufficient clinical details. low pump flow: clinical reported the pump was getting lower flows than spec, and claimed it was due to high after load. the product was not returned for investigation, and no event logs were provided. the cause of the low pump flow could not be determined due to lack of product or event logs returned. device history review: the device passed all the post sterile inspection checks.
5035 | 1220648-2025-31697 | 23159420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159420&pc=OZD | 02/08/2025 | 10/27/2025 |  | 09/26/2025 |  |  | 02/08/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025574344 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing when after over three weeks there was a loss of optical signal. the placement signal loss did not prompt pump explant. the pump remained on for support despite the malfunction. the patient was successfully weaned from the pump and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product or data logs were returned for evaluation. clinical details noted that patient was experiencing a few instances of suction alarms that clinical team attributed to decreased volume in days leading up to placement signal not reliable (psnr) alarm. it was not noted if psnr alarm resolved for remainder of case. the root cause of the optical signal issue cannot be definitively determined as no product or logs were returned and limited clinical details were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | B1 adverse event was inadvertently reported as ¿yes.¿ b2 outcomes attributed to adverse event was inadvertently reported as death.
5036 | 1220648-2025-46919 | 23159537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159537&pc=OZD | 08/29/2025 | 01/09/2026 |  | 09/26/2025 |  |  | 08/29/2025 | 09/26/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026693249 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that they encountered pump flow during impella cp support. the incident pump was inserted but maintained a flow of just 0.6 l at all times. several repositioning attempts failed to maintain a proper pump flow. in every position of the pump, 0.6 l was shown on the automated impella controller (aic). physician and nurse changed the flow from auto to several p-levels but 0.6 l were shown at all times. after exchange of the pump and the aic, there were no issues that occurred. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Correction: d4, e4. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5037 | 1220648-2025-31698 | 23159556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23159556&pc=OZD | 04/15/2025 | 09/26/2025 |  | 09/26/2025 |  |  | 04/16/2025 | 09/26/2025 |  |  |  |  | 09/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589725 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support was ambulating when a placement signal not reliable alarm occurred. the team monitored motor current. the pump was eventually successfully weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. after 16 days on support, placement signal not reliable (psnr) alarms occurred during patient ambulation and placement signal (ps) was lost. the doctor believes optical sensor damage with motor current (mc) and flows still present and accurate. the device was not returned. data log review showed ps trending down of the course of 24 hours until railing to 3000mmhg. 5s parameters were within nominal ranges with signal-to-noise ratio (snr) and contrast decreasing after ps failure. the cause of the optical signal issue was sensor silicone wear due to the data log review indicating known sensor wear metrics of ps trending down over the course of 24 hours until railing to 3000mmhg with nominal 5s parameters. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5038 | 1220648-2025-46824 | 23140279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23140279&pc=OZD | 08/25/2025 |  |  | 09/25/2025 |  |  | 08/25/2025 | 09/25/2025 |  |  | 09/30/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2024401770 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 78 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ischemia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that the patient presented to the emergency room with shortness of breath symptoms. the patient was found to have a significant clot burden in his right atrium and imaging indicated there was a possible thrombus in the pulmonary artery (pa). the patient¿s right ventricle was strained, and a reduced systolic function was observed. additionally, observed left ventricle septal flattening was consistent with the patient¿s volume and pressure overload. the initial ejection fraction was determined to be approximately 40%, and a tricuspid annular plane systolic excursion of 1.2 cm. the patient became progressively short of breath and computed tomography (ct) was performed to confirm the pa thrombus. the patient was taken to the operating room for an open chest thromboembolectomy of the main pa. despite the successful procedure, the patient¿s cardiac function progressively weakened. after returning to the intensive care unit (icu), the patient began to decompensate with vasoactive intravenous drips actively being increased. the patient was then cannulated for veno-arterial extracorporeal membrane oxygenation (va ecmo). repeat imaging revealed severely reduced bi-ventricular function and worsening septal dyskinesis. the decision was then made to take the patient to the cardiovascular laboratory (cvl) to assess coronary function and to place an impella cp for left ventricular unloading in the presence of va ecmo. in the cvl, the patient¿s lateral circumflex femoral artery was accessed, and prior to impella cp insertion the physician decided to assess the coronary arteries. the physician noted there was some bridging in the left anterior descending artery, but timi 3 flow also noted distal to the lesion and the decision was made to not intervene at that time. the impella cp was implanted, and support initiated without complications. the patient was released back to the icu for rest and recovery. upon arrival to the icu on va ecmo and impella cp support, the physician reviewed the placement of the impella cp. some imaging indicated that the impella was quite shallow, however, all waveforms were appropriate, with no signs of hemolysis, and no suction events. the physician therefore made the decision not to attempt to reposition the device. the patient was able to be weaned off of intravenous epinephrine drop, however when milrinone was started, the patient experienced increased ectopy, and the medication was discontinued. the following day it was reported that the va emco flows had decreased to 3l, and with the decrease of flows the patient¿s blood lactate results increased from 3.2 mmol/l to 4.0 mmol/l. the patient also became hypotensive. repeated imaging revealed the impella shallow at 2.5 cm, but the physician declined to reposition the pump. it was noted that the patient was placed on continuous renal replacement therapy the previous night. additionally, a left lower extremity ultrasound was performed, and it revealed limited flows distal to the impella insertion site. the medical team consulted a vascular surgeon to determine to escalate the patient to an impella 5.5 or to add a distal perfusion catheter to reestablish flows down the leg. the decision was made to escalate the patient to an impella 5.5. the impella 5.5 was implanted without issue, and the impella cp was explanted by a vascular surgeon. it was noted that the patient was hemodynamically stable after arriving back at the icu. it was reported that on the first day of impella 5.5 with va ecmo support, the patient was noted for potentially having north south syndrome. the impella flows were reduced, and it was noted that the patient¿s arterial blood gas results and patient color improved. imaging was performed to confirm placement of the impella 5.5 was optimal. it was noted that prior to impella 5.5 placement, the patient¿s ejection fraction was determined to be 40-45%, with noted moderate right ventricular systolic dysfunction. a head computed tomography (ct) revealed small acute or subacute middle cerebral artery infarcts. additionally, it was noted that the patient may have small bowel ischemia on ct, however the bowels were still moving. it was also noted that the pa thrombi were significantly reduced. the following day, the patient was removed from sedation, and it was noted that the patient had no gag reflexes, cough, or withdrawal from pain. the next day it was noted that repeat echocardiography from the previous day confirmed impella was well placed. it was also noted that the patient had a severely reduced left ventricular systolic function, and the ejection fraction was unable to be calculated. later the same day, the decision was made to withdraw care. the patient¿s care was withdrawn and the patient expired. this complaint involves two impella devices: pump 1: impella cp with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. he pump position was questionable per imaging, but the physician elected not to reposition as waveforms were appropriate, and there was no hemolysis nor suctioning. the patient had increased ectopy with medication changes. the patient became hypotensive and positioning showed the pump in questionable positioning. the physician again elected to not reposition the pump during echocardiogram. ultrasound of the left lower extremity showed limited distal flows to the impella insertion site. care was escalated with the removal and replacement of the impella cp with an impella 5.5. patient survived. |  |  |  |  | Additional narrative added to b5. correction to b5: patient outcome was incorrectly reported as survived in the initial report 1220648-2025-46824. the patient expired. added mrn number to related report in h10. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). h.10 related report number was added. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. ischemia, arrhythmia: in order to make a cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): no device was returned. pump was measured to be 3.2 cm and 2.5 cm the next day. the cause of the positioning issue was unable to be determined due to insufficient clinical details. | Medical safety reviewed the event. the event with the impella cp describes shallow device positioning (x2) where the physician made the decision not to reposition as there was no hemolysis or suction. subsequently distal limb perfusion concerns were noted with the impella cp, so the decision made to escalate to impella 5.5 axillary. thereafter escalation to the impella 5.5, the patient became hemodynamically stable. the event occurring with the impella cp is a serious injury and should be reported. the following day after start of impella 5.5 mcs, the ct scan revealed acute and subacute mca infarct; physicians unconcerned; patient expired two days later. there is not enough information dissociate the impella 5.5. however, the infarct is most likely unrelated to the impella 5.5 due to the fact that preecmo insertion there was identified clot in right side of the heart; presumed pulmonary embolism. the patient's underlying condition likely contributed most to an outcome of death. the patient had severely reduced left ventricular systolic function, and care was withdrawn which led to the patient expiring. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5, which is associated with the demise outcome.
5039 | 1220648-2025-46825 | 23141814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23141814&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614220 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 47 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Atrial Fibrillation | Inflammation | Ischemia | Renal Failure | Tachycardia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the patient had previously been supported by an impella cp where no issues were encountered, and the medical team decided to escalate the patient to an impella 5.5 for continued unloading and planned for coronary revascularization. during the impella 5.5 implant, a right axillary graft was placed, and it was reported that there was difficulty advancing the pump through the graft into the artery and fluoroscopy was used to guide the pump placement. after the impella cp was successfully weaned and explanted, the femoral-to-femoral bypass which had been pro-actively placed for limb perfusion was removed and closed with two angioseals. it was reported that the patient had an episode of ventricular tachycardia and two episodes of atrial fibrillation with rapid ventricular response requiring defibrillation and cardioversion. each shock converted back to normal sinus rhythm with perfusing pressure and pulsatility. the patient was transfused with cell saver blood and two units of packed red blood cells due to blood loss during graft placement. an echocardiogram was completed and the impella 5.5 position was reviewed. the medical team decided to not reposition the device and continue support and the patient¿s care plan was for revascularization and volume replacement. the following day, the medical team reported that repositioning of the pump would be attempted when the physician was available. additionally, it was noted that the patient was placed on continuous renal replacement therapy and dialysis. on the third day of support, vascular surgery noted that they continued to monitor the impella site and believed there to be a venous issue at the impella access site. the impella 5.5 was conservatively being weaned, and one unit of packed red blood cells were transfused. the following day, it was noted that the patient¿s arm swelling slightly improved and pulses were confirmed. the arm progressively improved during support. the patient required re-intubation. the impella 5.5 was explanted via surgical wound closure. pocket exploration and washout were performed, graft cut down to anastomosis and oversewn. the wound was closed with the drain left in place. the pump was successfully weaned and hemostasis was achieved. the pump was explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
5040 | 1220648-2025-46826 | 23141884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23141884&pc=PYX | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025587970 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support. it was reported that purge system open alarm was noted on the automated impella controller (aic). the nurse performed two purge cassette changes trying to resolve the alarm but it remained. upon further examination of the purge sidearm and blue plug, the nurse noted purge fluid leaking from the top and bottom part of the blue plug. the team at the account site was going to attempt to perform a purge side-arm bypass but noted that purge sidearm tubing came out of gray connection and was no longer attached to the blue power plug. at this time, the doctor was notified of the alarm and need for explant or replacement of device. the decision was made to proceed with an explant. the patient was brought down to the catheterization lab for explant. the intensive care unit staff was unsure of what led to the issue as they did not recognize an event that led to a leak of the purge sidearm. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5041 | 1220648-2025-46827 | 23141958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23141958&pc=OZD | 08/26/2025 | 11/05/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026714736 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | Yes | Inaccurate Flow Rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was placed but explanted after the coronary artery bypass graft and day two of the support due to purge alarming for purge flow and saturation. the team had attempted to troubleshoot by adding tissue plasminogen activator to the purge. the pump was not replaced after explant as the patient was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | H6 medical device problem code (a141102) was inadvertently omitted on the initial manufacturer device report number. additional information was provided by user medwatch (b)(4). b1, b2, and e4 was updated accordingly. the investigation was completed. no product was returned. after 2 days on support, purge pressure high alarms occurred with elevated pump flows. tpa was added with unknown resolve to the issue as the pump was explanted. data log review showed pump turned off 14 hours prior to purge pressure high alarms, noted time when pt underwent cabg procedure. on pump restart, purge pressure is elevated. purge pressure continues to rise with lowering purge flows for the next 14 hours before alarms occur. no mc spikes outside of p-level changes. flows at p-8 intermittently rise above the expected 4.3-4.9l/min for impella 5.5. high or saturated pump flow - the cause of the saturated pump flow was biomaterial build up due to intermittent saturated flows and no mc spike noted. high purge pressure - the cause of the high purge pressure was biomaterial buildup due to the gradual rise in purge pressure over 14 hours with lowering purge flows and no mc spike noted. the device history review was completed and passed all post sterile inspection checks.
5042 | 1220648-2025-46829 | 23142311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23142311&pc=OZD | 08/27/2025 | 09/24/2025 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713981 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 a patient was going for a possible washout due to bleeding. the patient went into atrial fibrillation (af) overnight with three boluses to contain. the patient was still in af but was controlled on 1mg of amiodarone. the patient was less pulsatile but has since regained pulsatory. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. major bleed: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. paroxysmal. atrial fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5043 | 1220648-2025-46828 | 23142324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23142324&pc=OZD | 08/28/2025 | 09/24/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671766 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 88 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella 5.5 placed to support with ongoing cardiogenic shock and cardiomyopathy. the pump was placed but was repeatedly repositioned due to suction despite being in the appropriate position. the pump was explanted after 29 days and was replaced by the left ventricular assist device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump suction: after reviewing logs, multiple suction alarms were observed. 1 motor current (mc) spike was observed but the parameters resolve afterwards and the suction started prior to the spike. the cause of pump suction cannot be determined due to insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks.
5044 | 1220648-2025-46831 | 23142682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23142682&pc=OZD | 08/26/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574188 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 115 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient admitted in with a past medical history and new arrest and cardiopulmonary resuscitation. the impella cp was placed and allowed for support but it was explanted with hematuria and consideration of possible hemolysis. the patient's urine was pink tinged. plasma-free hemoglobin was sent out and lactate dehydrogenase and lactate were on a downward trend. additional information was received. hemolysis was never confirmed. the pump was discarded once removed. the pump was not removed due to hematuria. it was removed due to the significant improvement of the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. mechanical interaction with blood (hematuria): data log was not provided or could not be retrieved from the remote server. the cause of the mechanical interaction with blood issue was not determined without returned product investigation and sufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5045 | 1220648-2025-46832 | 23142921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23142921&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707146 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 118 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | Difficult to Insert | Pumping Stopped | Device Damaged by Another Device | Device Sensing Problem | Difficult to Advance | Hemodynamic instability | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | PROTEK. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient collapsed during a firefighting training exercise and cardiopulmonary resuscitation was immediately initiated, and the patient was taken to the emergency room. the patient was then cannulated for veno-arterial extracorporeal membrane oxygenation (va ecmo) and then was urgently rushed to the cardiac catheterization laboratory to have an angiogram, swan, and impella cp placed. a high-risk catheterization was performed to treat a blockage in the left main coronary artery (lmca) via the left common femoral artery. angiography confirmed the lmca blockage, and the patient received an impella cp for temporary heart support. a penumbra catheter was used to remove the blood clot blocking the lmca and stents were placed in both the left anterior descending and left circumflex arteries. the patient was stabilized, a distal perfusion catheter was connected to the arterial line of the va ecmo circuit for further support, and the patient was transferred to the cardiac care unit in critical stable condition. the patient remained on impella cp support for five days without issue, and the medical team made the decision to escalate the patient to an impella 5.5 for increased support. the patient was brought to the operating room for an impella 5.5 implant and impella cp explant. the initial impella 5.5 selected for use had ¿placement signal not reliable¿ codes that occurred during the set up of the device. multiple attempts were made to resolve the issue, including attaching the impella 5.5 to a different automated impella controller unit, but ultimately the set-up was aborted and a new impella 5.5 was used. the second impella 5.5 pump was successfully primed and set up without issue. left ventricle access obtained and the impella 5.5 insertion attempt was made but resistance was met at the aortic valve with the impella cp in place. after successfully crossing the valve, it was reported that the impella 5.5 would only advance approximately two centimeters across the valve. the physician made the decision to not use the ¿double barrel technique¿ and direction to advance the impella 5.5 to two centimeters by company support after the impella location was measured at one centimeter with transesophageal echocardiogram. as result, the impella 5.5 was pulled back across the aortic valve during the pullback of the impella cp. attempts were made to recross the aortic valve with the impella 5.5, however when the impella 5.5 recrossed, it rotated 180 degrees with the outlet cage and motor housing sitting in the left ventricle, and the inlet cage in the aorta. the impella 5.5 was completely removed, left ventricle access was reobtained with a guidewire and the impella 5.5 was successfully reinserted. the guidewire was removed, and impella 5.5 support was successfully initiated. it was reported that after impella 5.5 support was initiated, the patient¿s right ventricle showed show further signs of decompensation and failure, and the decision was made to add an additional circuit to va ecmo circuit for protek cannulation (vp ecmo). the impella cp was then successfully removed, and the impella cp access site was closed. the patient then began having flow and suction issues on both va/vp ecmo and impella 5.5. the impella 5.5 displayed pump saturation and flows decreased from a mean of min of 3.5 to a max of 5 liters per minute (l/min) to 0.1 l/min of flow. to attempt to prolong pump failure due to the pump saturation and low flows, the medical team decided to add tissue plasminogen activator to the purge solution and lower the p-level of the pump. additionally, it was noted that the protek cannula of the vp ecmo was possibly too deep and the patient received volume as needed for the suction events on both vp ecmo and the impella 5.5. the following day, the impella 5.5 had a pump failure. the impella 5.5 was explanted, and the patient was closely monitored for a repeat impella 5.5 implant if needed. the patient remained on va/vp ecmo and was noted to be in very critical condition with a very guarded prognosis. this complaint involves two impella 5.5 pumps: pump 1: impella 5.5 with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). this report specifically addresses the second impella 5.5. a separate report was submitted for the first impella 5.5. refer to section h.10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H.6 codes a141204, a140508, and a150206 were reported incorrectly on the initial report that was submitted. code a150205 has been added. h.6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned. difficult to deliver/advance: clinical details reported impella 5.5 insertion was met with resistance at the aortic valve (av) and would only advance approximately two centimeters across the valve. there was difficulty placing the pump in the right position as evidenced by clinical summary. the cause of the difficulty to deliver pump issue was determined to be use related as evidenced by the device interaction (previous impella cp pump) causing the impella 5.5 pump not to advance past two centimeters. device interaction: the impella 5.5 was pulled back across the av during pullback of the impella cp. the cause of the device interaction issue was determined to be use related as evidenced by the pump being pulled back during manipulation of another device. saturated flow: clinical details reported the patient was put in reverse trendelenburg which immediately led to suction occurring on both extracorporeal membrane oxygenation (ecmo) circuits as well as the impella 5.5 showing signs of pump saturation. flows immediately went from a mean of 3.5 to 5 liters per minute. data logs were reviewed and motor current (mc) spike was observed on (b)(6) 2025. during that first mc spike, the purge pressure started to rise gradually, and the pump flows became saturated showing higher flows with increased motor current due to biomaterial ingestion causing added resistance to motor. the saturated flows along with high purge pressure continued till end of case for next approximately 21 hours until high motor current pump stop. the left ventricular pressure (lvp) was spiking upward or trending high decoupling from aorta-placement signal (ao-ps) due to ingestion. no other abnormalities were seen. the cause of saturated flows were due to biomaterial ingestion based on log review. optical sensor issue: as per clinical, left ventricle (lv) signal systolic and diastolic values were both higher than placement signal. data logs were reviewed and mc spike was observed on (b)(6) and lvp spiked up from 30 mmhg to >200 mmhg decoupling from ao-ps. during that first mc spike, the purge pressure started to rise gradually and the pump flows became saturated showing higher flows with increased motor current due to biomaterial ingestion causing added resistance to motor. apart from the lvp decoupling and trending high from ao-ps, all other 5s parameters were in normal range. real time logs showed initially prior to mc spike on (b)(6) the signal to noise ratio, contrast were slowly trending down but were up back in normal range slowly later. also, the raw optical sensor showed a downward drop/trend in relevance to mc spike (ingestion). no other abnormalities or hard failure of optical signal was seen during the case. the cause of lv optical signal being high than ao-ps was due to biomaterial ingestion based on log review. clinical symptom: - hemodynamic instability: the cause of injury is not determined due to insufficient clinical information. device history review: the pump passed all post sterile inspection checks.
5046 | 1220648-2025-46833 | 23143234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143234&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709917 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 118 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient collapsed during a firefighting training exercise, cardiopulmonary resuscitation was immediately initiated, and the patient was taken to the emergency room. the patient was then cannulated for veno-arterial extracorporeal membrane oxygenation (va ecmo) and then was urgently rushed to the cardiac catheterization laboratory to have an angiogram, swan, and impella cp placed. a high-risk catheterization was performed to treat a blockage in the left main coronary artery (lmca) via the left common femoral artery. angiography confirmed the lmca blockage, and the patient received an impella cp for temporary heart support. a penumbra catheter was used to remove the blood clot blocking the lmca and stents were placed in both the left anterior descending and left circumflex arteries. the patient was stabilized, a distal perfusion catheter was connected to the arterial line of the va ecmo circuit for further support, and the patient was transferred to the cardiac care unit in critical stable condition. the patient remained on impella cp support for five days without issue, and the medical team made the decision to escalate the patient to an impella 5.5 for increased support. the patient was brought to the operating room for impella 5.5 implant and impella cp explant. the initial impella 5.5 selected for use had ¿placement signal not reliable¿ codes that occurred during the set up of the device. multiple attempts were made to resolve the issue, including attaching the impella 5.5 to a different automated impella controller unit, but ultimately the set-up was aborted and a new impella 5.5 was used. the second impella 5.5 pump was successfully primed and set up without issue. left ventricle access obtained and the impella 5.5 insertion attempt was made but resistance was met at the aortic valve with the impella cp in place. after successfully crossing the valve, it was reported that the impella 5.5 would only advance approximately two centimeters (cm) across the valve. the physician made the decision to not use the ¿double barrel technique¿ and direction to advance impella 5.5 to two cm by company support after the impella location was measured at one cm with transesophageal echocardiogram. as result, the impella 5.5 was pulled back across the aortic valve during the pullback of the impella cp. attempts were made to recross the aortic valve with the impella 5.5, however when the impella 5.5 recrossed, it rotated 180 degrees with the outlet cage and motor housing sitting in the left ventricle, and the inlet cage in the aorta. the impella 5.5 was completely removed; left ventricle access was reobtained with a guidewire and the impella 5.5 was successfully reinserted. the guidewire was removed, and impella 5.5 support was successfully initiated. it was reported that after impella 5.5 support was initiated, the patient¿s right ventricle showed further signs of decompensation and failure, and the decision was made to add an additional circuit to va ecmo circuit for protek cannulation (vp ecmo). the impella cp was then successfully removed, and the impella cp access site was closed. the patient then began having flow and suction issues on both va/vp ecmo and impella 5.5. the impella 5.5 displayed pump saturation and flows decreased from a mean of min of 3.5 to a max of five (liters per minute) l/min to 0.1 l/min of flow. to attempt to prolong pump failure due to the pump saturation and low flows, the medical team decided to add tissue plasminogen activator to the purge solution and lower the p-level of the pump. additionally, it was noted that the protek cannula of the vp ecmo was possibly too deep and the patient received volume as needed for the suction events on both vp ecmo and impella 5.5 the following day, impella 5.5 had a pump failure. the impella 5.5 was explanted, and the patient closely monitored for a repeat impella 5.5 implant if needed. the patient remained on va/vp ecmo and was noted to be in very critical condition with a very guarded prognosis. this complaint involves two impella 5.5 pumps: pump 1: impella 5.5 with serial number: (b)(6). pump 2: impella 5.5 with serial number: (b)(6). this report specifically addresses the first impella 5.5. a separate report was submitted for the second pump. refer to section h.10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 code 1903 was reported incorrectly on the initial report that was submitted. component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. no product was returned but the event logs were returned on 20-oct-2025. optical sensor issue: clinical details noted that during setup and prep of the impella 5.5, there was a peak signal-to-noise ratio alarm with dashed out placement signal. they attempted to reseat the white cable multiple times as well as switching out the console. the pump was replaced. data log review showed signal-to-noise ratio (snr) around 7000, contrast at approximately 28%, gain at approximately1.4, and lamp at approximately 85% in the first minute of data logs. after switching consoles, snr was approximately 7000, contrast at approximately 32%, gain at approximately 1.9, and lamp at 100%. placement signal was out of range at 3000. the log review of automated impella controller revealed that optical calibration was unsuccessful since g0=0 even after multiple attempts. the cause of the optical sensor issue was eeprom corruption since the log review showed unsuccessful optical calibration with g0=0. device history review: the pump passed all post-sterile inspection checks.
5047 | 1220648-2025-46830 | 23143256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143256&pc=OZD | 08/26/2025 | 09/23/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711274 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 123 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient went into ventricular fibrillation (vf)/ventricular tachycardic (vt) rhythm on arrival to the catheterization lab. cardiopulmonary resuscitation (cpr) was started and the physician decided to place the impella cp. the impella was successfully placed in the right femoral artery and started. the patient was shocked multiple times, but expired in the procedural area prior to performing intervention. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. ventricular fibrillation/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device is not available for return by the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed.
5048 | 1220648-2025-46834 | 23143355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143355&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690243 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 63 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had been brought to the cardiac catheterization laboratory for left heart catheterization, and during the procedure, the patient became hypotensive and required vasopressors. the physician then elected to place the impella cp to help stabilize the patient. the impella cp was inserted according to the manufacturer's instructions for use, and support was initiated. a thrombectomy was performed in the left anterior descending and left circumflex arteries, a balloon angioplasty was then performed and nipride was administered. following the procedure, the patient was transferred to the intensive care unit for continued monitoring and care. the patient was later intubated for airway protection. the following day, a stroke alert was called due to due to loss of reflexes and no fixed pupils, and remained acidotic. the patient was made do not resuscitate. the patient then expired while on impella cp support. | [5 empty] | Based on the information available, it is unlikely that the stroke was related to the use of the impella device; however, a device-related contribution cannot be definitively ruled out. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5049 | 1220648-2025-46835 | 23143505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143505&pc=OZD | 08/26/2025 | 09/24/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684874 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had shockwave being used during a high-risk percutaneous coronary intervention (hrpci) with the support of an impella cp. during the percutaneous coronary intervention, there was a loss of signal. the optical signal loss did not prompt intervention. at the conclusion of the hrpci, the pump was explanted. no harm or injury was noted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary- impella cp sn (b)(6) returned for investigation. data logs are not available. optical sensor issue: clinical details note that there was a psnr alarm and loss of ps waveforms at the same time that shockwave was being used on lcx. the exact distance between shockwave and the sensor is unknown. data log review showed a loss of ps along with drop in snr near zero. there was also a decrease in contrast but the rest of the optical parameters remained fairly stable. psnr alarmed and ps was lost for the rest of the case. the product was returned and evaluated. the snr was near zero, contrast was below spec, and lamp was ~60%. gain, light, and cavity length were within normal limits. the pump passed laser continuity testing. the sensor was removed and inspected. the sensor face appeared to be damaged. the cause of optical sensor issue was a damaged sensor likely due to shockwave interaction since clinical details noted that ps was lost while using shockwave and the returned product had a damaged sensor. device history lot- device lot: 1931464. device history batch- subcomponent lot: n/a. device history review- the pump sn (b)(6) passed all post-sterile inspection checks.
5050 | 1220648-2025-46836 | 23143610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143610&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699755 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for unloading during percutaneous coronary intervention. the pump was placed per instructions for use and the patient remained hemodynamically stable throughout the case. the impella cp was weaned and removed. the perclose was locked down. there was some oozing remaining so the physician placed one angio seal. the patient was sent to the intensive care unit without any groin issues or bleeding. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | B.7 additional information was added as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the cause of the bleeding was not determined since the product was not returned and insufficient clinical details provided. the device history review was completed and the pump passed all the post sterile inspection checks.
5051 | 1220648-2025-46837 | 23143735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143735&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026715968 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 76 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in non-st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient presented to the emergency room with chest pain and wheezing that had been occurring for two days. right bundle branch block was noted on electrocardiogram and there was concern for pulmonary edema with new onset of congestive heart failure. the patient was brought to the cardiac catheterization laboratory for angiography. the procedure revealed an occluded right coronary artery and extensive, diffuse disease in the left anterior descending artery (lad). during the planned angioplasty of the lad, the vessel thrombosed, and the patient's left ventricular end-diastolic pressure became critically high at 44 mmhg, and the left ventriculogram showed a low ejection fraction. the decision was then made to abort the procedure, place an impella cp and transfer the patient to another hospital for a surgical workup. the planned percutaneous insertion of the impella pump was unsuccessful as it was reported that the impella would not cross the iliac artery. after multiple attempts to cross, the decision was made to place intra-aortic balloon pump due to patient¿s deteriorating condition. the patient¿s outcome is unknown. | [5 empty] | Investigation summary. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition because the pump was unable to pass the vasculature due to resistance in femoral artery and impella would not cross iliac artery because of patients deteriorating condition despite multiple attempts and the insertion was aborted. hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. device history lot: device lot: 1955124. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5052 | 1220648-2025-31670 | 23143918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143918&pc=OZD | 02/05/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 02/05/2024 | 09/25/2025 |  |  | 11/30/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434176 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a position in aorta alarm appeared during impella cp support. this indicates an optical sensor issue. the placement of the device was verified via echocardiogram, and the decision was made to disable the alarms. decision made to withdraw care. patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. product was not returned. data logs were returned and investigated. patient's health was deteriorating. there was a motor current spike observed and after 19 hours the "placement signal unreliable" alarms were recorded which suggests sudden biomaterial ingestion. placement was confirmed to be correct by echocardiogram by the doctor suggesting that it was a false alarm. per the clinical and data log review, the root cause of the optical signal alarm issue was biomaterial ingestion related to false "pump in aorta" alarm. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5053 | 1220648-2025-46838 | 23143932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143932&pc=OZD | 08/26/2025 | 09/24/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703379 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Atrial Fibrillation | Hematoma | Hemolysis | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during impella cp support a hematoma superior laterally to the impella and swan site (left groin) was noted. manual pressure was held by the medical doctor. the hematoma was soft post hold. changing to the dressing was noted. the patient was hemodynamically unstable during an episode of atrial fibrillation and was treated with amiodarone. inotropes were increased. the interdisciplinary team discussed viability and the next steps. plan was pending. no hematoma was noted and reposition given hemolysis and shown deep on echocardiogram. the patient's family made the decision to withdraw care and the patient passed away overnight. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | Product complaint # (b)(4). investigation summary: no product was returned. adverse site access event (hematoma): the cause of the hematoma was unable to be determined due to insufficient clinical details. hemolysis: no product was returned. after reviewing logs, suction alarms were observed. the cause of hemolysis cannot be determined due to insufficient clinical details provided. paroxysmal atrial fibrillation: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
5054 | 1220648-2025-31671 | 23143965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23143965&pc=OZD | 02/05/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 02/05/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024415027 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the console intermittently displayed sporadic placement signal numbers. despite this, the flows remained unchanged, and the patient was stable. pump position was confirmed, and the placement signal was disabled. no patient harm was associated with the issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the console logs are consistent with complaint and show optical placement signal was intermittently invalid. the console was returned and testing of the console did not reveal any irregularities, and optical bench parameters were within specification. there was no issue related to this device. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5055 | 1220648-2025-31672 | 23144084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144084&pc=OZD | 12/13/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 12/13/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576788 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that their impella cp¿s placement signal was still deactivated. the pump position had been confirmed via echocardiography. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. the data log shows a downward trend in the placement signal without any reported abnormalities on 5s optical signal parameters. several suction alarms were observed, and the impella position in the ventricle was checked before disabling the placement signal monitoring. additionally, a motor current trend was noted during the placement signal issue. the cause of the optical signal issue was most likely patient condition related, based on data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5056 | 1220648-2025-46839 | 23144087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144087&pc=OZD | 08/23/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711489 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient with worsening heart failure was presented to the hospital with low blood pressure and shortness of breath. to stabilize the patient¿s condition, the patient was placed on pressors and inotropes and then an impella 5.5 was then placed in the right axillary artery to help rest and unload the patient¿s failing heart while they were evaluated for a heart transplant. on the thirteenth day of support, it was reported that the previous day the patient had suction events, and the p level of the impella had to be decreased to p-4. the patient became hypotensive during this time and required administration of llvophed and a 500 cubic centimeter bolus of lactated ringer's solution. a transthoracic echocardiogram was performed and revealed the pump to be in good position and the patient¿s right heart had good function. additionally, it was reported that the suction events ultimately led to the removal of the impella 5.5. three days prior to the explant, the patient experienced an embolic cerebrovascular accident. suction events were noted to occur the following two days after the stroke occurred, which were treated with volume given to the patient. ultimately, due to stroke history, the decision was made to remove the impella for suction events believed to be caused by clot formation. imaging did not demonstrate a clot formation in the left ventricle or on the device, and no clot was observed at pump removal. the patient¿s outcome at end of procedure is unknown. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. no product was returned. stroke: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hypotension: the cause of the injury was not determined due to insufficient clinical details. pump suction: no product was returned. patient had suction alarms which were resolved with volume. after reviewing logs, suctions alarms were observed. the cause of pump suction was most likely patient condition related since pump suction alarms were resolved after volume was given. device history review: the device passed all post sterile inspection checks. | The device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5057 | 1220648-2025-31673 | 23144139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144139&pc=OZD | 03/20/2023 | 09/25/2025 |  | 09/25/2025 |  |  | 03/20/2023 | 09/25/2025 |  |  | 12/31/2024 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298871 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp support ongoing when the team observed signal issues in which the pump was improperly alarming for the pump in the aorta. the pump remained on support. the patient eventually expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for review. data was reviewed, which showed false impella position in aorta alarms two days into support. at the time of the alarm, the motor current pulsatility was decreasing and the placement signal remained stable showing strong aortic pressures. there is no evidence of biomaterial interacting with the pump. thus, the root cause of the alarm was most likely patient condition. all pertinent information available to abiomed has been submitted at this time.
5058 | 1220648-2025-46840 | 23144268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144268&pc=OZD | 08/26/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711302 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered a placement signal not reliable alarm during impella cp support. the physician stated that a placement signal not reliable alarm occurred 10-15 minutes after the case was started. hemodynamics and support remained stable. additionally, the pump was discarded inadvertently by the staff post procedure. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. data logs are not available. optical sensor issue: the cause of the optical sensor issue was not determined since product/logs were not returned and insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5059 | 1220648-2025-31675 | 23144337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144337&pc=OZD | 07/26/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 07/26/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486988 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and there were placement signal not reliable alarms. the audio was disabled, and placement was confirmed via echocardiography. support continued and the staff monitored. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. insufficient data logs were returned for evaluation from the date of optical failure. clinical details noted placement signal not reliable (psnr) alarm was noted in the field. the root cause of the optical signal issue cannot be determined with no product or data logs returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5060 | 1220648-2025-31676 | 23144369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144369&pc=OZD | 11/29/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 12/10/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547855 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during insertion, the shaft of the motor connection bent at the artificial blood vessel anastomosis site. the implant was successful, but the motor waveform reading went completely flat. support with the implanted pump continued despite the resulting optical signal issues. there was no report of patient harm resulting from the failure. the patient passed away on support after seven days. there was no allegation of device association with the patient's passing. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. log review showed multiple suction and impella position unknown alarms within the first 30 minutes of the pump running. there are no abnormalities with the motor current in this time. the optical parameters were stable throughout the case. there are some small motor current spikes throughout the rest of the case, but no significant spikes indicating large biomaterial ingestion. product was returned and evaluated. the 5s parameters were all in-spec. the signal-to-noise ratio was near the low end of the spec at ~2800. the pump also passed the laser continuity test. a catheter kink was found right before the motor housing as noted in the clinical details. the 5s parameters did worsen when bending at the kink. the pump passed all electrical testing related to the motor current cables. there was a small amount of biomaterial found on the underside of the impeller. the pump was run in a bucket of water for 15 minutes. there was slightly saturated flows, but there was biomaterial on the returned pump. otherwise, the pump ran without issue. the product evaluation showed the pump's functionality was not affected by the catheter kink or biomaterial. based on the clinical details and data log review, it is unclear what the reported issue was. the root cause of the optical signal issue was not determined since insufficient clinicals details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5061 | 1220648-2025-31677 | 23144391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144391&pc=PYX | 11/09/2023 | 09/25/2025 |  | 09/25/2025 |  |  | 11/09/2023 | 09/25/2025 |  |  | 02/28/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024348386 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp flex support ongoing when the placement signal (ps) was lost. the signal loss did not cause the team to explant the pump. the patient remained on support for 5 days until the family made choice to withdraw care and allow the patient to expire. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs were viewed, and there were placement signal unreliable alarms on (b)(6) where the signal-to-noise ratio goes to 0, lamp steps up, gain increases, contrast decreases, and light decreases. this pattern of the 5s parameters is indicative that the optical fiber most likely broke. due to no product returned or no information provided in the clinical or patient updates of how this breakage could have occurred, the location of the break is unknown. the root cause of the optical signal issue is most likely due to a damaged fiber line. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5062 | 1220648-2025-31678 | 23144461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144461&pc=PYX | 08/20/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 08/20/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529806 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella 5.5 and rp pump to support a patient. the chest was to be left open for post-operative washouts, and the 5.5 was placed surgically with transesophageal echocardiography guidance. the next day, when placing another line (venous for dialysis), the rp lost its optical signal but remained on for support and caused no reported harm or injury. the pump remained on support despite signal loss. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed, and the "placement signal not reliable (opm invalid)" alarm did trigger on (b)(6). just before the alarm, there is a spike in motor current, drop in motor speed, automatic drop in p-level, and drop in placement signal. this is followed by the central venous pressure railing out for the remainder of the case along with the signal-to-noise ratio (snr) dropping to zero, lamp increasing, contrast dropping quite low, gain dropping slightly, and light remaining constant. this combination of changes in optical parameters is indicative of an optical sensor break. this also aligns with the clinical timeline that a dialysis catheter was being inserted from the femoral artery. it is very possible the catheter interacted with the pump, causing the motor current spike and automatic p-level drop, breaking the sensor. returned product was investigated and the pump failed 5s testing as snr was zero and cavity length was 32,768 nm. the pump did, however, pass the laser continuity test proving that there was no damage to the fiber optic cables. the optical sensor was imaged and there appeared to be damage on its face. the root cause of the optical signal issue was determined to be a damaged sensor likely due to an interaction with the dialysis catheter. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5063 | 1220648-2025-31679 | 23144551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144551&pc=PYX | 06/17/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 06/17/2024 | 09/25/2025 |  |  | 09/30/2024 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023270455 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm triggered during impella rp support. the flows were no longer being calculated, and the motor current was observed to be constant. intervention was not required, and support was maintained with the implanted pump. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or logs were returned for evaluation. clinical details noted the motor current was consistent when the placement signal not reliable alarm was noted and flows were lost. the root cause of the placement signal issue cannot be determined, and no logs or product was returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5064 | 1220648-2025-31680 | 23144618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144618&pc=PYX | 12/23/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 12/23/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540878 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella rp flex support, and, during support, there were placement signal issues. the position was verified. the patient was actively decompensating when the issue occurred. widespread systemic clotting was also noted. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs confirm placement signal not reliable alarms. they show suction alarms with a decreasing central venous pressure before going out of range. data logs also show the signal-to-noise ratio trending downwards. contrast was also trending downwards, and lamp stepped up to 100. gain and light were stable. clinical mentioned it was a bipella case and mentioned that patient was actively decompensating. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5065 | 1220648-2025-31681 | 23144659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144659&pc=OZD | 04/08/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 04/08/2024 | 09/25/2025 |  |  | 12/31/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453418 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 support for a patient, the placement signal was lost. the other values were in range, and the pump was working well. the pump's position was confirmed via echocardiography. neither the console nor the pump were replaced as a result. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product and data logs were not provided for this investigation. as per clinical details, pump lost placement signal during use. the cause of the optical signal issue was not determined since product and logs were not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5066 | 1220648-2025-46842 | 23144660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144660&pc=OZD | 08/26/2025 | 09/24/2025 |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  | 01/31/2026 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504126 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed on a patient for ventricular tachycardia ablation, however during support decision was made to remove the cp due to evidence of hemolysis and hematuria. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. hematuria - the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood - after 1 day on support, impella cp was removed due to hemolysis as evidence by hemolyzing labs and hematuria. pt noted to be scheduled for high risk vr ablation. the cause of the mechanical interaction with blood was not determined due to no product returned, no data logs recovered, and insufficient clinical details. h6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. e4 added selection as it was omitted from the initial report that was submitted. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5067 | 1220648-2025-46843 | 23144721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144721&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716170 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 96 KG |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted as the pump was unable to pass from the innominate to the aorta. an impella cp was implanted instead. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ the investigation has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had stenosis in the aorta preventing successful delivery of impella. the device history review was completed and the pump passed all the post sterile inspection checks.
5068 | 1220648-2025-31682 | 23144778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144778&pc=PYX | 08/06/2023 | 09/25/2025 |  | 09/25/2025 |  |  | 08/06/2023 | 09/25/2025 |  |  | 06/30/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024363976 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp placed for support. the rp had lost placement signal after two days of support. audio was disabled. an aorta rip and chest bleeding were also noted. platelets, fresh frozen plasma (ffp), blood, kcentra, and factor 7 were given to help with the bleeding. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. according to the clinical, six hours after beginning impella rp support, a placement signal not reliable alarm appeared. no attempts were made to resolve the issue, and the patient continued on support. no product was returned for a visual inspection. data log analysis confirmed multiple placement signal not reliable alarms throughout support. motor current was trending down along with flow as the placement signal is trending up. the optical sensor appears to be going negative in these suction conditions, as the central venous pressure and signal-to-noise ratio trended down. there does not appear to be a hard failure of the optical sensor. based on the data log analysis, the most likely cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5069 | 1220648-2025-46845 | 23144795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144795&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026710686 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support the patient admitted in for postcardiotomy cardiogenic shock/ low cardiac output syndrome. during implant, the team attempted placement of impella 5.5 but the case was aborted due to patient anatomy. the pump was inserted but unable to make a turn from the axillary artery to the aorta. | [5 empty] | The investigation of the reported failure mode has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had difficult vascular anatomy preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5070 | 1220648-2025-46844 | 23144813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23144813&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725509 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 96 KG |  | White | Malfunction | No | Yes | Device Sensing Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | Followup | [8 empty] | It was reported that after implantation of an impella cp, the aorta sensor immediately failed. the patient continued on impella cp support for a total of 67 hours until they were successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The root cause of the optical signal issue is due to eeprom corruption. the root cause of the delivery issue is due to the patients condition.
5071 | 1220648-2025-31683 | 23145216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23145216&pc=OZD | 01/21/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 01/21/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590509 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that while observing support via impella connect, the placement signal decoupled. the systolic left ventricle placement signal was no longer aligned with the systolic aorta placement signal. further details state that the patient was unstable and care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. the data logs show that the 5s parameters are within specification. there are suction alarms throughout the entire case, and the p-level is changed very frequently. this patient is most likely experiencing a volume issue. the root cause of the optical signal issue is most likely due to the patient¿s condition as they were in suction conditions the whole case causing the discrepancy in the waveforms. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5072 | 1220648-2025-31684 | 23145438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23145438&pc=OZD | 01/25/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 01/25/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1268745 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal low alarm appeared during automated impella controller support. the pump continued to provide circulatory support and was not removed as a result of the alarm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the data logs confirm the clinical description as placement signal trends downward as the case proceeded. the console was tested but the failure mode was not reproduced. the console's 5s were well above passing parameters so that would not have caused the optical signal behavior. the root cause of the optical signal issue was not determined since the failure mode was not reproduced. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5073 | 1220648-2025-46841 | 23145861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23145861&pc=PYX | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  | 10/31/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024410394 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 98 KG | Non Hispanic | White | Death | Yes | Yes | Backflow | Pumping Stopped | Failure to Read Input Signal | Malposition of Device |  |  |  | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock with low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that the patient was admitted for a coronary artery bypass graft and aortic valve replacement using a mechanical valve and root enlargement. post-procedure, the medical team was unable to wean the patient from cardiopulmonary bypass (cpb) due to the right ventricle (rv) becoming distended and the patient becoming hemodynamically unstable. of note, the left ventricle appeared to be functioning normally with no significant abnormalities. in addition to cpb, the patient was supported by an intra-aortic balloon pump and several vasoactive medications. the medical team successfully implanted an impella rp, and support was initiated. it was noted that there was immediate decompression of the rv with improved flows and pulmonary artery (pa) pressures. attempt to wean from cpb was again unsuccessful. the anesthesiologist observed pink, frothy sputum from the endotracheal tube, a sign of severe pulmonary congestion. the medical team decided to cannulate the patient for veno-arterial extracorporeal membrane oxygenation (va ecmo). while waiting for the va ecmo team, the rv became distended again and the impella rp flows decreased. the physician became concerned with impella rp placement and elected to reposition the device with fluoroscopy. during an attempt to reposition the pump, it dislodged from the pa. the physician attempted to recross the valve without success. the cannula was explanted, flushed, and re-prepped. during re-priming, there were alarms for ¿impella stopped¿, ¿retrograde flow¿, and ¿impella defective.¿ the decision was made to replace the impella rp with a new impella rp. the second impella rp was prepped, primed and inserted with no issues. support was initiated, and the patient was then cannulated for veno-arteriovenous extracorporeal membrane oxygenation (vav ecmo), however the patient continued to decompensate. the physician reported ongoing difficulty maintaining hemodynamic support during the transition from cardiopulmonary resuscitation to vav ecmo. it was noted at that time that the patient had received a massive transfusion protocol of 13 units of packed red blood cells and 13 liters of fluid replacement. mean arterial pressures remained in the low-40s despite maximal support. after the prolonged procedure, a multidisciplinary team determined that no further treatment would be successful. therefore, resuscitative efforts were stopped, and the patient expired. this complaint involves two impella rp pumps: pump 1: impella rp with serial number (b)(6). pump 2: impella rp with serial number (b)(6). this report specifically addresses the first pump. a separate report was submitted for the second pump. refer to section h.10 for the related report number. | [5 empty] | Investigation summary: no product was returned. device in wrong position: clinical details note that while attempting to reposition to fix cannula position, the pump dislodged from the pulmonary artery. the cause of the positioning issues was not determined since insufficient clinical details were provided. pump stop/retrograde pump flow: clinical details note that during re-priming, the automated impella controller (aic) alarmed "impella stopped. retrograde flow." data log review confirms impella stopped - reverse flow alarms but there are no pump metrics at the same time. based on clinical details, the alarms most likely occurred while the pump was outside of the body. the cause of the pump stop was not determined since product was not returned and insufficient clinical details were provided. pump not detected: clinical details note that during re-priming, the aic alarmed ¿impella defective.¿ data log review confirms an impella defective alarm. the cause of the pump detection issue was not determined since product was not returned and insufficient clinical details were provided. device history lot ==> device lot: 1750590 device history review ==> the pump sn (b)(6) passed all post-sterile inspection checks. | Although the patient's presenting condition may be more likely to have impacted the event, the first impella rp pump cannot be excluded as a possible contributing factor in the patient's demise. the device was not returned for evaluation. it was reported that the device was explanted and discarded by the facility and is therefore unavailable for return to the manufacturer. should additional information be received, a supplemental report will be filed.
5074 | 1220648-2025-46847 | 23146028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23146028&pc=PYX | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  | 10/31/2025 |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024410397 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock with low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that the patient was admitted for a coronary artery bypass graft and aortic valve replacement using a mechanical valve and root enlargement. post-procedure, the medical team was unable to wean the patient from cardiopulmonary bypass (cpb) due to the right ventricle (rv) becoming distended and the patient becoming hemodynamically unstable. of note, the left ventricle appeared to be functioning normally with no significant abnormalities. in addition to cpb, the patient was supported by an intra-aortic balloon pump and several vasoactive medications. the medical team successfully implanted an impella rp, and support was initiated. it was noted that there was immediate decompression of the rv with improved flows and pulmonary artery (pa) pressures. attempt to wean from cpb was again unsuccessful. the anesthesiologist observed pink, frothy sputum from the endotracheal tube, a sign of severe pulmonary congestion. the medical team decided to cannulate the patient for veno-arterial extracorporeal membrane oxygenation (va ecmo). while waiting for the va ecmo team, the rv became distended again and the impella rp flows decreased. the physician became concerned with impella rp placement and elected to reposition the device with fluoroscopy. during an attempt to reposition the pump, it dislodged from the pa. the physician attempted to recross the valve without success. the cannula was explanted, flushed, and re-prepped. during re-priming, there were alarms for ¿impella stopped¿, ¿retrograde flow¿, and ¿impella defective.¿ the decision was made to replace the impella rp with a new impella rp. the second impella rp was prepped, primed and inserted with no issues. support was initiated, and the patient was then cannulated for veno-arteriovenous extracorporeal membrane oxygenation (vav ecmo), however the patient continued to decompensate. the physician reported ongoing difficulty maintaining hemodynamic support during the transition from cardiopulmonary resuscitation to vav ecmo. it was noted at that time that the patient had received a massive transfusion protocol of 13 units of packed red blood cells and 13 liters of fluid replacement. mean arterial pressures remained in the low-40s despite maximal support. after the prolonged procedure, a multidisciplinary team determined that no further treatment would be successful. therefore, resuscitative efforts were stopped, and the patient expired. this complaint involves two impella rp pumps: pump 1: impella rp with serial number (b)(6). pump 2: impella rp with serial number (b)(6). this report specifically addresses the second impella rp. a separate report will be submitted for the first impella rp. refer to section h.10 for the related report number. | [5 empty] | Although the patient's presenting condition may be more likely to have impacted the event, the first impella rp pump cannot be excluded as a possible contributing factor in the patient's demise. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5075 | 1220648-2025-46849 | 23146376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23146376&pc=OZD | 08/26/2025 |  |  | 09/25/2025 |  |  | 08/26/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681628 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 68 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support for a percutaneous cardiac intervention (pci) following left heart catheterization and selective coronary angiography. the impella cp was placed without complications, and following a successful pci without complications. the patient was admitted to the intensive care unit (icu) still in cardiogenic shock. after arriving to the icu, the patient¿s impella site dressing was found to be saturated with blood. the dressing was changed, and no further bleeding or oozing was noted. the patient¿s partial thromboplastin time (ptt) was elevated with a result of 200 seconds, and the decision was made to hold heparin until the ptt decreased to therapeutic targets. there was additional oozing noted later the same evening, and the patient¿s dressing was changed again. additionally, the medical team was advised to verify angle matching and forward suture tension. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5076 | 1220648-2025-31685 | 23146595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23146595&pc=PYX | 03/12/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 03/12/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025587561 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp flex pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal unreliable were observed. despite this issue, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the failure mode 'optical signal issue' was changed to 'placement signal issue' against the dp sensor. no product was returned for analysis. the data logs show that 7 hours after start, the placement signal begins to drift down and the flows drift up before the placement signal goes out of range and the flow calculation is disabled. the root cause of the placement signal issue is due to a dp sensor drift as evidenced by log pattern. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5077 | 1220648-2025-46850 | 23147438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23147438&pc=OZD | 07/31/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698591 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 67 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp endured cardiogenic shock with a "possible" relationship to the impella device used: ¿upon arrival to the cicu, patient was lethargic. his lab were pertinent for a lactic acid of 7>7.9, given worsening lactate, he was started on dobutamine¿. the patient recovered with no sequelae. additionally, the patient endured a right common femoral artery (cfa) thrombus with a "possible" relationship to the impella device used: ¿we then performed right cfa angiogram through 4f microsheath in the cfa which revealed mobile clot in the right cfa extending into the right profunda and superficial femoral artery. subsequent thrombectomy was performed via left femoral access (vascular surgery assisting) using pounce thrombectomy system. successful thrombectomy x2. distal runoff of entire superficial femoral artery with patent vessels down to plantar arch. palpable dorsalis pedis and posterior tibial pulses post-thrombectomy. impella was eventually removed and pt found to have r femoral artery clot, for which vascular surgery did a thrombectomy and patient was started on lih.¿. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | G.2 added selection that was inadvertently omitted from the initial report that was submitted. the root cause of the cardiogenic shock/thrombosis was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
5078 | 1220648-2025-46851 | 23148039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23148039&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698839 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was placed to support a patient admitted for high-risk percutaneous coronary intervention. during implantation, the patient experienced ventricular tachycardia requiring defibrillation. repositioning of the pump was attempted under fluoroscopy but the pump remained under the papillary muscle. later, an echocardiogram confirmed the device was in the correct position and the case continued. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the positioning issue was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
5079 | 1220648-2025-46853 | 23148816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23148816&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704335 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed bleeding at the access site. the patient was transfused six units of packed red blood cells and a femstop was applied to obtain hemostasis, and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned. major bleed: the cause of major bleed cannot be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks.
5080 | 1220648-2025-46854 | 23148843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23148843&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653470 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during attempted implantation of an impella 5.5 the pump could not advance through the vasculature. several unsuccessful attempts were made before implantation was aborted. an impella cp was implanted instead. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ the investigation has been completed. no product was returned for investigation. the cause of the delivery issue is determined to be patient condition because the impella 5.5 was unable to be delivered due to anatomical challenges and decided to place an impella cp. the device history review was completed and the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5081 | 1220648-2025-46856 | 23149059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149059&pc=OZD | 08/27/2025 | 09/24/2025 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026707167 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella 5.5 placed to support the patient post coronary artery bypass graft procedure. during implant, the impella 5.5 could not be delivered due to patient anatomy so an impella cp was placed. reportable malfunction despite no harm or injury from delivery issue. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature. device history review: the device passed all the post sterile inspection checks.
5082 | 1220648-2025-46857 | 23149062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149062&pc=OZD | 08/25/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026719080 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 147 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp for the support of a post cardioverted acute myocardial infarction with cardiogenic shock patient. the patient was tachycardic and intervention is unknown. it was further reported that there was a placement signal issue with no harm to the patient. later it was reported that the patient expired with no additional information. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. clinical symptoms (tachycardia): the root cause of the injury was unable to be determined due to insufficient clinical details. optical signal issue: data log shows the signal not reliable (snr) was low about 500 at the beginning of the case and decreased to around 250 after the pump started. later during the case, the placement signal trended upward and was lost while running on a steady p-level, with snr reported as zero. lamp values were less than 100%. no known pump's optical failure was confirmed based on data log review. the cause of the optical signal issue was not determined since the product was not returned. device history review: the device passed all the post sterile inspection checks.
5083 | 1220648-2025-46859 | 23149078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149078&pc=OZD | 08/27/2025 | 10/30/2025 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026697086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, insertion of the long peel away was attempted but could not be inserted fully due to patient vascular anatomy, so the short peel away sheath was advanced. the impella was inserted and turned on but was unable to be passed distal to the aorto-iliac bifurcation. implantation of the pump was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the delivery issue was determined to be patient condition as the patient had difficult anatomy preventing successful delivery of impella. the device passed all post sterile inspection checks.
5084 | 1220648-2025-46860 | 23149080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149080&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693885 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the team declined coronary artery bypass graft (cabg) for the patient and the cp was placed for the hrpci. during the pci the patient arrested and had cpr (cardiopulmonary resuscitation) and was chemically coded and intubated. the pump supported for 4 days until patient made care and comfort only and expired. no allegation made against the pump causing the death. physician mentioned multiple times that this patient would not have made it off the table without impella . | [5 empty] | Correction d4. the investigation of the reported failure mode has been completed. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5085 | 1220648-2025-46852 | 23149081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149081&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705221 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 130 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Tachycardia | Ventricular Fibrillation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient in cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient developed a large hematoma from right ulnar catheter site. additionally, the patient experienced ventricular tachycardia and the pump generated diastolic suction alarms. later, the patient experienced ventricular fibrillation and expired while on impella support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | No product was returned. the cause of the hematoma cannot be determined since insufficient clinical details were provided. the root cause of the tachycardia and ventricular fibrillation was unable to be determined due to insufficient clinical details. the cause of pump suction was most likely patient condition related since suction was resolved by giving fluid boluses. the device passed all post sterile inspection checks.
5086 | 1220648-2025-46861 | 23149084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149084&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026706188 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when inserting the impella cp the catheter kinked, making it not possible to deliver the impella. the impella cp placement was aborted. there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had calcification and there was slight tortuosity preventing successful delivery of impella. product damage (catheter kink) : the cause of the injury was determined to be patient condition as the patient had calcification and there was slight tortuosity preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5087 | 1220648-2025-46862 | 23149099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149099&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026712002 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during implantation of an impella cp the catheter would not advance around the aortic arch. the pump was removed and new impella cp was placed successfully. no patient harm was reported. | [5 empty] | The pump was returned for investigation. no physical damage was noted on the returned pump. it was reported that the pump encountered heavy resistance at the aortic arch while being loaded on the 0.018 guidewire. the physician reported being unable to remove the wire, so both the device and the wire were pulled together. no defect was found on the returned product the cause of the delivery issue was unable to be determined as insufficient clinical details were provided. the device passed all post sterile inspection checks. | B.7 additional information was added that was inadvertently omitted from the initial report that was submitted. d.3 manufacturer name should not of contained numbers behind it on the initial report that was submitted. d.9 the device was returned and the date was added. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 code a150207 was reported incorrectly on the initial report that was submitted. investigation codes were updated due to the device being returned. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5088 | 1220648-2025-46863 | 23149107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149107&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706484 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 59 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during support of the impella 5.5 the patient developed a clot in an unknown location that needed thrombectomy to remove it from the lower left extremity. the patient was later weaned successfully and the survived. | [5 empty] | Product complaint # (b)(4). the investigation for the reported limb ischemia and thrombosis has been completed. the root cause of the limb ischemia and thrombosis could not be determined. h6 type of investigation code b01 has been added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the pump and purge cassette returned for investigation and no issues were found with the returned product upon inspection. ischemia/ thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5089 | 1220648-2025-46864 | 23149118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149118&pc=OZD | 08/27/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713897 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 81 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported they encountered delivery issues during implant attempt of impella 5.5. the patient was taken to the operating room for 5.5 upgrade. however, due to patient's small anatomy, impella 5.5 would not pass despite exchanging for stiffer 0.025 platinum plus. as a result, the decision was made to place impella cp via axillary site. this case is for the second pump (5.5). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. delivery issue: the cause of the delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to small patient anatomy and decision to place axillary cp made. device history review: the device passed all the post sterile inspection checks.
5090 | 1220648-2025-46865 | 23149120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149120&pc=PYX | 08/27/2025 | 02/18/2026 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026693635 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella rp developed bleeding and a hematoma. the site dressing was changed, and the hematoma was mashed out. additionally, doppler pulses were lost in the lower extremities. the patient was treated with albumin and heparin gtt, and distal perfusion sheaths were placed. a fasciotomy was performed on the right lower extremity with partial success. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ the cause of the hematoma/bleeding is most likely use related since patient was moving and reported bleeding and a small hematoma to left groin. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
5091 | 1220648-2025-46866 | 23149125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23149125&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026711235 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp lost doppler pulses in the lower extremity. a fasciotomy was performed to the right lower extremity. additionally, the patient experienced atrial fibrillation with rapid ventricular response and ventricular tachycardia. the patient was treated with epinephrine. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. the root cause of the ischemia/tachycardia/paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks.
5092 | 1220648-2025-46846 | 23150387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23150387&pc=OZD | 08/27/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026613371 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 on (b)(6) 2025, there was an attempt made late last night to remove the device given the concern for sepsis that was identified. patient was moved to the operating room and intubated for airway protection. lines were initially removed but the patient did not tolerate wean of the device and it was ultimately kept in place. the patient remained intubated for most of the day and was extubated shortly before the current update. they are ultimately stable but given the septic risk the plan has changed to likely place a durable ventricular assist device for the patient. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. sepsis: the root cause of the sepsis was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5093 | 1220648-2025-31686 | 23151581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151581&pc=OZD | 11/02/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 11/02/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204377 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal not reliable alarm occurred on the automated impella controller. the alarm was disabled, and motor current was monitored. patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. product was returned. data logs confirmed the issue. while the failure mode symptoms seen in the field were not reproduced while running a known-good pump, the 5s parameters show a failing contrast. the root cause of the placement signal not reliable alarms and oscillating contrast was the defective optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5094 | 1220648-2025-31687 | 23151583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151583&pc=OZD | 11/02/2024 | 09/25/2025 |  | 09/25/2025 |  |  | 11/02/2024 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528916 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 was getting placement signal not reliable alarm. the alarm was disabled, and motor current has been monitored. patient expired while on support (this is reported on sister record for aic) the complainant reported that an impella 5.5 pump was implanted to support a patient . during support, placement signal unreliable were observed. despite this issue, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. there is no problem found against the pump. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5095 | 1220648-2025-46867 | 23151876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151876&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702343 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Priming Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered high purge pressure, pump stop and air in purge alarm during impella cp support. the physician confirmed a handling problem by the user. the purge system drew air during a transport and there was an ¿air in the purge system¿ alarm. after that the purge system stopped. this alarm was also only confirmed multiple times. the purge system was deaired. then, the pump stopped. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | G.1 added manufacturing site information as it was inadvertently omitted from the initial report that was submitted. h.6 code a1414 was entered incorrectly on the initial support that was submitted. | The device was discarded. the investigation was completed. after 1 day on support, purge pressure high and pump stopped. pp 1072mmhg, pf<1.0ml/h. a short time before there was an ¿air in the purge system¿ alarm and the system was deaired. md stated handling problem by the user. the purge system drew air during a transport. purge pressure high - data logs confirmed purge pressure high and blocked purge alarms. purge was not running for at least 90 minutes, purge pressure above 1000mmhg upon purge restart up, and pump stop 1 hour later. the cause of the purge pressure high was biomaterial build up that was caused by lack of purge flow from air in purge. pump stop - data log review showed pump stop without mc spike. purge was not running for at least 90 minutes, purge pressure above 1000mmhg upon purge restart up, and pump stop 1 hour later. the cause of the pump stop was biomaterial build up due to the gradual rise in mc after the purge issues occurred. air in purge system - data log review showed air in purge system alarms trigger a period of no purge flow. when alarms resolved after 90 minutes, purge pressure high alarms triggered. the cause of the air in purge system was an adverse event caused by the user due to the md noting a handling problem by the user. device history report was completed and passed all post sterile inspection checks.
5096 | 1220648-2025-46868 | 23151878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151878&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102527 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) alarmed continuously when plugged in. a different aic was used to support patient management. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5097 | 1220648-2025-46869 | 23151882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151882&pc=OZD | 08/27/2025 | 10/28/2025 |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026714591 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 86 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered placement signal issue and bleeding during impella cp support. during percutaneous coronary intervention (pci) procedure there was an optical sensor failure that was not resolved. pci of the left anterior descending artery was successful after 1 stent. this vessel supplied what turned out to be a chronic rca (right coronary artery). the rca was aborted and ultimately the drips were weaned off. upon field rep arrival there was some discussion on whether to remove or leave in impella due to some ooze post operation. ultimately since the electrocardiogram improved, impella was explanted and perclose was deployed. | [5 empty] | The investigation for the access site bleed and optical signal issue has been completed. data logs were reviewed. signal-to-noise ratio dropped to zero, contrast dropped, gain and lamp stepped up, and light remained relatively stable. these patterns in the optical signal metrics are indicative of sensor damage. returned product was inspected and there were no visual abnormalities such as damaged pins or catheter kinks. the pump's 5s parameters were all still out of specification when connected to a console, but the pump passed laser continuity. the optical sensor face was removed from the pump, and imaging confirmed evidence of a crack on the surface. this aligns with the data log review. based on the clinical details that a percutaneous coronary intervention procedure was being done, the patterns in the optical 5s parameters on data logs, and the damage seen to the sensor face on returned product, the cause of the optical sensor issue was determined to be a damaged sensor. the root cause of the access site bleed and optical signal issue could not be determined due to lack of clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5098 | 1220648-2025-46870 | 23151884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151884&pc=OZD | 08/27/2025 |  |  | 09/25/2025 |  |  | 08/27/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676993 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 83 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient implanted with an impella 5.5 had a fever and oozing from the access site. systemic anticoagulation was altered and the patient was transfused two units of platelets and one unit of packed red blood cells to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the cause of the bleeding is most likely patient condition since all access sites are oozy including the impella axillary site. the cause of the fever was not determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
5099 | 1220648-2025-46871 | 23151885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151885&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627712 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 102 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced arrhythmias and pump suction and placement alarms. proper pump placement was verified via echo. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A.4 added as it was inadvertently omitted from the initial report that was submitted. b.1 add selection as it was inadvertently omitted from the initial report that was submitted. b.2 required intervention was entered incorrectly on the initial report that was submitted. a new selection was added. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. the device history review was completed and the device passed all post sterile inspection checks. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no product was returned. suction and placement suction low alarms were present along with verified placement. after reviewing logs, suction alarms and trend in placement signal was observed. the cause of pump suction cannot be determined since insufficient clinical details were provided.
5100 | 1220648-2025-46872 | 23151887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151887&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716277 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support the patient admitted in for a cabg, with known cardiac and systemic comorbidities. the pump was placed but in the operating room when cannulas removed and surgical sites checked for bleeds the patient went into ventricular fibrillation and was reversed after 1 shock. impella dropped to p-3 and dobutamine to 1. urine was reported to be pink and more fluids were given. urine turned back to yellow by the end of the case. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. arrhythmia: the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5101 | 1220648-2025-46873 | 23151894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151894&pc=OZD | 08/28/2025 | 12/01/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635938 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient with acute myocardial infarction complicated by cardiogenic shock. the pump was placed but the access site was found to be bleeding. the team had to troubleshoot and intervene by the placement of sutures, gentle pressure, adjustment of the angle, and giving 2 units of rbc. the pump was explanted after 24 hours. patient survived event. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.4 added selection as it was omitted from the initial report that was submitted. g.1 revised manufacturing site name as it was entered incorrectly on the initial report that was submitted. added manufacturer site address street line 1, city, postal code, fax number and country code as they were omitted from the initial report that was submitted. h.11 added instructions for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ investigation summary: the investigation has been completed. no product was returned for investigation. major bleed: bleeding from right femoral access site was noted. the cause of the bleeding is determined to be use issue because hemostasis was achieved by manipulation of catheter at site, placing additional sutures and gentle compression were noted. device history review: the device passed all the post sterile inspection checks.
5102 | 1220648-2025-46874 | 23151901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151901&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710387 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 141 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support the patient admitted in with a non-st-segment elevation myocardial infarction and taken to the operating room for coronary artery bypass graft and a bentall procedure. the pump was supporting when the patient went into ventricular fibrillation that required shocks. the team confirmed by echocardiogram that pump was appropriately positioned and the team discussed a rp flex but, the patient expired before that right sided support could be placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. arrhythmia: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5103 | 1220648-2025-46875 | 23151918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151918&pc=OZD | 08/25/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662776 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Embolism/Embolus | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon explantation of an impella 5.5 a thrombus was found in the inominant. the patient survived. | The complainant had an impella 5.5 pump placed to support a patient admitted in cardiogenic shock/acute myocardial infarction. the pump was supporting while the patient awaited a heart transplant. the pump was explanted at the time of transplant on day 10. after explant there was an embolic cerebrovascular accident that the team presumed to be from the impella. the cerebrovascular accident was at a site of a prior stroke. the patient survived. |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 revised as an incorrect selection was made on the initial report that was submitted. b.5 information was added as previously entered information was reported incorrectly on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. stroke : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. embolism : the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks.
5104 | 1220648-2025-46876 | 23151923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151923&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716024 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had oozing from the impella site. upon completion of the case, a hematoma was noted in the right groin. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 required intervention was incorrectly selected on the initial report that was submitted. a new selection was added. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. hematoma: the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. major bleed: oozing around the insertion site was noted. all best practices were performed and the impella repositioning sheath was inserted with correct angle matching technique. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
5105 | 1220648-2025-46877 | 23151928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151928&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1882177 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant was attempting to transfer the patient from automated impella controller (aic) to aic after being sent from community to tertiary center. during attempted bag & cassette change the screen froze with ¿please wait¿ and the progress bar not progressing. after a period of time the aic would begin to alarm ¿please complete¿ however the screen remained frozen in this alarm state. aic 12058 was powered off and re-started in between the two separate attempts to complete this transfer with the same result. after the failed second attempt, transport was completed with the original aic that case support had started with. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5106 | 1220648-2025-46879 | 23151932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151932&pc=OZD | 08/28/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698031 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 92 KG |  |  | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed for the support of a patient who had arrested/ventricular tachycardia in the community before being brought to the hospital. while on the support there was alarms for ps and "in aorta" but the pump was observed to be in the left ventricle. this did not prompt explant. the pump was also showing suction alarms while on extracorporeal membrane oxygenation as well. the patient condition deteriorated, blood was given and an abdominal bleed was suspected. after just 6 hours the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned. data log review showed low placement signal (ps) and suction alarms occurring above p-2. no motor current (mc) spikes noted. ps trended down over the course of the case. optical signal issue - the cause of the optical issue was patient condition related due to the trending down of the ps as well as the noted decompensation of the patient. pump suction - the cause of the pump suction was patient condition related due to a lack of volume. the lack of volume is a combination of parallel extracorporeal membrane oxygenation (ecmo) support, noted abdominal bleed, and patient decompensation. patient code - bleeding: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5107 | 1220648-2025-46878 | 23151933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151933&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026651691 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support the patient admitted in with st-segment elevation myocardial infarction. the pump was placed and the patient transferred to the tertiary center for escalation and coronary artery bypass graft (cabg). the patient had ischemia at the access site which was relieved by removal of the 14fr sheath and repositioning the pump. the cabg was ongoing and thrombus was seen in the aorta that could have been from pump. when the team tried to restart after surgical mode there was a failure to start, but they were able to restart and support proceeded until pump successfully weaned. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5108 | 1220648-2025-46881 | 23151936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151936&pc=OZD | 08/28/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723821 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 71 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp placed to support the patient in inferior st-segment elevation myocardial infarction. while on support, the patient experienced ventricular fibrillation arrest after a stent and the patient received four rounds of cpr. return of spontaneous circulation was achieved. the case continued. care was withdrawn and the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5109 | 1220648-2025-46882 | 23151938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151938&pc=OZD | 08/28/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703470 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella 5.5 was attempted to be placed, but unable due to the patient anatomy so 5.5 insertion was aborted. the impella cp was successfully placed axillary. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of delivery issue is determined to be patient condition because 5.5 was unable to pass through the anatomy and insertion aborted. cp successfully placed axillary. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5110 | 1220648-2025-46883 | 23151939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151939&pc=OZD | 08/11/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710517 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was a gastrointestinal consult for hematochezia. there were suction events related to low preload. resuscitation efforts were noted. a new m1 stroke occurred; the patient was taken to neuro interventional radiology for thrombectomy. the purge retainer began to leak. the decision was made to explant the pump. the patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation was completed. thrombosis/stroke: on (b)(6) 2025, it was reported that the patient had an embolic stroke. they had a successful thrombectomy to manage the issue. the root cause of the injuries were unable to be determined due to insufficient clinical details. major bleed: it was reported on the first day of support that gi was consulted for hematochezia. no further details were provided. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. pump suction: it was reported that the pump was experiencing suction on (b)(6) 2025. it was believed to be related to preload, so resuscitation efforts were made. no further details were provided. data logs were reviewed, and the suction conditions on the first day of support were confirmed. there were many triggers of the suction alarm #14 (low pulsatility algorithm) as well as position unknown alarm #33. a closer look at the lt log during this period confirmed low pulsatility in the ps, which improved with time. the rt log of a specific suction trigger suggested that the alarms were triggering due to a combination of a trend down in mcmax values, as well as relatively low ps values. there was a mc spike with ms deviation on (b)(6), however, it was unlikely related to the low flows considering that suction alarms resolved between the evening of (b)(6) and the afternoon of (b)(6). product was not returned for investigation. based on the clinical details provided that low preload was suspected at the time of suction alarms and data log review confirmed suction conditions and low pulsatility of mc/ps, the cause of the low pump flows was determined to be patient condition related. purge leak - pump: it was reported that a leak was noted from the purge sidearm retainer on (b)(6) 2025. all purge related metrics and pump performance remained unchanged, but the decision was made to explant the pump regardless. data logs were reviewed and there were no relevant purge alarms. also, purge pressure/flows were generally stable throughout the case until (b)(6). at this time, there appeared to be the start of a slight trend up in purge flows and down in pressures. however, the values remained within the expected range. product was not returned for investigation. since product wasn't returned, the origin of the purge leak could not be determined. however, based on the clinical details provided that the leak was noted to be from the purge sidearm retainer, the cause of the purge leak was determined to be a damaged sidearm. the device history review was completed and passed all post-sterile inspection checks.
5111 | 1220648-2025-46885 | 23151942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151942&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026705310 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Unspecified Infection | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support placed for the patient admitted in with prior coronary artery bypass graft and st-segment elevation myocardial infarction. the pump was placed and was supporting when they observed was suction and poor pump positioning. the patient experienced tachycardia with the position. the pump was repositioned and tachycardia resolved. while on the support the patient had an infection presumed to be from the pacer, this prompted the pacer removal and temporary new pacer placed. a hematoma was reported at the old pacer site and (b)(4) units of red blood cells were given. the pump was later explanted and the patient survived. | The pump was explanted on day 16 with placement signal loss. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information: b5, h6. investigation of the reported failure mode has been completed. the pump was not returned for investigation. data log shows several placement signal (ps) low alarms, with ps readings dropping close to zero without significantly impacting optical signal parameters or pump performance. several suction alarms were present at that time. as per the clinical details, there was some unconfirmed positioning issue during the placement signal and suction events, where the pump was found to be deep and was also found to have been pulled into the aorta during repositioning. optical sensor issue: the cause of the optical signal issue was not determined, as the product was not returned and sufficient clinical information was not provided. pump suction: the cause of the suction issue was not determined, as the product was not returned and sufficient clinical information was not provided. device in wrong position: the cause of the positioning issue was not determined, as the product was not returned and sufficient clinical information was not provided. clinical symptoms (tachycardia, unspecified infection): the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms(hematoma): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5112 | 1220648-2025-46886 | 23151945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151945&pc=OZD | 07/01/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695269 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 47 YR |  | Female | 83 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed's study database that a patient implanted with an impella cp experienced ventricular arrhythmia. synchronized cardioversion was performed with mild improvement, and an amio bolus was transfused. the patient then experienced an atrial flutter. again, amio was transfused and the patient was converted back to normal sinus rhythm. later, the pump was explanted and the patient survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. the root cause of the arrhythmia/paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks.
5113 | 1220648-2025-46887 | 23151950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151950&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703635 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 75 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding and a hematoma at the impella site. manual pressure was held and a suture was placed to obtain hemostasis. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter phone number as it was inadvertently was omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. hematoma/ major bleed: the cause of the bleeding is determined to be a use issue because the bleeding was controlled by placing a suture and achieved hemostasis. device history review: the device passed all post sterile inspection checks.
5114 | 1220648-2025-46888 | 23151953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151953&pc=OZD | 08/28/2025 |  |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703372 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced heparin-induced thrombocytopenia. no intervention was noted. the pump was explanted a day later, and the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.2 required intervention was previously selected incorrectly on the initial report that was submitted. a new selection was added. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the investigation has been completed. no product was returned for investigation. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks.
5115 | 1220648-2025-46889 | 23151957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151957&pc=OZD | 08/28/2025 | 09/25/2025 |  | 09/25/2025 |  |  | 08/28/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026723098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the there was high purge pressure and a kink during impella cp with smart assist support. it was reported that impella purge blocked alarms were encountered at bedside and that product was kinked due to a turn. there were no further impella related concerns or alarms. the patient remained stable throughout support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the pump was not returned for investigation. data log shows a sudden rise in purge pressure with a noted drop in purge flow, followed by a purge system block alarm. the purge pressure then suddenly returned to normal within two minutes, indicating a possible purge line kink during the case. clinical details also confirm that the kink was due to patient movement (turning) during the purge system block alarm. high purge pressure: the cause of the high purge pressure issue was most likely a kinked purge line due to patient movement, based on data log review and provided clinical details; however, damage to the purge lumen could not be verified without the returned product. product damage (purge tubing kink): the cause of the purge tubing kink was most likely related to patient movement; however, damage to the purge lumen could not be verified without the returned product. device history review: the device passed all the post sterile inspection checks.
5116 | 1220648-2025-46890 | 23151960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151960&pc=OZD | 08/29/2025 | 10/07/2025 |  | 09/25/2025 |  |  | 08/29/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026647770 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported patient had a 5.5 pump support ongoing for the patient admitted in with cgs and cardiomyopathy. the pump was supporting when after 8 days there was a rise in the purge/purge system blocked which prompted the addition of lysis/tpa. this troubleshooting did not resolve the issue and the pump stopped. the pump was explanted after the 8 days and was not replaced. | [5 empty] | Correction d6b and e4. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. the pump plus the pc was returned for evaluation. data logs were not recovered. the pump was returned cut. the returned pump was found to have biomaterial wrapped around the impellor blade and in the purge lumen. pump stop - the cause of the pump stop was blood ingress into the motor as the purge line was found to have blood in it near the motor housing. purge system blocked - the cause of the purge system blocked could not be determined as the pump was returned cut and reproduction testing could not be completed, the data logs were not recovered, and insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5117 | 1220648-2025-46891 | 23151962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23151962&pc=OZD | 08/29/2025 |  |  | 09/25/2025 |  |  | 08/29/2025 | 09/25/2025 |  |  |  |  | 09/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026693836 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 77 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted to support a patient admitted with extensive cardiac history. a high purge pressure alarm occurred which resulted in adding tissue plasminogen activator to the purge. there was also a suction alarm which was treated by adding volume. ultimately, the family withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for evaluation. the cause of the high purge pressure was due to biomaterial ingestion based on returned data log analysis and clinical details. the cause of the suction alarms could not be determined as limited clinical details were provided. the pump passed all post-sterile inspection checks.
5118 | 1220648-2025-46793 | 23131819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23131819&pc=OZD | 08/25/2025 | 09/23/2025 |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699727 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 94 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient with ventricular fibrillation (vf) arrest and cardiac history who was brought into the hospital in acute myocardial infarction complicated by cardiogenic shock. the patient was sent to the operating room (or) for coronary artery bypass graft and the patient vf arrested in the operating room again and was put on impella cp support. after 25 hours, the pump was explanted due to limb ischemia. they were unable to dopple distal pulses. the pump was explanted and computed tomography imaging was ordered post impella cp removal. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia: the root cause of the ischemia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5119 | 1220648-2025-46794 | 23131882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23131882&pc=OZD | 08/25/2025 | 09/23/2025 |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026699468 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp placed for a high-risk percutaneous coronary intervention procedure. the pump had a high purge pressure alarm. no kinks were noted in the purge line. the alarm resolved itself after around 30 minutes with no intervention done to fix the issue. no issues with pump flow or motor current were noted. the high purge pressure did not prompt the device to be explanted. the patient survived and the pump was discarded post procedure. | [5 empty] | The investigation of the reported failure mode has been completed. no product returned. purge pressure high (hpp) - data log review showed purge pressure ramp up over the 20 minutes after pump start, briefly recovering, before ramping back to hpp alarms over 8 minutes. no motor current (mc) spike seen with pump and purge flows also decreasing at that time. the cause of the purge pressure high alarms was biomaterial build up due to the 20 minute purge pressure build up with lowering pump and purge flows. clinical details note no heparin was used. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5120 | 1220648-2025-46797 | 23132216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23132216&pc=OZD | 08/25/2025 | 09/23/2025 |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 94 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient admitted in for high-risk percutaneous coronary intervention. while on support, the patient was sent to the intensive care unit for recovery and groin management due to a hematoma. the case continued. the device was explanted and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5121 | 1220648-2025-46798 | 23132568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23132568&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694740 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. prior to placement of the impella peel-away sheath, two precloses were placed. while sequentially dilating, the 12 french dilator apparently caught on one of the perclose sutures. the physician continued to place the peel-away sheath and impella without issue occurring from the pump, but hematoma formation occurred. while on support, the patient developed a hematoma and was transferred to an operating room (or) to urgently evacuate the hematoma and repair vessel bovine patch angioplasty. while in the or, the patient received one unit of packed red blood cells. the patient peel-away sheath was removed in or and hemostatic control via vessel loops in surgical cutdown was noted. the patient was stable while exiting the or. the case continued. the device was explanted and the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. hematoma : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5122 | 1220648-2025-46795 | 23132596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23132596&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716257 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while rounding, the patient was found to be in supraventricular tachycardia (svt). the patient was awake and alert with no reports of palpitations. an amiodarone bolus and drip were ordered. the patient was intermittently self converting back to a baseline heart rate in the low 100's. it was reported that dobutamine was decreased this morning. additionally during the support, there was epistaxis that prompted anticoagulation pause. the pump remained on while awaiting a left ventricular assist device. additional information was received. the svt was treated with cardioversion from an external defibrillator and intravenous amiodarone. the plan is for the patient to go for a lvad. it is believed that the patient has a clotting disorder that they attributed to the epistaxis. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. arrhythmia: i the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. epistaxis: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5123 | 1220648-2025-46799 | 23132892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23132892&pc=OZD | 08/25/2025 | 11/06/2025 |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695573 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 105 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received: the patient had hematuria as the urine test was positive for red blood cells (rbcs). | The user facility reported that a patient on impella cp support had red color urine. the nurse was to send a urine analysis and a lactate dehydrogenase test to help determine the cause. |  |  |  |  | Following field update with additional information: b5. the investigation of the reported failure mode has been completed. peri-procedural adverse event(hematuria): the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5124 | 1220648-2025-31652 | 23133024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133024&pc=OZD | 03/04/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 03/04/2024 | 09/24/2025 |  |  | 12/31/2025 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438131 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. the pump was supporting along with protek duo when after 5 days there was signal loss. the placement signal not reliable did not prompt the team to explant the pump. pump remained on for support until the patient was successfully weaned and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. upon review of the data on the device it is apparent that the console is the potential cause of the issue. please refer to the complaint (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5125 | 1220648-2025-46800 | 23133035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133035&pc=OZD | 08/25/2025 | 10/24/2025 |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633161 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that at the beginning of the impella cp procedure there was an impella failure. the easy guide lumen was left in place and the pump started out of the body which lead to an immediate pump stop. the impella was exchanged. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. clinical details report that the pump was started outside the body with the red easyguide lumen still in place. the case was aborted. data logs were reviewed, and the pump stop was confirmed. 2 minutes after the data log started, the pump was attempted to be started and a pump stop alarm #115 triggered. at this time, the ps was reading 0 mmhg, indicating the pump was still outside the body. there were no attempts to restart, and the pump was disconnected. the pump was returned with the red easyguide lumen still in the cannula. closer visual inspection confirmed a deformation/kink of the red lumen near the outflow cage, further supporting that the lumen was still in the pump when it was started. the pump passed electrical resistance testing of the motor cables. the lumen was removed, and the pump was able to be started without an issue in a bucket of water. based on the clinical details provided, data log review, and returned product, the cause of the pump stop was determined to be a user error.
5126 | 1220648-2025-46801 | 23133242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133242&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599653 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that an impella cp was inserted without issues and percutaneous coronary intervention was completed. there was no flow to distal limb. external femorofemoral bypass was placed for distal limb perfusion. the patient was sent to the intensive care unit (icu) in grave condition. upon arrival at the icu, the patient went into uncontrolled ventricular tachycardia followed by ventricular fibrillation. the family had made the patient do not resuscitate and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. additional information was received. the doctor placed an antegrade sheath on the right femoral side and retrograde sheath on the left side. the device is not available to be returned as it was discarded. | [5 empty] | The investigation of the reported failure mode has been completed. ischemia/ tachycardia/ ventricular fibrillation the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5127 | 1220648-2025-46802 | 23133305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133305&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703558 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 113 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. at some point during the intervention, the patient had an episode of ventricular fibrillation and was shocked once but never lost consciousness. percutaneous coronary intervention (pci ) was completed and the impella was weaned and removed as planned. the patient is stable and survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5128 | 1220648-2025-46803 | 23133536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133536&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710494 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient experienced a cardiac arrest that was witnessed by a bystander. cardiopulmonary resuscitation was administered for 30 minutes before return of spontaneous circulation was achieved. the patient was brought to the hospital and placed on veno-arterial extracorporeal membrane oxygenation. the patient was taken to the cardiac catheterization laboratory where the left anterior descending coronary artery spontaneously dissected. a stent was placed, and the patient was brought to the operating room for impella 5.5 placement and an additional venous return line was placed for the veno-arterial-venous ecmo (vav ecmo). after the patient was transferred back to the intensive care unit on vav ecmo and impella 5.5 support, the patient required additional volume, crystalloid, and blood. it was noted that the patient was bleeding significantly from the back of their head, and as a result systemic anticoagulation was withheld. the device was subsequently weaned and explanted successfully. the explant was not performed earlier than planned as a result of the complication. the patient¿s arterial cannula was removed, and the patient remained on veno-veno extracorporeal membrane oxygenation support. additional information was received. the product is not available to be returned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5129 | 1220648-2025-46804 | 23133615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133615&pc=OZD | 08/22/2025 |  |  | 09/24/2025 |  |  | 08/26/2025 | 09/24/2025 |  |  | 10/31/2025 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024414150 |  | 000000000001000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an initial impella 5.5 pump placed to support a patient in heart failure. later, they had to redo mitral valve replacement and coronary artery bypass graft. at that time, the impella 5.5 was replaced with another impella 5.5. while on support on the second impella 5.5, there was an infection at the insertion site which was treated with antibiotics. later, it was reported that liver failure worsened and the patient expired on support. no further information provided. | [5 empty] | Correction: b2 date of death and d6b explant date. the investigation of the reported failure mode has been completed. infection : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5130 | 1220648-2025-46805 | 23133758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133758&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708096 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female | 62 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient experienced impella access complications noted as bleeding and oozing at the right groin insertion site. the physician tightened the perclose and re-sutured the site in the catheter lab. packed red blood cells (prbc's) and platelets were required for the bleeding and oozing, despite activated clotting time of 363. manual pressure was held. on day 2 of support, it was noted that the patient required an additional three units of prbc's due to groin ooziness overnight. the site was re-sutured again, reducing the ooziness. while the complication was not successfully managed, it did not lead to pump removal. the patient is stable post event. | [5 empty] | The investigation of the reported failure mode has been completed. major bleed: the cause of the injury is most likely use related since the physician tightened the perclose and re-sutured the site and the ooziness subsided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5131 | 1220648-2025-46806 | 23133986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23133986&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  | 10/31/2025 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423610 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp and experienced an optical sensor error. the pump was disconnected and reconnected as directed. however, the placement signal showed 0/0/0. no placement waveform was displayed on the automated impella controller (aic). the wire was removed and pump was placed in auto. there were no aorta/left ventricle or pump flows noted. the decision was made to keep the pump in place and proceed with the procedure. the procedure was completed successfully and the pump was removed without issues afterwards. this report represents the pump. another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | Due to a lack of product or data logs returned, the root cause of the optical signal issue cannot be determined. the impella pump passed all post sterile inspection checks.
5132 | 1220648-2025-46807 | 23134172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23134172&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1731304 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp and experienced an optical sensor error. the pump was disconnected and reconnected as directed. however, the placement signal showed 0/0/0. no placement waveform was displayed on the automated impella controller (aic). the wire was removed and the pump was placed in auto. there were no aorta/left ventricle or pump flows were noted. the decision was made to keep the pump in place and proceed with the procedure. the procedure was completed successfully and the pump was removed without issues afterwards. this report represents the automated impella controller. another report was submitted for the pump. refer to section h.10 for the related report number. | [5 empty] | A.4 is unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5133 | 1220648-2025-46808 | 23134393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23134393&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720627 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 81 KG | Non Hispanic | White | Death | Yes | Yes | Crack | Difficult to Open or Close | [5 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction/cardiogenic shock was placed on impella cp support. the team in the intensive care unit noticed that the tuohy borst was broken. a technician came and re-tightened it and it appeared to be locked down. the patient continued on impella support, but a decision was made to withdraw care and the patient expired. | [5 empty] | H.6 added medical device problem code. the investigation has been completed. difficult to open/close catheter anchor or touhy: clinical details reported the tuohy borst was broken. tech re tightened and appeared to be locked down now but there was fear to twist any further. the cause of the issue was not established as no product was returned and limited information was provided. product damage (pump crack): the intensive care unit team said the tuohy borst was broken. the cause of the issue was not established as no product was returned and limited information was provided. hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5134 | 1220648-2025-46811 | 23134534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23134534&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679916 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 105 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Clinical details noted while on support, patient experienced pink urine and slight drop in hemoglobin. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the physician called with a new placement signal not reliable alarm. upon review on impella connect, the placement signal waveform was still present on the automated impella controller (aic) and looked to be in good condition. the customer service center was called who also agreed that the waveform still looked good. they suggested pushing in the white plug in to the aic to see if that would change it but it did not. therefore, the audio for the alarm was disabled. additionally, initially a hematoma was present and then blood started oozing from the groin site. pressure was given and heparin was held. however, ultimately a suture was placed around the repositioning sheath and angle matching was corrected to better reflect a steeper angle. |  |  |  |  | Investigation summary: no product was returned for evaluation. only event logs were retuned. hematuria: the cause of the injury was not determined due to insufficient clinical details. access site adverse event: the cause of the bleeding at access site was user related as suture and angle matching not initially applied helped resolve bleeding per clinical details. optical signal issue. data logs were reviewed and lt logs showed "placement signal not reliable" alarms #1036 and #130 on first day of pump support. alarms resolved within 24 hours of triggering. no sustained alarms noted. no hard failures of the optical parameters were observed. upon start-up of pump, snr drifts down over 10 hours before stabilizing for remainder of case. contrast follows similar pattern with drifting up over same time period before stabilizing for remainder of case. additionally, placement signal drifts up then drops suddenly to within normal limits and remains stable for remainder of case. the cause of the optical signal issue could not be definitively determined as no hard failures of the optical sensor were observed and no product was returned for evaluation. device history review: per qn #(b)(4), pump sn (b)(6) from lot #1915040 did not pass post-sterile inspection due to product label legibility. pump sn was reworked and re-inspected to pass post-sterile inspection. per qn #(b)(4), pump sn (b)(6) from lot #1915040 did not pass post-sterile inspection due to hair in packaging, pump was reworked and sent for re-sterilization. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5135 | 1220648-2025-31653 | 23134879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23134879&pc=PYX | 07/09/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 07/09/2024 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025480454 |  | 1000323 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella rp along with an impella cp and extracorporeal membrane oxygenation (ecmo) as bipella support for a patient experiencing heart failure. the rp flex was to replace the ecmo. after four days of the support, the optical signal was lost on the rp, but this did not cause the team to explant the pump. the pump remained on for support until the patient was successfully weaned and the device explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. during multiple placement signal not reliable events, the data log shows several motor current spikes and suction, which resulted in negative cvp values and out-of-range optical signal values. the cause of the optical signal issue was most likely biomaterial. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5136 | 1220648-2025-30747 | 23134994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23134994&pc=OZD | 07/02/2025 | 07/28/2025 |  | 09/24/2025 |  |  | 07/02/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610584 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that multiple complications were encountered during impella 5.5 support. eleven days into support, the patient experienced some runs of ventricular tachycardia. magnesium was given to help treat the condition. six days later, some oozing was noted at the impella access site. the issue resolved with no noted intervention required. after another eleven days of support, the staff reported that they believed the placement signal reading was not accurate. they expressed no concerns over the discrepancy and continued support. roughly four days later, the staff noted that the pump was too deep but they were unable to pull the device back. no hemolysis was noted at the time and support continued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5137 | 1220648-2025-31654 | 23135146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23135146&pc=PYX | 01/31/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 01/31/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025600627 |  | 1000323 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an impella rp was implanted in a patient due to heart failure. the patient was transferred into the tertiary center for care escalation to an intra-aortic balloon pump. the rp pump was supporting when after three days there was optical signal loss. the pump remained on for support despite the signal loss. the patient was then successfully weaned from the pump and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log indicates several motor current spikes during the case run, which coincided with impacted placement signals and pump flow. also, 2 hours after the motor current spike event, where placement signal not reliable (psnr) 1051 was triggered, the placement signal was drifting downward, while the pump flow was drifting upward with stable cvp. none of the 5s optical parameters were inspected during the case run. this trend was seen multiple times, returning to normal. the optical signal issue failure mode was change to placement signal issue since the dp sensor was responsible for the reported psnr issue. the cause of the optical signal issue was most likely biomaterial, based on data log review. all pertinent information available to abiomed has been submitted at this time.
5138 | 1220648-2025-46813 | 23135206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23135206&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702648 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 74 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was noted as high risk percutaneous coronary intervention was planed to be implanted with an impella cp for mechanical circulatory support. during impella cp implant, the surgeon placed a 14 french sheath. however, while trying to advance the impella cp, it would not pass through the iliacs. the decision was made to abort the impella left femoral artery which was also heavily calcified. they proceeded with coronary stenting. the sheath was pulled at the end of the case and closed with perclose. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. | The investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of delivery issue is determined to be patient condition sue to calcification in patient's iliac vessels and the pump was unable to be delivered and the insertion was aborted. device history review: the device passed all the post sterile inspection checks.
5139 | 1220648-2025-46812 | 23135340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23135340&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703332 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 107 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | Hemodynamic instability | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during implant, the patient went in and out of ventricular tachycardia with loss of pulsatility multiple times. however, eventually this arrhythmia was not self-limiting and the patient had to be shocked. flows from impella dropped at that point and the patient was dependent on temporary pacer. with decreased flows, the patient became hemodynamically unstable and required cardiopulmonary resuscitation. in total, the patient received 20mg ivp epinephrine, six amps of bicarbonate, and three shocks. return of spontaneous circulation was attained. additionally, the patient was started on continuously renal replacement therapy. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist, section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was returned for investigation. renal failure/ hemodynamic instability: the cause of the clinical symptom was not determined due to insufficient clinical details. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5140 | 1220648-2025-46814 | 23135412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23135412&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725455 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 97 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient's urine turned red. the impella p-level was dropped. echocardiogram showed the impella in good position and not under any structures or up against the wall. the next day, the patient¿s urine color improved. the patient continued with impella support and the impella was explanted as planned. the patient is stable post explant. | [5 empty] | Product complaint # (b)(4). investigation summary - peri-procedural adverse event(hematuria): the cause of the injury was not determined due to insufficient clinical details. device history lot - device batch: 1963645. device history batch - subcomponent lot: na. device history review - the pump (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5141 | 1220648-2025-31655 | 23135535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23135535&pc=OZD | 01/09/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 01/09/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541856 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, the placement signal low alarm persisted throughout support. despite these issues, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. field data logs are unavailable. the pump was returned and evaluated. the 5s parameters of the pump showed no abnormalities. there was no ingested biomaterial around the impeller or near the sensor. the pump was run in the pressurized flow test loop for 4 hours. no abnormalities were seen with the 5s parameters during this test run. the root cause of the optical signal issue was not determined since field data logs are unavailable. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5142 | 1220648-2025-46815 | 23135588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23135588&pc=OZD | 08/25/2025 | 10/30/2025 |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026717612 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 54 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced a cerebral vascular accident, thrombus, and minor bleed. after implanting the pump, the patient was unable to move the left side of their body. one to two hours earlier, the patient was extubated, speaking and moving all extremities. a computed topography scan confirmed the patient had a stroke. a clot was surgically removed. a small bleed was also noted with no intervention specified for it. the physician was concerned and is uncertain about the reason the patient had a stroke. due to the increased risk of intracranial bleeding and need for heparin with the impella, the impella was successfully weaned from the pump and the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | No product was returned for investigation. the root cause of the thrombosis/stroke was unable to be determined due to insufficient clinical details. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks.
5143 | 1220648-2025-31658 | 23136331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136331&pc=OZD | 12/22/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 12/22/2024 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578598 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support had a placement signal alarm. impella was in good position. the patient had appropriate flow and placement signal audio was disabled. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. on the day of the complaint, log review showed the aortic placement signal adopt a widened waveform (looking ventricular) with impella position unknown alarms. shortly after, a placement signal adjustment was applied, which shifted the placement signal down below zero. this was followed by suction and then the placement signal monitoring was disabled. for the next two weeks, there are times that the placement signal intermittently goes in and out of range. gain, light, and lamp are stable throughout the case. contrast is low and the signal-to-noise ratio has a wide range with the minimum under spec. given the log review, it appears that the issue is not a hard sensor failure. however, the root cause of the optical signal issue related to the pump was not determined since product was not returned and insufficient clinical details were provided. its apparent that console may also be the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5144 | 1220648-2025-31656 | 23136365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136365&pc=OZD | 12/13/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 12/13/2024 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508138 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support. while on support, placement signal monitoring was going in and out, 3-point fixation was appropriate. there were no kinks or problems in the line, sometimes it¿s good and sometimes it¿s unreliably high. it was reported that it appears to be a wiring issue, not an optical sensor issue. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were provided. the root cause of the optical signal issue could not be determined since the product/logs were not returned and insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5145 | 1220648-2025-31659 | 23136479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136479&pc=PYX | 04/18/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 04/18/2024 | 09/24/2025 |  |  | 10/31/2025 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024418526 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal not reliable alarm was present after patient on impella rp flex support¿s bed was lifted. motor current and flows were without any issues. the pump was eventually weaned and explanted. there was no patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs show that 230 hours into support, the optical sensor begins to drift over a period of 6 hours before triggering placement signal not reliable (psnr) alarm and going out of spec (from around -40 to -90 mmhg). on the returned 5s parameters, lamp was at 100 and signal-to-noise ratio was low. light, contrast, and gain were in normal ranges. psnr alarm was reproduced when plugged into the lab console. the pump passed the laser continuity test. pressure testing in the uson at 50 mmhg for 50 minutes showed the optical sensor to be drifting at around the same rate as in the field, confirming the optical sensor issue. the root cause of the optical signal issue was most likely sensor silicone wear. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5146 | 1220648-2025-46796 | 23136514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136514&pc=OZD | 08/24/2025 |  |  | 09/24/2025 |  |  | 08/24/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682546 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 90 KG | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, there were active suction alarms. additionally, there was a hematoma to the right groin. an ultrasound was completed and there was no visualization of an active bleed. the patient received two units of packed red blood cells that night and an additional unit in the morning. it was noted that hemostasis was achieved through manual pressure. care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. hematoma: patient had hematoma in right groin. the cause of the hematoma was not determined due to insufficient clinical details. pump suction: no product was returned. patient had suction alarms. volume given. not resolved. after reviewing logs, suction alarms were observed. the cause of pump suction cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5147 | 1220648-2025-31657 | 23136567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136567&pc=OZD | 02/06/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 02/06/2024 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1446476 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) had an optical signal issue. the aic cable/plug were reported to be at fault. no further information was provided. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the automated impella controller issue was a defective blue receptacle. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5148 | 1220648-2025-46809 | 23136618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136618&pc=PYX | 07/11/2025 |  |  | 09/24/2025 |  |  | 07/11/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561594 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella rp and impella cp device for bipella mechanical circulatory support (mcs) and subsequently experienced complications requiring intervention. the patient would eventually expire. the patient had a complaint of general malaise and weakness for one week. the patient confessed running out of medications three to four days prior and also had chest pain. however, the patient did not seek help. the patient was admitted as st-segment elevation myocardial infarction and an impella rp flex was placed. the patient was began having some ventricular tachycardia. milrinone drip was stopped, and vasopressin drip was started. while on the unit, the patient continued to decompensate requiring additional pressors. the patient was taken back to the catheterization lab for diagnostic left heart catheterization and possible impella cp placement. the decision was made to place impella cp for additional support. levophed was given 33 mcg/kg/min, vasopressin: 0.04 u/min, epinephrine: 0.8 mcg/kg/min, and heparin: 800 u/hr. the 6fr sheath was exchanged for impella 14 fr sheath after serial dilation. the impella cp was advanced over the wire into left ventricle without complications and started in auto to 3.1 l/min flow. the patient was in atrial fibrillation and cardioversion was performed twice to normal sinus rhythm. impella peel-away sheath was removed; repositioning sheath was advanced. suction alarms sounded, pulsatility lost, and the patient's pressure dropped. an angiogram revealed the impella cp kinked at groin and in the left ventricle. the impella cp was pulled back and straightened without difficulty but remained in suction. repositioning was attempted several times. an echocardiogram showed adequate placement with no sign of effusion or any other problems causing suction. however, the decision was made to replace impella due to possible kink/damage from exchanging the repositioning sheath. impella cp was removed, and a new impella 14 fr sheath was advanced into left femoral artery to advance the second impella cp pump into left ventricle without any complications. the second impella cp pump was slowly ramped up to p-4 without any complications. however, there was suction alarm at performance level p-5 noted. echocardiogram showed a small left ventricle and dilated right ventricle. normal saline bolus for volume was started. the patient went asystole and cardiopulmonary resuscitation was started. however, despite all efforts, the patient was unable to survive this event. | [5 empty] | Medical safety officer reviewed the event and determined the following: regarding the impella rp, there is no reason to believe the arrhythmic event is related to the impella rp flex. the underlying condition of the patient appears to have contributed to this. however, dissociation cannot be made, and the arrhythmic event is being reported as a serious injury. there is an identifiable connection, as the impella rp device was present at the time of the arrhythmic event. during an unsuccessful attempt to place the first impella cp, the device kinked folded on itself, and the patient had an episode of atrial fibrillation. another cp was successfully implanted. regarding the replacement or second impella cp, there was suction event; the patient may have needed more support per the note. the patient went into asystole, and this, however, is unrelated to the impellas. conservatively the report is being submitted as death for the (second) impella cp. although, the patient's underlying condition is likely the reason for the patient going into asystole requiring cardiopulmonary resuscitation. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella rp flex report is one of three devices associated with the event story. refer to the related manufacturer report numbers as noted in section h10 for the medical device report on the first impella cp and the second impella cp, which is associated with the demise outcome. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding arrhythmic event, in order to make a root cause determination, all essential clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6: investigation: type, findings and conclusion codes and h6: component code were updated accordingly based on the completed investigation. related medical device reports: this impella rp flex report is one of three devices associated with the event story. refer to the related manufacturer report numbers as noted in section h10 for the medical device report on the first impella cp and the second impella cp, which is associated with the demise outcome.
5149 | 1220648-2025-46810 | 23136621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136621&pc=OZD | 07/12/2025 |  |  | 09/24/2025 |  |  | 07/12/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633340 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 67 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella rp and impella cp device for bipella mechanical circulatory support (mcs) and subsequently experienced complications requiring intervention. the patient would eventually expire. the patient had a complaint of general malaise and weakness for one week. the patient confessed running out of medications three to four days prior and also had chest pain. however, the patient did not seek help. the patient was admitted as st-segment elevation myocardial infarction and an impella rp flex was placed. the patient was began having some ventricular tachycardia. milrinone drip was stopped, and vasopressin drip was started. while on the unit, the patient continued to decompensate requiring additional pressors. the patient was taken back to the catheterization lab for diagnostic left heart catheterization and possible impella cp placement. the decision was made to place impella cp for additional support. levophed was given 33 mcg/kg/min, vasopressin: 0.04 u/min, epinephrine: 0.8 mcg/kg/min, and heparin: 800 u/hr. the 6fr sheath was exchanged for impella 14 fr sheath after serial dilation. the impella cp was advanced over the wire into left ventricle without complications and started in auto to 3.1 l/min flow. the patient was in atrial fibrillation and cardioversion was performed twice to normal sinus rhythm. impella peel-away sheath was removed; repositioning sheath was advanced. suction alarms sounded, pulsatility lost, and the patient's pressure dropped. an angiogram revealed the impella cp kinked at groin and in the left ventricle. the impella cp was pulled back and straightened without difficulty but remained in suction. repositioning was attempted several times. an echocardiogram showed adequate placement with no sign of effusion or any other problems causing suction. however, the decision was made to replace impella due to possible kink/damage from exchanging the repositioning sheath. impella cp was removed, and a new impella 14 fr sheath was advanced into left femoral artery to advance the second impella cp pump into left ventricle without any complications. the second impella cp pump was slowly ramped up to p-4 without any complications. however, there was suction alarm at performance level p-5 noted. echocardiogram showed a small left ventricle and dilated right ventricle. normal saline bolus for volume was started. the patient went asystole and cardiopulmonary resuscitation was started. however, despite all efforts, the patient was unable to survive this event. | [5 empty] | B7 added information that was omitted from the initial report that was submitted. e1 added zip code extension as it was omitted from the initial report that was submitted. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. investigation summary: hemodynamic instability: the cause of the injury was not determined due to insufficient clinical details. pump suction: clinical details reported that suction alarms occurred at p-5, and fluid was administered but was unsuccessful. the product was not returned for investigation. data log review confirmed the suction conditions throughout the case, as well as multiple "position in aorta" alarms about 23 minutes after starting the pump. placement signal started at very low magnitudes (roughly 30-40 mmhg) which triggered placement signal low alarms, but became more pulsatile around the same time that the positioning alarms started to occur. flow levels were lower than spec when placement signal was showing low magnitude/pulsatility. when position in aorta alarms started to trigger, placement signal increased in magnitude and pulsatility, and flows began to increase to spec. clinical details reported that at the time of suction alarms, echocardiogram showed small left ventricle (lv) and dilated right ventricle (rv), and normal saline (ns) bolus for volume was administered. the pump was turned off around 36 minutes in. the cause of the suction was most likely related to patient condition, as the data logs showed very low pulsatility and, clinical reported the patient having a small lv and dilated rv, and administering ns bolus for the patients volume issues. device history review: the pump passed all post sterile inspection criteria. | Medical safety officer reviewed the event and determined the following: regarding the impella rp, there is no reason to believe the arrhythmic event is related to the impella rp flex. the underlying condition of the patient appears to have contributed to this. however, dissociation cannot be made, and the arrhythmic event is being reported as a serious injury. there is an identifiable connection, as the impella rp device was present at the time of the arrhythmic event. during an unsuccessful attempt to place the first impella cp, the device kinked folded on itself, and the patient had an episode of atrial fibrillation. another cp was successfully implanted. regarding the replacement or second impella cp, there was suction event; the patient may have needed more support per the note. the patient went into asystole, and this, however, is unrelated to the impellas. conservatively the report is being submitted as death for the (second) impella cp. although, the patient's underlying condition is likely the reason for the patient going into asystole requiring cardiopulmonary resuscitation. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. related medical device reports: this impella cp report is one of three devices associated with the event story. refer to the related manufacturer report numbers as noted in section h10 for the medical device report on the impella rp flex and the first impella cp.
5150 | 1220648-2025-31660 | 23136631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136631&pc=OZD | 02/28/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 02/28/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612835 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 91 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella cp support. position was confirmed via echocardiography, and the alarm subsided at p7. support continued uninterrupted. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and evaluated. pump was returned for investigation. no physical abnormalities were observed. the pump passed the laser continuity test for the optical line. the pump read all 5s parameters within the range. the pump was further tested in a pressure environment to check if the optical sensor could hold the pressure, and it passed the pressure test. the pump was run per specifications in a bucket of water. data log shows the signal-to-noise ratio (snr) drop to below 2500 from 6000 after 30 minutes of case run while running on steady p-level p7, with a loss in placement signal and placement signal not reliable (psnr) alarm. this trend of snr and placement signal returned to normal at the end of the case while the pump was set to a lower p-level. imc logs were not available to review any other console-related abnormalities. no abnormalities were reported on the previous pump utilized on the same console. the cause of the optical signal issue was not determined since no abnormalities were reported on returned product and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5151 | 1220648-2025-46818 | 23136672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136672&pc=OZD | 08/25/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026674609 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the insertion site. manual pressure was held and a femstop was placed to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the bleeding was determined to be patient condition because of the poor native vasculature and unable to use perclose. the device passed all post sterile inspection checks.
5152 | 1220648-2025-46816 | 23136736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136736&pc=OZD | 07/16/2024 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505090 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening heart failure with an "unknown (undetermined)" relationship to the impella device used: palliative care was consulted. the patient was made do not resuscitate, but the family still hoped the patient would improve enough to get to percutaneous coronary intervention. the patient was weaned from the ventilator. the patient was also extubated but only tolerated for about two hours and then had worsening hemodynamics and respiratory status. the family was offered reintubation, but the family agreed to not reintubate and the patient passed away. | [5 empty] | The investigation of the reported failure mode has been completed. cardiac failure : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed.
5153 | 1220648-2025-46817 | 23136762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136762&pc=OZD | 08/23/2025 |  |  | 09/24/2025 |  |  | 08/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527639 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Device Contaminated During Manufacture or Shipping | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation. no patient harm was reported. | [5 empty] | The cause of the contamination was not established as no product was returned and insufficient details were provided. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. the pump passed all post sterile inspection checks | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5154 | 1220648-2025-46819 | 23136784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136784&pc=OZD | 08/26/2025 |  |  | 09/24/2025 |  |  | 08/26/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER U |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1301514 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a console swapped due to a reported "damage" in which there was unknown liquid flows out behind the purge cassette. the console was replaced without any harm or injury noted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5155 | 1220648-2025-46820 | 23136785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136785&pc=OZD | 08/26/2025 |  |  | 09/24/2025 |  |  | 08/26/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026692779 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient enduring ventricular tachycardia during impella 5.5 support. support proceeded, medication was administered, and the device was repositioned. | [5 empty] | D6b explant date updated. the investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5156 | 1220648-2025-46822 | 23136788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136788&pc=OZD | 08/26/2025 | 11/07/2025 |  | 09/24/2025 |  |  | 08/26/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026700316 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 112 KG |  |  | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Unspecified Infection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device and then escalated to an impella 5.5 device for mechanical circulatory support (mcs). during this time, the introducer appeared to be clotting off requiring intervention. also, subsequently the patient was suspected of having sepsis which was later unconfirmed. there was a "small infection" that was treated and later noted again, the patient experienced multiple frequent runs of torsades de pointes and would eventually expire. the patient was found to be in cardiogenic shock and was started on levophed and placed in bilevel positive airway pressure. catheterization was completed and diseased left main was noted. the decision was made to place an impella cp device, which was successfully completed and 17 days later the patient was transferred for higher level care. the decision was made to postpone escalation to an impella 5.5 some time after the percutaneous coronary intervention (pci). the patient was taken back to the catheterization lab for high-risk pci which was completed. two days later the optical sensor stopped giving a signal and the purge flow became blocked. a decision was made to explant the impella cp in the operating room with vascular cutdown and have cardiothoracic surgery on standby for an impella 5.5 placement if the patient required more support. it was noted the impella cp peel away sheath remaining in place and appeared to be clotted off. during explant and after, the patient did not require more support and vitals were stable. the decision was made not to place the impella 5.5 device, and the patient was extubated and sent to the intensive care unit. however, later this day, the patient decompensated requiring reintubation and high dose pressors. the impella 5.5 was emergently placed, and the patient was taken back to the unit on impella 5.5 mcs. the following day, it was noted the patient was possibly septic and had an increase in pressor dose. lactate climbed overnight to 3.2 but was now trending down at 2.2. the impella 5.5 device was running at p-9 and maintained at p-9 until pressor support could be weaned. the patient remained sedated due to agitation, but sedation was reduced as tolerated. discussions were made regarding strategies around revascularization of the left anterior descending artery if weaning is unsuccessful. impella support continued. four days later, however, it was noted that the patient had multiple frequent runs of torsades de pointes. pressors have continued to wean down. the patient was awake and able to talk this day. the following day, pressors continued to be weaned, and the patient was still having runs of torsades/ventricular tachycardia, although less frequent. the impella 5.5 device was reduced to support level p-8 and continued to wean as noted the following day. of note, there was no swan-ganz catheter in place, and the central line was not set to transduce. all labs remained normal, and the team was working toward left ventricular assist device work up. however, the following day, the patient was reintubated as every time the impella 5.5 device was weaned down the patient went into ventricular tachycardia. pressors were started after reintubation. the patient was kept sedated. the plan was to remain as is and try to extubate the following day if tolerated. over the next two days it was noted the impella ran at high levels with an attempt to wean pressors as tolerated. an attempt to clear the lungs was made to extubate. however, the patient¿s status remained unchanged. urine output was adequate, and it was noted the patient was administered one unit of packed red blood cells. now day 12 of support the patient was noted looking septic. this day the patient expired. due to the continuous decline, the family decided to withdraw all care. additional information was subsequently received and noted the following: the cause of death as not due to the pump. sepsis was not confirmed but rather a small infection that was treated. the cause, assumed, was the impella cp being placed with questionable sterility. and lastly, the arrhythmia presented when support was reduced. the cause of this was not stated and the reason for transfusion was not state. however, there was no continuous bleed. | The complainant had a impella cp pump placed to support the patient admitted in with cardiogenetic shock. the pump was placed and was supporting when the patient sat up in bed twice last night and broke their optical pressure sensor. the alarm type reported was a high purge pressure/purge system blocked. the purge system is now blocked and troubleshooting did not solve the issues. the impella cp was replaced with the impella 5.5 however during impella cp the peel away sheath still remaining in place and appearing to be clotted off. support continued on the impella 5.5 and the patient later expired (b)(4). |  |  |  |  | B5 event description, b6 test/lab data and b7 medical history were updated to include complete available details. based on the additional information received, h6 health effect - clinica/impact coding was updated/repopulated accordingly. medical safety review. medical safety reviewed the event. the first event describes access site thrombus requiring intervention. this is a serious injury attributed to the introducer. the event involving the impella cp was purge issues; the purge pressure rose, and the patient had sat upright and damaged the purge sensor. the optical signal was unreliable. the pump was explanted. during impella cp explant and after, the patient did not require more support and vitals were stable. the decision was made not to place the impella 5.5 device, and the patient was extubated and sent to the intensive care unit. however, later this day, the patient decompensated requiring reintubation and high dose pressors. the impella 5.5 was emergently placed, and the patient was taken back to the unit on impella 5.5 mechanical circulatory support. the following day, it was noted the patient was possibly septic and days thereafter the patient would experience runs of torsades de pointes. performance level was reduced and every time the device was weaned down the patient experienced the ventricular tachycardia. the cause was not provided, unnoted. additional information noted the sepsis was not confirmed. however, a small infection was treated and the cause, assumed, was the original impella cp being placed with questionable sterility. the patient condition remained unchanged and eventually the patient expired; care was withdrawn. regarding the impella cp device, there was a device malfunction, but it did not contribute to the death outcome. the cp was explanted and patient continued support with the impella 5.5 device. additionally, there is no information that indicates or suggests that the patient expired from the small infection. the event involving the impella cp is a serious injury type of reportable event. the impella 5.5 was in at the time of the patient's demise. conservatively the death should be reported as the impella 5.5 device was in use at the time of expiration. there is no information that indicates or suggests that the patient expired from arrhythmic events. the arrhythmia presented with support was reduced. the patient's underlying condition contributed most to the death outcome. due to the continuous decline in the patient's condition from the cardiac event, the family decided to withdraw all care and this led to the patient expiring. related medical device reports: this impella cp report is one of three devices associated with the event story. refer to the related manufacturer report numbers as noted in section h10 for the medical device report on the impella introducer and impella 5.5 which is associated with the demise outcome. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5157 | 1220648-2025-31662 | 23136962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136962&pc=OZD | 04/05/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 04/05/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621570 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella 5.5 support. the alarm was disabled and support continued uninterrupted. there was no noted patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. clinical details report that a placement signal not reliable (psnr) alarm triggered on (b)(6) 2025. the pump was not removed due to the complication. data logs were reviewed, and a single psnr alarm was confirmed the night of (b)(6) 2025. at this time, contrast began to oscillate between +3200/-3200. after 13 minutes, the alarm resolved and there were no further optical signal issues for the rest of the case. product was not returned for investigation. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5158 | 1220648-2025-31661 | 23136986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136986&pc=OZD | 04/02/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 04/02/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625680 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Unknown | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp support ongoing. the pump lost signal and had an unreliable placement signal alarm. the team effectively troubleshot the issue by the straightening of the white cable. the pump remained on for support without any harm or injury. the pump was eventually weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for investigation. no physical damage was noted on the returned product. the pump passed the laser continuity test, and all the 5s optical parameters were within specification range. the pump passed the water bucket test and pressure test on the optical system. the pump was run according to the specifications during the test. the data log was not returned for investigation. according to the clinical report, intermediate failures of placement and lv metrics occurred with a placement signal not reliable alarm. this issue was resolved after the connector was firmly plugged in. the cause of the optical signal failure was most likely air gap from between the automated impella controller and the patient cable plug or within the middle of the red handle ferrule, based on given clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5159 | 1220648-2025-46823 | 23136988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23136988&pc=OZD | 08/26/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 08/26/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713290 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had placed a 5.5 after the prior intra-aortic balloon pump was causing an ischemic leg in the acute coronary syndrome / acute myocardial infarction patient. the 5.5 was placed but after 4 days there were purge alarms. the pp rose and tpa was added to the purge, and the purge blockage was alarming as well. the pump was explanted and patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product returned. high purge pressure/purge system blocked. based on the clinical details and data log analysis the root cause of the high purge pressure and purge system blocked is most likely due to biological material ingestion. device history review: the device passed all the post sterile inspection checks.
5160 | 1220648-2025-31663 | 23137033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23137033&pc=OZD | 04/09/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 04/09/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547104 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient placed on the impella 5.5 pump. at the time of insertion there was thought to be optical sensor damage. the damage did not prompt any intervention. the pump remained on for support without harm or injury noted. the pump was eventually weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data log was provided for the console. the case start was not recorded on the console from the provided imc log and lt log. we are assuming that the pump case start was done on a different console or that insufficient logs were provided. the optical sensor offset difference was 2000 between the calibration and factory values. the provided lt data log has all the 5s parameters within specification range. the placement signal was at 3000 from the beginning and intermittently dropped to 1000 with an active placement signal not reliable (psnr) alarm. as per the clinical assessment, the optical sensor was damaged during the implant. without the returned product, it was difficult to verify the damage, and the returned log did not confirm any known trends of pump's optical sensor damage failure. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5161 | 1220648-2025-31664 | 23137049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23137049&pc=OZD | 04/25/2025 | 09/24/2025 |  | 09/24/2025 |  |  | 04/25/2025 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617614 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed for a patient. the pump was placed, but the optical signal failed. the pump remained on for support despite the placement signal loss. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned product was inspected and the pump passed light continuity testing. the 5s parameters on the returned product showed very low signal-to-noise ratio (snr) and gain and contrast were also out of normal range. the data logs confirm placement signal not reliable (psnr) alarm and show placement signal going out of range to 3000mmhg. at the same instant snr drops to zero, lamp and gain increase, and contrast decreases. clinical mentioned intravascular lithotripsy (ivl) was estimated to be about 20mm away from the catheter and that psnr alarm occurred after first irradiation. per the instructions for use, this is considered to be a use related issue. the root cause of the optical signal issue was determined to be due to a damaged sensor as evidenced by 5s parameters on returned product and clinical detail of ivl position in relation to catheter. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5162 | 1220648-2025-31665 | 23139967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23139967&pc=OZD | 05/23/2023 | 09/24/2025 |  | 09/24/2025 |  |  | 05/23/2023 | 09/24/2025 |  |  | 10/31/2024 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275567 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp was implanted under cardiopulmonary resuscitation, and no placement signal and left ventricle (lv) metrics were displayed, as well as no flow calculated. unplugging/ plugging the white connector cable did not resolve the issue. the automated impella controller was changed, which brought back the impella flow calculation, but the placement signal and lv signal were still not available. the impella cp was not removed because the patient¿s condition was critical. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the issue was not determined since product/data was not returned. all pertinent information available to abiomed has been submitted at this time.
5163 | 1220648-2025-31666 | 23139986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23139986&pc=OZD | 03/20/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 03/20/2024 | 09/24/2025 |  |  | 11/30/2025 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435690 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | EXTRA CORPOREAL MEMBRANE OXYGENATION | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient paced on the imeplla cp and extra corporeal membrane oxygenation (ecmo). after fourteen days of support. the pump lost placement signal. impella p-level dropped to p-3 and ecmo flows increased. care was withdrawn, and the patient was put on comfort care. the patient eventually expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no problem was found with the pump. it is apparent the automated impella controller was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5164 | 1220648-2025-31667 | 23140012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23140012&pc=OZD | 06/19/2024 | 09/24/2025 |  | 09/24/2025 |  |  | 06/19/2024 | 09/24/2025 |  |  |  |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490816 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal on an impella 5.5 was intermittent. motor current was normal, and echocardiography was performed. staff confirmed the impella was in an appropriate position. the pump was eventually weaned and explanted, and no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for evaluation. data logs were reviewed, and period low signal-to-noise ratio (snr) was observed. no issues with motor current were noted. clinical details noted placement signal was occasionally reading high and off the automated impella controller (aic) screen. no problem was found with the pump to cause the optical signal issue. upon review of the data logs, it is apparent that the console (aic) is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5165 | 1220648-2025-31668 | 23140021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23140021&pc=OZD | 12/11/2023 | 09/24/2025 |  | 09/24/2025 |  |  | 12/22/2023 | 09/24/2025 |  |  | 07/31/2025 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381688 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the implanted impella cp pump had a persistent discrepancy between the position waveform and the arterial line. the device positioning was subsequently verified via echocardiography and support was maintained. it was noted that the patient passed away during support; however, there was no allegation or evidence of association between the patient's passing and the impella device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs show placement signal (ps) low alarms starting six hours into support and continuing throughout. upon investigation of the pump, it passed the laser continuity test. ps low alarm did not reproduce on the lab console. the optical parameters were within normal ranges. the root cause of the optical signal issue was most likely patient condition related as the patient's condition was deteriorating. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5166 | 1220648-2025-31669 | 23140066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23140066&pc=PYX | 03/17/2023 | 09/24/2025 |  | 09/24/2025 |  |  | 03/17/2023 | 09/24/2025 |  |  | 11/30/2024 |  | 09/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024286331 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there was an optical signal failure. no patient harm or interruption to support was reported. the patient was successfully weaned from the pump and the device was explanted. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. field lt log from two different consoles were reviewed and persistent suction and placement signal not reliable alarms triggered throughout support. cvp is erratic with negative values which occurs with suction on rp optical. no spikes in motor current. 5s parameters were stable throughout support. signal not reliable dipped at the times of of the issue. the returned pump was connected to an automated impella controller (aic) and placement signal unreliable was unable to be reproduced. 5s parameters were normal. the pump was visually inspected and biomaterial was observed in the cannula and several cannula kinks were observed. the cause of the optical signal issue was suction due to biomaterial in the cannula. cannula kink also observed. no sensor issues were observed. refer to es2020-082 for additional biomaterial pattern details. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5167 | 1220648-2025-46760 | 23119282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119282&pc=OZD | 08/23/2025 |  |  | 09/23/2025 |  |  | 08/23/2025 | 09/23/2025 |  |  | 02/28/2026 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479688 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | Increase in Pressure | Pumping Stopped |  |  |  |  | Arrhythmia | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [35 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported on the second day of impella 5.5 support the pump had a purge blocked alarm and saturated flows. the medical team added tissue plasminogen activator to the purge solution to attempt to resolve the issue. the following day, the pump abruptly stopped functioning. it was noted there was possible damage to the pump from cross clamping. the impella 5.5 was pulled back across the aortic valve into the aorta and the patient began to decompensate. the decision was made to take the patient urgently to the operating room for impella cp insertion and impella 5.5 removal. the impella cp was successfully inserted, and support was initiated. the patient remained on the pump for an additional day, and with heart function recovered the patient was successfully weaned and the impella cp was explanted. | It was additionally reported that the patient experienced arrhythmia which caused hemodynamic instability on the third day of support. the patient was treated with a bolus of amiodarone, and an intravenous drip was started, which converted the patient into a normal sinus rhythm. |  |  |  |  | Correction to h6 health effect - clinical code: e2343 replaces incorrect code e2403 from the initial medwatch report. correction to h6 health effect - impact code f26 was erroneously reported in initial medwatch. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. | Device status: the impella was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 is being updated to include related device information not included in the initial report. h6. health effect - clinical code e2343 and e0601 added for b5 information. health effect - impact code f2303 added for additional b5 information. component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. medical device problem code a141102 was not included in the initial mdr and is now correctly added. investigation summary: pump sn (b)(6) and a purge cassette returned on 09/02/2025. clinical details report that a purge system blocked alarm triggered and pump flows became saturated on (b)(6) 2025. the following day, the pump stopped. data log review confirmed the report. on the (b)(6), purge pressure rose over a 50 minute period around 16:00. at the time of the onset of the purge pressure rise, there was no spike or variation in the motor current (mc)/ motor spike (ms), and there were no long purge bag changes leading up to the event. several hours later, mc started to trend up and exhibit some spikes with ms deviations. this persisted through the end of the case, when high mc pump stop alarm #13 began to trigger on and off. returned product was inspected and there were no kinks noted in the catheter, patient cable, or purge cassette tubing. the pump was soaked, and there was evidence of biomaterial in the purge gap and near the impeller. the pump passed electrical resistance testing of the motor cables. the pump was connected to a console and run in a bucket. initially, pump stop alarms reproduced, but the pump was ultimately able to be restarted. purge pressure gradually trended up during the run, ultimately triggering the high purge pressure alarm. pump flows were saturated at p-9 for the entire run. a window was then cut in the catheter near the motor housing, and there was evidence of blood ingress. the catheter was cut at this location, but fluid did not expel, suggesting the blockage was more proximal. the red handle was opened, and blood had ingressed through the purge line all the way into the end of the patient cable. high purge pressure: data log review confirmed a 5-hour trend up in the purge pressure without a mc spike. the mc subsequently trended up hours later and began to spike, resulting in high mc pump stops. returned product did have evidence of biomaterial in the purge gap and reproduced high purge pressure, however, there was also evidence of blood ingress into the purge line all the way up to the end of the patient cable. there were no signs of kinks in the purge line (patient cable, catheter, purge cassette,). the patterns noted on data log review and evidence of biomaterial at the purge gap suggest that biomaterial was related to the high purge pressure. however, it couldn't be definitively established as the cause because the data log trend does not align with blood ingress noted on returned product without any purge line kinks. saturated flow: data log review confirmed a trend up and spikes in mc with speed deviations several hours after purge pressure became high. testing of the returned device confirmed evidence of biomaterial in the purge gap and reproduction of saturated flows. there was also evidence of blood ingress into the purge line. therefore, it was determined that the cause of the saturated flows was likely blood ingress and/or biomaterial buildup, however, the cause could not be determined as data logs and returned product were inconclusive. pump stop: data log review confirmed a trend up and spikes in the mc with speed deviations several hours after purge pressure became high. the mc spikes triggered a high mc pump stop. testing of the returned device confirmed that there was still evidence of biomaterial in the purge gap and blood ingress into the purge line, but there was no evidence of kinks in the catheter/purge cassette. therefore, it was determined that the cause of the pump stop was likely blood ingress into the motor, however, the cause of the blood ingress could not be determined as data logs and returned product were inconclusive. arrythmia: clinical details report arrhythmias overnight on (b)(6) 2025. pressor support was increased to troubleshoot. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. hemodynamic instability: clinical details report hemodynamic instability overnight on (b)(6) 2025. pressor support was increased to troubleshoot. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. device history lot: device lot: 1804011. device history review: pump sn (b)(6) passed all post-sterile inspection checks.
5168 | 1220648-2025-46761 | 23119288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119288&pc=OZD | 08/23/2025 |  |  | 09/23/2025 |  |  | 08/23/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025560122 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 91 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Unspecified Nervous System Problem | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed for the patient admitted in with cardiac history and new cardiac arrest with acute myocardial infarction/cardiogenic shock. the pump was placed and sheath removed. when the reposheath was placed and another sheath for access the limb became ischemic. there was no flow down the limb. the team transferred the patient to tertiary center for care escalation and new impella access when pupils were fixed and care was withdrawn and patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. neurological impairment : the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5169 | 1220648-2025-46762 | 23119289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119289&pc=OZD | 08/23/2025 |  |  | 09/23/2025 |  |  | 08/23/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574168 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. on the third day of support, the physician stated that they believed the patient was in septic shock, and a source of infection was not provided. additionally, the patient experienced a large amount of bleeding at the impella access site. as a result of the bleeding, the impella cp was removed by a vascular surgeon. the patient survived impella explant, however patient procedure outcome is unknown. | [5 empty] | Device status: the impella was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5170 | 1220648-2025-46763 | 23119300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119300&pc=OZD | 08/23/2025 | 11/13/2025 |  | 09/23/2025 |  |  | 08/23/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026682629 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on the second day of support, it was reported that the patient experienced oozing at the impella access site. as a result, heparin was discontinued until the oozing resolved. four days later, it was noted that the patient experienced intermittent ectopy, and the p level was decreased from p6 to p4. the following day, the patient decompensated and became anxious. the patient requested to start comfort care. the decision was made to withdraw patient care, and the patient expired. | [5 empty] | Investigation summary: no product was returned for investigation. major bleeding: groin site and ij site are oozing. the cause of the bleeding is determined to be use issue due improper anticoagulation management because hemostasis was achieved by stopping heparin. arrhythmia: in order to make a root cause determination, additional clinical would have to be provided. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. device history lot: device lot: 1934628. device history review: device with sn: (b)(6) passed all post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5171 | 1220648-2025-46764 | 23119311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119311&pc=OZD | 08/23/2025 |  |  | 09/23/2025 |  |  | 08/23/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665229 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 105 KG | Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Retroperitoneal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). while on the unit, a ¿slight¿ hematoma was noted and managed with pressure and femostop. however, the hematoma and bleeding worsened. the patient was taken to the catheterization lab for the percutaneous coronary intervention (patient was a surgical turn down for coronary artery bypass graft surgery) with the intent to leave in the impella cp device. successful thrombectomy of the right coronary artery was completed. the physician utilized the side closure technique to attempt to minimize the hematoma and bleeding and replaced the repositioning sheath. the patient was returned to the unit on the impella mcs. however, the bleeding continued, and a retroperitoneal bleed was suspected. additionally, an alarm triggered for the impella in the ventricle, and the support level was reduced to p-2. the patient was brought back to the catheterization lab to remove the impella cp device, repair the site and placed the patient on venoarterial extracorporeal membrane oxygenation. the patient received ¿multiple¿ transfusions. survival at explant and the outcome of the patient because of the event were unnoted. however, there was no report or indication of a demise outcome. | [5 empty] | Patient-specific information a6: race is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for the impella cp with smartassist system: section: general patient care considerations - ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer - ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, clinical symptoms retroperitoneal hemorrhage, the cause of the injury was not determined due to insufficient clinical details. access site adverse event hematoma, major bleed, the pump was not returned for investigation. data log review was not required for this case run. the cause of the access site bleeding issue was not determined, as the pump was not returned for investigation and sufficient clinical information was not provided. device in wrong position: pump was not returned for investigation. data log review was not required for this case run. the cause of the positioning issue was not determined, as the pump was not returned for investigation and sufficient clinical information was not provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5172 | 1220648-2025-46765 | 23119318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119318&pc=OZD | 08/23/2025 |  |  | 09/23/2025 |  |  | 08/23/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026724410 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced device suction unresolved and left ventricle thrombus leading intervention and device explantation. the patient presented to catheterization lab for inferior st-segment elevation myocardial infarction and required intubation for respiratory distress. at this time the patient went into cardiac arrest and cardiopulmonary resuscitation along with defibrillation and pressors utilized. the decision was made to place an impella cp and started in auto. after 20min resuscitation efforts, return of spontaneous circulation was achieved. cardiothoracic surgeon present and evaluated the patient for extracorporeal membrane oxygenation (ecmo) versus impella 5.5. the decision was made to take the patient to the operating room. impella was flowing 3.3l at performance level p-9. upon arrival at the operating room, it was decided to cannulate for ecmo, which slowly started up to 2.5lpm and the impella cp was turned down to p-6 in which suctioning immediately started. albumin 250cc were given and impella flow was turned to p-2 for suction events. pressors were turned off. final activating clotting time was 170seconds. the patient was transferred to the unit waiting for transport to an outside facility. once on the unit suction alarms at p-2 were noted with central venous pressure of 3mmhg. the patient was off all pressors with only albumin and amino infusing. albumin 500cc boluses total was given and central venous pressure was now 13mmhg. one unit of packed red blood cells was administered prior to transport. sedation medications were off. and it was still, however noted, suction at p-2 with mean arterial pressure of 110 mmhg. impella giving 0.6l and ecmo 2.7l. bilaterial groin sites were stable. impella site soft without hematoma or bleeding. the patient awaited transport for continuation of care. the patient continued, however, to have suction at p-2 despite the albumin and packed red blood cells administration. with point of care ultrasound, the impella cp device was pulled back from 85cm to 86cm. the impella cp device was then able to be turned up to performance level p-4 flowing at 1.5l and ecmo flow of 4l. bilateral groin sites were noted with no bleeding or hematoma, and doppler pulses at bilateral feet. overnight, impella was running at p-4 and was turned up to p-8 by the next morning. echocardiogram was completed and showed left ventricle thrombus and an ejection fraction of 5%. the impella device was approximately 4cm across the valve. the patient underwent chest/head computed tomography scan that showed an acute versus old infarction on head scan and required a magnetic resonance imaging to confirm. computed tomography scan of the chest showed bilateral pleural effusions. the patient also had an ultrasound of the right upper extremity as pulses were lost in the right arm. further update noted the patient woke up and followed commands. doppler pulses were present in bilateral lower extremity. the patient had ample urine output overnight without hematuria, and bilateral groins were stable without oozing or hematoma. the impella cp device was eventually explanted without issue. the groin site was stable and the patient continued venoarterial ecmo at 3.8lpm, milrinone and therapeutic on bivalrudin. repeat echocardiogram this day in the morning showed laminar thrombus attached to inferior apical wall remained measuring 2.1x1.4cm. the status of the patient following the event was unnoted. | [5 empty] | Patient-specific information a4: weight and a5: ethnicity is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows for the impella cp with smart assist system. section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, clinical symptoms (thrombosis/thrombus), in order to make a cause determination, all of the clinical details would need to be provided. the root cause of the injury could not be determined due to insufficient clinical details. pertaining to the pump suction, the pump was not returned for investigation. data log shows suction alarms while running at p9, around 4 hours into the case. the pump was then reduced to p-2, with persistent suction alarms for about 5 hours. the pump placement signal showed a decreasing trend during the early case run until the suction alarm resolved. some pump flow improvement was noted after resolution of the suction alarm, and the pump was then set to p4. there was no motor current spike or positioning alarm reported during the case run. clinical details suggest low left ventricular ejection fraction (lvef) and thrombus in the lv. extracorporeal membrane oxygenation (ecmo) was started after the pump was implanted. the pump p-level was reduced to p-6 after ecmo was initiated, and the pump was noted to be in suction condition. the pump p-level was further reduced to p2, and blood was administered due to the ongoing suction alarm. the cause of the suction alarm was most likely related to the patient condition, based on data log review and provided clinical details. h6. investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5173 | 1220648-2025-46766 | 23121057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121057&pc=OZD | 08/24/2025 | 10/28/2025 |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617952 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient with st-segment elevation myocardial infarction. while on support, it was noted that the impella cp was removed due to bad positioning causing hemolysis. the patient outcome is unknown. | [5 empty] | The investigation for the reported hemolysis and positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and positioning issue could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5174 | 1220648-2025-31615 | 23121294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121294&pc=OZD | 08/09/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 08/09/2024 | 09/23/2025 |  |  | 03/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356971 |  | 0550-0004 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was noted. pump position was verified by echocardiography. despite the signal issue, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump and data logs were not provided for investigation. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5175 | 1220648-2025-31616 | 23121432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121432&pc=OZD | 05/30/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 05/30/2023 | 09/23/2025 |  |  | 06/30/2024 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234905 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.0 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues; however, flow rates, motor current, and purge values remained within expected ranges. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump and log was not returned for investigation. the root cause of the issue was not determined since product was not returned. all pertinent information available to abiomed has been submitted at this time.
5176 | 1220648-2025-31617 | 23121473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121473&pc=OZD | 01/24/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 01/24/2023 | 09/23/2025 |  |  | 02/29/2024 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023195896 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.0 pump was implanted for circulatory support. during support, the pump exhibited low placement signal issues; however, flow rates, motor current, and purge values remained within expected ranges. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. the data logs showed characteristic of a durability sensor wear failure. no optical fiber or sensor issue was observed. the root cause of the optical signal issue was most likely sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5177 | 1220648-2025-31618 | 23121553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121553&pc=OZD | 02/07/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 02/07/2023 | 09/23/2025 |  |  | 09/30/2024 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023262473 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, the pump exhibited placement signal issues. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs were reviewed and show deterioration of the optical signal and confirm the placement signal unreliable alarms. there are periods during support that the signal returns to normal values. the 5s parameters are not showing a hard failure of the sensor. the root cause of the optical signal issue is unable to be determined. all pertinent information available to abiomed has been submitted at this time.
5178 | 1220648-2025-46768 | 23121582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121582&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562128 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 102 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced a "placement signal low" alarm. proper pump position was confirmed and the pump's audio was disabled. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5179 | 1220648-2025-46769 | 23121657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121657&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683670 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient was admitted to an outside hospital with shortness of breath and was diagnosed with a non-st-elevation myocardial infarction. a cardiac catheterization revealed multivessel disease and a left ventricular end diastolic pressure of 20 mmhg. the patient was transferred to the complainant for surgical consultation. the morning prior to impella cp insertion, the patient experienced atrial fibrillation with rapid ventricular response, which progressed to ventricular fibrillation and eventual cardiac arrest. extracorporeal life support was considered. however, due to the prolonged duration of the code, the decision was made to proceed with impella insertion and percutaneous coronary intervention. revascularization of the right coronary artery was successful. the ostial of the circumflex artery was treated with a balloon angioplasty. however, it was noted there was extensive atherosclerosis in the circumflex and left anterior descending arteries (lad). a chronic total occlusion was identified in the lad artery but was not treated during the procedure. the patient was taken back to the intensive care unit for rest and recovery. after the transfer, it was noted that the patient experienced a hematoma at the impella access site around the sheath. the patient was transfused with a unit of blood as a result of the hematoma. it was noted that the patient was taking anticoagulants and antiplatelet medications, and it was also noted that the deep placement of the repositioning sheath may have contributed to the hematoma. it was noted later the same day, the patient expired while on impella cp support. | [5 empty] | Although the patient's presenting condition may be more likely to have impacted the event, the impella cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5180 | 1220648-2025-46770 | 23121728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23121728&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661766 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 developed a thrombus and hypoxia during support. a thrombus on the left ventricle was noted on the daily echocardiogram, while the patient was on heparin drip. additionally, the patient had a vasovagal episode and became hypoxic. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the cause of the respiratory failure was not determined due to insufficient clinical details. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
5181 | 1220648-2025-46773 | 23122092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122092&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026716177 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Fracture | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella 5.5 encountered a pump stop and catheter product damage. the pump stops resolved without intervention. upon removing the device through the graft, the catheter snapped into two parts. the patient was taken to surgery to remove the other portion of the catheter in graft. impella support continues. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this is one of two related reports. this report represents the pump and another report was submitted to represent the automated impella controller. refer to section h.10 for the related report number.
5182 | 1220648-2025-31619 | 23122170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122170&pc=OZD | 03/16/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 03/16/2024 | 09/23/2025 |  |  | 06/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374505 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support the patient experienced placement signal issues. the cable was checked and there were no kinks noted, and the white light was working. the pump was checked via tte and was in good position. the issue was unable to be resolved therefore the placement signal was muted. the pump was not explanted and no further information was provided. patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. about 10 hours after implant, there were no waveforms visible. the placement signal issue did not resolve after troubleshooting. echo was done and good position was confirmed. no product or data logs were returned. the root cause of the optical signal issue was not determined as no product was returned and insufficient clinical information was provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5183 | 1220648-2025-46772 | 23122201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122201&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1598659 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 66 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient implanted with an impella 5.5 encountered an automated impella controller error alert. the error resolved without intervention. the impella support continued. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. this is one of two related reports. this report represents the automated impella controller and another report was submitted to represent the pump. refer to section h.10 for the related report number. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5184 | 1220648-2025-46771 | 23122314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122314&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026705182 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that delivery issues were encountered during implantation of an impella 5.5 due to the patient's tortuous anatomy. implantation was aborted. no patient harm was reported. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. | No product was returned for investigation. the cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to patients tortuous anatomy and the insertion was aborted. the device passed all post sterile inspection checks.
5185 | 1220648-2025-31620 | 23122533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122533&pc=OZD | 03/19/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 03/19/2023 | 09/23/2025 |  |  | 10/31/2023 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154382 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide support for over two months. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data log showed placement signal not reliable alarms. data also showed 5s parameters were within specification. the product was not returned for review. based on clinical description & data analysis, the root cause of the optical signal issue was sensor wear. the root cause of the optical signal issue was sensor silicone wear. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5186 | 1220648-2025-46289 | 23122544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122544&pc=OZD | 08/11/2025 | 10/07/2025 |  | 09/23/2025 |  |  | 08/11/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675556 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Consumer | Health Professional | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan had a cp pump and automated impella controller (aic) support being initiated to support the patient in need of pci for the ami/cgs patient. the pump stopped and was unable to be restarted successfully. in troubleshooting they also wished to replace the aic but there was no backup/not in patient use console available. the pump was explanted and pci performed without hemodynamic support of impella. patient survived. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of the 2 report, where this report represents the pump and the other report is for aic. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
5187 | 1220648-2025-46302 | 23122547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122547&pc=OZD | 08/10/2025 | 10/29/2025 |  | 09/23/2025 |  |  | 08/10/2025 | 09/23/2025 |  |  | 01/31/2026 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457844 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 49 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced access site complication and suction events, prior to the impella placement, an intra-aortic balloon pump was successfully attempted due to a pump failure. the decision was made to switch to an impella cp, which was successfully placed without incident in the right groin. the patient received 8000 u heparin intravenous-procedurally. activated clotting time was 252 seconds at the end of the case. the patient was transferred to the unit on impella mcs. when on the unit, the support level was dropped from p-9 to p-7 to break the suction. the patient was administered 250 milliliters of albumin bolus and was responsive to the fluid bolus. suction events stopped and was resolved. additionally, the patient experienced groin site oozing. the dressing was removed, and pressure was held over the insertion site until oozing stopped. a large, firm hematoma was also noted in the right groin (notedly not present in the catheterization lab at time of departure). labs revealed a drop in hemoglobin and the patient received one unit of packed red blood cells. distal pulses were confirmed present by doppler. the patient continued on mcs for an additional five days before being successfully weaned and the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device was not returned; therefore, an evaluation/analysis of the device was not possible. however, event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical reported oozing from the groin site on the first day of the case, which was resolved with manual pressure. the product was not returned for investigation. the cause of the bleeding and hematoma was not determined due to insufficient clinical details and no product returned. regarding, the pump suction, clinical reported suction events occurring on the first day, which were resolved after administration of fluid bolus. the product was not returned for investigation. data logs show that at the time of the suction alarms, placement signal was exhibiting a slight downward trend. there were no motor current spikes outside of p-level changes. the cause of the suction was most likely patient related, because data logs showed a downward trend in placement signal, and the clinical stated the suction resolved when volume was administered. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5188 | 1220648-2025-46303 | 23122551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122551&pc=OZD | 08/09/2025 | 09/09/2025 |  | 09/23/2025 |  |  | 08/11/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639809 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 121 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Patient Device Interaction Problem | [5 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had hemolysis concerns, and arrythmia. the pump had continuous suction with flows of zero the patient was anticoagulated and there was no evidence of clot. the pump was exchanged for a 5.5. the 5.5 was successfully placed and the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Correction: g4. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. hemolysis: the pump was returned cut so a hemolysis test could not be performed. data logs showed suction alarms occurring at the time hemolysis was reported. at this time, placement signal was trending down slightly, but there were no motor current spikes to indicate the biomaterial observed in investigation was ingested during the case. the cause of the hemolysis was not determined due to insufficient clinical details and insufficient product was returned, leading to insufficient testing. pump suction/low pump flow: data log review at this time revealed that there were suction alarms triggered and pump flows on the lower end at p-1. suction alarms were triggered during this time as there was a lack of motor current (mc) pulsatility, but no spikes in motor current were seen. the p-level is increased to p-4 but flows of almost 0l/min remained. there weren't any abnormalities in the placement signal at this time, or signatures of a clot ingestion. in physical product investigation, a kink was observed on the cannula near the motor housing and biomaterial was observed around the impeller. the cause of the suction and low pump flow was most likely the cannula kink, because the cannula likely folded when the pump was readjusted and pulled back, which is what led to the suction alarms and flows of 0l/min. device history review: the device passed all the post sterile inspection checks.
5189 | 1220648-2025-31621 | 23122571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122571&pc=OZD | 10/25/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 10/27/2023 | 09/23/2025 |  |  | 06/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368161 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. data logs were reviewed, and placement signal was downtrending from the start till the end of the case. several placements signal low, impella position unknown alarms and suction alarms were found with likely poor patient condition. the 5s parameters had no major impact during this failure. the root cause of the optical signal issue was most likely patient condition related per data log review and clinical review. all pertinent information available to abiomed has been submitted at this time.
5190 | 1220648-2025-31622 | 23122634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122634&pc=OZD | 04/02/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 04/02/2023 | 09/23/2025 |  |  | 02/28/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310347 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump exhibited placement signal issues but continued to provide adequate support. the device was successfully weaned and explanted, and no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the log shows that there is an upward trend in placement signal and change in motor current pulsatility as well as saturated flow. the returned 5s were not within specification. biomaterial was observed around the impeller and on the sensor. the pump was soaked in tergazyme and afterwards the 5s were re-tested and were within specification. the cause of the optical signal issue was biomaterial. the catheter was found to be separated at the kink protector exposing the coil underneath. the cause of the product damage was catheter damage at the transition from inside the kink protector. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5191 | 1220648-2025-46775 | 23122641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122641&pc=PYX | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  | 11/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024434878 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 112 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Inflammation | Renal Failure | Sepsis | Asystole | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support following worsening right ventricular function status post pulseless electrical activity arrest and quadruple coronary artery bypass graft. after impella rp flex support was initiated, pressor support was able to be lowered. however, the patient was also placed on continuous renal replacement therapy (crrt). the following day, the patient remained on crrt, and suction alarms were noted. additionally, the same day, the patient was taken off amiodarone and was noted to be back in atrial fibrillation. it was also noted that the patient¿s mixed venous oxygen saturation dropped. it was additionally noted that there was a concern for an infectious process versus inadequate support, and the patient was already being treated for sepsis. the following day, it was noted that the patient became increasingly septic. the patient went into asystole and return of spontaneous circulation (rosc) was achieved. the family made the patient do not resuscitate (dnr). shortly after, the patient went into asystole again. care was withdrawn and the patient expired. | [5 empty] | Based on the information available, it is unlikely that the sepsis was related to the use of the impella device. however, a device-related contribution cannot be definitively ruled out. although the patient's presenting condition may be more likely to have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5192 | 1220648-2025-31623 | 23122770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122770&pc=OZD | 02/03/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 02/03/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547525 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient acute myocardial infarction and cardiogenic shock. during therapy, the pump experienced an optical sensor issue, with the difference between the aortic and placement signal readings exceeding acceptable limits. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned for investigation. no physical abnormalities were noted on the returned product except for the cosmetic catheter kink at the kink protector. the pump passed the laser continuity testing for the optical line and all the 5s parameters were within limits. the pump also passed the pressure testing for the optical sensor. the data log showed a high placement signal for about 5 minutes with low pump flow during the very start of the case, which was resolved and remained around 50 throughout the case. the cause of the optical signal issue was not determined since no abnormalities were found on the returned product, and the placement signal high issue was resolved during the case. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5193 | 1220648-2025-31624 | 23122836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122836&pc=OZD | 01/08/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 01/09/2024 | 09/23/2025 |  |  | 10/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418722 |  | 2000204 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced persistent placement signal issues. despite this, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was returned for evaluation. biomaterial was found on impeller and sitting on sensor face. data logs were reviewed and there is no indication of a sudden ingestion of biomaterial on sensor face. the root cause of the optical issue resulting in a differential in placement signal from clinical metrics was likely due to biomaterial on the sensor face. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5194 | 1220648-2025-31625 | 23122859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122859&pc=OZD | 01/22/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 01/22/2024 | 09/23/2025 |  |  | 09/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403937 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the placement signal was noted to be lower than the arterial line pressure. pump position was verified by echocardiography. despite the signal issue, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was returned for evaluation with catheter cut right before motor housing of pump, so no optical testing was able to be performed. no data logs are available for this pump. the root cause of the optical signal issue cannot be determined as the pump was cut and optical tests were unable to be conducted and there are no data logs available for this pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5195 | 1220648-2025-31626 | 23122928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23122928&pc=OZD | 04/07/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 04/07/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026618438 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a cp pump placed to support a patient. the pump was placed and was supporting with placement signal loss. the pump remained on for the 90minutes and was then explanted. no patient harm reported. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. as impella was started, lv waveform did not appear. motor current appropriate, ao placement signal appropriate, flows appropriate. no evidence of waveforms being disabled. procedure continued until pt weaned off for closure. the pump was not returned. data logs were not recovered. the cause of the optical signal issue was not determined since product was not returned, data logs were not recovered, and insufficient clinical details. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5196 | 1220648-2025-31627 | 23123015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23123015&pc=OZD | 05/13/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 05/13/2024 | 09/23/2025 |  |  | 12/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025443378 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella 5.5 implantation the placement signal was lost. placement signal was disabled to avoid alarms. the impella functioning was normal, and the patient was hemodynamically stable. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was not returned for investigation. data logs show that the 5s parameters were in spec for about 30min until they suddenly are lost. placement signal not reliable (psnr) immediately goes out of spec, snr drops to 0, lamp spikes to max 100, contrast increases in amplitude, light decreases, and gain increases. there is a psnr alarm at this time. 5 minutes later the pump is started. the clinical states that the psnr alarm occurred after verifying the pump's position. the optical parameters never recover. this matches the pattern of a damage optical fiber line. the root cause of the optical signal issue was most likely a damaged optical fiber line but the cause of the damage is unknown because product was not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5197 | 1220648-2025-31628 | 23123060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23123060&pc=OZD | 04/25/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 04/25/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599993 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 had placement signal not reliable alarm. no patient harm reported. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: clinical details note placement signal not reliable alarms during support. the pump was not removed due to this issue. data log review shows the placement signal intermittently going out of range. when the placement signal was out of range, the contrast was seen oscillating between 3000/-3000. no product was returned. upon review of the logs, it is apparent that the console is the potential cause of the issue. please refer to 25-44075-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5198 | 1220648-2025-31629 | 23123457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23123457&pc=PYX | 07/20/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 07/22/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025502656 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient was on impella rp support. while on support, placement signal and flow not displaying and are incorrect on the automated impella controller (aic). it was noted that less than 24 hours after implant, the signal went out. confirmed that the sensor of the flex was going out. chest xray was obtained and placement of the device was confirmed. no report of patient harm. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella rp flex was returned for evaluation. upon visual inspection, no abnormalities were noted on pump. no biomaterial was present on impeller. pump passed electrical testing and optical parameters were found to be within specification. pump was run in flow loop and dp sensor drift was reproduced. data logs were reviewed and a sudden increase in placement signal with a drop in pump flows. motor current remained stable and cvp remained stable throughout case. clinical details noted that after loss in placement signal, pump was able to run without issue for an additional 6 days. the root cause of the placement signal issue was due to sensor drift of the dp sensor. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5199 | 1220648-2025-31630 | 23123640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23123640&pc=OZD | 07/21/2025 | 11/11/2025 |  | 09/23/2025 |  |  | 09/11/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026691801 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male | 75 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Death | Other |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The medical safety officer reviewed and determined the following: the patient was already being treated by extracorporeal membrane oxygenation (ecmo) and was surgically treated. impella was implanted for venting support and operated without an issue. conservatively reporting due to the pump being used at the time of the patient's demise. there was no issue reported and no known device malfunction with venting support. underlying disease is the likely cause of the patient's demise. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. this file was reassessed on 11sept2025 and was determined to be reportable for death. the medical safety officer reviewed and determined the following: the patient was already being treated by ecmo and was surgically treated. impella was implanted for venting support and operated without an issue. conservatively reporting due to the pump being used at the time of the patient's demise. there was no issue reported and no known device malfunction with venting support. underlying disease is the likely cause of the patient's demise. a5 race and ethnicity was unknown. | The root cause of the unspecified allegation -death was not determined due to lack of sufficient clinical details.
5200 | 1220648-2025-31631 | 23123646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23123646&pc=PYX | 07/25/2025 | 10/24/2025 |  | 09/23/2025 |  |  | 07/25/2025 | 09/23/2025 |  |  | 11/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024434405 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella rp pump was implanted in a patient with prior coronary artery bypass grafting and heart failure. the complainant reported that there was a placement signal unreliable alarm on the impella rp. a chest x-ray was ordered, and the rp indicated to be in a good position. in addition, it was reported that the patient was maxed on vasopressor support and heavy ventilator support along with continuous renal replacement therapy (crrt). a decision was made by the family to withdrawal care and the patient expired. | [5 empty] | D6b: corrected the explant date. | G3: corrected date. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. data logs show a 1054-placement signal not reliable alarm (bad snr) with signal not reliable (snr) dropping to zero, lamp at 100%, contrast decreased to 10, gain increased to 2.6, and light decreased to 2.4, which is indicative of a severe air gap. no issue was reported with pump flow. according to the clinical details, the placement signal not reliable alarm appeared during the patient's transfer to the icu. the cause of the optical signal failure was most likely air gap from between the automated impella controller (aic) and the patient cable plug or within the middle of the red handle ferrule. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5201 | 1220648-2025-46777 | 23123941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23123941&pc=OZD | 08/25/2025 |  |  | 09/23/2025 |  |  | 08/25/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026711585 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that implantation of an impella 5.5 pump failed due to the patient's anatomy. the patient had a rare anatomical condition where the right subclavian has a sharp turn on the left side of the aortic arch. alternative access sites were not viable. the patient was supported by extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation has been completed. the cause of the delivery issue was determined to be patient condition as the patient had rare anatomical condition having right subclavian with a sharp turn on the left side of the aortic arch preventing successful delivery of impella. the device history review was completed and the pump passed all the post sterile inspection checks.
5202 | 1220648-2025-31632 | 23124203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23124203&pc=OZD | 07/24/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 07/24/2025 | 09/23/2025 |  |  | 01/31/2026 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454437 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that a patient with history of cardiogenic shock and heart failure was implanted with an impella 5.5 pump for circulatory support. it was reported that the placement signal on the pump disappeared during patient support. leading up to the event, it was noted that a .035 wire had gotten temporarily stuck in the motor prior to insertion. it was believed that the optical sensor may have been damaged by the .035 wire, resulting in the lost reading. the patient remained on support; however, the family made the decision to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the day after impella insertion, during a lhc, the 0.035 wire became stuck in the motor of the 5.5 per staff reporting. when they were able to release the wire, the placement signal disappeared. clinical team suspect that the optical sensor was damaged by the 0.035 wire and now is not able to be read. no product was returned. the data logs confirm loss of placement signal during the procedure and a drop in snr to zero. the 5s parameters show failure pattern characteristic to an optical fiber break. the root cause of the optical signal issue was most likely a damaged optical fiber as characterized from the 5s parameters. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5203 | 1220648-2025-46776 | 23124635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23124635&pc=OZD | 08/25/2025 |  |  | 09/23/2025 |  |  | 08/25/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1563183 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | False Alarm | Backflow | Failure to Pump |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had a console being setup for patient support. the console alarmed with high motor current, pump stop, retrograde flow and purge alarms. the automated impella controller was replaced and support proceeded. there was no patient harm. | [5 empty] | Complaint product experience coding was updated to better describe the reported event. therefore, h6 medical device problem coding was revised accordingly. related medical device reports: this automated impella controller device report is one of two devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella cp. | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Product complaint # (b)(4). b5 is being updated to include related device information not included in the initial report. the revised b5 provides a complete event narrative. h6. medical device problem code false alarm was not included in the initial mdr and is now correctly added. h.10 related report number was added.
5204 | 1220648-2025-31633 | 23124941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23124941&pc=OZD | 06/01/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 06/01/2023 | 09/23/2025 |  |  | 01/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. it was reported that there was an ongoing placement signal low alarm and oozing during impella support. repositioning was attempted, however the issue remained. the alarm was silenced. the oozing was managed with dressing change. care was eventually withdrawn, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root causes of the optical signal issue could not be determined as insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5205 | 1220648-2025-46779 | 23125354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125354&pc=OZD | 08/24/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684843 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 101 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support a patient admitted in with acute myocardial infarction complicated by cardiogenic shock. a hematoma developed at the end of the implant procedure. the hematoma was held manually for 30-40 minutes. a femstop was placed. eventually, the patient was transferred from an impella cp to an impella 5.5. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. hematoma: the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5206 | 1220648-2025-46780 | 23125460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125460&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698066 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp pump placed to support the patient admitted in with acute myocardial infarction complicated by cardiogenic shock. after implant, it was noted that the groin site was oozing. manual pressure was held. on the same night, it was noted that oozing continued and a femostop was used. additionally, the patient was hypotensive. the case continued and the device was explanted. the patient survived. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. hypotension : the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5207 | 1220648-2025-46783 | 23125538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125538&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026712275 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 65 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 exhibited oozing from the access site along with bleeding from the nose and mouth. heparin has withheld and the purge was swapped to bicarb. the patient was transfused red blood cells. later, care was withdrawn and the patient expired. | [5 empty] | The cause of the bleed is most likely patient condition related since patient is having heavy nose bleeds, bleeding from the mouth, and increased output from his jp drain at the impella site. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5208 | 1220648-2025-46784 | 23125539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125539&pc=OZD | 08/01/2025 | 10/29/2025 |  | 09/23/2025 |  |  | 08/01/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678168 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia (vt). the patient was treated with amio to resolve the vt. the patient had low hemoglobin, hematocrit, and platelet values, and was transfused two units of packed red blood cells. systemic heparin was withheld. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia/hypotension was not determined due to insufficient clinical details. the device passed all post sterile inspection checks.
5209 | 1220648-2025-46781 | 23125542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125542&pc=OZD | 08/24/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695400 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed successfully on patient with suspected mixed shock. the team disable placement signal due to actual diastolic pressure below 30. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. optical sensor issue - data log review showed placement signal (ps) low alarm occurring at pump start with ps readings of 71 mmhg systolic and 25mmhg diastolic. the cause of the optical signal issue was patient condition related due to low ps readings seen due to the patients actual diastolic readings being low. device history review: the device passed all the post sterile inspection checks.
5210 | 1220648-2025-46785 | 23125543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125543&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026703342 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 92 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Pleural Effusion | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 was being turned for care when the graft on the aorta ¿popped off¿. the patient's chest cavity filled with blood and the patient coded. the patient was taken to surgery for intervention where it was confirmed that the anterior half of the graft placed on the aorta was completely torn from the body. the tissue of the aorta appeared healthy and the anastomosis was reinforced. after repositioning the pump, the patient became hypotensive, diaphoretic, and hemodynamically unstable requiring high dose pressors and va ecmo. a chest x-ray showed a new pleural effusion requiring re-intubation. the patient's chest was opened at the bedside where blood and clots were seen at the aortic graft site. the large clots were evacuated from the chest wall. ecmo was removed, and the patient was transfused eight units of packed red blood cells, eight units of fresh frozen plasma, two units of platelets, and two units of cryoprecipitate. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5211 | 1220648-2025-46786 | 23125560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125560&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a hematoma at the access site. a femstop was applied to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation into the reported event has been completed. no device was received for evaluation. the cause of the hematoma is most likely use related since upon arrival patient had hematoma at impella site. the pump passed all the post sterile inspection checks. | A4 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5212 | 1220648-2025-46787 | 23125570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125570&pc=OZD | 08/24/2025 |  |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026725332 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 55 KG |  |  | Death | Yes | No | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | Tachycardia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced atrial fibrillation with suction alarms and placement signal low alarms. the patient had a do not resituate order. an echocardiogram was performed and showed the pump positioned deep, the left ventricle appeared collapsed, and the right ventricle was not moving. the physician attempted to reposition the pump but was unsuccessful. the impella was turned off and the patient expired. | [5 empty] | No product was returned for evaluation. the root cause of the cardiovascular insufficiency was unable to be determined due to insufficient clinical details. the cause of the pump in wrong position could not be determined as no clinical details were provided as to how the pump came out of position. the cause of the pump suction was patient condition related as clinical details noted lv was collapsed and pump was deep when suction alarms were triggered. the cause of the optical signal issue was patient condition related per clinical details noting patient condition deteriorating. the pump passed all post-sterile inspection checks. | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5213 | 1220648-2025-31634 | 23125577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125577&pc=OZD | 07/25/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 07/25/2024 | 09/23/2025 |  |  | 11/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431099 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was implanted and exchanged for a impella 5.5 pump to support a patient experiencing acute myocardial infarction and cardiogenic shock. the complainant reported placement signal issues on the pump. echocardiogram confirmed that there was ideal positioning of the impella 5.5 and that the patient¿s hemodynamics were stable. the placement signal monitoring was disable. the patient was sent to icu; however, the patient experienced ventricular tachycardia and was shocked more than 20 times. decision was made to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. partial data logs were returned but the logs at the time of failure are missing. therefore, the root cause of the optical signal issue was not determined since product was not returned and insufficient data logs were returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5214 | 1220648-2025-46788 | 23125651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125651&pc=OZD | 07/28/2025 | 10/24/2025 |  | 09/23/2025 |  |  | 07/28/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690931 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Thrombocytopenia | Aortic Valve Insufficiency/ Regurgitation | Mitral Valve Insufficiency/ Regurgitation | Tricuspid Valve Insufficiency/ Regurgitation | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was escalated to impella 5.5 support and had low hemoglobin, supraventricular tachycardia, hypotension, heparin induced thrombocytopenia, tricuspid, aortic, and mitral valve insufficiency/regurgitation, optical signal issue, pump suction, retrograde pump flow, high purge pressure, and purge system blocked alarms. they required blood transfusions, medication, pump repositioning, and purge cassette change. | [5 empty] | The root cause of the reported tachycardia, thrombocytopenia, high purge pressure and pump suction has been completed. clinical reported purge system blocked at the end of mitral clip placement procedure. clinical reported tissue plasminogen activator (tpa) added to purge which resolved issue. the returned pump was inspected and no biomaterial was observed; no physical abnormalities related to these issues were observed. the pump was run in the lab and high purge pressure and purge system blocked issues did not reproduce. the data logs confirm high purge pressure and purge system blocked alarms. there was no motor current spike before high purge pressure. purge pressure increased while purge flow decreased for 10 minutes until alarms were triggered. issue was sustained for about 3 hours before parameters returned to normal values. the root cause of the issue was likely biomaterial build up as evidenced by no reproduction on returned product and clinical mention of tpa resolving the issue after initiation. the root cause of the tachycardia, thrombocytopenia, and pump suction was not determined as limited clinical information was provided and insufficient data logs were returned. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: correct impella catheter position ¿catheter angled toward the left ventricular apex away from the heart wall and not curled up or blocking the mitral valve¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
5215 | 1220648-2025-46778 | 23125699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125699&pc=OZD | 08/25/2025 |  |  | 09/23/2025 |  |  | 08/25/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704366 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed an impella 5.5 to support a patient awaiting a transplant. the patient had been admitted in with known cardiomyopathy, valve disease, and ejection fraction of 10%. the placement signal had been gradually lowered to be consistently less than 30mmhg less than the non-invasive blood pressure. the position of the impella was verified and volume status was appropriate. no alarms or issues related to sensor drift were noted. the impella flow and motor current remained stable/unchanged. brief suction alarms were noted while ambulating. later that day, lower placement signal was noted and appeared maybe to have migrated just a bit. the next day, the placement signal appeared to be drifting. the team was aware and was monitoring. a brief suction event occurred overnight while the patient was vomiting. the plan was for an echocardiogram in the morning to assess the position. three days later, the patient's nausea was slightly improving and the impella remained stable. two days later, the patient was experiencing more nausea and vomiting. suction events occurred overnight due to the vomiting. no real cause for it is know. gastroenterologist was notified. the next day the patient's condition was unchanged/stable. a couple of suction events occurred again overnight. albumin was given and there were no alarms since. over the next several days. the patient continued to show improvement with nausea/vomiting but was still experiencing suction events with movement. lasix drop was started. the patient ended up having a near syncopal episode. the patient was put in a cardiac chair and was taken back to his room where he was stabilized. bivalirudin was paused due to a nosebleed. the patient remains stable. | [5 empty] | No product was returned for investigation. the cause of the epistaxis was not determined due to insufficient clinical details. the root cause of the suction is most likely due to patient movement. the root cause of the optical signal issue cannot be determined. the pump passed all post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Correction: d6b explant date was inadvertently omitted. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5216 | 1220648-2025-31635 | 23125749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23125749&pc=PYX | 05/11/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 05/11/2024 | 09/23/2025 |  |  | 10/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024413913 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that while delivering support to a patient an impella rp generated differential sensor drift throughout the night. no harm to the patient was reported. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. rp flex was returned for evaluation. upon visual inspection, biomaterial was observed around base of impeller blade and purge gap. electrical testing was performed and sensor lines were within range. biomaterial was removed from pump and pump was run in pressure loop set-up. sensor drift was reproduced when run in loop. data logs were reviewed and lt logs from field show placement signal sensor drift throughout case. no alarms were triggered during case related to placement signal. additionally, a motor current spike was observed on day 8 of support. a speed deviation, drop in pump flow, and increase in purge pressure with a decrease in purge flow were also observed at time of motor current spike. clinical details noted that rp flex flows had decreased over the first night of use and automated impella controller (aic) was displaying 0.0l/min but no change was noted with patient's hemodynamics and was still being supported by the pump. clinical team re-zeroed the pump and flows returned back to normal range. the root cause of the placement signal issue was due to sensor drift of the dp sensor on the rp flex. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5217 | 1220648-2025-31636 | 23126008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126008&pc=OZD | 11/05/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 11/05/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534900 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support. while on support, placement monitoring was disabled due to an unreliable placement signal. echocardiography was performed since a placement signal not reliable alarm occurred. motor current was reviewed and issues without placement signal were identified with the nurse. the placement signal alarm occurred while audio was off. the impella was eventually successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed and the placement signal not reliable alarms (psnr) were confirmed on 11/5-7. immediately after support is started, the pump runs at p-9, pump flows are low and placement signal not reliable triggers. during psnr, signal-to-noise ratio and contrast are affected while light and gain are not, indicating the optical signal is being diminished by environmental factors. this is supported when pump flows are reduced to p-8 and the psnr resolves, with pump flows are within specification again. however, they attempt to turn back up to p-9 and psnr triggers again and pump flows return to being low. on 11/7, psnr resolves as pump flows show improvement. the team then starts to wean the pump, and all optical signal issues, suction alarms, and pump flows resolve. returned product was investigated and the pump's 5s parameters were within specification, and the pump passed the laser continuity test. additionally, the pump was run in a bucket in the lab overnight, and all 5s parameters were stable/within specification. clinical details support the data log review findings based on the report of improved hemodynamics from case start to 11/7. from clinical details, data log review, and returned product, the root cause of the optical signal issue was determined to most likely be patient condition. the root cause wasn't able to be verified as the pump couldn't be uson tested. in the middle of the night when the pump was being run in the lab, there was a transient loss of optical signal data for 1 minute. however, the acute issue resolved on its own and is unrelated to the failure mode seen in the field as it is a console issue, not a failure of the pump. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5218 | 1220648-2025-46789 | 23126011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126011&pc=OZD | 08/25/2025 |  |  | 09/23/2025 |  |  | 08/25/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679132 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the pump was placed in a less-than-ideal location due to the patient being prone to ventricular tachycardia. therefore, the pump was positioned toward the posterior wall to avoid stimulation with the impella pigtail. due to the position, it was advised to the medical team to use p-6 or lower, but support was continued for two days at p-7. on the third day of support, the patient¿s blood potassium level rose, and although there was no hemolyzed urine or anemia, renal damage was suspected due to possible hemolysis, and continuous renal replacement therapy was initiated. the p level of the impella was lowered to p-5 to p-6 and support was continued. it was noted that the patient was stable. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary product & data were not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. device in wrong position: the cause of device in wrong position was unable to be determined as limited clinical details were provided and no product was returned for review. renal failure: the cause of the injury was not determined due to insufficient clinical details. device history review. pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5219 | 1220648-2025-31637 | 23126013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126013&pc=OZD | 12/23/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 12/23/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582105 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following implant, an impella cp pump triggered a placement signal not reliable alarm. hemodynamics were confirmed and the functionality of the pump remained unaffected. the alarm and placement monitoring were subsequently disabled for the course of planned support. the pump was successfully weaned and explanted. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed, and the placement signal not reliable (psnr) alarm was confirmed. from case start, signal-to-noise ratio (snr) was below 3500 and gradually worsened until it hit zero. contrast increased in value and amplitude throughout the case. lamp, gain, and light were unaffected. there were no abnormalities that would suggest patient condition or damage to the pump initiated the diminished optical metrics. no known hard failure patterns of the pump were observed in the data logs. other pump performances on the same console were reviewed and a complaint already existed with similar symptoms. snr gradually varied throughout the case, and the snr and contrast in rt logs during case start looked identical to this one. returned product was investigated and the pump passed the laser continuity test. it was then plugged into the console, 5s parameters were all within specification, and it was run for four hours without any abnormality. upon review of data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5220 | 1220648-2025-31638 | 23126135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126135&pc=OZD | 09/06/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 09/06/2023 | 09/23/2025 |  |  | 03/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346927 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | On proactive call, the nurse complainant informed that the placement signal low alarm was ongoing. the nurse stated echocardiography confirmed placement. pump was not removed due to the placement signal issue. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were returned for investigation. the root cause of the optical signal issue was unable to be determined as no product or logs were returned for analysis. all pertinent information available to abiomed has been submitted at this time.
5221 | 1220648-2025-31639 | 23126156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126156&pc=OZD | 10/20/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 10/20/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503005 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient placed on impella 5.5 support. the 5.5 was placed and after four days, the optical signal was seen to be malfunctioning. the loss of placement signal did not prompt any intervention or explant. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs were evaluated and an optical signal issue pattern of oscillating contrast was observed. no problem was found as the root cause of the optical signal issue, and it is apparent the console was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5222 | 1220648-2025-31640 | 23126173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126173&pc=OZD | 11/01/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 11/01/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562165 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella 5.5 pump began to trigger placement signal position in aorta alarms during patient support. positioning was verified via x-ray and the functionality of the pump remained unaffected. the alarm was disabled and support with the implanted pump continued. the pump was eventually weaned and explanted. there was no patient harm reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were investigated and the position in aorta alarms were confirmed on 11/1. during this time, all 5s parameters are within specification and stable. pump flows are also within specification, however, they're extremely unstable. the final position in aorta alarm triggers early in the morning of 11/2, and from that point forward pump flows start to stabilize as well. clinical details report that x-ray confirmed that the pump was in good position when the alarms were triggering and that positioning monitoring was never disabled. additionally, the patient was on venous-arterial extra corporeal membrane oxygenation and three units of blood were administered during the evening of 11/1. finally, it was confirmed that the console in the field was using software v12.2. product was not returned for investigation. based on clinical details and data log review, the root cause of the optical signal issues was determined to most likely be patient condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5223 | 1220648-2025-31641 | 23126187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126187&pc=OZD | 11/25/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 11/25/2024 | 09/23/2025 |  |  | 06/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359723 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed an impella 5.5 pump to support a patient. the patient had arrested and was placed on the 5.5 pump. the pump remained on despite a loss of optical signal. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. log review showed placement signal going out of range at this time. signal-to-noise ratio dropped to 0, light dropped to 0.5, gain dropped to 1.0, lamp was stable at 100, and contrast was oscillating between 2000/-2000. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5224 | 1220648-2025-31642 | 23126694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126694&pc=PYX | 04/13/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 04/13/2023 | 09/23/2025 |  |  | 01/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024299933 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump, as part of bipella support, was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. it was reported that there was a damaged sensor on the impella rp. there were no flow calculations appearing on the automated impella controller (aic). there was good hemodynamic support by the impella rp; however, the patient suffered from influenza and multiorgan failure due to late admission after myocardial infarction. a decision was made to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for investigation. the cause of the placement signal issue was not determined since product was not returned and no data logs were provided for review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5225 | 1220648-2025-31643 | 23126847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23126847&pc=OZD | 10/11/2024 | 10/10/2025 |  | 09/23/2025 |  |  | 10/11/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-0000-US |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | UNKNOWN |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the automated impella controller (aic) had lost placement signal issues while patient was on support. the controller was held for testing. the aic was not reported to have been removed during support. no further information was provided. | [5 empty] | D4: added device information. | D4: udi # info is unknown this mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data analysis: the console was last used on 8/17 before the complaint date of 10/11 (imc ic9123.pdf). the console was then used again on 10/30. it is unlikely this is the malfunctioning console given the two month gap but efforts to locate the malfunctioning console were unsuccessful. device analysis: the console performed as expected. root cause: there was no issue found during the case. the console functioned/operated to specification. repair: please perform a functional check on the console (0042-7316). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5226 | 1220648-2025-46790 | 23127248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23127248&pc=OZD | 08/24/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704453 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision was made to place impella cp on a patient with severe microvascular disease and left ventricular end-diastolic pressure 38. the impella was inserted per best practices, however was shocked for tachycardia arrhythmia, but was not intubated or coded. fluids were given to optimize. the patient recovered well and labs/hemodynamics looked great. decision to explant was made. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. tachycardia: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5227 | 1220648-2025-46791 | 23127282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23127282&pc=OZD | 08/24/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 08/24/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026706392 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 62 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5 on (b)(6) 2025, the patient experienced lower extremity weakness last night when getting back to bed. code stroke was called and cat scan revealed new right cerebellar territory stroke. the patient does have a history of strokes. not a tissue-type plasminogen activator (tpa), candidate per team. the patient was offered transfer to main campus for clot retrieval interventions, and he declined. echocardiogram performed and no clot visualized around impella catheter. team is attributing stroke to low flow and low cardiac output, state and not to impella usage. partial thromboplastin time goals increased and patient showing improvements in motor function on my rounding this morning. currently it is unclear if this will alter plans for his scheduled left ventricular assist device tomorrow morning. the heart failure feels that it should be delayed but surgery feels it should not, and he would do better with proceeding. | [5 empty] | Correction e4: the investigation of the reported failure mode has been completed. no product was returned for investigation. stroke: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5228 | 1220648-2025-31644 | 23127311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23127311&pc=PYX | 07/23/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 07/23/2024 | 09/23/2025 |  |  | 04/30/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024338006 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. it was reported that the placement signal was absent with no active alarm. motor current and flows were monitored. while on support, the patient had a cerebral study that indicated the patient did not have any brain activity. the patient was made dnr and the family withdrew care. the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. the root cause of the placement signal issue could not be determined as no product or logs were returned for analysis and limited information was provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5229 | 1220648-2025-31645 | 23127730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23127730&pc=OZD | 07/22/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 07/22/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2025486795 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. during support, a placement signal not reliable alarm occurred. the pump positioning could no longer be monitored via the waveform, but the motor waveform and purge conditions remained normal. support continued with the utilized pump and automated impella controller (aic) despite the noted issue. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console logs from the reported day of event are consistent with complaint and show that 7 days after pump start, a ¿placement signal not reliable¿ (psnr) alarm (#1036, due to opm signal invalid) presented. the next day, the psnr condition momentarily self-resolved, which coincided with brief interruption of communication from optical bench, but soon the communication was restored, and psnr condition (and alarm) recurred. the alarm resolved after pump was unplugged 2 days later. ltlog data indicates that during that time optical contrast had erroneous values between -3000 and +3000, and snr was close to 0. this pattern is consistent with loss of light on the optical bench detector, due to optical bench malfunction. console ic5659 was returned to jp fs for investigation and repair. its optical bench was replaced and the issue did not recur. repair: replace optical bench assembly (0042-3015 or equivalent), and perform functional check of the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5230 | 1220648-2025-31646 | 23128883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23128883&pc=PYX | 10/25/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 10/25/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540760 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. it was reported that after four days of support, the pump was alarming for 'placement signal not reliable.' the pump was not explanted. imaging indicated the pump was in an appropriate position. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. due to lack of product and data logs returned, the root cause cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5231 | 1220648-2025-46792 | 23128887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23128887&pc=OZD | 08/25/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 08/25/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711111 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 76 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hematuria | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient transferred from community to tertiary center after admission with st-segment elevation myocardial infarction and coronary microvascular disease the intra-aortic balloon pump was not enough support. during the cp support there was urine color change from possible bladder injury and there was both a subdural hematoma and retroperitoneal hematoma. the patient did get blood products for the intervention but after 6 days care was withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. subdural hematoma: the root cause of the subdural hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. hematoma: the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. hematuria: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5232 | 1220648-2025-31647 | 23128997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23128997&pc=OZD | 04/11/2025 | 10/24/2025 |  | 09/23/2025 |  |  | 04/11/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402264 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella rp pump was implanted for circulatory support. during support, the rp had signal issues and was removed and replaced. the console was returned as well for investigation despite no aic replacement. there was no reported harm or injury to the patient. | [5 empty] | H4 device manufacture date updated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned logs confirmed issue. this console was tested after arriving at fs. reproduced the reported symptoms. the issue may be related to the specific pump type since no consistent issues were found with the returned console. the root cause for the invalid placement signal, prior to insertion into the patient, was not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5233 | 1220648-2025-31648 | 23129735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23129735&pc=OZD | 03/10/2025 | 09/23/2025 |  | 09/23/2025 |  |  | 03/10/2025 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2026631242 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention (hrpci). the pci was successfully supported by the pump despite the optical signal intermittently malfunctioning. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. intermittent placement signal not reliable (psnr) alarms occurred. pump remained in used until case end following successful pt wean off. the pump was returned. returned data logs confirmed psnr alarm. the signal not reliable (snr) was low with spike to placement signal and contrast was pulsatile to snr being low. no other issues with other 5s parameters. the returned pump was unable to reproduce psnr alarms and was run for over 10 minutes while exhibiting nominal parameters. the cause of the optical signal issue was not determined since the returned device functioned/operated to specification during investigation and did not reproduce psnr alarms despite reported issues from the field. for additional details about the relation of the automated impella controller (aic) in the complaint, please refer to 25-42195-2. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5234 | 1220648-2025-31649 | 23129766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23129766&pc=OZD | 03/04/2023 | 09/23/2025 |  | 09/23/2025 |  |  | 03/04/2023 | 09/23/2025 |  |  | 12/31/2024 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291808 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with impella 5.5 for care escalation from an impella cp and extracorporeal membrane oxygenation (ecmo). the team had difficulty in placement but was able to successfully deliver the 5.5. it was reported that there was loss of placement signal; however, this did not result in explantation of the 5.5. the pump remained on support for the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the cause of the optical signal issue was not determined since neither product nor data logs were returned. all pertinent information available to abiomed has been submitted at this time.
5235 | 1220648-2025-31650 | 23129779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23129779&pc=OZD | 05/21/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 05/21/2024 | 09/23/2025 |  |  | 03/31/2025 |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486795 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm occurred during impella cp support. pump positioning could no longer be monitored via the waveform, but the motor waveform and purge conditions remained normal. support continued with the implanted pump despite the noted issue. there was no patient harm resulting from the failure. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the product was returned and inspected. no abnormalities were found with laser continuity, biomaterial, kinks in catheter, or bond issues. no abnormalities related to optical signal were found after running the pump at p-9 for 10 minutes. upon review of the data logs and pump, no problem was found with the pump, and it is apparent the console was the potential cause of the issue. please refer to 24-29642-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5236 | 1220648-2025-31651 | 23129790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23129790&pc=OZD | 11/22/2024 | 09/23/2025 |  | 09/23/2025 |  |  | 11/22/2024 | 09/23/2025 |  |  |  |  | 09/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED , INC. |  |  | US |  | 2025541224 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant in the eu had a cp pump placed for a high-risk percutaneous coronary intervention (hrpci). the pump had lost placement signal during the support. there was no harm or injury and the event did not cause any intervention. the pump was weaned and explanted per plan and after the failure of the hrpci. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump and data log were not provided for this case. according to the clinical notes, the psnr alarm was triggered on the day of implant, and the pump continued to run for another 77 hours. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5237 | 1220648-2025-46728 | 23108380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23108380&pc=OZD | 07/22/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 07/22/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654146 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 98 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | Increase in Pressure |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered flow saturation, high purge pressure, thrombus, suction, and bleed during impella support. during start of procedure, floseal/bioglue was applied to graft anastomosis for bleeding control and bulb drain inserted into pocket prior to closure. the patient was later on p9 with flows at 5.5 and rvad 4.5l. then, the patient's purge pressure spiked to around 1080 with purge fluid (pf) less than 1. tissue plasminogen activator (tpa) at 2mg/50ml was started through the purge and was reviewed with the nurse. during support, the patient also encountered several instances of suction. additionally, patient showed signs of flow saturation with an increasing motor current (mc). the impella was removed from the right axillary and a clot was seen around the impeller/motor housing. the patients outcome is unknown. | [5 empty] | There investigation was completed. no product was returned for analysis. high purge pressure: the logs confirm purge pressure high alarms and show a gradual rise in purge pressure over about 2 hours before alarm was triggered. there were no motor current spikes outside p-level changes before high purge pressure. motor current values also trended upwards and pump flow became saturated. clinical mentions tpa was added to the purge and clot observed upon explant. the cause of the issue was most likely biomaterial as evidenced by the gradual rise in purge pressure with effects on pump flow and motor current and by the clot visualized upon explant. high or saturated pump flow : the data logs confirm saturated pump flow and show a gradual rise in purge pressure and flow saturation with pump flowing above the expected range for p9 at 6.5l/min. there were no motor current spikes outside p-level changes before saturated pump flow. there were high purge pressure alarms and purge system blocked alarms after a gradual rise in purge pressure and motor current. clinical mentions tpa was added to the purge and clot observed upon explant. the cause of the issue was most likely biomaterial as evidenced by the gradual rise in purge pressure with effects on pump flow and motor current and by the clot visualized upon explant. pump suction: several suction events during support. the data logs confirm suction alarms during support. the logs do not show any motor current spikes. suction alarms occurred at different points during support with some instances self-resolving. the cause of the issue was not determined as no product was returned for analysis and limited information was provided. access site leak: access site bleeding around sheath reported. floseal/bioglue applied to graft anastomosis for bleeding control. limited information was provided to determine cause of access site leak. the cause of the issue was not determined as limited clinical information was provided. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. a device history review was completed and showed the device passed all post sterile inspection checks.
5238 | 1220648-2025-46730 | 23109502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23109502&pc=OZD | 06/23/2025 |  |  | 09/22/2025 |  |  | 06/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681618 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Fever | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support after developing a chronic driveline infection with resistant pseudomonas that began to involve central left ventricle assist device (lvad) components. the patient was brought to the operating room to explant the lvad and implant an impella 5.5. it was reported that after impella 5.5 support was initiated, there were alarms observed for flow saturation and then the pump abruptly stopped functioning. the decision was made to replace the impella 5.5. upon the explant of the pump, a piece of muscle was found entrained in the motor housing. a new impella 5.5 was placed successfully. during use of the second pump, the patient experienced a fever that began on the 8th day of support. on the 44th day of support, it was noted that the patient still had a sternal infection the medical team was attempting to resolve prior to placing a new lvad. on the 46th day of support, blood cultures were still noted to be positive for bacterial growth, and that the impella 5.5 p level had to be increased for a low mixed venous oxygen saturation level. on the 56th day of support, the p level had to be decreased due to suction alarms. the patient received albumin as a result of the suction alarms to increase volume. the following day, it was noted that the patient was receiving antibiotics for an ongoing infection. on the 64th day of support, it was reported that there were more suction alarms. as a result, the p level of the pump was lowered. three days later, the patient was taken to the operating room to remove the impella 5.5 and place a new lvad. during the procedure, the medical team was unsuccessful in placing the lvad due to anatomical concerns. a new impella 5.5 was placed in addition to a right ventricular assist device (rvad) as bridge to explore alternative lvad options or possible heart transplant. two days later, the rvad was explanted, and it was noted that the same day the patient was taken for computed tomography. it was reported that following surgery, the patient experienced a hemorrhagic stroke. the following day, on the third day of the latest impella 5.5 pump support, the patient¿s care was withdrawn and the patient expired. this complaint involves three impella 5.5 devices: pump 1: impella 5.5 with serial number(b)(6), pump 2: impella 5.5 with serial number (b)(6), pump 3: impella 5.5, with serial number (b)(6). this report specifically addresses the second impella 5.5 with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h.10 for the related report numbers. | [5 empty] | Medical safety reviewed the event. the event describes a pump stop, which resulted in replacement of the device. this serious injury is attributable to a piece of muscle was found entrained in the motor housing. replacement pump was in for at least 64 days. the patient tested positive for bacterial growth and also there was pump suction. this is a serious injury attributable to the pump's association to the infection. the second impella 5.5 pump was replaced. the patient underwent device removal of the impella 5.5 and placement of a left ventricular assist device which was unsuccessful. therefore, a third impella 5.5 device was placed. the next day after start of this support, it was reported the patient had a hemorrhagic stroke post-surgery. in this case, it is unknown which device procedure, if any, contributed to the hemorrhagic stroke and therefore the impella 5.5 pump 3 cannot be dissociated; this pump was in place at the time of the event. however, it is noted care was withdrawn which appears to have led to the demise outcome. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical devices reports for the other two impella 5.5 device in which one is associated with the demise outcome. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H11 as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. nvestigation summary: the investigation has been completed. infection: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the injury was not determined due to insufficient clinical details. pump suction: no product was returned. speed was dropped from p9 to p7 and then to p5 due to suction events. after reviewing logs, suction events was observed. no motor current spikes were outside of p-level changes. placement signal is variable. the cause of pump suction cannot be determined due to insufficient clinical details and no product was provided. device history review: the device passed all post sterile inspection checks.
5239 | 1220648-2025-46732 | 23109685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23109685&pc=OZD | 08/13/2025 | 09/12/2025 |  | 09/22/2025 |  |  | 08/13/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704103 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 91 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp experienced ventricular fibrillation requiring defibrillation. the patient experienced a lack of distal limb perfusion so the impella cp was explanted. soon after, the family decided to withdraw care, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5240 | 1220648-2025-46733 | 23110036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23110036&pc=PYX | 07/23/2025 |  |  | 09/22/2025 |  |  | 07/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625189 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 149 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death |  |  |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the decision was made to move patient support from an impella 5.5 to impella rp flex for a patient. however, the impella rp flex did not improve the patient's condition. the patient arrested on the table before the impella rp flex was placed requiring multiple shocks, return of spontaneous circulation was noted but blood pressure was never recovered. the impella rp flex was placed and the patient was taken to the intensive care unit. multiple high dose vasopressors were given and a second dose of cyanokit was given which did not help the patient hemodynamically. per physician, the patient was not an extracorporeal membrane oxygenation candidate. the family ultimately withdrew care on this critically ill patient later that evening. pumps were not saved and the patient expired. | [5 empty] | The impella device was not returned, and the investigation has been completed. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. a device history review was done and passed all the post sterile inspection checks. | Medical safety officer reviewed the event and noted the following: physiologically depleted patient in cardiogenic shock on multiple devices. there is no documentation that the ectopy was of clinical significance given the documentation at hand. the event is being filed as death because the device (in this case the impella 5.5) cannot be excluded, although the impella devices are highly unlikely associated based on the patient's condition. as the ectopy is more likely to come from the impella 5.5 versus the impella rp flex, the impella 5.5 is death type of reportable event, and the impella rp is serious injury. change in reportability: after review of the case by medical safety, it was determined the impella rp flex is a serious injury type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. additional correction made for the report was concomitant device impella 5.5 was added to section 10 accordingly. note: h6: device problem/component/investigation coding remains unchanged. related medical device reports: this impella rp flex report is one of two devices associated with the event. refer to the related manufacturer report number in h10 for the report on the impella 5.5.
5241 | 1220648-2025-46734 | 23110296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23110296&pc=OZD | 08/26/2025 |  |  | 09/22/2025 |  |  | 08/26/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595415 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 83 KG | Non Hispanic |  | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp exhibited amber urine and potential vessel damage. additionally, the silver latch that connects the pump to the automated impella controller (aic) was broken. no patient harm was reported. this report represents the pump. another reported was submitted to represent the aic. refer to section h.10 for the related report number. | [5 empty] | No product was returned for investigation. the cause of the hematuria was not determined due to insufficient clinical details. the cause of the vascular dissection was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the device passed all post sterile inspection checks. | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5242 | 1220648-2025-46735 | 23111155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23111155&pc=OZD | 08/22/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678215 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 91 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | IMPELLA RP FLEX. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted in with cardiac history and prior proteck duo had bipella support of an impella 5.5 and an impella rp flex placed. the pump was supporting when after 22 days, the pump was explanted with hemolysis concerns that were treated by explant and dialysis. the impella rp flex had already been explanted and the patient remained on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of the reported failure mode has been completed. hemolysis/mechanical interaction with blood: the data logs for the 5.5 show no related abnormalities and no positioning issues were indicated. the cause of the issue was determined to be traced to another device as issue occurred after rp flex implant and ingestion and resolved post-rp flex explant. please see mwr 18092025-0001935455 for an investigation into the rp flex device history review: the device passed all the post sterile inspection checks.
5243 | 1220648-2025-31603 | 23111370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23111370&pc=OZD | 03/06/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 03/06/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599846 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 115 KG |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump escalated to the 5.5 pump for a patient. the 5.5 was supporting for over two weeks when the pump's optical signal was lost. it was also noted two weeks after implant that the patient developed spontaneous retroperitoneal bleed and multi-organ failure, both unrelated to the device/use. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. device implanted (b)(6) 2025 as escalation of care. after 13 days of care, placement signal not reliable (psnr) alarm triggered and optical sensor signal cut out. the pump was not returned. data log review showed psnr alarm beginning on (b)(6) 2025 at 23:00. immediately prior to the alarm, the 5s parameters were checked and showed oscillating contrast. there was no issue found during the case related to the pump. upon review of the data logs and complaint history of the automated impella controller (aic), the oscillating contrast issue linked to complaint console confirmed from reference complaint. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5244 | 1220648-2025-31604 | 23111422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23111422&pc=OZD | 03/07/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 03/07/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574136 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, the pump was having intermittent impella in aorta alarms mixed with impella position unknown notifications. position was verified. placement signal alarms were disabled. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. there was no product returned. the data logs show that the signal not reliable varies while the placement signal spikes, lamp steps up, while gain and light are stable. this occurs during a p-level changes. the placement signal would rapidly spike while the motor current would lose pulsatility. the alarm was correctly met the triggering criteria. echo confirmed the position in the correct place and the impella position in aorta alarms were intermittently triggering leading to infer that the alarm was falsely triggered. however, the triggering of these alarms is not known. due to lack of product returned and sufficient clinical details provided, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5245 | 1220648-2025-31605 | 23111467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23111467&pc=OZD | 04/24/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 04/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025605566 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that false impella in aorta alarms appeared during patient support. the staff was walked through disabling the placement signal monitoring and support continued uninterrupted. decision made to withdraw care, patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. after 5 days on support, false position in aorta alarms occur despite an echcardiogram which verified positioning. the device used for another 12 days without another reported incident. the pump was not returned. data logs were recovered. data log review showed no known abnormality patterns linking this issue to the automated impella controller (aic). all impella position in aorta alarms occurred within the same hour before placement monitoring was disabled. placement monitoring disabled was turned off within 1 hour and no impella position in aorta alarms occurred again. the cause of the optical signal issue was not determined since product was not returned, insufficient clinical details, with data log patterns that did not link the problem to the aic and did not fit any known problems. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5246 | 1220648-2025-31606 | 23111613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23111613&pc=PYX | 11/25/2024 | 09/22/2025 |  | 09/22/2025 |  |  | 11/25/2024 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025561565 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp flex support. while on support, placement signal unreliable alarm appeared. audio alarms were silenced. patient expired while on support. | [5 empty] | Data logs were reviewed, and it was found that the placement signal not reliable alarms were optical signal related, so the failure mode was changed from placement signal to optical signal issues. data logs are only available for the first 3 days of support. from case start, pump flows struggle to achieve specification for p-9 (3.7-4.2 l/min). the motor current is quite unstable while running at p-9, there are 23 triggers of the suction alarm during the first 5 hours of support, and when placement signal metrics are active, cvp is a negative value. these are all indications of environmental factors altering the optical signal. clinical details report that placement signal not reliable began triggering during heart manipulation when closing the patient's chest, which could have also played a role in the noise in the optical signal early in the case. product was not returned for investigation. based on clinical details and data log review, the root cause of the optical signal issue was determined to most likely be patient condition. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5247 | 1220648-2025-46739 | 23112618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23112618&pc=OZD | 08/13/2025 | 10/30/2025 |  | 09/22/2025 |  |  | 08/13/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026704103 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 91 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp experienced ventricular fibrillation requiring defibrillation. the patient experienced a lack of distal limb perfusion so the impella cp was explanted. soon after, the family decided to withdraw care, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported limb ischemia and ventricular fibrillation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and ventricular fibrillation could not be determined.
5248 | 1220648-2025-46738 | 23112769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23112769&pc=OZD | 07/29/2025 |  |  | 09/22/2025 |  |  | 07/29/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622556 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hematoma | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported impella cp support was being escalated to impella 5.5 support for a patient admitted in with acute myocardial infarction/ cardiogenic shock. the impella 5.5 was supporting despite hematoma, hematuria and hemolysis being reported. the patient had been treated with albumin and red blood cell infusions. the pump remained on for 69 hours when it was explanted. the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
5249 | 1220648-2025-46736 | 23113341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23113341&pc=OZD | 08/23/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026663314 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, a patient¿s right hand became cool and dusky overnight. initially, there was concern that the impella may be occlusive in the artery, but it was then realized that the catheter is significantly smaller than her axillary artery. an ultrasound of the extremity was performed and showed pulsatile flow in the distal brachial artery. the team was leaning towards this being some type of thrombus embolic event. vascular surgery was consulted for recommendations. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of the reported failure mode has been completed. thrombosis/ ischemia : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5250 | 1220648-2025-46737 | 23113483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23113483&pc=OZD | 08/23/2025 | 10/22/2025 |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688237 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a 5.5 pump placed to support the patient who had care escalation from the cp and extracorporeal membrane oxygenation the 5.5 was supporting when there was positional issues and pump stopped after 3 days. the pump was replaced and support proceeded on a 2nd 5.5 pump. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported pump stop has been completed. data log shows a sudden pump stop with a 119 impella stopped alarm. the pump could not be restarted, and a motor current spike of 30,000 was observed. a slight increase in purge pressure was noted at the time of the pump stop. the root cause of the pump stop issue was electrical open lines inside the red handle, most likely due to unintentional user error, based on returned product investigation. h6 type of investigation code b01 has been added.
5251 | 1220648-2025-46729 | 23113509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23113509&pc=OZD | 06/16/2025 |  |  | 09/22/2025 |  |  | 06/16/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681620 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 79 KG | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Flow Rate | Pumping Stopped | [5 empty] | Unspecified Tissue Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support after developing a chronic driveline infection with resistant pseudomonas that began to involve the central left ventricle assist device (lvad) components. the patient was brought to the operating room to explant the lvad and implant an impella 5.5. it was reported that after impella 5.5 support was initiated, there were alarms observed for flow saturation and then the pump abruptly stopped functioning. the decision was made to replace the impella 5.5. upon the explant of the pump, a piece of muscle was found entrained in the motor housing. a new impella 5.5 was placed successfully. during the use of the second pump, the patient experienced a fever that began on the 8th day of support. on the 44th day of support, it was noted that the patient still had a sternal infection that the medical team was attempting to resolve prior to placing a new lvad. on the 46th day of support, blood cultures were still noted to be positive for bacterial growth, and that the impella 5.5 p level had to be increased for a low mixed venous oxygen saturation level. on the 56th day of support, the p level had to be decreased due to suction alarms. the patient received albumin as a result of the suction alarms to increase volume. the following day, it was noted that the patient was receiving antibiotics for an ongoing infection. on the 64th day of support, it was reported that there were more suction alarms. as a result, the p level of the pump was lowered. three days later, the patient was taken to the operating room to remove the impella 5.5 and place a new lvad. during the procedure, the medical team was unsuccessful in placing the lvad due to anatomical concerns. a new impella 5.5 was placed in addition to a right ventricular assist device (rvad) as bridge to explore alternative lvad options or possible heart transplant. two days later, the rvad was explanted, and it was noted that the same day the patient was taken for a computed tomography. it was reported that following surgery, the patient experienced a hemorrhagic stroke. the following day, on the third day of the latest impella 5.5 pump support, the patient¿s care was withdrawn and the patient expired. this complaint involves three impella 5.5 devices: pump 1: impella 5.5 with serial number (b)(6). pump 2: impella 5.5 with serial number (b)(6). pump 3: impella 5.5, with serial number (b)(6). this report specifically addresses the first impella 5.5 with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h.10 for related report numbers. | [5 empty] | Medical safety reviewed the event. the event describes a pump stop, which resulted in replacement of the device. this serious injury is attributable to a piece of muscle that was found entrained in the motor housing. a replacement pump was in for at least 64 days. the patient tested positive for bacterial growth and also there was pump suction. this is a serious injury attributable to the pump's association to the infection. the second impella 5.5 pump was replaced. the patient underwent device removal of the impella 5.5 and placement of a left ventricular assist device which was unsuccessful. therefore, a third impella 5.5 device was placed. the next day after start of this support, it was reported that the patient had a hemorrhagic stroke post-surgery. in this case, it is unknown which device procedure, if any, contributed to the hemorrhagic stroke and therefore the impella 5.5 pump 3 cannot be dissociated. this pump was in place at the time of the event. however, it is noted care was withdrawn which appears to have led to the demise outcome. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report numbers as noted in section h10 for the reports for the other two impella 5.5 devices in which one is associated with the demise outcome. | The impella device was not returned, and the investigation has been completed. the pump passed all post-sterile inspection checks. no product was returned for evaluation. soft tissue injury: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. saturated pump flow: clinical details noted that upon pump activation, pump flow was saturated. additionally, a piece of muscle was observed in the motor housing on explant. images from the field showed biomaterial in the outflow. data logs were reviewed and approximately 1 minute into support when p-level is increased, a small motor current spike then fully saturated flows is observed. pump flow remained fully saturated for approximately 12 minutes before p-level was turned down then turned off. the cause of the saturated pump flow was most likely due to biomaterial ingestion based on data log review showing a motor current spike prior to saturated flows and clinical details noting biomaterial on explant. pump stop: clinical details reported a pump stop soon after the pump was inserted and start. saturated pump flow noted prior to the pump stop. data logs were reviewed and saturated pump flow was observed during support but no pump stop alarms were noted during pump run on the logs, only an alarm indicating reverse flow of the pump after the pump was stopped at time of console shutdown. pump flows were fully saturated for approximately 12 minutes before p-level was turned down to p-0 and turned off. the cause of the pump stop could not be definitively determined as no pump stop alarms were present on the returned data logs and limited clinical details were provided. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
5252 | 1220648-2025-46742 | 23113589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23113589&pc=OZD | 08/24/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 08/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681970 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella cp for mechanical circulatory support. after the pump was implanted, a small, stable hematoma was noted at the groin site. a femostop was in place with no additional pressure. the patient became febrile; cultures negative and cat scan completed. antibiotics started for fever and were later escalated | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. hematoma: the root cause of the hematoma was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. fever: the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5253 | 1220648-2025-46740 | 23113776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23113776&pc=OZD | 08/23/2025 | 09/22/2025 |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619143 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 120 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented with anterior st-elevation myocardial infarction and the decision was made for an impella cp to be placed. the right femoral artery (rfa) was accessed and the impella cp was placed without issues. during support, it was reported that the patient had developed a rfa hematoma. the nurse immediately applied pressure to the site, which she manually compressed out. no bleeding or bruising was noted after. additionally, it was noted that there were multiple suction alarms throughout the evening and into the morning requiring multiple fluid boluses. the physician was please with left ventricle recovery and the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. hematoma: the cause of the hematoma was not determined due to insufficient clinical details. pump suction: after reviewing the logs, suction alarms along with a trend in purge and placement signals were observed.. the cause of pump suction was most likely patient condition related since hemostasis was achieved after giving multiple fluid boluses. device history review: the device passed all post sterile inspection checks.
5254 | 1220648-2025-46743 | 23115281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23115281&pc=OZD | 08/24/2025 |  |  | 09/22/2025 |  |  | 08/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694054 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 86 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was implanted to support a patient with st elevated myocardial infarction. while on support the patient experienced bleeding at the right femoral artery access site. hemostasis was achieved by holding manual pressure, adding a perclose, placing a femstop, and transfusing blood products. systemic heparin was withheld. additionally, there were suction alarms which were resolved by adjusting the pump's p level. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The product was not returned for investigation. the cause of the bleeding was not determined due to insufficient clinical details. the cause of the suction alarms was most likely an adverse event related to patient condition, as the placement signal is seen trending downward at the time of the alarms, and clinical reported blood products being given, which was followed by an increase in placement signal, but it continued to decrease as the case progressed. the device passed all post sterile inspection criteria.
5255 | 1220648-2025-31607 | 23116449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23116449&pc=OZD | 03/24/2023 | 09/22/2025 |  | 09/22/2025 |  |  | 03/24/2023 | 09/22/2025 |  |  | 08/31/2024 |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272457 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient with acute myocardial infarction and cardiogenic shock. during use, the optical sensor appeared to fail, disabling the placement monitoring function. despite this, the impella pump continued to provide effective hemodynamic support, and its operation was not otherwise affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for investigation. data logs were returned and evaluated. data logs showed a steady downward trend in placement signal indicative of patient condition despite no changes in the 5s parameters throughout the run confirming that there was no optical system defect. per the given clinical information, the root cause of the optical signal issue was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
5256 | 1220648-2025-31608 | 23116502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23116502&pc=OZD | 06/15/2023 | 09/22/2025 |  | 09/22/2025 |  |  | 06/15/2023 | 09/22/2025 |  |  | 11/30/2024 |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. following shockwave therapy, the pump lost the aortic placement signal. hemodynamics remained stable and pump functionality was unaffected. the device was subsequently weaned and explanted without issue. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned for investigation. no physical abnormalities were observed. data log shows the placement signal not reliable alarm along and all 5s parameter was down which indicate hard failure of the sensor. placement signal was lost immediately after the shockwave device used. the cause of the optical signal issue was a damaged sensor due to the shockwave device interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5257 | 1220648-2025-31610 | 23116596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23116596&pc=OZD | 09/30/2023 | 09/22/2025 |  | 09/22/2025 |  |  | 09/30/2023 | 09/22/2025 |  |  | 04/30/2025 |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient with acute myocardial infarction and cardiogenic shock. during use, the device experienced optical signal loss. despite this, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the clinical detailed that the placement signal went out upon shockwave pulsing. the pump was not returned. the data logs were reviewed and confirmed placement signal not reliable alarm. the root cause of the issue was a damaged sensor due to interaction with the shockwave device. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5258 | 1220648-2025-31611 | 23116621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23116621&pc=OZD | 10/24/2023 | 09/22/2025 |  | 09/22/2025 |  |  | 10/24/2023 | 09/22/2025 |  |  | 06/30/2025 |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024382494 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient with acute myocardial infarction and cardiogenic shock. during use, the device experienced optical signal loss. despite this, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product returned and insufficient data logs available. the root cause of the optical signal issue was not determined due to no product nor sufficient logs. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5259 | 1220648-2025-46746 | 23117556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117556&pc=OZD | 08/22/2025 |  |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660907 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Patient Device Interaction Problem | [5 empty] | Hematoma | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing for weeks for the patient admitted in with known cardiac history. during support there was this known alarm but still had a waveform. now the waveform is not visible. it was further reported that the was having runs of ventricular tachycardia. additional the patient was noted to have a hematoma in the hip in which anticoagulation was held. there was also some concerns for hemolysis with pfhgb 60 and urine yellow. the patient remains in stable condition and on support. | [5 empty] | D6b. device explant date was added as additional information was provided. the investigation was completed and no product returned. optical signal issue - after 5 days on support, psnr alarms occurred. waveforms eventually became not visible. no change to hemodynamics, flow, and mc. pump remained in use. data log review showed psnr alarm occurred with ps railing to 3000mmhg before returning after 18 hours. no ps issues for the remainder of the case. during the ps blackout, snr drops to zero, contrast barcodes, gain decreases, light decreases, lamp increases to 100. upon review of the data logs and cross functional review, it is apparent that the console is the potential cause of the issue. mechanical interaction with blood (hemolysis) - after 24 days on support, elevated pfhb noted in patient. urine remained yellow. pfhb remained elevated for the remainder of the case, when labs were taken, with no clinical details relevant to hemolysis. red urine never noted. the cause of the mechanical interaction with blood was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. vt- the causes of the ventricular tachycardia were not determined due to insufficient clinical details. hematoma - the cause of the hematoma was not determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5260 | 1220648-2025-46747 | 23117569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117569&pc=OZD | 08/23/2025 |  |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676289 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 163 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 the patient had an episode of ventricular tachycardia required direct current cardioversion and amiodarone infusion. final measurement 3.9. medical doctor aware of shallow placement and attempted to reposition but was causing ectopy, decision was made to leave shallow. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). no product was returned. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: no product was returned. echo showed impella at 5.9cm. the cause of the positioning issue was unable to be determined due to insufficient clinical details.
5261 | 1220648-2025-46748 | 23117584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117584&pc=OZD | 08/24/2025 | 11/25/2025 |  | 09/22/2025 |  |  | 08/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608268 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male | 68 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered air in purge and optical signal issues during impella 5.5 support. it was noted that there was air in the purge system, however, alarm was solved independently. additionally, there was intermittent but persistent impella in aorta alarm despite pulsatile but damped motor current. echo confirmed correct left ventricle (lv)-aorta impella positioning. there was no patient harm. | [5 empty] | D.4 revised primary udi number as it was entered incorrectly on the initial report that was submitted. e.1 added initial reporter address line 2 as it was omitted from the initial report that was submitted. g.1 added manufacturer site postal code as it was entered in the incorrected field on the initial report that was submitted. investigation summary: the investigation was completed. no product returned for evaluation. optical signal issue: clinical details noted intermittent but persistent "impella in aorta" alarm despite pump in proper position. data logs were reviewed and "impella in aorta" alarms were noted. no hard failures of the optical sensor or motor current spikes to indicate possible ingestion were observed. the cause of the optical signal issue could not be definitively determined as no product and limited clinical details were provided. air in purge: clinical details noted that an "air in purge" was triggered on support and was resolved independently. data logs were reviewed and confirm "air in purge" alarms noted. a de-air procedure was completed to resolved the "air in purge" alarms. no purge related procedures were performed prior to "air in purge" alarms on that day. the cause of the air in purge issue could not be determined as no product was returned for evaluation. device history review: the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information provided: the impella device was discarded by the customer; therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
5262 | 1220648-2025-46745 | 23117598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117598&pc=OZD | 08/24/2025 |  |  | 09/22/2025 |  |  | 08/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026700788 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 106 KG | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella 5.5 was implanted in a patient needing a heart and kidney transplant. after four weeks of impella 5.5 support a high purge pressure and purge system blocked alarm occurred. tissue plasminogen activator was added to the purge and was running for one hour before the surgeon decided to replace the impella. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The pump and a purge cassette were returned for investigation. the root cause of the high purge pressure is due to biological material buildup. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5263 | 1220648-2025-46750 | 23117661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117661&pc=OZD | 08/27/2025 |  |  | 09/22/2025 |  |  | 08/27/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194748 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the selector knob of an automated impella controller was not working. | [5 empty] | There was no patient involved in the reported event. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5264 | 1220648-2025-46751 | 23117665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117665&pc=OZD | 08/27/2025 |  |  | 09/22/2025 |  |  | 08/27/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1559426 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure on the automated impella controller. the battery was empty and impossible to recharge. there was no clinical procedure involvement. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5265 | 1220648-2025-46752 | 23117693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117693&pc=OZD | 08/15/2025 |  |  | 09/22/2025 |  |  | 08/22/2025 | 09/22/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316955 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an "unstable position sensing signal" alarm occurred on an automated impella controller (aic) during impella support. the aic was replaced to continue patient support. no patient harm was reported. | [5 empty] | . the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the console was powered on and a pump was driven at a flow level of 9 for approximately 3 hours. the only alarms occurring in this time were yellow suction alarms, to be expected when testing at high flow levels. root cause: no issues were found with (b)(4), as the reported failure mode was confirmed to have occurred only on the initial console, (b)(4).
5266 | 1220648-2025-46744 | 23117767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23117767&pc=OZD | 08/23/2025 |  |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658743 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had quadruple coronary artery bypass graft and after this procedure an impella 5.5 was placed. on day two of impella 5.5 support, the impella had suction alarms. an echocardiogram was performed and inlet was noted to be pointed toward the septum. the medical doctor came in and repositioned the impella 5.5 and suction alarms were resolved. on day five of impella 5.5 support, the patient had runs of ventricular tachycardia that would break themselves with no intervention. one unit of blood was administered. it was further reported that the patient improved hemodynamically heart wise and impella 5.5 was removed as planned. the patient survived. | [5 empty] | The investigation of the reported failure mode has been completed. the product was not returned for investigation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: the cause of the positioning issue was unable to be determined due to insufficient clinical details. pump suction: data log review confirmed suction alarms occurring on the first day of the case. at this time, the placement signal decreased in magnitude, triggering the suction alarms. however, it resolves after about 46 minutes as a result of it being repositioned. there were no spikes in motor current observed. the cause of the pump suction was most likely improper positioning, as clinical reported that the pump was improperly positioned at the time of suction, which was confirmed by the drop in placement signal seen in the data logs which resolved for the remainder of the case. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5267 | 1220648-2025-46753 | 23118077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118077&pc=OZD | 08/31/2025 |  |  | 09/22/2025 |  |  | 08/31/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1280703 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that automated impella controller (aic) battery failure issues were encountered when the aic was turned on. patient harm was not reported. | [5 empty] | Patient data was not provided. section a was left blank. console logs were reviewed and confirmed the reported failure. the returned aic was tested on a lab bench where the failure mode was successfully reproduced. battery 1 was found to be not charging. after replacing the original batteries with known good ones, no battery-related alarms or issues were observed, and both the power battery manager and batteries operated within specifications. furthermore, the aic was run with a known good pump and purge cassette for an hour without any issues. the battery kit was replaced and a full functional check was performed on the aic before returning the device to the customer.
5268 | 1220648-2025-46755 | 23118087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118087&pc=OZD | 09/04/2025 |  |  | 09/22/2025 |  |  | 09/04/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1367614 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that upon starting the automated impella controller (aic) a "system self check failed" occurred. a replacement aic was used for patient support. | [5 empty] | The reported event did not occur during patient use. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5269 | 1220648-2025-46756 | 23118095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118095&pc=OZD | 07/02/2025 | 01/30/2026 |  | 09/22/2025 |  |  | 09/15/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1450701 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an automated impella controller (aic) was pulled from use. this occured during patient use. the procedure was successfully completed with a new console. there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5270 | 1220648-2025-31609 | 23118291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118291&pc=OZD | 08/16/2023 | 09/22/2025 |  | 09/22/2025 |  |  | 08/16/2023 | 09/22/2025 |  |  | 05/31/2025 |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354625 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient with acute myocardial infarction and cardiogenic shock. during use, the device exhibited an ¿impella in aorta¿ alarm. echocardiography was performed, and pump position was verified. despite the alarm, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. the data logs were returned, and 22 false alarms were found. the cause of the optical signal issue was most likely patient condition related due to software with false iimpella position in aorta alarm observed. all pertinent information available to abiomed has been submitted at this time.
5271 | 1220648-2025-31612 | 23118300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118300&pc=OZD | 10/27/2024 | 09/22/2025 |  | 09/22/2025 |  |  | 10/27/2024 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541816 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient with acute myocardial infarction and cardiogenic shock. during use, the device experienced placement signal not reliable alarms. despite this, the impella continued to provide effective hemodynamic support, and its function was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product not returned. data logs show sudden placement signal not reliable alarm. this matches the pattern for damaged fiber. the root cause of the optical signal issue is most likely a damaged optical fiber line but the exact cause and location of the damage cannot be determined due to no product returned and lack of clinical information. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5272 | 1220648-2025-46749 | 23118458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118458&pc=OZD | 08/24/2025 |  |  | 09/22/2025 |  |  | 08/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026703210 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 72 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was implanted to support a patient with st elevated myocardial infarction. while on support, a large left ventricle thrombus was found on an echocardiogram. no intervention was performed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | D.4 revised primary udi number as it was previously entered incorrectly on the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." the investigation has been completed. the root cause of the thrombosis was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5273 | 1220648-2025-46754 | 23118495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118495&pc=OZD | 09/04/2025 |  |  | 09/22/2025 |  |  | 09/04/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1647627 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during a clinical procedure the automate impella controller (aic) incorrectly detected a pump plug in the ¿start case" menu. the aic was replaced to continue patient care. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. | The console was returned for testing, and the reported failure mode was not reproduced. the console successfully detected the pump and purge cassette, as well as the purge disc. during the investigation, it was found that the rotary encoder (push button) was inoperable. further testing of the console revealed that the root cause was a defective power battery manager (pbm) printed circuit assembly (pca). replacing the old pbm with a known good unit resolved the issue, which was not related to the reported complaint. further the console was then tested with a known good pump and purge cassette, and it functioned as expected. root cause: the root cause for "aic incorrectly detected pump plug in the ¿start case menu" was due to an impellatronic pca malfunction. the impellatronic printed circuit assembly, the power battery manager, and the rotary encoder were replaced. then a full functional check was performed before returning the device to the customer. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5274 | 1220648-2025-46757 | 23118759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118759&pc=OZD | 08/23/2025 |  |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720199 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 99 KG | Non Hispanic | Asian | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | High Blood Pressure/ Hypertension | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was reported that the patient was unstable. hypotensive refractory to volume and escalation of vasoactive. additionally, there was expanding hematoma at site for which 1 unit red blood cells were given. pressure held. transthoracic echocardiogram showed impella inlet to be quite shallow, 2cm and oriented toward pap muscle. it was noted that given orientation of inlet and angled aortic takeoff, attempts to reposition may place impella in worse spot and patient dependent on flow. however, the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Correction b1, d4. the investigation of the reported failure mode has been completed. clinical symptoms (high blood pressure/ hypertension): the cause of the injury was not determined due to insufficient clinical details. device in wrong position: the pump was not returned for investigation. data log review was not required for this case run. the cause of the positioning issue was not determined, as sufficient clinical details were not provided and the product was not returned for investigation. access site adverse event (hematoma): the cause of the access site adverse event was not determined, as the product was not returned and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5275 | 1220648-2025-46758 | 23118762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118762&pc=OZD | 08/23/2025 |  |  | 09/22/2025 |  |  | 08/23/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2025552284 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 110 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. placement signal (ps) initially good but migrated to in aorta with dampened motor current (mc) . echo called to bedside stat and physician assisted in advancing impella back in position under echo guidance. final measurement at 3.2 cm across valve once slack removed. ps back to appropriate with pulsatile mc. patient only on amio gtt, hypertensive. awake and alert on nasal cannula. plan to place arterial line. received 16,000 heparin and groin is quite oozy. pending coags but keeping heparin gtt on hold for now. impella cp successfully weaned and explanted and patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5276 | 1220648-2025-46759 | 23118767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23118767&pc=OZD | 08/24/2025 |  |  | 09/22/2025 |  |  | 08/24/2025 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691947 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 85 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient presented with disorientation, acute renal failure, and hypotension. septic shock was of concern which caused myocarditis. the patient was taken to the catheterization lab for right heart catheterization after intubation and increased drips. the patient would undergo percutaneous coronary intervention (pci) and the impella was placed pre-pci for hemodynamic support. after the pci, two stents to the left anterior descending and circumflex arteries, the impella sheath was secured and locked. dressing was applied. however, the patient still has some oozing that was successfully managed with manual pressure and femostop. the patient was transferred to the unit on impella mcs. update approximately 2.5 hours later noted an echocardiogram was performed and showed the impella 3.2cm in parasternal long axis and 3.6cm from 4 chamber view. unable to obtain a good view as the patient¿s heart anatomy is rotated posterior. care team noted this is a ¿last ditch effort¿ for the patient. approximately 23 hours later, however, care was withdrawn, and the patient expired. no further details surrounding the event was noted. | [5 empty] | E1 added zip code extension as it was omitted from the initial report that was submitted. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the minor bleed, the cause of the injury was not determined since product was not returned and insufficient clinical details provided. h6 investigation type, findings and conclusion codes and component code were updated accordingly based on the completed investigation. | The patient experienced access site bleeding requiring minimal action to resolve. there is no reason to believe that this minor bleeding event is associated with the demise outcome given the information at hand. however, the impella cp device was present at the time of the outcome and no alternative cause as the likely cause of the adverse event outcome was provided. therefore, conservatively, the report is being submitted as death for the impella cp. although, the patient's underlying condition likely played a primary role (team noted ¿last ditch effort¿ for the patient) and the care was withdrawn which appears to have led to the demise outcome. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5277 | 1220648-2025-31613 | 23119102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119102&pc=OZD | 09/12/2024 | 09/22/2025 |  | 09/22/2025 |  |  | 09/12/2024 | 09/22/2025 |  |  | 02/28/2026 |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514574 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, there were placement signal issues. the pump was successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was a damaged sensor due to the shockwave device interaction. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5278 | 1220648-2025-31614 | 23119114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23119114&pc=OZD | 09/24/2024 | 09/22/2025 |  | 09/22/2025 |  |  | 09/24/2024 | 09/22/2025 |  |  |  |  | 09/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella cp pump experienced intermittent placement signal alarms during patient support. repositioning was attempted, but the issue remained. the alarms were disabled and support was maintained with the implanted pump. the pump was successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. all pertinent information available to abiomed has been submitted at this time.
5279 | 1220648-2025-46664 | 23105881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105881&pc=OZD | 08/21/2025 |  |  | 09/21/2025 |  |  | 08/21/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026698076 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that a patient presented with chest pains and shortness of breath. electrocardiogram showed inferior st depression and elevation in avl. decision for impella cp was made. during support there was a notification of ongoing suction alarm at p5. ps waveforms decoupled. additionally, the patient had hypotension and ventricular tachycardia/ventricular fibrillation arrest. return of spontaneous circulation achieved. the patient also had large amount out of chest tube right before code. volume was replaced and pressors weaned off thru the night. heparin was held overnight. the impella was successfully weaned and explanted. | [5 empty] | Correction: d4, e4. the investigation of the reported failure mode has been completed. no product was received for evaluation. hypotension: the cause of the injury was not determined due to insufficient clinical details. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of bleed was not determined due to insufficient clinical details. pump suction: suction alarms present. after reviewing logs, suction alarms were observed along with multiple sudden motor current spikes and drop in flow and trend in placement signal. the cause of pump suction was most likely biomaterial ingestion per the log review. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5280 | 1220648-2025-46666 | 23105885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105885&pc=OZD | 08/21/2025 | 10/20/2025 |  | 09/21/2025 |  |  | 08/21/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703818 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | False Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that the patient was rushed to the hospital with st-segment elevation myocardial infarction. veno-arterial extracorporeal membrane oxygenation was introduced first, and percutaneous coronary intervention was performed. considering the diameter of the lower limb vessels and the neurological prognosis, an intra-aortic balloon pump (iabp) was selected. subsequently, electroencephalogram measurements showed no problems, so the decision was made to replace the iabp with an impella. while on impella support the impella position in aorta alarm kept sounding but the impella was positioned appropriately. it was further reported that the patient expired. the cause of death was myocardial infarction. there were no problems with impella operation. | [5 empty] | Correction b1, e4. the investigation of the reported failure mode has been completed. the impella device was returned for investigtaion. optical signal issue/false alarm buzzer: the data logs show that there were 1124 triggers of "impella position in aorta" and 368 "impella position unknown" alarms throughout support with 5s parameters within specification. there were no spikes in motor current and flows remained in the expected range. a log was viewed during 8/21 which is the first day where the alarms begin to trigger. these alarms were accurately triggered but were on the verge of not triggering as the motor current modulation was only slightly below the alarm criteria. additionally, the placement signal can be seen to be spiking in different areas and there are also frequent instances of it losing pulsatiltiy before the alarms were triggered. due to lack of information regarding imaging to confirm pump position and what was occurring with the patient, the root cause of the optical signal issue/false alarms cannot be determined. device history review: the device passed all the post sterile inspection checks. | Correction e4. the investigation of the reported failure mode has been completed. the impella device was returned for evaluation. optical signal issue/false alarm buzzer: the pump was returned and there was no biomaterial noted, the 5s were within specification and the pump passed laser continuity. the data logs show that there were 1124 triggers of "impella position in aorta" and 368 "impella position unknown" alarms throughout support with 5s parameters within specification. there were no spikes in motor current and flows remained in the expected range. due to lack of information regarding imaging to confirm pump position and what was occurring with the patient, the root cause of the optical signal issue/false alarms cannot be determined. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. h6 type of investigation code b19 added.
5281 | 1220648-2025-46667 | 23105887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105887&pc=OZD | 07/10/2025 |  |  | 09/21/2025 |  |  | 07/10/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687444 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 77 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient admitted for cardiogenic shock was implanted with an impella cp. during support the patient experienced tachycardia, borderline blood pressure, and had increased o2 requirements the physician decreased the impella support from p3 to p2. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A5 and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5282 | 1220648-2025-46665 | 23105890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105890&pc=OZD | 07/03/2025 |  |  | 09/21/2025 |  |  | 07/03/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613014 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Difficult to Remove | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Pulmonary Edema | Thrombosis/Thrombus | [8 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with bridge to transplant was implanted with an impella 5.5 for mechanical circulatory support. patient experienced atrial fibrilation which was treated with medication. additionally patient had bleeding in the lungs and pulmonary edema. concerns about some infection. broadening antibiotics. there were positional suction alarms are occurring in the last 24 hours. clam shell chest closure has prevented any windows for echo to view impella but waveforms suggest positional suction, repositioned bedside under transthoracic echocardiogram. the patient being weaned to p-5 overnight however, increased pulmonary edema. increased to p-6 and patient doing better. the physician attempted to explant at bedside. pump was stuck so had to go to operation room to remove. opened the pocket more and found pump was stuck on the rib. physician was able to remove successfully without having to open chest. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was not received, only data logs were returned and an investigation was completed. data logs showed flow was fluctuated during the case corresponded with clinical narrative that triggered auto suction response & suction alarms. suction issue: product was not returned for review. based on clinical description & data review, the cause of suction issue was most likely patient condition related since the suction alarms were only triggered when the patient had hypotension, or arrhythmias during support. difficult/unable to remove issue: the cause of removal issue was unable to be determined as limited clinical details were provided and no product was returned for review. arrhythmia/major bleed/thrombosis/infection: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pulmonary edema: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided.
5283 | 1220648-2025-46669 | 23105891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105891&pc=OZD | 07/22/2025 |  |  | 09/21/2025 |  |  | 07/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691695 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp and an impella rp exhibited pink tinged urine. the impella rp was explanted and the patient continued on impella cp support until the next day when the patient was escalated to an impella 5.5. no patient harm was reported. | [5 empty] | A4 is unknown. the root cause of hematoria cannot be determined since the product was not returned and insufficient clinical details were provided.
5284 | 1220648-2025-46668 | 23105892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105892&pc=OZD | 07/11/2025 | 09/15/2025 |  | 09/21/2025 |  |  | 07/11/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639600 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered optical signal issues during impella cp support. the pump was placed without issues, however, during shockwave at start of procedure, the placement signal not reliable alarm went off. there were no issues with pump flow and no further information was mentioned on the signal issue. | [5 empty] | B.7 added information that was omitted from the initial report that was submitted. stroke should not of been reported on the initial report that was submitted. d.9 added device returned and date as it was omitted from the initial report that was submitted. e.1 added zip code extension as it was omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.11 the initial report should of stated that the device was returned for evaluation. investigation summary: the pump was returned for investigation. optical signal issue: the impella cp was returned for evaluation. upon visual inspection, no biomaterial was present on returned pump. the pump passed laser continuity testing. the returned pump optical parameters were consistent with a damaged sensor face. data logs were reviewed and long term log showed a sudden increase of placement signal to 3000 with a "placement signal not reliable" alarm. real time log at time of alarms showed optical parameters at that time had sudden drop in signal to noise ratio, step down in contrast, increase in lamp, and slight increase in gain, all consistent with damage to optical sensor face. clinical details noted that the peak signal to noise ratio alarm occurred when using a shockwave device during percutaneous coronary intervention procedure. the cause of the optical signal issue was due to a damaged sensor per returned product, data log analysis, and clinical details noting that shockwave was being used at time of alarm. the distance between the optical sensor and shockwave was unknown per the clinical details. per impella cp instructions for use, "use with shockwave intravascular lithotripsy (ivl) catheter at a distance of less than 20 mm between the optical sensor and ivl device may interfere with or damage the impella optical sensor." device history review: the pump passed all post-sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5285 | 1220648-2025-46670 | 23105894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105894&pc=OZD | 07/22/2025 |  |  | 09/21/2025 |  |  | 07/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694959 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on the second day of impella 5.5 support there were alarms for high purge pressure. the medical team was advised to check the purge line for any kinks, none were found. the medical team discussed the use of tissue plasminogen activator in the purge fluid; however, it was not confirmed if the medical team performed this action. additionally, on (b)(6) 2025, the patient experienced runs of ventricular tachycardia and the following day the patient experienced pulseless electrical activity arrest. as a result, the patient was re-intubated, and it was noted that the patient had questionable neurological status and poor prognosis. of note, there were no impella related issues observed during the episodes of arrythmia or arrest. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. after explanting the device, it was noted there was a loss of the right arterial line waveform with concern for thrombus in the right arm. it is unknown if any interventions were performed to resolve the issue, it was noted that the patient received clotting factors for post operative bleeding which may also have contributed to the suspected thrombus. | [5 empty] | H.6. medical device problem code a24 was incorrectly reported in initial report that was submitted. component code and type of investigation, investigation findings, and investigation conclusions codes were updated according to investigation results. investigation summary: no product was returned. high purge pressure: on (b)(6), there was high purge pressure (hpp) and no kinks were noted. there was possible tissue plasminogen activator use. the data showed that about nine hours into the case the purge pressure began to rise and trigger hpp and purge system blocked alarms. the rise time was about 10 minutes and there were no motor current spikes or long cassette or bag changes leading up to this. the purge solution was changed from bicarb to heparin about 11 hours after it had been trending up, but it still remained high for about another 30 hours until it started to trend down and it was changed back to bicarbonate. the purge pressure remained stable for about 279 hours until it started to gradually trend up again. based on the data log analysis and clinical information, the root cause of the high purge pressure is most likely due to biological material buildup. arrythmia: it was reported that the patient had an episode of pulseless electrical activity. in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombosis: it was reported that there was concern for thrombus down the right arm. in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5286 | 1220648-2025-46671 | 23105895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105895&pc=OZD | 07/28/2025 |  |  | 09/21/2025 |  |  | 07/28/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677215 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Renal Impairment | Hemodynamic instability | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on the fifth day of support, it was reported that there were self-resolving suction alarms from the impella. throughout the support of the patient, the medical team reported daily, ongoing, continuous suction alarms. the patient was given fluids and the medical team decided to continue impella support without lowering the p level of the pump and declined to verify the pump¿s position. the patient¿s hemodynamics worsened during support. the medical team was encouraged to assess the right ventricular function if they determined that impella placement and the patient¿s volume were satisfactory. the medical team disabled the alarm for suction and continued support. additionally, it was noted that there was no hemolysis observed in the patient due to suction issues. on the 23rd day of support, the patient was approved for a durable left ventricular assist device (lvad). an echocardiogram was ordered the same day and impella position was confirmed. continuous suction alarms persisted. on the 27th day of support, continuous renal replacement therapy was initiated and the lvad was placed on hold. two days later, the patient was determined to no longer be a candidate for an lvad due to decreasing kidney function. the patient was placed on palliative care. the patient¿s care was then withdrawn, and the patient expired. | [5 empty] | Based on the information available, the impella 5.5 cannot be excluded as a possible contributing factor to the patient¿s outcome. investigation summary: pump suction: clinical details state that there were continuous suction alarms throughout the case. on the day of the first suction report, placement was not checked. the next report, they gave 500 bolus without further issue. two days later, it was noted that suction was occurring above p7. suction continued throughout the rest of the case but there were no more details about troubleshooting. no product was returned. data log review confirmed the continuous suction alarms throughout the case with more frequent alarms while the pump was above p7. the suction alarms and slightly low pump flow began on (b)(6) 2025. the placement signal (ps) was fairly stable throughout the case until the last few days when it began to trend downward. there were no spikes/abnormalities in motor current or purge pressure. while the pump was above p-7, there were occasional drops in ps at the time of suction alarms but it was not continuous. also, around this time, there were occasional impella position unknown alarms. without more clinical details describing troubleshooting attempts, it is difficult to establish a cause of the suction alarms. the cause of the suction was not determined since insufficient clinical details were provided and no product was returned. hemodynamic instability/acute kidney failure: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. the pump sn (b)(6) passed all post-sterile inspection checks.
5287 | 1220648-2025-46031 | 23105923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105923&pc=OZD | 07/29/2025 |  |  | 09/21/2025 |  |  | 07/29/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684961 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced a vascular injury that required intervention after the impella cp was explanted. the patient was initially implanted with an impella cp device (serial number (b)(6)), however within the first twenty minutes of support a high purge pressure alarm occurred which could not be resolved. the first impella cp pump was replaced with a second impella cp device (serial number (b)(6)). after 1 day of support, the medical team decided to escalate the patient to an impella 5.5 (serial number (b)(6)). upon explanation of the second impella cp, it was reported that the patient experienced a vascular injury at the impella cp access site after a perclose device failed. a covered stent was placed to resolve the bleeding. it was reported during impella 5.5 support, on (b)(6) 2025, there was an impella positioning issue which caused the patient to experience ectopy. the medical team was advised to reposition the impella, which resolved the ectopy. it was additionally noted that the patient¿s care was withdrawn, and the patient expired. this complaint involves three impella devices: pump 1: impella cp, with serial number (b)(6). pump 2: impella cp, with serial number (b)(6). pump 3: impella 5.5, with serial number (b)(6). this mdr specifically addresses the second impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Product status: the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. investigation summary: peripheral artery dissection: the root cause of vascular injury cannot be determined since the product was not returned and insufficient clinical details were provided. device history review: device (sn: (b)(6)) passed all post sterile inspection checks.
5288 | 1220648-2025-46673 | 23105929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105929&pc=OZD | 07/29/2025 |  |  | 09/21/2025 |  |  | 07/29/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026691839 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had a nosebleed and that they experienced pump suction. it was reported that heparin was stopped due to a nosebleed. additionally, suction alarms were encountered overnight, and as a result lasix drip was stopped and no alarms appeared since. patient is stable post | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of the 2 report, where this report represents the pump and the other report is for aic. | H10 added related report number as it was omitted from the initial report that was submitted. the investigation has been completed. clinical symptoms(hemorrhage/bleeding): the cause of the injury was not determined due to insufficient clinical details. pump suction: the pump was not returned for investigation. data log showed a few suction instances during the case run. one occurred on 30-jul-2025 while running at a steady p-level, with the placement signal becoming flattened and low in amplitude, and a drop in motor current and flow for about two minutes. there was no confirmed positioning issue noted in the data log. the cause of the suction issue was not determined, as sufficient clinical information was not provided and the product was not returned for investigation. device history review: the device passed all the post sterile inspection checks.
5289 | 1220648-2025-46674 | 23105951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105951&pc=OZD | 07/30/2025 |  |  | 09/21/2025 |  |  | 07/30/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694333 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was successfully supported with a impella cp through emergent coronary artery bypass graft. the catheter was shallowed out in the left ventricle (lv) under echocardiogram guidance before the procedure. the catheter was cross clamped in the lv and surgical mode was used. no soft insets were available for the aortic clamp, which lead to a leak from the aorta. while repositioning the impella cp back into the lv, the physician pulled the impella cp out of the lv. attempts were made to recross the lv and all were unsuccessful. the groin site was closed with one perclose and the chest was left open. | [5 empty] | The cause of the positioning issue was determined to be an unintentional use error as the doctor pulled the pump out of the left ventricle during repositioning. the cause of the heart valve incompetence cannot be determined as insufficient clinical details were provided. the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion. ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
5290 | 1220648-2025-46676 | 23105952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105952&pc=OZD | 07/31/2025 |  |  | 09/21/2025 |  |  | 07/31/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697189 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Nerve Damage | Pain | Thrombocytopenia | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while awaiting surgery for heart transplant. on the fourth day of impella support, it was reported that the patient gained 5 pounds overnight, increased shortness of breath, and had a decrease in mixed venous oxygen levels. as a result, the patient was given an epinephrine drip and diuretic medications were adjusted. the patient was monitored and when they did not respond to interventions, the patient was cannulated for veno-arterial extracorporeal membrane oxygenation for increased hemodynamic support. it was also noted that the patient continued to report numbness in their right hand distal to the impella 5.5 insertion site. on the 16th day of support, it was reported that the patient had an elevated white blood count and was in consistent pain. blood cultures were drawn, and the patient¿s heart transplant status was delisted to status 7, temporarily inactive. the following day, the patient remained as status 7 on the transplant list due to acute decompensation. three days later, the patient was again listed as active status 1 on the transplant list. on the 23rd day of support the patient experienced heparin induced thrombocytopenia (hit). as a result, heparin was discontinued and angiomax was started for anticoagulation therapy. it was additionally noted that the patient¿s lactate dehydrogenase tripled after the impella p level was increased due to low mixed venous oxygenation levels. the patient was transfused with one unit of packed red blood cells. the following day the patient was determined to no longer be a candidate for heart transplant following the hit complication. the patient remained on impella 5.5 support until successfully bridged to durable left ventricle assist device with post op right ventricle assist device support. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: no product was returned for investigation. thrombocytopenia/ compression/ pain: the cause of the clinical symptom was not determined due to insufficient clinical details. infection: in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the infection was unable to be determined due to insufficient clinical details. device history lot: device lot: 1932218. device history review: device with sn: (b)(6) passed all post sterile inspection checks. the with sn# (b)(6) from batch no: 1932218 was found to have a loose particle and failed 100% visual inspection checks. rework of this unit has been successfully completed; unit was sent back to sterilizer. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5291 | 1220648-2025-46064 | 23105955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105955&pc=OZD | 07/31/2025 |  |  | 09/21/2025 |  |  | 08/01/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1133147 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 63 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Decrease in Pressure | Failure to Read Input Signal | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a purge disc error alarm that was resolved by replacing the automated impella controller. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5292 | 1220648-2025-46677 | 23105956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105956&pc=PYX | 08/24/2025 |  |  | 09/21/2025 |  |  | 08/24/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026675604 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had an impella rp flex (and cp and ecmo) implanted to support a patient admitted for st elevated myocardial infarction and concerns of microvascular decompression and a ventricular septal defect. the patient developed hematuria that prompted p level adjustment and echo imaging for pump positioning. the team explanted the impella rp flex and biomaterial was observed on the pump. the patient remained on ecmo and impella cp support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the returned product was investigated. biomaterial was seen in the outflow cage. no biomaterial was seen around the impeller blades. no other abnormalities were seen. thrombosis : clinical details mention that a tissue was found on the tip of rp flex proximal to the pigtail on the distal end of the outflow cage upon explant. the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria : according to clinical details, the care team was aware of the hematuria in the patient during impella support. the cause of the injury was not determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5293 | 1220648-2025-46678 | 23105957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105957&pc=OZD | 08/02/2025 |  |  | 09/21/2025 |  |  | 08/02/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657442 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Arrhythmia | Hematoma | Pain | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. prior to impella 5.5 support, the patient had an intro-aortic balloon pump (iabp) placed and was supported for 13 days. the medical team decided to escalate the patient from the iabp to the impella 5.5 for an advanced option workup. it was reported that on the second day of support; the patient experienced an episode of ventricular tachycardia (vt) during the previous night. dobutamine was turned off and the vt did not recur, and the impella position was confirmed with echocardiography. the following day it was reported that the patient had another episode of vt and to resolve they adjusted the amiodarone medications. of note, there were no impella related issues reported, and a plan was made for a formal echocardiogram to verify impella position. the impella position was confirmed to be well placed within the left ventricle. on the 21st day of support, the patient was noted to have hematoma in their right arm, located below the cutdown site where the impella 5.5 was inserted. no interventions were reported to resolve the hematoma, and a week later it was noted to have improved. the following day, on day 29 of support, the site was noted to be stable but swollen. on day 31 of support, the site was still noted to be swollen but was improving. on day 32 of support, the patient complaint of pain in the right arm and it was noted that the hematoma was resolving. the nerve pain in the arm persisted over the next several days, and a nerve block was ordered to resolve the pain. on day 46 of support, the patient experienced ectopy overnight. the impella was found to be too deep, and the pump was repositioned. the following night, there were suction alarms that occurred, and the echocardiography showed the impella deep. the impella was repositioned, and impella flows improved from 3.7 to 4.2 l/min. the next day it was reported that the purge arm was wet, and there appeared to be an intermittent purge leak. the following day the purge side arm was still noted to be leaking. the source of leak was unable to be located, however pump flows and pressures remained unchanged and there was no impact on patient support noted. the patient remains on impella support as of the writing of this report. | [5 empty] | Investigation summary. pain in extremity: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. arrhythmia: in order to make a root cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the cause of the hematoma was not determined since the product was not returned and sufficient clinical information was not provided. purge leak - pump: as per the clinical details, the pump sidearm was found to be leaking during the field run. data log review was not required for this case. the cause of the purge leak pump issue was not determined since the product was not returned for investigation and sufficient clinical information was not provided. device history review: the pump (b)(6) passed all post sterile inspection checks. | Product complaint #: (b)(4). investigation summary: analysis summery: pain in extremity: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. arrhythmia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma: the cause of the hematoma was not determined since the product was not returned and sufficient clinical information was not provided. purge leak: pump: as per the clinical details, the pump sidearm was found to be leaking during the field run. data log review was not required for this case. the cause of the purge leak pump issue was not determined since the product was not returned for investigation and sufficient clinical information was not provided. device history lot: device batch: 1911719. device history batch: subcomponent lot: na. device history review: the pump sn (b)(6) passed all post sterile inspection checks.
5294 | 1220648-2025-46680 | 23105958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105958&pc=PYX | 08/02/2025 |  |  | 09/21/2025 |  |  | 08/02/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026652200 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 81 KG | Non Hispanic | White | Death | Yes | Yes | Inaccurate Flow Rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support (mcs) and experienced placement signal issues, increased pump flow with elevated motor current and bleeding requiring intervention. however, the patient would subsequently expire. the patient had a redo sternotomy as well as aortic and mitral valves replacement and redo coronary artery bypass graft surgeries. the patient failed to decannulate in the operating room requiring crash onto central veno arterial extracorporeal membrane oxygenation (va ecmo). the right ventricle continued to struggle. right aorta thrombus was present on echocardiogram. the risks of the rp flex placement with thrombus as well as exchange of indwelling lines were discussed with physicians. however, the physicians decided to continue with rp flex placement. the impella rp flex was placed and started at performance level p-2. ecmo flows were stable. impella was increased to p-4, flows greater than 1.5. the patient remained hemodynamically stable throughout procedure. the impella rp flex was secured and dressed. next the physician placed a new pulmonary artery catheter via left internal jugular. impella rp flex position was confirmed prior to leaving catheterization lab for the intensive care unit. the following day the nurse called the emergency call center due to the placement signal changes and artifact following ecmo decannulation in the operating room. the patient returned to the unit and changes were noted in the placement signal. motor current and flows remained the same as prior to the operating room and have been consistent. there were no changes to medications or the patient¿s condition. there were no alarms as well. the nurse requested local team support. the next day, an update noted the patient¿s condition remained unchanged and ¿bleeding has stopped.¿ heparin was off with no current plans to restart. the patient continued to decompensate with kidney function declining to the point of dialysis. two days later another call was made to the emergency service center for an increased pump flow and elevated motor current gaining over 100 ma. the care team was advised of a likely pump stop and reminded of the known right aortic thrombus present at implant. the patient continued to require pressers and lactate dehydrogenase continued to increase. patient was noted for possible comfort measures only and would expire in the evening of this day. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
5295 | 1220648-2025-46506 | 23105964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105964&pc=PYX | 08/15/2025 |  |  | 09/21/2025 |  |  | 08/15/2025 | 09/21/2025 |  |  | 01/31/2026 |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025461499 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella rp flex device for mechanical circulatory support (mcs) experienced ventricular fibrillation and would subsequently expire the following day. the patient arrived with right coronary artery st-segment elevation myocardial infarction. the patient underwent a right heart catheterization and thrombectomy was performed. intra-aortic balloon pump (iabp) and temporary transvenous pacing wire were placed and the decision was made to place the patient on an impella rp flex. the patient, intubated and sedated, was sent to the unit on impella rp flex and iabp mcs and settled. overnight, the patient went into ventricular fibrillation and self-converted. an x-ray at bedside for reconfirmation of correct placement of rp flex was made. impella rp flex waveforms and flow were noted as appropriate, and pulses were dopplerable. urine output was clear and copious. cardiothoracic surgery was consulted, and potential plan was made to escalate from the iabp to an impella 5.5. the following day, the patient remained intubated and sedated, the impella access site dressing was noted to have ¿old¿ blood, but hemostasis was intact. additionally, it was noted the patient ventricular fibrillation arrested overnight and it appears 1 shock was delivered. the patient's status changed to do not resuscitate. the family decided not to escalate in care and was determining if care would be continued. labs and hemodynamics showed poor perfusion. the decision was subsequently made to place the patient on comfort measures only, the care was withdrawn, and the patient would, therefore, expire. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding, minor bleed, the cause of the injury was not determined since product was not returned and insufficient clinical details provided. pertaining to the ventricular fibrillation, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5296 | 1220648-2025-46592 | 23105965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105965&pc=OZD | 08/12/2025 |  |  | 09/21/2025 |  |  | 08/12/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658382 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 pump was placed to support a patient with known cardiomyopathy and other systemic comorbidities. during support the patient experienced arrythmia which was managed with medical therapy. additionally, the patient suffered a nose bleed which prompted systemic anticoagulation of angiomax to be put on hold. impella support continues. | [5 empty] | No product was returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the cause of the epistaxis was not determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5297 | 1220648-2025-46266 | 23105966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105966&pc=OZD | 08/07/2025 |  |  | 09/21/2025 |  |  | 08/08/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569906 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to impella cp insertion, the patient presented to the emergency room complaining of chest pain and an electrocardiogram revealed a st elevation myocardial infarction. left heart catheterization was performed and revealed mitral valve disease and the patient became hypotensive. the patient was then started on norepinephrine and an impella cp was implanted and the patient was transferred to a large hospital in the network for further management and a possible coronary artery bypass graft. it was reported that on the second day of support the p level was weaned to p4, and following the drop in p level the patient went into atrial fibrillation and the mitral valve prolapsed. the p level was turned back up to p6 and normal sinus rhythm was established. on (b)(6) 2025 it was reported that the patient continued to have episodes of atrial fibrillation, and the medical team discussed changing the amiodarone drop to dobutamine in attempt to resolve the arrythmia and if unsuccessful then the patient may need cardioversion. the following day, the patient was taken to the catheterization laboratory for a high-risk percutaneous cardiac intervention due to decompensation and requirements of high dose pressor and inotrope support. it was noted that the patient was cardioverted twice that morning for ventricular tachycardia and once the previous day for atrial fibrillation. after the patient¿s procedure, it was noted that the impella cp had to be repositioned as it was found to be deep. the impella was measured at 5.27 cm in the ventricle and pulled back to 3.53 cm. on (b)(6) 2025 it was noted that the patient remained in atrial fibrillation and was non-pulsatile. it was reported later that night, the purge system on the impella cp required a bypass due to a leak. on (b)(6) 2025, the medical team decided to escalate the patient to an impella 5.5. the patient was brought to the operating room and the right axillary artery was prepped for impella 5.5 access. during insertion of the 5.5, resistance was met in the axillary artery and the impella 5.5 could not be advanced. the impella was removed, and the axillary artery was shockwaved and a balloon inserted and inflated to open the vessel. the 5.5 was still unable to pass through the artery. the artery was shockwaved again and the 5.5 was still unable to be advanced through the patient¿s artery due to the patient's anatomy. the decision was then made to abort the impella 5.5 insertion and to insert a new impella cp into the axillary artery. the cp was able to be delivered with no resistance. two days later, the patient experienced ventricular tachycardia and was successfully defibrillated to resolve. later that day, the decision was made to withdraw the patient¿s care. the patient¿s care was withdrawn, and the patient expired. this complaint involves 3 pumps: pump 1, impella cp with serial number (b)(6). pump 2, impella 5.5 with serial number (b)(6). pump 3, impella cp with serial number (b)(6). this mdr specifically addresses pump 1, impella cp with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Medical safety review was completed and noted the following: the patient expired while on impella cp support. the patient implanted with an impella cp pump 1 and experienced arrhythmic events requiring support level adjustment, possible change in medication, defibrillation, and pump repositioning. this device's purge system was bypassed due to system leak. the impella cp pump 1 was explanted and an impella 5.5 was unsuccessfully attempted. however, another impella cp pump 2 was successfully placed with no loss of support. the failure to deliver the impella 5.5 is a product malfunction and the first impella cp device event resulted in serious injury and should be reported. while on the second impella cp, the patient continued to have ventricular tachycardia (vt) runs that were successfully defibrillated; however, the patient subsequently expired as care was withdrawn. given the information at hand, there is not enough information to clearly dissociate the demise out from the second impella cp device. however, although this dissociation cannot be made, it is the care that was withdrawn that appears to have led to the patient's demise; vt runs were successfully defibrillated. medical safety officer reviewed the event and noted the same in the second impella cp device cannot be dissociated for the outcome and the first impella cp pump and impella 5.5 is a serious injury and product malfunction respectively type of reportable event. coding - h6: medical device problem coding a24 was removed as one of the codes describing the event and should, therefore, be disregarded as captured on the initial medwatch report. related medical device reports: this impella cp device report is one of three devices associated with the event story. refer to the related manufacturer report number as noted in section h10 for the medical device report on the impella 5.5 and second impella cp, which is associated with the demise outcome. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5298 | 1220648-2025-46681 | 23105968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105968&pc=OZD | 08/15/2025 |  |  | 09/21/2025 |  |  | 08/15/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695318 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 82 KG | Non Hispanic | Black Or African American | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Inflammation | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced ventricular tachycardia (vt) requiring treatment with an amio bolus. impella suction alarms occurred over night so the p-level was lowered. the following day albumin was given to improve flows. additionally, swelling was noted on the patient¿s right arm with stable right radial waveforms. the patient continued to experience vt events requiring defibrillation. later, care was withdrawn and the patient expired. | [5 empty] | Product complaint # (b)(4). investigation summary : data review: data logs showed pump ran without issue during support. there were suction alarms triggered in the first 4 days of support then were resolved. there were no mc spikes, abnormal ps or purge issues observed during support. pump suction issue: the cause of suction issue was most likely patient condition related since the suction was improved after albumin transfusion. tachycardia: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. epistaxis/inflammation: the cause of the injury was not determined due to insufficient clinical details. device history lot : device lot #: 1937661. device history batch : subcomponent lot#: n/a. device history review : pump sn: (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5299 | 1220648-2025-46682 | 23105969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105969&pc=OZD | 08/08/2025 |  |  | 09/21/2025 |  |  | 08/08/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622863 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device via the right axillary artery for mechanical circulatory support (mcs). support level was increased to p-8, and the patient was sent to the unit on mcs. the next day, overnight, positioning and suction alarms along with some ectopy were experienced. the pump was repositioned at bedside under echocardiogram guidance measuring at 4.9cm with good waveforms. impella mcs continued with possible plan for left ventricular assist device (lvad) workup. now four days later, the patient was noted as doing well despite right hand pain managed with exercise with pain medication. later that night, suction alarms triggered related to the patient moving. impella mcs continued with potential explant planned for the next day. overnight the patient was maintained at support level p-3, and the following day the device was explanted with a survived patient outcome. it was noted the patient went home on milrinone and possible lvad workup, placement at a later time. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. event logs were received and analyzed. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the arrhythmia, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. regarding, pain, the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. pertaining to the pump suction, after reviewing the logs, suction alarms along with a trend in purge was observed. the cause of pump suction was most likely patient use issue related per the review of logs. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | Patient-specific information section a4: weight is unknown. device status: the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5300 | 1220648-2025-46570 | 23105970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105970&pc=OZD | 08/15/2025 |  |  | 09/21/2025 |  |  | 08/18/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026673798 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male | 85 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Multiple Organ Dysfunction Syndrome | Thrombocytopenia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant report that during impella 5.5 support, a patient experienced heparin induced thrombocytopenia. heparin was discontinued and argatroban was started for anticoagulant therapy. the same day, it was noted that the patient developed anisocoria and signs of cerebral hemorrhage were observed. the patient¿s status was updated to do not resuscitate, and the decision was made not to implement aggressive treatment. the patient died of multiple organ failure. the physician noted that the patient¿s pre-existing condition was severe, and there was a high risk of bleeding due to hit. it was additionally noted that there was no direct relationship to the impella, and the impella was functioning stably until it was removed. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. stroke: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia/ organ failure: the cause of the injury was not determined due to insufficient clinical details. device history review: the pump passed all the post sterile inspection checks. | Although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5301 | 1220648-2025-46561 | 23105971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105971&pc=OZD | 08/12/2025 |  |  | 09/21/2025 |  |  | 08/12/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629954 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 87 KG | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. after the impella 5.5 was implanted, the patient had suction alarms during the day shift and volume was given. six day later, the patient had a ventricular tachycardia run with which required shocking x 1 during turning. additionally, the patient experienced bleeding on day nine of support on the right lower leg. heparin was put on hold due to that. the plan is for bronch, exchange lines, and to draw blood cultures. care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint #(b)(4). investigation summary: tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the bleed was most likely patient condition related because it occurred at a non-impella site. pump suction: the cause of the suction was most likely patient related, as the suction alarms occurred at times where the placement signal had low pulsatility and a downward trend, and they resolved when p-level was lowered.
5302 | 1220648-2025-46683 | 23105972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23105972&pc=OZD | 07/26/2025 |  |  | 09/21/2025 |  |  | 07/26/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026703548 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a "placement signal low" alarm on an automated impella controller (aic) supporting an implanted impella 5.5 pump. a transthoracic echocardiogram was performed to confirm that the impella was in proper position and the patient's hemodynamics did not correlate with the placement signal. the audio was disabled for the placement signal alarm. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The cause of the placement signal low alarm was patient condition related as evidenced by issue resolution after volume was given with no identified log abnormalities related to the optical sensor. the pump passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5303 | 1220648-2025-46684 | 23106020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106020&pc=OZD | 08/02/2025 |  |  | 09/21/2025 |  |  | 08/02/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667560 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the patient experienced oozing at the impella access site. an additional suture was placed at the access site, heparin was withheld, and lidocaine with epinephrine was injected around the site to resolve the bleeding. additionally, it was noted that the patient was transfused with blood products, quantity of units transfused was not provided. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5304 | 1220648-2025-46685 | 23106021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106021&pc=OZD | 08/03/2025 |  |  | 09/21/2025 |  |  | 08/03/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026658785 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 92 KG | Non Hispanic | White | Death | Yes | No | Difficult to Advance | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 pump was placed to support a patient admitted in for acute myocardial infarction/cardiogenic shock. ventricular access was attempted multiple times, and the standard jtip wire was switched for glidewire advantage, but the doctor was unable to cross the valve. the pigtail catheter was switched for a straight flush catheter to obtain ventricular access. the impella 5.5 was attempted to be placed across aortic valve but kept getting stuck on edge of the prosthetic valve. the doctor manually manipulated pump multiple times and the impella was placed into position after much difficulty. the patient was bleeding from the chest tubes from previous cardiac surgery where the chest was left open, so heparin was withheld and the patient was transfused four units of red blood cells and two units of fresh frozen plasma. a ct scan showed an infarction involving the right parietal-occipital region and the patient¿s white blood cell count was elevated. antibiotics were administered and the swan was moved. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5305 | 1220648-2025-46686 | 23106022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106022&pc=OZD | 05/26/2025 |  |  | 09/21/2025 |  |  | 05/26/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557907 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 86 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | Nerve Damage | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 pump was placed to support a patient with known cardiomyopathy listed for transplant of both heart and kidney. the patient reported nerve pain in the arm which was treated by medical therapy. the pump experienced intermittent optical signal issues so the placement signal was recalibrated. later, the patient received organ transplants and the impella was explanted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5306 | 1220648-2025-46687 | 23106023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106023&pc=OZD | 08/11/2025 |  |  | 09/21/2025 |  |  | 08/13/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665950 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered pump suction during impella 5.5 support. it was reported that there were occasional suction events due to the pump being positioned shallow and facing the mitral valve. later, intermittent suction alarms occurred. pump placement was confirmed, the patient was transfused additional volume, and the pump's p-level was reduced from p-5 to p-4. the patient remains on impella support. | [5 empty] | D.6b explant date was added as it was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. pump suction: no product was returned. after reviewing the logs, suction alarms along with a trend in purge and placement signals were observed. the cause of pump suction was most likely patient condition related per the review of logs. device history review: the device passed all post sterile inspection checks. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5307 | 1220648-2025-46688 | 23106024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106024&pc=OZD | 07/26/2025 |  |  | 09/21/2025 |  |  | 07/26/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026697052 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 74 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Swelling/ Edema | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient had an impella 5.5 placed for care escalation and bridge to heart transplant. during support the patient developed a new arrythmia, an arm clot, edema, and bleeding that required a transfusion of red blood cells. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the bleed was not determined since product was not returned and insufficient clinical details provided. the root cause of the arrythmia and thrombosis was unable to be determined due to insufficient clinical details. the cause of the edema cannot be determined as insufficient clinical details were provided. the pump passed all the post sterile inspection checks.
5308 | 1220648-2025-46689 | 23106095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106095&pc=OZD | 08/15/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-0000 |  |  | ABIOMED, INC. |  |  | US |  | 1316957 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant indicated that a patient classified as undergoing high-risk surgery was implanted with an impella cp device to provide mechanical circulatory support. the complainant also noted that an alarm for an unstable position sensing signal was triggered during the period of impella support. subsequently, the aic was substituted with the ic5660, and support for the patient was maintained. no patient harm was reported. | [5 empty] | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. data analysis: review of the imc logs confirms the complaint. the pump is detected when plugged into the console and eeprom reads are initiated to determine the various pump parameters. within a few seconds, optical sensor calibration begins. eleven seconds later a ¿calculator placement signal 1048¿ error appears. approximately 54 seconds later, there are log entries indicating that calibration has timed out and stopped, followed immediately by ¿case start: optical sensor calibration 4 g0=0¿ (calfail.png). the ¿placement signal not reliable¿ (psnr) alarm occurs a few moments later (psnr.png). pump flow was initiated but discontinued after 6 minutes. it does not appear that there was any attempt to disconnect and re-connect the pump to this console; rather, the decision was made to move this pump to a second console, (b)(4) . the logs for that console were obtained and compared to those of (b)(4) , and no notable differences in sequence of steps for pump calibration were noted. the pump was, however, correctly initialized on the alternate console. the rt logs for (b)(4) show that for the length of time the pump was connected, contrast, lamp and gain were within the specifications listed in pm procedure 0042-7309. the optical sensor and placement signal, however, show invalid values (32760 mmhg and 3000 mmhg, respectively) for the duration of time that the pump was connected. device analysis: the console was powered on and a pump connected. there were no issues with pump initialization or calibration. pump flow was started at level 9 and allowed to continue for approximately 3 hours. no psnr or other unexpected alarms occurred during this test. the console was opened and thoroughly inspected for damage, loose cables/connections, etc. and no issues were found. eight total attempts were made to reproduce the issue by turning on the console, connecting a pump, waiting several minutes for an alarm, and then repeating this process. no errors or alarms ever occurred. root cause: the root cause of the psnr alarm was intermittent out-of-range placement signal data sent by the optical bench to the 4-in-1 module, resulting in unsuccessful pump sensor calibration. the cause of the optical bench transmitting out-of-range data could not be determined, due to the inability to reproduce the issue. | The device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5309 | 1220648-2025-46692 | 23106096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106096&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  | 10/31/2025 |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419196 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 68 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had placed a 5.5 pump for post-cardiotomy cardiogenic shock due to coronary artery bypass graft performed in the operating room. due to bleeding concerns, anticoagulation had not been initiated. during the support, embolic cva was observed. the family made the choice to harvest organs and withdraw care. the pump had supported for 11 days. | [5 empty] | Corrections: d2, e4. the investigation of the reported failure mode has been completed. stroke :the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5310 | 1220648-2025-46691 | 23106097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106097&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708012 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | Yes | Crack | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had a cp for the high-risk cardiac procedure. the pump supported for the pci but was explanted after just 24 hours with an access site ooze and hematoma treated by femostop and dressing with manual hold. the ooze was from pump damage observed by team. the patient was successfully weaned. | [5 empty] | B.1 product problem was selected incorrectly on the initial report that was submitted. b.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. e.4 added selection as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name should of been left blank on the initial report that was submitted. h.6 medical device problem code (a0404) was reported incorrectly on the initial report that was reported. investigation summary: access site adverse event/hematoma: clinical details noted oozing from the impella site and hematoma. the cause of the issue was not determined as limited clinical information was provided and device was not returned. device history review: the pump passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5311 | 1220648-2025-46690 | 23106098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106098&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713762 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 108 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | Material Puncture/Hole | Patient Device Interaction Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that decision was made to place a patient on impella cp for hemodynamic support. however, it was noted that during single access stick the physician poked a hole in the impella catheter causing purge fluid to leak from the impella catheter. purge alarms started and attempts to patch the hole with a tegaderm with no success. the physician elected to proceed without replacing new pump. left anterior descending artery was wired. pros left anterior descending artery (lad) was ballooned shockwave used to treat lad-left main (lm). it was noted that motor currents began to spike and impella was restarted multiple times until it would no longer. stents were imaged with intravascular ultrasound and post dilated. impella cp was removed and patient was left on small doses of levo and dobutamine for blood pressure. peel away sheath was removed and preclose both broke off. 2 more attempted and failed. 8f angioseal deployed and manual compression held to successfully close right femoral artery with no evidence of hematoma or bleeding. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections: d9, e4, h6. the investigation of the reported failure mode has been completed. the impella cp and purge cassette were returned for investigation. a hole was observed on the catheter near the 85 cm mark. no other evidence of physical damage was noted on the returned product. the attempt to prime was able to reproduce the purge leak from the hole on the catheter. the impeller was inspected and found to have biomaterial blocking the purge gap. the pump was not able to start in a bucket of water and displayed a pump stop alarm with motor current around 1400. all three electrical motor cable lines passed resistance testing. a small cut was made on the catheter near the motor, and blood was noted in the purge lumen. data log shows purge system open alarm and open line characteristics on the pressure and purge flow after the start of the case. the pump was only utilized for one hour, and a gradual rise in motor current was observed, followed by several pump stop-motor current high alarms. as per the clinical details, the impella catheter was accidentally punctured, causing fluid leaks and purge alarms. attempts to patch the catheter were unsuccessful, and the impella eventually stopped working after multiple restarts. the impella cp was removed. product damage (hole in device): the cause of the product damage issue was unintentional use error during the procedure, based on clinical details. purge leak - pump: the cause of the purge leak-pump issue was a damaged purge lumen in the catheter due to unintentional use error during the procedure, based on clinical details. pump stop: based on the returned product evidence, purge leak from the damaged purge lumen allowed blood to ingress and clot in motor housing. the cause of the pump stop issue was blood ingress due to unintentional use error. device history review: the device passed all the post sterile inspection checks.
5312 | 1220648-2025-46694 | 23106099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106099&pc=PYX | 08/21/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  | 01/31/2026 |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025462218 |  | 000000000001000323 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient experienced atrial fibrillation (afib). cardioversion was discussed for the afib but no confirmation of intervention was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of the reported failure mode has been completed. no product was returned for investigation. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5313 | 1220648-2025-46693 | 23106100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106100&pc=OZD | 08/22/2025 | 12/01/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026713178 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility had placed a 5.5 pump for the support of a patient with known cardiac history and elective cabg. the pump was placed but the lv was perforated. the pump had been placed via the direct approach. the perforation was repaired and the pump was instead placed via the axillary. | [5 empty] | The investigation for the reported cardiac perforation and positioning issue has been completed the root cause of the injury was not determined due to insufficient clinical details. the root cause of the positioning issue was unable to be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections: b1 product problem was inadvertently omitted on the initial manufacturer device report. d6b was inadvertently omitted on the initial manufacturer device report. e4 was inadvertently omitted on the initial manufacturer device report. h6 medical device problem code a01 was inadvertently omitted on the initial manufacturer device report. h11 dhr statement inadvertently omitted on the supplemental manufacturer device report. device history review: the device passed all the post sterile inspection checks.
5314 | 1220648-2025-46695 | 23106102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106102&pc=OZD | 08/22/2025 | 10/14/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2026710431 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant indicated that a patient classified as high-risk for percutaneous coronary intervention had an impella cp device implanted for mechanical circulatory support. the pump support continued when the pump flows decreased. furthermore, at the time of explantation, the impella cp catheter was kinked at the cannula just before the outflow cage. alarms for reduced impella flows were activated initially when the kink was not observed. the kink became visible towards the conclusion of the procedure, with no alarms sounding at that time. | [5 empty] | The investigation of the reported failure mode has been completed. the pump was returned. event logs were not provided. low pump flow: clinical reported that low flow alarms were triggering at the beginning of the case, and later on, a kink was observed on the cannula. it's also noted that the patient had a small ascending aortic arch, which is what caused the cannula to kink. the product was returned for investigation, and a small kink was observed on the cannula near the outlet cage. a flow test was run on the pump in a bucket of water, and low flow was not reproduced. data logs were not returned for this case, so the low flow alarms that triggering during the case were not confirmed. although the low flow was not reproduced, a cannula kink doesn't typically reproduce low flows in investigation, because the bucket does not apply forces on the cannula like it would in a patient's anatomy. the cause of the low pump flow was most likely due to the cannula kink that was reported in clinical details and observed during product investigation. product damage (cannula kink): clinical reported that the pump kinked at the cannula before the outlet cage, and said it was due to the patients small aortic arch. the product was returned for investigation, and a small kink was observed on the cannula near the outlet cage. the cause of the cannula kink was most likely patient condition, as the clinical report states the patient had smaller anatomy which caused the cannula to kink. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5315 | 1220648-2025-46696 | 23106104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106104&pc=DQX | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | WIRE, GUIDE, CATHETER |  |  | IMPELLA |  |  | DQX |  |  | GUIDE WIRE 0.018" PTFE COATED |  |  | ABIOMED, INC. |  |  | US |  | 9452080 |  | 0052-3005 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient classified as high-risk for percutaneous coronary intervention received an impella cp device for mechanical circulatory support. during the preparation of the pump for delivery, it was noticed that the .018 guidewire appeared to be "bent." the team proceeded to replace the wire, and no harm or injury was reported. | [5 empty] | The device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | A.4 added as it was omitted from the initial report that was submitted. b.7 added information as it was omitted from the initial report that was submitted. d.1,d.2a and d.2b were revised as they were previously entered incorrectly on the initial report that was submitted. d.4 model number was revised as it was previously entered incorrectly on the initial report that was submitted. lot number and expiration date were added as they were omitted from the initial report that was submitted. h.4 date was added as it was omitted from the initial report that was submitted. the investigation of the reported failure mode has been completed. product damage (guidewire kink): clinical details were reported that an 0.018 guidewire had a bent tip when it was pulled from its plastic packaging. the product was not returned for investigation. the cause of the guidewire kink could not be determined due to insufficient clinical details and no product returned. device history review: the 0.018" guide wire passed all post sterile inspection criteria.
5316 | 1220648-2025-46697 | 23106105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106105&pc=OZD | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | 2000154 |  |  | ABIOMED, INC. |  |  | US |  | 2026707610 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant noted that a visible leak was observed from the purge filter while the patient was receiving support from the impella cp device. there were no alarms activated on the console during this time. the leakage occurred during the exchange of purge fluid and deairing, with visible leakage noted from the check valve of the yellow luer lock. no patient harm was reported. the pump was successfully weaned off and subsequently explanted. | [5 empty] | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. clinical details note that there was visible leakage from the purge filter. also, during purge fluid exchange and de-airing, there was visible leakage out of check valve of yellow luer lock. the cause of the purge leak - pump was not determined since product was not returned and insufficient clinical details were provided. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5317 | 1220648-2025-46698 | 23106106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106106&pc=OZD | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-0000 |  |  | ABIOMED, INC. |  |  | US |  | 1222703 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Network Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was implanted with an impella cp device to provide mechanical circulatory support. while in the catheterization lab, a white alarm appeared stating "unexpected controller-turn off - change to back up controller if situation won't change." the alarm ceased without any manual intervention. furthermore, establishing a connection to impella connect was unsuccessful, although the blue light on the connect module was illuminated. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5318 | 1220648-2025-46700 | 23106159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106159&pc=OZD | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | IMPELLA CP SMART ASSIST |  |  | ABIOMED, INC |  |  | US |  | 2026690104 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was fitted with an impella cp device to provide mechanical circulatory support. on the second day of support, the pump dislodged; however, it was successfully repositioned across the valve into the left ventricle (lv). there were no adverse effects observed from the pump being situated in the aorta. furthermore, there was an occurrence of arrhythmia during the support, which was addressed through medical management. it is noted that no defibrillation was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. no product was returned. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5319 | 1220648-2025-46699 | 23106160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106160&pc=OZD | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2026710377 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant indicated that a patient experiencing acute myocardial infarction and cardiogenic shock was implated with an impella cp device to provide mechanical circulatory support. while the impella cp was in operation, the patient exhibited persistent capillary ooze, which proved challenging to manage. by the following morning, the patient's hemoglobin levels had decreased, and the treatment team was unable to identify a solution to address the bleeding; consequently, they opted to remove the pump and administer packed red blood cells to the patient. the purge fluid was altered to d5 with bicarbonate, and an echocardiogram was conducted, confirming that the pump was positioned correctly. the pump was extracted with manual assistance, utilizing femstop support. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be use issue due patient movement because of the turning episodes. patient had frequent bouts of diarrhea which necessitated frequent turns to change linens and after one turning episode the patient had a persistent capillary ooze. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
5320 | 1220648-2025-46701 | 23106161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106161&pc=OZD | 08/21/2025 | 10/23/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMART ASSIST |  |  | OZD |  |  | IMPELLA CP SMART ASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2026701401 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant indicated that a patient classified as high-risk for percutaneous coronary intervention received an impella cp device for mechanical circulatory support. the easy guide lumen did not pass through the impella while positioning the 0.018 guide wire, which resulted in an incorrect pump placement. subsequently, the pump was replaced with another impella cp pump. following this change, everything functioned properly. no patient harm was reported. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. unable to deliver or advance: the pump was returned for investigation. the red lumen was not attached to the pump and was also not found in the returned package. the pump was inspected and no damage was found. the pump functioned as per specifications when tested in a bucket of water. the cause of the delivery issue was not determined, as sufficient product was not returned for investigation and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5321 | 1220648-2025-46702 | 23106162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106162&pc=OZD | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2026701917 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant indicated that a patient identified as high-risk for percutaneous coronary intervention had an impella cp device implanted for mechanical circulatory support. access to the left ventricle was secured, and the pump was positioned within the ventricle. optimal positioning was achieved, and the patient remained in auto mode at p9 for the duration of the procedure, with flow rates around 3.8 liters per minute (lpm). the left anterior descending artery (lad) and left main (lm) artery were treated using shockwave therapy and two stents. during the shockwave treatment of the proximal lad, the placement signal was lost. however, both flow rates and motor current remained stable. throughout the procedure, the patient began to express discomfort due to pain in the left leg. pulses were still detectable but weak, and both legs exhibited reduced temperature. after completing the stenting of the lad and lm, contrast was injected through the side arm of the peel-away sheath, which indicated limited to no flow in the leg. in light of the patient's discomfort and concerns regarding ischemia, the impella device was gradually weaned and ultimately removed. the right coronary artery exhibited significant stenosis; nevertheless, the physician deemed it essential to remove the device to restore flow to the leg and to plan for stenting of the right coronary artery at a subsequent date. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation of the reported failure mode has been completed. the device was returned for investigation. ischemia: clinical details report that the patient experienced pain in the impella access site leg during a pci procedure. post-procedure, contrast injection showed limited flow down the limb, so impella was removed. the root cause of the injury was unable to be determined due to insufficient clinical details. optical sensor issue: clinical details report that while shockwave was being used during a pci procedure, the psnr alarm triggered. mc and flows remained stable. the procedure was completed, and impella was weaned and explanted. data logs were reviewed, and midway through the case, snr dropped to zero, contrast dropped, and gain/lamp increased. the psnr alarm triggered and remained active for the rest of the case. returned product was investigated and there were no visual abnormalities. when the pump was connected to a console, all 5s parameters were out of specification, but the pump passed the laser continuity test. this indicates that the issue was isolated to the sensor face, supporting data log and clinical details. clinical details report the complication occurred during shockwave use, data log optical signal signatures align with a damaged sensor, and this was supported by the status of returned product. the impella cp ifu states: "use with shockwave coronary intravascular lithotripsy (ivl) catheter at a distance of less than 20 mm between the optical sensor and ivl device may interfere with or damage the impella optical sensor." based on the clinical details provided, signatures noted in data logs, and review of returned product, the cause of the optical sensor issue was determined to be a damaged sensor, though the cause could not be established as it was not clear whether the devices were operated within a 20 mm distance at the time of the damage. device history review: the device passed all the post sterile inspection checks.
5322 | 1220648-2025-46706 | 23106163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106163&pc=OZD | 08/22/2025 | 10/03/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669154 |  | 000000000001000482 |  | Device was returned to manufacturer |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant in the eu had a 5.5 pump support ongoing for just under 24 days when the pump stopped. the team attempted to troubleshoot by aic replacement, but when this failed, the pump was explanted. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Correction g1 manufacturing site name, address, and fax number were inadvertently omitted on the initial manufacturer device report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5323 | 1220648-2025-46705 | 23106165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106165&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670706 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient admitted in with cardiac history. the cp was placed to allow for pci with support. the patient did get coded during that time and survived to transfer from lab to intensive care unit, where optical signal was noted to be unreliable. the pump supported but due to critical nature the patient expired after 10 hours. | [5 empty] | Corrections: e4. the investigation of the reported failure mode has been completes. no product returned. optical signal issue - within 6 hours of implant, placement signal not reliable (psnr) alarms occurred. placement signal (ps) waveform remained visible, but low. data log review showed ps trends down with flows over the 10 hours of the case, independent of p-level. motor current (mc) spike occurs at 0240 on (b)(6) 2025, 15 minutes before case end, where mc spikes to 1418ma and pump positioning noted to be in ventricle. 5s parameters from case end show no abnormalities to suggest sensor related issues. the cause of the optical signal issue was patient related due to data log review showing trending down ps and flows with the patient in critical condition on ventilator support and ultimately expiring on support. the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5324 | 1220648-2025-46707 | 23106166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106166&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026665243 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented with chest pains and became hypotensive. decision was made to place on extracorporeal membrane oxygenation with decision to place impella cp for venting. it was noted that the cp was primed however the physician was unable to gain access due to vessel size and calcification. procedure was aborted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had calcification and small vessel size preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks.
5325 | 1220648-2025-46708 | 23106167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106167&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688073 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 82 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented for mitra clip procedure. the procedure was unsuccessful and a physician was consulted for mitral valve regurgitation procedure, however the patient did not tolerate procedure hemodynamically and continued to rely on escalating does of pressors and inotropes. the physician opted to place intra-aortic balloon pump (iabp) and reassess patient¿s overall condition, however the patient continued to deteriorate and decision was made to escalate to impella 5.5. it was noted of a successful implantation of impella 5.5. axillary artery accessed via surgical cutdown. vessel loops were applied to control bleeding and graft was attached. axillary sheath inserted and secured with clamps. left ventricle (lv) was accessed. patient entered ventricular tachycardia (vt) rhythm and required multiple defibrillations to return to sinus. patient did not tolerate arrhythmia and deteriorated briefly requiring increased doses of pressors. impella was loaded onto wire and advanced into lv. impella positioning was ensured with transesophageal echocardiogram (tee) showing device at approximately 4.8cm. impella was slowly started and ultimately achieved p7 however flows were noted <1.5l. patient¿s map was noted >110. pressors were weaned and flows gradually increased. waveforms were fully separated and intermittent suction was noted. aic showed low volume to lv despite cvp showing 30. tee showed severely hypokinetic lv and right ventricle movement. decision was then made to place rp flex. upon completion of rp insertion, attention was then returned to impella 5.5. at this time, iabp resumed at 1:1 and flows were noted at p6 with approx 3.2l flows. graft was clamped and sheath was removed. graft was then trimmed and reposition sheath advanced into position. catheter noted at 41cm. sheath secured with sutures x3. access site was then closed with sutures and staples. during support it was noted that the patient¿s pressure dropped. advanced cardiovascular life support was activated and several rounds of cardiopulmonary resuscitation and medication administration were performed, however the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. ventricular fibrillation the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5326 | 1220648-2025-46709 | 23106168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106168&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026667327 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 103 KG |  |  | Death | Yes | Yes | False Alarm | Restricted Flow rate | Malposition of Device |  |  |  |  | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | [11 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that decision was made to escalate a patient from a cp to a 5.5. it was noted that the patient was prepped and draped. the 5.5 was inserted and initially the motor current was flat and impella in aorta alarm was sounding, but correct placement was verified with transesophageal echocardiogram (tee). shortly after, extracorporeal membrane oxygenation was ramped down and patient was decannulated. throughout the case, patient required shocks out of ventricular tachycardia(vtach) and ventricular fibrillation, both before and after 5.5 placement. otherwise, case went without incident. during support impella in aorta alarm intermittently sounded for very brief periods, but self resolveed. additionally, the patient had an episode of 33 seconds of v-tach. he was shocked x1 and became hypotensive so 500 cc bolus given. impella in aorta alarm had been occurring. placement was a little deep. there were also ongoing suction alarms. during night shift, patient was turned for bathing and coded, cpr initiated for 4 hours. care was ultimately withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Medical safety reviewed the event. the impella 5.5 related events occurring same day and after one to two days of implant was noted and approximately five days later, the patient was being repositioned coded and expired. it is unknown what occurred with the repositioning of the patient that led to the patient coding. given some unknown details and the impella 5.5 device was supporting the patient at the time of the demise, it cannot be dissociated from the outcome. however, given the patient's condition, this may have contributed to the demise outcome as well. additionally, the patient was on extracorporeal membrane oxygenation (ecmo). impella 5.5 was started. initially the motor current was flat and impella in aorta alarm was sounding, but correct placement was verified with transesophageal echocardiogram. the team was getting a false position in the aorta because ecmo was running high and the impella running. note: ecmo takes away the pulsatility and impella needs the pulsatility to give the positioning alarm. therefore, when pulsatility is lost, there will be a false alarm. further, as noted the patient was on ecmo and impella 5.5 at the same time which can compete with volume and cause ventricular tachycardia. it does not appear the patient had arrhythmia prior to implant. ecmo could have also contributed to the event/competed for volume. corrections: b.7 added information that was omitted from the initial report that was submitted. d.10 concomitant medical products and therapy dates was erroneously omitted from the initial and has been updated/populated accordingly. e.1 added initial reporter phone number as it was omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.6 complaint file codes for medical device problem code, health effect - clinical code and health effect - impact code were added to help better describe the reported event. note: h6 component code and h6 investigation type, findings and conclusion coding remain unchanged as initially reported. | G1 manufacturing site name should of been left blank on the initial report that was submitted. h6 medical device problem code a150202 was removed. type of investigation, investigation findings, conclusion codes and component codes were updated accordingly based on the completed investigation. investigation summary: the investigation has been completed. hypertension: clinical details reported that after an episode of ventricular tachycardia, the patient became hypotensive and 500 cubic centimeters bolus was administered. the product was not returned for investigation. the cause of the injury was not determined due to insufficient clinical details. pump suction/false alarm: on the fifth day, clinical details reported suction alarms triggering at p-6. data log review confirmed suction and positioning alarms began triggering at the beginning of the case. both 'impella position unknown' and 'impella position in aorta' alarms were triggered in addition to suction. clinical details reported the 'position in aorta' alarms and a flat motor current, but the pump position was verified. suction was also triggering at this time. motor current can go flat when there's a lack of pressure difference between the left ventricle and the aorta. the patient also received fluids the next day, which improved the patients pulsatility, implying the patient could have been experiencing volume issues. during this timeframe, signal to noise ratio drops but remains within spec. placement signal monitoring was disabled for the first time about three hours into the case, suspending the alarms. when it was turned back on, 'impella position unknown' alarms began triggering again. the placement signal monitoring was disabled two more times during the case, suspending alarms temporarily, but they would return whenever it was turned on again. on the fifth day of care, data log review confirmed the suction alarms at p-6, as well as an overall trend down in ps, snr rising slightly, and motor current dropping despite the p-level remaining unchanged. the patient was also re-cannulated on extracorporeal membrane oxygenation on this day. the cause of the pump suction and false alarms were most likely related to the patients condition, as seen by the down trending placement signal during suction events, flat motor current despite proper positioning, and the issues resolving when the patient received fluids. device history review: the pump passed all post sterile inspection criteria.
5327 | 1220648-2025-46711 | 23106170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106170&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709838 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was escalated to the impella 5.5. during 5.5 support it was noted that the impella was in good position wire was removed and impella started. then the patient went into ventricular fibrillation and shocked twice. additionally the patient was taken back to interventional radiology for coil x 2 to right and mid hep colon unrelated to impella, however in ir patient received 4 packed red blood cells (prbc), patient began running into constant suction alarms, stat echocardiogram (echo) performed impella seen to be directed towards mitral valve, impella repositioned under echo with surgeon, one more prbc to be given, ivc measurement 1.5cm, heparin esd on hold until bleeding resolved, surgeon to start systemic anticoagulation. the patient continued on support and the 5.5 was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed: the cause of the major bleed was not determined due to insufficient clinical details. pump suction: after reviewing logs, multiple suction alarms were noted. the cause of pump flow cannot be determined since insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
5328 | 1220648-2025-46714 | 23106172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106172&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  | 02/28/2026 |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501617 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Ventricular Fibrillation | [10 empty] | Disability | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient (pt) presented for diagnostic cath. initial access via right femoral difficult. left femoral had right upper extremity to left lower extremity bypass graft with low flow. severe lesion o plcx. the physician proceeded with balloon inflations and during second inflation pt experienced pulsless ventricular tachycardia/ventricular fibrillation (vtach/vfib) cardiac arrest. multiple defibrillations delivered unsuccessfully and code blue called, cardiopulmonary resuscitation (cpr)(advanced cardiovascular life support (acls) initiated by team, intubated by anesthesia. right femoral access exchanged for impella sheath and impella cp implanted with active cpr. impella flows 3.4-3.6 lpm with refractory vtach/vfib, shocks delivered and rose obtained. physician continued with percutaneous coronary intervention and delivered x1 des to lcx. image right fem distal to leg for perfusion, no flow visualized past introducer sheath. introducer was removed and inserted repo sheathed, imaged again and slow flow past sheath observed. vascular called to bedside to assess. physician imaged left fem and no flow observed to the common femoral. team decided on plan to attempt to wean impella, however pt had vtach/vfib on p4, defib x1 with response and impella increased to p9 by team. decision to leave impella in. during support it was also noted of right femoral site that was redressed for slow ooze, angle matched and forward tension re-applied, no oozing now. ultimately the patient was made a do not resuscitate and expired on support. | [5 empty] | Correction e4, h6. the investigation of the reported failure mode has been completed. ischemia/ ventricular fibrillation: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5329 | 1220648-2025-46713 | 23106173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106173&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694752 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was percutaneous coronary intervention who turned into unstable patient and required emergent impella support. patient was sent to intensive care unit with bleeding groin site and local staff were unable to control bleeding. femstop was placed and slowed bleeding but large hematoma and bleeding were still present. the physician time undressed/repositioned the impella sheath at a 45 degree angle match as well as foward tension sutures which stopped the bleeding. the patient also received 2 units packed red blood cells 1 unit of fresh frozen plasma and 1 unit of platelets. the cp was successfully weaned and explanted. | [5 empty] | Correction e4 the investigation of the reported failure mode has been completed. no product was returned for investigation. hematoma/ major bleed: bleeding was reported in icu. femstop was placed and slowed bleeding but large hematoma and bleeding were still present. the cause of the bleeding is determined to be use issue because the issue was resolved by repositioning and placing forward tension sutures. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5330 | 1220648-2025-46712 | 23106174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106174&pc=OZD | 08/22/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661322 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 60 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented non-st-segment elevation myocardial infarction. decision for impella cpsa insertion in the setting of acute myocardial infarction/cardiogenic shock (ami/cgs). the patient (pt) had successful implantation of impella cpsa in the setting of ami/cgs. pt also received 1 stent placed in the left main. decision was made to leave pt on impella support due to her increased edp and reduced ef. it was noted that at the end of the initial case it was noticed that there was substantial bleeding coming from the impella insertion site. pt had received a stent during the initial procedure and was placed on aggrastat immediately at the conclusion of the case. pressure at the site was initiated. discussion was had and a fem stop was placed and hemostasis was achieved. further discussion continued and based on the patients vital signs being stable the decision was made to explant the impella device. the impella was explanted without complications and 2 proglide closures were used. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5331 | 1220648-2025-46715 | 23106175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106175&pc=OZD | 06/11/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611367 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the team received a r long-term outcome and quality indicator (loqi) database reported upon subject 420-1084 for ae #1 of thrombocytopenia. the relationship listed as possibly related. the loqi states: the patient developed thrombocytopenia. admission platelets 82k/ul, nadir 11 k/ul (on (b)(6) 2025 (prior to impella)), nadir during impella support 20 k/ul (on (b)(6) 2025). it was later revealed that the patient had heparin-induced thrombocytopenia. the patient was switched to bivalirudin. patient survived. | [5 empty] | Correction: e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5332 | 1220648-2025-46716 | 23106177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106177&pc=OZD | 08/19/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/19/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515157 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented with atrial fibrillation with rapid ventricular response/acute decompensated heart failure. was admitted for suspicion of cardiogenic shock. decision made to place impella 5.5. the impella 5.5 was successfully placed. during support it was noted that the patient had multiple run of ventricular tachycardia (vt). the physician attempted to reposition pump however, unable to clear the papillae muscles, vt improved after reposition. additionally, they were holding heparin due to nose and mouth bleeding. last echo showed impella measuring 6.3 cm. interventional cards tried to pull back but not able to. also noted was platelets 30. heparin-induced thrombocytopenia panel sent in addition to transfusing platelets. the patient also had arrhythmias that resolved with amiodarone. the 5.5 was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.7 added information that was omitted from the initial report that was submitted. e.1 added initial reporter phone number and zip code extension as they were omitted from the initial report that was submitted. e.4 added selection as it was omitted from the initial report that was submitted. h.11 added instruction for use as they were omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the investigation has been completed. no product was returned. arrhythmia: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. thrombocytopenia, epistaxis: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): no product was returned. the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5333 | 1220648-2025-46717 | 23106178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106178&pc=OZD | 02/21/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/18/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609907 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Renal Failure | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had placed a cp to support the patient admitted in with known cad/pad in (b)(6) 2025. the pump was placed and supported for over 8 days until weaned and explanted rather than escalated to the 5.5 pump. the loqi database later in (b)(6) 2025, reported upon hemolysis and renal failure with relationship listed as "possibly" related. the patient survived and was medically managed. | [5 empty] | Data review: data logs showed pump ran without issue during support. there were positioning & suction alarms triggered during the case but resolved. there was a mc spike corresponding with the pump in wrong position likely causing an interaction between the pump and the aortic valve, abnormal ps or purge issues observed during support. product was not returned for review. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. renal failure/hematuria: the cause of the injury was not determined due to insufficient clinical details. h6. investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. medical device problem code was updated from a24 to a01. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5334 | 1220648-2025-46718 | 23106180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106180&pc=OZD | 08/12/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/17/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678076 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a 5.5 pump support placed for the patient admitted in with multivessel disease and low ef undergoing coronary artery bypass graft. the support was ongoing when there was bleeding on day 2 of the support which required blood products and re-exploration. the pump remained on and was still supporting when there was new vt that required cardioversion. the pump supported for 7 days and was weaned and explanted. | [5 empty] | B.3 revised as date was reported incorrectly on the initial report that was submitted. b.7 added information that was inadvertently omitted from the initial report that was submitted. e.1 added initial reporter phone number and zip code extension as they were inadvertently omitted from the initial report that was submitted. e.4 added selection as it was inadvertently omitted from the initial report that was submitted. g.1 manufacturing site name should have been left blank on the initial report that was submitted. g.3 the date was reported incorrectly on the initial report that was submitted. the date should have reflected 12-aug-2025. the investigation was completed and no product was returned for investigation. bleeding: chest exploration for bleeding and it was noted that the bleeding was primarily form adipose tissue. the root cause of access site bleeding is determined to be patient condition because of the bleeding from the non-impella site. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review was completed and the device passed all post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5335 | 1220648-2025-46719 | 23106184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106184&pc=OZD | 08/21/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/21/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026724283 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 67 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp experienced lactate dehydrogenase and alanine aminotransferase elevated, urine pink-tinged, no suction alarms on the impella reported. decision to explant cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections e4. the investigation of the reported failure mode has been completed. no product was returned for investigation. data logs were reviewed and the logs showed some suction alarms on first day of support. additionally, on first day of support a motor current spike with speed deviation was observed. the cause of the hemolysis was most likely due to biomaterial ingestion based on ingestion signature found in returned data log analysis. device history review: the device passed all the post sterile inspection checks.
5336 | 1220648-2025-46720 | 23106185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106185&pc=OZD | 08/21/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/21/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653895 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Injury | Yes | Yes | Malposition of Device | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that on day two with impella cp support the patient was transferred to another facility. upon arrival, the patient¿s labs were hemolyzed on p9 and were resent on p2. impella was slightly deep on echocardiogram. the team noted no pulse on impella leg, and it was cool. vascular consulted. ultrasound showed common femoral artery disease occluded. after team discussion with the physician the impella was lowered to p2, and the patient was hemodynamically stable. the impella was removed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D6b explant date added correction e4 the investigation of the reported failure mode has been completed. product was not returned for review. data review: data logs showed pump ran without issue during support. there were positioning alarm triggered in early of the case but resolved quickly. mechanical interaction with blood: the cause of mechanical interaction with blood was unable to be determined as limited clinical details were provided and no product was returned for review. positioning issue: the cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5337 | 1220648-2025-46721 | 23106186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106186&pc=OZD | 08/15/2025 | 11/06/2025 |  | 09/21/2025 |  |  | 08/15/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702480 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 86 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented for a scheduled percutaneous coronary intervention (pci) of the right coronary artery (rca). they had just started the procedure and got the guide in place and the patient coughed. when she coughed it caused a dissection at the ostia of the rca. patient started to deteriorate quickly and the decision was made to place impella cp. the cp was successfully placed and the pci and right heart catheterization were performed. during cp support it was noted that flows were left on p2 due to the inability to go any higher without suction alarms. additionally, the right groin impella site was oozing. manual pressure applied, lido with epi injected, mattress suture placed and neptune patch applied. the patient continued on support for a total of 125-43 hours and was successfully weaned and explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4. the investigation of the reported failure modes has been completed. major bleed: the product was not returned for investigation. the cause of the bleeding was not determined due to insufficient clinical details and no product returned. pump suction: data log review confirmed the occurrence of suction during the case. on the first day, the flows were slightly below spec at p-4, getting about 1.5-2 l/min. the cause of the pump suction was most likely patient related, as clinical reported the patient couldn't withstand higher flows than p-4 and required venoarterial extracorporeal membrane oxygenation (va ecmo) support. this was further confirmed in the data logs by the downward trending placement signal at the time of suction, and the increase in flows after placing the patient on ecmo. device history review: the device passed all the post sterile inspection checks.
5338 | 1220648-2025-46722 | 23106187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106187&pc=OZD | 08/07/2025 |  |  | 09/21/2025 |  |  | 08/07/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697216 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 125 KG | Non Hispanic | White | Injury | Yes | No | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant stated that an impella cp pump was implanted to support a patient admitted for cardiogenic shock. while on support, the patient¿s urine culture was positive and antibiotics were administered. in response to a concern for bleeding of unknown origin, the patient was transfused two units of packed red blood cells. additionally, there was a left groin preclose failure requiring a third perclose. the patient was found to have left common femoral extravasation so manual pressure was held and blood was transfused. the covered stent was unsuccessful and the patient decompensated. later, the patient expired. | [5 empty] | The root cause of the infection was unable to be determined due to insufficient clinical details. the cause of the bleeding is most likely use related since left groin preclose failed. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5339 | 1220648-2025-46723 | 23106188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106188&pc=OZD | 08/06/2025 |  |  | 09/21/2025 |  |  | 08/06/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687930 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 83 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was placed to support a patient admitted for cardiomyopathy. while on support, the patient experienced atrial fibrillation. additionally, anticoagulation was temporarily paused after the patient experienced a bloody bowel movement. no further episodes of bloody stools were reported. the patient suffered a stroke with no intracranial hemorrhage noted. a follow up brain ct was scheduled and the patient was noted to be alert, awake and oriented. at this time, the patient is still on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5340 | 1220648-2025-46724 | 23106189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106189&pc=OZD | 08/05/2025 |  |  | 09/21/2025 |  |  | 08/05/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682579 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5 on (b)(6) 2025. the patient had positive blood cultures for e coli and is receiving antibiotics (abx) currently. also, on 8/5 the patient had a couple of episodes of v-tach last night when nursing staff attempted to turn patient. no impella related issues overnight. the rn reports that the patient had gastrointestinal bleed (gib) yesterday, endoscopy done at bedside, ulcers noted. heparin gtt stopped at that time. no further gib symptoms noted; restarted heparin gtt this morning. on (b)(6) heparin stopped due to gi bleed. 2 units prbcs overnight. an endoscopy/colonoscopy done, ulcers shown on coloscopy, gi fixed with fibrin glue. the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 code f1905 was entered incorrectly on the initial report that was submitted. the investigation has been completed and no product was returned for investigation. major bleeding: gastrointestinal bleed was noted and the treatment was not provided in the clinical details. the cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. infection/ tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. the device history review was completed and the pump passed all post sterile inspection checks.
5341 | 1220648-2025-46725 | 23106335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106335&pc=OZD | 08/21/2025 | 09/21/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658357 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 93 KG | Non Hispanic | Asian | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient placed on impella 5.5 for mechanical circulatory support had episodes of ventricular tachycardia (vt). initially the vt was attributed to the impella positioning being too close to the septum and right ventricle strain. the patient continued to have runs of vt before going into vt arrest, requiring cpr with defibrillation prior to return of spontaneous circulation. medication was administered. there was worsening right ventricle dysfunction and the patient was deemed to not be a transplant candidate. denied ecmo also. care was withdrawn and the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product was returned. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. cardiac arrest: the cause of the injury was not determined due to insufficient clinical details. positioning issue: impella position being close to the septum. the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5342 | 1220648-2025-46726 | 23106336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106336&pc=OZD | 08/22/2025 | 10/22/2025 |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026704133 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 79 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient taken down for left ventricle assisted device (lvad) placement during procedure team was unsuccessful on reimplanting new lvad due to anatomical concerns. decision made to re implant a new 5.5 and right ventricle assisted device to explore alternative lvad options vs orthotopic heart transplantation. the pump was primed 5.5 inserted under transesophageal echocardiogram guidance wirelessly without difficulty. pump slowly ramped up from p2 to p8 while coming down on cardiopulmonary bypass machine. right ventricle failure concerns arose and rvad cannulation performed from right ventricle to pulmonary artery centrally. impella measured at 4.49 cm, patient to remain open with the hopes to bring back for alternative dvad techniques. it was note that the patient had a hemorrhagic stroke post surgery. ultimately care was withdrawn and the patient expired. | [5 empty] | Medical safety reviewed the event. the event describes a pump stop, which resulted in replacement of the device. this serious injury is attributable to a piece of muscle was found entrained in the motor housing. replacement pump was in for at least 64 days. the patient tested positive for bacterial growth and also there was pump suction. this is a serious injury attributable to the pump's association to the infection. the second impella 5.5 pump was replaced. the patient underwent device removal of the impella 5.5 and placement of a left ventricular assist device which was unsuccessful. therefore, a third impella 5.5 device was placed. the next day after start of this support, it was reported the patient had a hemorrhagic stroke post-surgery. in this case, it is unknown which device procedure, if any, contributed to the hemorrhagic stroke and therefore the impella 5.5 pump 3 cannot be dissociated; this pump was in place at the time of the event. however, it is noted care was withdrawn which appears to have led to the demise outcome. related medical device reports: this impella 5.5 device report is one of three devices associated with the event. refer to the related manufacturer report number as noted in section h10 for the medical device reports for the other two impella 5.5 devices. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | H.10 related report numbers were not added to the initial report submitted and were added. the investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
5343 | 1220648-2025-46727 | 23106337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23106337&pc=OZD | 08/22/2025 |  |  | 09/21/2025 |  |  | 08/22/2025 | 09/21/2025 |  |  |  |  | 09/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682492 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 76 KG | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was being escalated from a cp to a 5.5, however was having a aortic dissection repair first. the 5.5 was inserted into sidearm and advanced through the atrioventricular (av) with direct visualization. finished anastomosing distal end of graft. cross clamp removed and impella started. position of 5.5 was confirmed. it was directed away from mitral. position was mid ventricular. post op echocardiogram showed severe aortic insufficiency (ai) and surgeon was concerned she may need an aortic valve replacement (avr). impella was put in surgical mode and pulled back into the aorta while the valve was assessed. echo showed mild to moderate ai and surgeon decided not to intervene on her av. 0.35 wire and diagnostic catheter used to re-access the left ventricle( lv). wire exchanged for 0.018 wire. impella backloaded and placed back into lv and turned up to p-7. impella was alarming with impella position unknown as well as some suctioning. flows were also lower than normal for the p-level. position was checked multiple times. surgeon used a prefabricated aortic graft with sidearm and it was determined due to the size of her aorta, the impella outflow was inside the sidearm graft. flows were less than optimal and the surgeon placed a tourniquet around the proximal aspect of the sidearm graft, in order to redirect flow, and tunneled the tourniquet outside. this improved impella flow and her ef improved to 35%, to where the surgeon chose to keep the impella in place rather than remove. alarms were disabled for purposes of management. during 5.5 support, a cardiac computed tomography of the head also showed acute subarachnoid hemorrhage. ultimately care was withdrawn and the patient expired. | [5 empty] | Correction e4. the investigation of the reported failure mode has been completed. no product or data logs were returned for evaluation. optical signal issue: clinical details noted that "impella position unknown" alarms were triggered with suction alarms. pump was repositioned out aortic sidearm graft and flows improved, but placement signal alarms remained disabled. the cause of the optical signal issue could not be definitively determined as no product or data logs were returned for evaluation. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5344 | 1220648-2025-46645 | 23104026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104026&pc=OZD | 08/21/2025 |  |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  | 07/31/2025 |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2024380572 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 61 KG | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Deformation Due to Compressive Stress | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was placed to support a patient who had been admitted in as a bridge to transplant. on day 31 the team noted the pump had lost signal. the pump was monitored and later explanted after 49 days of support when the transplant occurred. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.1 revised as selection on the initial report that was submitted was incorrect. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.6 an additional medical device problem code was added. the investigation has been completed. there was no product returned. optical signal issue: the clinical details stated that about a month into support there were placement signal not reliable (psnr) alarms with absent aorta and left ventricle placement signal (ps). a kink in the catheter was noted and was thought to be a potential contributing cause for the psnr. the data logs showed that on the 21st, the day of the reported issue, the ps becomes unreliable and triggers psnr and controller error alarms as the 5s parameters show that the signal-to-noise ratio goes to 0, contrast barcodes, light decreases, gain rails, and the lamp remains maxed out at 100%. there were no changes in flows or motor current. there was no product returned. based on the clinical details and data log analysis, the root cause of the optical signal issue is most likely from a damaged fiber line. product damage (catheter kink): the clinical details stated there was a kink on the catheter that could be contributing to the psnr. due to lack of clinical details and no product returned, the product damage (catheter kink) cannot be determined. device history review: the pump passed all post sterile inspection checks.
5345 | 1220648-2025-46646 | 23104028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104028&pc=OZD | 08/21/2025 | 09/19/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683386 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | False Alarm | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp support placed for the support of a patient admitted in storm and underwent ablation. the pump did remain on for 6 days and during that time the pc was replaced for the purge leakage. additionally, there was an alarming for "in aorta" which prompted the team to assess position with imaging. the pump remained on for 2 more days and was then explanted. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. false alarm: data logs were not available for investigation. since neither product nor data logs were available for investigation, the cause of the false alarm could not be determined. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5346 | 1220648-2025-46028 | 23104244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104244&pc=OZD | 06/30/2025 |  |  | 09/20/2025 |  |  | 06/30/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026644172 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a small hematoma developed around the access site. manual pressure was applied to resolve the condition and the site remained properly dressed with gauze and clear dressings. a small amount of drainage was noted, but no intervention was deemed necessary. support continued uninterrupted. the following day, the patient developed hematuria and hemolysis. the patient was escalated to impella 5.5 support in response to the condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5347 | 1220648-2025-46648 | 23104245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104245&pc=OZD | 08/21/2025 | 10/29/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711286 |  | 000000000001000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Contamination | Restricted Flow rate | [5 empty] | Tachycardia | Asystole | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support the purge fluid was not exiting out of purge cassette. line was checked for kinks, d5w bag spike manipulated, purge cassette was removed and reinserted twice. some fluid was exiting tubing, but minimal. not enough to see purge fluid exit the impella catheter. new purge cassette was opened from another kit remedied the issue. purge exiting line impella catheter. while attempting to optimize impella position, the patient went into complete heart block/asystole. cpr was started. the patient was awake with chest compressions, moaning and moving arms. temp pacing wire was inserted. return of spontaneous circulation obtained impella flows 3.4l/min. then the patient started going in and out of non-sustained ventricular tachycardia, with it becoming sustained ventricular tachycardia. amino bolus was administered, synchronized cardioversion was performed multiple times, some more successful than others. a new companion sheath was opened due to contaminated table and inserted. on (b)(6) 2023 the patient had suction alarms during nightshift, restarting vaso gtt resolved. levo discontinued. the physician spoke with family wanting to give him another day before making any decisions, the patient was showing improvement; however, the family withdrew care today and the patient expired. | [5 empty] | No product was returned for investigation. data logs were reviewed and confirm the complaint condition. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the asystole cannot be determined as insufficient clinical details were provided. the cause of the pump suction was determined to be patient condition related. the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5348 | 1220648-2025-46650 | 23104246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104246&pc=OZD | 08/21/2025 |  |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701747 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the graft was bleeding due to dislocation and this blood loss contributed to suction observed on the pump. additionally, during the support the pump was repositioned repeatedly. this did not prompt pump replacement or removal. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. pump suction: patient had immediate suction believed to be because of volume related but not confirmed. after reviewing logs, suction alarms along with sudden spike in placement signal along with drop to 0 in pump flow was observed. the cause of the pump suction was most likely patient condition related per the log review. device in wrong position (positioning issue): pump was in deep which need position to be optimized. the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5349 | 1220648-2025-46649 | 23104247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104247&pc=OZD | 08/18/2025 |  |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMEPLLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667410 |  | 000000000001000482 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Hospitalization | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu had a 5.5 and ecmo support placed to support the patient who went to the operating room for emergent coronary artery bypass graft. the pump was supporting when there was a stroke and diagnosis of thrombocytopenia. after 8 days the care was withdrawn. | [5 empty] | Correction d4. the investigation of the reported failure mode has been completed. no product was returned for investigation. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. stroke: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5350 | 1220648-2025-46651 | 23104249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104249&pc=OZD | 08/21/2025 | 11/10/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026710599 |  | 000000000001000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 while supporting a patient could not go past p5 level without suction. echocardiogram was done and the impella was pulled back. the impella looked mid-ventricle, but maybe it was against the septum. waveforms were still decoupled. there was discussion about a new impella because the flows were never good since insertion, and any type of adjustment was not fixing the issue, but the patient continued five more days with impella 5.5 support. | [5 empty] | Additional information: the investigation of the reported failure mode has been completed. pump suction/low pump flow: data logs were reviewed and the complications were confirmed. product was not returned for investigation. since limited clinical details were provided, data logs were inconclusive, and product wasn't returned for investigation, the cause of the low flows and pump suction couldn't be determined. device in wrong position: no further details were provided regarding how the pump got out of position. there were no improper position alarms in the available data logs. product was not returned for investigation. since limited clinical details were provided, data logs were inconclusive, and product wasn't returned for investigation, the cause of the positioning issue couldn't be determined. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5351 | 1220648-2025-46652 | 23104250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104250&pc=OZD | 08/26/2024 | 11/10/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482076 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Nerve Damage | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient experience nerve damage and pain that they encountered delivery issues attempting to place impella 5.5. the loqui adverse event noted right upper extremity (rue) neuropathy with relationship noted as probably. the loqi stated: "rue pain/neuropathy most likely 2/2 brachial plexus injury with r axillary impella placement." medication, gabapentin 300 3 times a day, was given to the patient with improvement in pain control. additionally, during the unsuccessful 5.5 attempt, an arteriotomy was performed sufficient to accommodate a 10 mm dacron graft. the impella 5.5 was placed into the graft and subsequently through the axillary and subclavian arteries according to manufacturer instructions over a wire, which had previously been placed into the left ventricle. the impella 5.5 would not make the turn from the subclavian into the innominate artery, likely due to the patient's redo sternotomy anatomy. the team tried several times and was concerned with either dissecting or perforating the artery. consequently, the wire and impella 5.5 were removed, and a new wire was placed through the graft down the aorta into the left ventricle followed by placement of an impella cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: b1, e4, g2. the investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition due to patient anatomy because impella 5.5 would not make the turn from the subclavian into the innominate artery, likely due to the patient's redo sternotomy anatomy despite multiple methods and the insertion was aborted. nerve damage/ pain: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5352 | 1220648-2025-46654 | 23104251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104251&pc=OZD | 08/21/2025 |  |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698845 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was inserted but subsequently aborted due to patient's vasculature. there was no patient harm reported, and the patient outcome was noted to be unknown. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary : unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had difficult vasculature preventing successful delivery of impella. device history lot : device lot : 1943288. device history review : the pump s/n: 608638 passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5353 | 1220648-2025-46655 | 23104253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104253&pc=OZD | 08/20/2025 | 10/20/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026670752 |  | 000000000001000080 |  | No |  |  | No |  |  | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed a new left ventricle thrombus shown in echo. the impella was weaned and explanted. the patient outcome was reported as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombosis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombosis could not be determined.
5354 | 1220648-2025-46656 | 23104254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104254&pc=OZD | 08/21/2025 | 09/19/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682090 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp pump placed to support the patient admitted in with post cardiotomy cardiogenic shock (pccs) & low cardiac output syndrome (lcos). upon starting the device, flows would not get above 1.5 liter per minute. the motor current was low and a suction alarm was present. it was explained to the surgeon that there was a possibility of a clot ingestion due to suspected hit and plt 60 prior to cardiovascular operating room arrival. the surgeon elected to remove the impella cp and insert new impella cp device. a clot was noted on the inlet once device was removed. | [5 empty] | Correction: b1, e4. the investigation of the reported failure mode has been completed. product was not returned for review. data review: data logs confirmed low pump flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. there were no motor current (mc) spikes, abnormal placement signal (ps) or purge issues observed during support. low flow/suction: the cause of low flow/suction issue was unable to be determined as limited clinical details were provided and no product was returned for review. thrombus: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5355 | 1220648-2025-46658 | 23104255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104255&pc=OZD | 08/21/2025 |  |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681216 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered difficulty during the insertion of the impella cp. the radiologist stented the left femoral artery at first. after that it was possible to implant the impella. it was not very easy because of massive kinking. after the impella was placed, intervention was made of the last remaining vessel. patient was stable all the time and the pump was removed in the same procedure. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. delivery issue: the cause of the delivery issue is determined to be patient condition due to tortuosity of the vessels and massive kinking. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5356 | 1220648-2025-46657 | 23104256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104256&pc=OZD | 08/21/2025 | 09/19/2025 |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084304 |  | 0042-0000-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a 5.5 support ongoing for now 39 days for the patient admitted in and awaiting heart transplant. the pump is one of the aic serial number noted in prior fda recall. the md has made choice to not replace the aic for this patient despite concerns of voltage and recall. reporting for possible risk. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5357 | 1220648-2025-46659 | 23104257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104257&pc=OZD | 08/21/2025 |  |  | 09/20/2025 |  |  | 08/21/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709329 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues during insertion of impella cp. due to surgically altered vascular anatomy, there were challenging conditions during the implantation of the impella cp. as a result, the case was aborted and the pump was exchanged for a new one. | [5 empty] | D.3 added manufacturer fax number as it was omitted from the initial report that was submitted. d.4 added model number as it was omitted from the initial report that was submitted. investigation summary: the investigation has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had altered vascular anatomy preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5358 | 1220648-2025-46661 | 23104258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104258&pc=OZD | 04/02/2025 | 01/21/2026 |  | 09/20/2025 |  |  | 08/28/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000-US |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During the investigation of a returned automated impella controller intermittent oscillating contrast was identified. there was no patient involved in the reported issue. | [5 empty] | The reported event did not involve patient use. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Investigation summary: data analysis: console logs for the reported date were not available. however, lt data logs confirmed the reported issue. the lt logs indicate that the unreliable placement signal persisted for approximately nine days during the case, with varying intensity. this correlated with low snr and oscillating contrast, which resulted in placement signal not reliable alarms (#1036). rt logs also showed oscillating contrast along with low snr values. further review of the lt logs revealed a controller error (#1045). device analysis: console ic6242 was not returned to field service for further investigation. root cause: the root cause of the placement signal not reliable alarm (low snr and oscillating contrast) could not be determined because the aic was not returned for analysis. the controller error (event #1045) was attributed to an optical bench firmware communication protocol anomaly, which caused misalignment of data packets received by the epc. the aic software operated as designed. device history lot: n/a. device history batch. subcomponent name: n/a. material number: n/a. lot number: n/a. serial number: n/a. device history review: q1) service history review - are there an qns associated with the complaint device¿s most recent service report? there were no nonconformances observed in most recent fsr. q2) subcomponent lot review - were there any other product malfunction complaints in this subcomponent lot related to this failure mode? n/a. q3) complaint history review - have there been any other complaints against this console? no. product history review summary : there were no issues related to complaint discovered when reviewing the product history of console ic6242.
5359 | 1220648-2025-46660 | 23104259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104259&pc=OZD | 08/22/2025 | 09/19/2025 |  | 09/20/2025 |  |  | 08/22/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657556 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with an impella cp pump had hemolysis overnight. the issue resolved with decreasing p level. the physician opted to remove the device anyway and place patient on va ecmo. | [5 empty] | Correction: e4 the investigation of the reported failure mode has been completed. the pump and the purge cassette were returned for investigation. mechanical interaction with blood (hemolysis): data log was not provided or was not retrieved from the remote server. according to the given clinical details, hemolysis resolved after changing to a lower p-level. no damage was noted on the returned product. the pump was soaked in warm water, and biomaterial was found on the impeller. the biomaterial could not be removed from the impeller, but the pump was then tested in a bucket of water without any major performance-related issues. the cause of the hemolysis issue was most likely ingested biomaterial related to the patient's condition, based on the given clinical information and returned product investigation. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5360 | 1220648-2025-46663 | 23104261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104261&pc=OZD | 09/01/2025 |  |  | 09/20/2025 |  |  | 09/02/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 1332595 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during preparation for a procedure, the automated impella controller (aic) generated a blue screen with an ekg rhythm along the bottom upon start up, and not the normal start up screen. the case was completed with another aic. no patient harm was reported. | [5 empty] | The console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. the reported failure mode was reproduced during testing. the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. before returning the console to the customer, the following actions will be performed: 1. replace the pbm 2. perform preventive maintenance 3. perform a full functional check. | The reported event did not involve patient use. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5361 | 1220648-2025-46662 | 23104262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23104262&pc=OZD | 08/29/2025 | 10/13/2025 |  | 09/20/2025 |  |  | 08/29/2025 | 09/20/2025 |  |  |  |  | 09/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 1658143 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 71 KG | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that they encountered an automated impella controller (aic) battery issue. it was reported that when the nurse unplugs the aic from red plug, the battery immediately goes from 100% to 50%. then, the nurse plugs aic back in and aic begins charging and eventually goes back to 100%. there was customer concern that the aic would potentially stop charging and cause a pump stop. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. this is 1 of the 2 report, where this report represents the aic and the other report is for pump. | D9 device return date has been updated.
5362 | 1220648-2025-31578 | 23092717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092717&pc=OZD | 02/24/2025 | 09/18/2025 |  | 09/19/2025 |  |  | 02/24/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596895 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an impella cp pump implanted in a patient undergoing a high-risk percutaneous coronary intervention. when shockwave was initiated, the placement signal on the impella cp stopped working; however, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data log were not provided for this case run. the pump was returned for investigation. no physical abnormalities were reported on the returned product; the pump passed the laser continuity test, and the alarm was reproduced during pump testing. the root cause of the optical signal issue was a damaged sensor, based on returned product review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5363 | 1220648-2025-31579 | 23092726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092726&pc=OZD | 03/06/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 03/06/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594244 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. the complainant reported that false placement signal low alarms and inaccurate aortic readings occurred. an echocardiogram indicated the impella was in perfect placement. the patient was successfully weaned from the pump and the device explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows no abnormalities related to all 5s optical signal parameters. an impella position in aorta alarm was reported on (b)(6), without any major trends observed in the placement signal and motor current. the imc log was not provided for the review console's software version. the cause of the optical signal issue was most likely patient condition related to false positioning alarm. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5364 | 1220648-2025-46602 | 23094410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23094410&pc=OZD | 08/20/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695282 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 137 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during removal of an impella cp, it was noted that there was a thrombus present near and around the outlet. the patient also was noted to have loss of distal pulses in right lower extremity (rle). the thrombus at the popliteal was removed and flow was restored to the rle. this was the site where the impella cp was placed. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | The investigation of the reported failure mode has been completed. no product was returned for investigation. thrombosis/ ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5365 | 1220648-2025-46603 | 23094848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23094848&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696856 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after removal of the impella cp, the impella leg demonstrated signs of ischemia. vascular surgery was consulted and was managing the leg. at this time, the medical doctor felt it was not in the best interest of the patient to proceed with ventricular assist device placement. the patient continued to decline and pressor support was initiated. due to the patients anatomy, the doctor felt the axillary was not appropriate and would like to preserve the left leg for bailout options due to previous complications with sternotomy procedures. it was also noted a thrombus/biomaterial was observed in the patient or on the device. no further information was provided. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. thrombosis/ ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional. interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
5366 | 1220648-2025-46604 | 23094972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23094972&pc=PYX | 08/20/2025 | 09/18/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025574572 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 91 KG | Non Hispanic |  | Injury | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on bipella support decompensated significantly overnight. the patient was experiencing significant hemolysis and suspicion of clot ingestion. a decision was made to remove the impella rp flex and cannulated for veno-arterial extracorporeal membrane oxygenation. additional information was received that the pump can not be sent back. it did however ingest a clot. | [5 empty] | A.6 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. hemolysis: the data logs show that about 38 minutes into the case there is a large step up in motor current (mc) with a decrease in pump and purge flows. lysis solution was added which seems to resolve the issue after about 3 hours. mc and flows remain steady for 17 hours while the pump is kept in. there are no positioning alarms noted. based on the clinical details provided and data log analysis, the root cause of the hemolysis is due to biological material ingestion. thrombosis the cause of the injury was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5367 | 1220648-2025-46605 | 23095071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23095071&pc=OZD | 08/25/2025 |  |  | 09/19/2025 |  |  | 08/26/2025 | 09/19/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465594 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | [8 empty] | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient underwent impella insertion for acute myocardial infarction. there were no problems with the insertion technique, and the patient was managed without any issues afterward. however, after a few days on support, the position waveform suddenly disappeared and immediately afterwards a controller error alarm was issued. the pump was used with an alternative automated impella controller, and the position waveform was subsequently restored. there was no patient harm. | [5 empty] | The investigation for the reported event has been completed. the reported failure mode was reproduced. it was found that the bulb power was abnormal (0.04 w). root cause: the placement signal issue and the controller error were caused by an optical bench lamp malfunction. service & repair: before returning this console (b)(4) back to the customer, fs was instructed to perform the following, replace the optical bench (0042-1012) as a precaution. perform a full functional check (0042-7316). | A.4 is unknown. the automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5368 | 1220648-2025-31580 | 23095244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23095244&pc=OZD | 08/03/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 08/03/2024 | 09/19/2025 |  |  | 01/31/2026 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464148 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. it was reported that during support, there was a sensor failure where no placement signal and no lv signal were noted. the pump was not explanted as a result of the event. the patient was stable and remained on support. the patient was successfully weaned from the pump and the device explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product or data logs were returned for evaluation. clinical details noted placement signal was lost on automated impella controller (aic) while pump remained in proper position and continued to support patient. the failure mode investigated was changed from "placement signal issue" to "optical signal issue" as the impella 5.5 has an optical sensor and no differential pressure sensor. the root cause of the optical signal issue cannot be determined as no product or data logs were returned for evaluation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5369 | 1220648-2025-46607 | 23095302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23095302&pc=OZD | 01/21/2025 |  |  | 09/19/2025 |  |  | 08/27/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544621 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pneumonia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported in a publication that there was a need for surgical exploration of the access site due to minor bleeding, and was treated by suture. additionally, there was pneumonia that was medically managed by antibiotics during the hospital stay. none of theses issues were known by the company representative at the time of pump use. the patient survived. | [5 empty] | A.4 - a.6 are unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ literature citation: glanowski, s., pietryka, b., okarski, m., niewiara, l., legutko, j., & kleczynski, p. (2025). dissection of epigastric artery after percutaneous coronary intervention with impella cp device. advances in interventional cardiology, 21(2), 278¿279. https://doi.org/10.5114/aic.2025.151707. | The investigation of the reported failure mode has been completed. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. infection : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5370 | 1220648-2025-46608 | 23095433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23095433&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601552 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 52 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced suction issues. albumin was given to resolve the issue. additionally, the patient had gastrointestinal bleeding. blood products and anticoagulant was stopped to achieve hemostasis. the patient also experienced atrial fibrillation with rapid ventricular response and was treated with an amiodarone bolus. later, the patient was successfully weaned from the pump and survived. | [5 empty] | No product was returned for investigation. the root cause of the paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the cause of the retroperitoneal hemorrhage was not determined due to insufficient clinical details. the cause of pump suction was most likely due to patient related per the log review and clinical. the device passed all post sterile inspection checks. | A.5 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5371 | 1220648-2025-46609 | 23095501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23095501&pc=OZD | 08/27/2025 |  |  | 09/19/2025 |  |  | 08/27/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161398 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that when the automatic impella controller was turned on there was a battery failure. it was not in use on a patient. | [5 empty] | Section a: no patient was involved. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5372 | 1220648-2025-46610 | 23096001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096001&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple patients treated with impella pumps experienced ischemia and strokes during support. the medical staff reported that they had two or three strokes and three ischemic arms from axillary impellas this year. two of the strokes occurred on patients with heartmate 3 devices, and one was a catastrophic stroke. the banner tucson lab found the partial thromboplastin time and anti xa results to be sub-therapeutic, and this has been corrected. patient information was not provided. | [5 empty] | Patient information was not provided. therefore, section a was left blank. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. ischemia/ stroke: in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the necessary materials were not returned for analysis so the serial number could not be identified to complete history review inspection.
5373 | 1220648-2025-46611 | 23096069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096069&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693703 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in for high-risk percutaneous coronary intervention. while on support, there were multiple hypotensive events. during impella explant, two perclose were deployed with oozing at the site. one angioseal 6 french was deployed in the right femoral artery with good hemostasis achieved. heparin was held. the device was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. major bleed : the cause of the injury is most likely use related since two perclose were deployed, one angioseal 6 french was deployed and heparin was held to achieve hemostasis. hypotension : the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
5374 | 1220648-2025-46612 | 23096242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096242&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697823 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 85 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and when the access site dressing was changed, small ooze was noted from the femoral site. one unit of packed red blood cells was administered. additionally, despite clinical improvements, the family made the decision to withdraw care and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned for investigation. major bleed: the cause of the bleeding is determined to be patient condition because of the bleeding from other sites. device history review: the device passed all the post sterile inspection checks. | A.6 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5375 | 1220648-2025-46613 | 23096668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096668&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720190 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient after an impella cp was inserted and once on support went into ventricular tachycardia/ventricular fibrillation and was shocked twice. the impella support continued until successful weaning and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. ventricular fibrillation/ tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5376 | 1220648-2025-31581 | 23096670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096670&pc=OZD | 03/09/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 04/03/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1717562 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 121 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 support ongoing for 26 days when the optical signal was lost on the automated impella controller (aic). the pump remained on for support despite the loss of signal. during the investigation, team made choice to investigate aic as well. there was no aic replacement noted, and the pump remained on for support. the decision was eventually made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the placement signal alarm was an optical bench malfunction. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5377 | 1220648-2025-46614 | 23096727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096727&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690225 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was inserted in the right femoral artery successfully. the patient had percutaneous primary coronary intervention while on impella support. a dissection of the left femoral artery from the sheath occurred. vascular surgeon was consulted to remove the sheath. the patient remained in stable condition. | [5 empty] | The investigation of the reported failure mode has been completed. peripheral artery dissection: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
5378 | 1220648-2025-46616 | 23096849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096849&pc=OZD | 08/20/2025 | 10/23/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026693531 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 80 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had one unit of blood administered which was believed to be due to a hematoma. the patient continued with impella 5.5 support until successful weaning. the patient survived. | [5 empty] | The investigation for the reported hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5379 | 1220648-2025-31582 | 23096930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23096930&pc=PYX | 03/15/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 03/15/2024 | 09/19/2025 |  |  | 11/30/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024433028 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. the complainant reported a placement signal issue with the impella rp that had appeared to be there since rp placement. no further information was provided. patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs show that about 1 hour into the case there are placement signal not reliable (psnr) and suction alarms. snr drops to 0, lamp increases, there is a spike in gain, and contrast decreases. light is stable. cvp spikes to 300 mmhg and does not resolve. the console was replaced and the problem did not resolve. lamp remained at 100, snr extremely low, and cvp out of spec at 300 mmhg. the root cause of the optical issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5380 | 1220648-2025-31583 | 23097131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097131&pc=PYX | 09/01/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 09/01/2024 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025534004 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported a patient was implanted with an impella rp flex implant for circulatory support. during support there were placement signal not reliable alarms. the patient's position remained unchanged and the cable was fully engaged. the staff monitored the patient and position and support continued. no further information was provided. patient expired despite being on support with a left ventricular assist device and triple vasopressors. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. clinical details noted placement signal not reliable (psnr) alarm was triggered with no changes in hemodynamics. data logs were reviewed and an alarm #1055 psnr (opm invalid bench signal) was triggered during support. at time of alarm, cvp goes out of range with raw electrical placement signal remaining stable indicating the dp sensor was not affected. additionally, at time of alarm, contrast becomes oscillating with a decrease in snr, light, and gain. lamp increases to 100%. no patient movement was noted at time of alarm. upon review of data logs, it is apparent that the console is the potential cause of the issue. . no problem was found with the pump. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5381 | 1220648-2025-46618 | 23097152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097152&pc=OZD | 08/10/2025 |  |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674463 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 14 YR |  | Male | 52 KG | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a cardiomyopathy patient. it was reported that the optical sensor stopped displaying the position waveform after about the second day of use. support was continued and there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # ==> (b)(4). impella cp sn (b)(6) + purge cassette returned for investigation. optical sensor issue - data log review showed psnr alarm from ps railing to 3000mmhg with snr dropping to zero, contrast barcoding, lamp remaining unchanged at 100, gain decreasing, and light decreasing. the pump was returned and inspected. pump had abnormal 5s parameters with an snr of 1 and a kink found in the catheter just below the motor housing. the pump failed the laser continuity test at this kink. the cause of the optical signal issue was a kink in the catheter causing material fracture of the optical fiber, due to an unintended user error. pump s/n (b)(6) passed all post sterile inspection checks.
5382 | 1220648-2025-31584 | 23097300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097300&pc=OZD | 12/10/2023 | 09/19/2025 |  | 09/19/2025 |  |  | 12/10/2023 | 09/19/2025 |  |  | 10/31/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406202 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal not reliable alarm was triggered following impella 5.5 repositioning. the alarm was subsequently silenced, and support with the implanted pump was maintained. the decision to withdraw care was made, and the patient eventually expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs available. the cause of the optical signal issue was not determined since product and data logs were not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5383 | 1220648-2025-31585 | 23097301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097301&pc=PYX | 07/25/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 07/25/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025563919 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient experiencing acute myocardial infarction and cardiogenic shock. it was reported that roughly three days into support, the pump began to experience placement signal failures. the staff disabled the audio and support with the implanted pump was maintained. the patient was listed for a heart transplant; however there was a sudden increased pressor requirement and instability that called for extracorporeal membrane oxygenation (ecmo) shock. the family decided against any further treatment and withdrew care. shortly after the decision was made to withdraw care, the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product/logs were returned. the fm was changed to "placement signal issue" as reported in the field since there is insufficient evidence of which sensor is involved. the root cause of the placement signal issue was not determined since product/logs were not returned and insufficient clinical details were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5384 | 1220648-2025-31586 | 23097323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097323&pc=OZD | 05/15/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 05/15/2024 | 09/19/2025 |  |  | 12/31/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025444258 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant placed an impella 5.5 for support of a patient. the placement signal was lost during the support. the decision made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump and data logs were not returned for investigation. as per clinical observations, 'placement signal not reliable' alarm was noted during the case. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5385 | 1220648-2025-31587 | 23097363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097363&pc=PYX | 04/15/2023 | 09/19/2025 |  | 09/19/2025 |  |  | 04/15/2023 | 09/19/2025 |  |  | 07/31/2023 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022124648 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the device's placement signal failing after patient was rolled during a procedure. patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the data logs were reviewed and a placement signal not reliable alarm triggered with ps spiking to 3000 mmhg for 1 hour and then the alarm resolved and placement signal (ps) returned to normal. throughout the run, ps was trending up and shifting down. the cause of the issue was not determined since product was not returned. all pertinent information available to abiomed has been submitted at this time.
5386 | 1220648-2025-46619 | 23097367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097367&pc=OZD | 08/21/2025 |  |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720647 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Degraded | Failure to Advance | Device Damaged by Another Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had an impella cp placed to support a patient admitted in with a plan for high risk percutaneous coronary intervention. it was reported that the 0.018 wire was stuck in the cannula and was unable to remove. therefore, the whole system had to be removed without a problem. it was also noted that the guidewire was unravelling. a new impella cp and 0.018 wire were used for a successful placement. there was no patient harm. | [5 empty] | Product complaint #(b)(4). investigation summary: device interaction or damage by another device: the guidewire and pump (b)(6) were returned for investigation. the returned guidewire was found to have a stretched coil and kink/fracture on the guidewire core (see pi-17604604875059152). no damage was observed on the returned pump. according to clinical information, the 0.018" guidewire could not be removed from the cannula, so both the pump and guidewire were removed together. the device interaction was with the guidewire; however, the cause of the guidewire damage was not determined due to insufficient clinical information. device history lot: device batch: 1958578. device history batch: subcomponent lot: na. device history review: the pump (b)(6) passed all post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5387 | 1220648-2025-31589 | 23097453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097453&pc=PYX | 06/09/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 06/09/2024 | 09/19/2025 |  |  | 06/30/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024369394 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella rp pump triggered a placement signal not reliable alarm during patient support. the audio on the alarm was disabled, and support with the pump continued. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs show that the 'placement signal not reliable' alarm appeared for 15 minutes, along with several suction alarms and a fluctuating placement signal trend during the time of the event. there was no trend noted related to hard failure of the optical signal 5s parameters. during the event, central venous pressure was reported to be negative. the cause of the optical signal issue was most likely patient condition related, based on data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5388 | 1220648-2025-46617 | 23097532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097532&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694648 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information reported the morning the patient had a trialysis catheter placed for continuous renal replacement therapy (crrt) and they thrashed in bed, which cause pump migration leading to placement signal low alarm. the physician adjusted pump under echocardiogram and resolved the issue. | The complainant had an impella cp placed to support a patient admitted in with acute myocardial infarction/cardiogenic shock. while on support, continuous renal replacement therapy (crrt) was started. the patient thrashed in the bed, which caused pump migration leading to a placement signal low alarm. the doctor adjusted the pump under echocardiogram and resolved the issue. crrt was continued. it was noted that the patient was very oozy from all her access sites. platelets dropped to 92 from 280. the clinical team was going to send a heparin-induced thrombocytopenia panel off later that day. heparin was off at that time due to bleeding. the patient was going to get three units of packed red blood cells with the first being administered at that time. despite bleeding, levophed had been weaned off and the patient's hemodynamics and labs looked great. the impella cp was explanted and the patient survived. |  |  |  |  | Additional information received b5 and h6 updated. the investigation of the reported failure mode has been completed. no product was received for evaluation. major bleed: patient bleed all the access sites. the cause of major bleed most likely use issue related since withholding heparin resolved the impella access site bleed. renal failure: the cause of the injury was not determined due to insufficient clinical details. device in wrong position (positioning issue): patient thrashed in bed, which cause pump migration leading to placement signal low alarm but was resolved by adjusting the pump under echo. the cause of the positioning issue was determined to be an unintentional use error as the pump was pulled out of position by patient movement. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5389 | 1220648-2025-31588 | 23097546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097546&pc=PYX | 07/25/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 07/25/2025 | 09/19/2025 |  |  | 09/30/2023 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023193945 |  | 0046-0035 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient post-op to cardiothoracic surgery. it was reported that there was an unreliable placement signal sensor during use of the device. support continued despite this issue. as there was no concern from the team. no further information was provided. patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the data logs were reviewed and confirm the negative cvp as well as the placement signal not reliable alarms. the alarm code specified that the optical sensor was out of range during the time period this alarm was triggered. it is also noted that there are multiple suction alarms during this time. the suction alarms and the placement signal unreliable alarms resolve at the same time. all of the 5s parameters show that the sensor is functioning, the signal is just dampened by an increase in noise as a result of the suction conditions. the root cause of the optical signal issue was most likely patient condition as the pump was experiencing suction the time of the event. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5390 | 1220648-2025-46621 | 23097621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097621&pc=OZD | 08/20/2025 | 09/18/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606211 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with coronary artery disease had an attempted impella placement. the pump was opened and attempted to be used, but was aborted due to the patient's cardiac anatomy. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. unable to deliver or advance (delivery issue): the cause of the deliver issue was determined to be patient condition as the patient had resistance met at femoral artery and coronary artery disease preventing successful delivery of impella. patient anatomy or condition prior to therapy: the cause of the issue was most likely patient condition related. device history review: the device passed all the post sterile inspection checks.
5391 | 1220648-2025-46623 | 23097641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097641&pc=OZD | 08/17/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675404 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp and ecmo were supporting a patient with acute myocardial infarction/cardiogenic shock. the patient experienced ventricular fibrillation requiring defibrillation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
5392 | 1220648-2025-46622 | 23097666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097666&pc=OZD | 08/18/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707182 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 was placed to support a patient with st elevated myocardial infarction. an optical signal issue prompted the team to adjust the pump position and review the blood pressure and placement signal correlation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
5393 | 1220648-2025-46625 | 23097818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097818&pc=OZD | 08/20/2025 | 10/29/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658297 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp exhibited bleeding at the access site. manual pressure was held for 20 minutes which did not seem to slow the bleeding. therefore, the impella cp was explanted. a contralateral femoral balloon tamponade was used in 5 minute increments to achieve hemostasis in response to bleeding during explantation. | [5 empty] | The cause of the bleeding was not determined since product was not returned and insufficient clinical details provided. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5394 | 1220648-2025-46626 | 23097824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097824&pc=OZD | 08/20/2025 | 09/18/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708681 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 74 KG | Non Hispanic | Asian | Injury | Yes | Yes | Inaccurate Flow Rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was saturated flow and thrombus during impella cp support. after the patient was brought to the intensive care unit, acute thrombus and saturated flow were found. as a result, the decision was made to remove and exchange the impella cp with a new one. | [5 empty] | Correction: b1, e4 the investigation of the reported failure mode has been completed. no product was received for evaluation. thrombosis: the root cause of the injury was unable to be determined due to insufficient clinical details. high/saturated pump flow: data logs were reviewed and the logs showed pump flows being saturated when pump was turned up to p-4 with flows greater than 2.5l/min. the log at time of saturated pump flow showed sudden loss of pulsatility for motor current and pump flow. no motor current spike was present at time of start of saturated flows. pump flows remained saturated for remainder of case. the cause of the saturated flow could not be determined as no product was returned and no motor current spike was present in data logs indicating possible ingestion. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5395 | 1220648-2025-46627 | 23097835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097835&pc=OZD | 08/20/2025 | 12/23/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026716784 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 143 KG | Non Hispanic | White | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had an impella cp pump placed to support the patient admitted in for high-risk percutaneous coronary intervention. during device insertion, there was significant bleeding from the peel-away sheath diaphragm due to pressure applied to the companion sheath. this was noted as necessary due to patient¿s height and length of catheter needed to cannulate coronary arteries. the guide catheter needed additional pressure to reach the coronary arteries. at the end of the case, the device was explanted and the patient survived. | [5 empty] | During complaint review by the engineering team, it was found that the impacted product should have been the introducer, and not the pump. the insertion kit is addressed under manufacturer report number (b)(4). | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5396 | 1220648-2025-46628 | 23097837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097837&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678516 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 87 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was placed to support a patient with acute myocardial infarction/cardiogenic shock. while on support, there was some bleeding from the graft site and low platelet count. platelets and one unit of packed red blood cells were transfused to achieve homeostasis. the cutdown was closed and a three point fixation was completed. anticoagulation was turned off, and the bleed resolved. later, purge flows began to drop and there was a slight increase in purge pressures. the doctors made the decision to change the purge fluid from bicarbonate to heparin. impella support continues. | [5 empty] | The cause of the access site bleed was determined to be patient condition related. the cause of the renal failure cannot be determined as insufficient clinical details were provided. the pump passed all post-sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5397 | 1220648-2025-46629 | 23097848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097848&pc=PYX | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026693657 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female | 112 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Device Sensing Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella rp support motor current spike with placement signal spikes and drops in flow occurred. it was suspected that the pump ingested a clot so the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | No product was returned for investigation. the cause of the placement signal issue was biomaterial ingestion due to the mc spike seen in the data logs that triggered a higher ps. the cause of the low flows was biomaterial ingestion due to the motor current spike that triggered a trending down of flows at the same p-level. the cause of the hemolysis was not determined due to no product returned, inconclusive data log review, and insufficient clinical details. the root cause of thrombosis was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
5398 | 1220648-2025-46630 | 23097911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097911&pc=OZD | 08/21/2025 |  |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713202 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 95 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Patient-Device Incompatibility | [5 empty] | Renal Failure | Sepsis | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 for mechanical circulatory support. originally the patient was admitted to an outside hospital with intent to place an impella cp, but was found to have triple vessel disease. two days after the pump was implanted, the patient coded overnight and required two rounds of cardiopulmonary resuscitation. dissection was found to the bilateral lower lobes via computed tomography. creatinine was elevated while urinary output decreased, and the patient was started on continuous veno-venous hemofiltration. the patient was found to be septic and was treated with antibiotics. the patient was placed on veno-arterial extracorporeal membrane oxygenation and then veno-venous extracorporeal membrane oxygenation after completion of a fasciotomy. the ventricular tachycardia (vt) did not respond to lidocaine, amiodarone, nor defibrillation at 360j, resulting in decreasing filling pressure and suction alarms. the patient was treated with albumin, packed red blood cells, and platelets. long vt run and possible suction alarms were noted. monomorphic vt continued the next day, along with a change in the patients status. the family withdrew support and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6 investigation: type, findings, and conclusion codes and h6 component codes were updated accordingly based on the completed investigation. the device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. the pump was not returned for investigation. clinical details indicate that on (b)(6) 2025, there was discussion of balancing flows and possible suction alarms with the left support higher than the right. on (b)(6) 2025, the patient experienced decreasing filling pressures and a suction alarm following unsuccessful defibrillation for ventricular tachycardia. the patient was treated with albumin, prbcs, and platelets. data logs were reviewed and suction alarms were confirmed on (b)(6) 2025 between 10 and 11 am while running on p-level 8, with a decreasing trend in pump flow and no major trend on the placement signal. there was also a non-reported suction instance on (b)(6) 2025. a motor current spike was reported on (b)(6) 2025 without impacting any of the pump's performance, and this instance was not reported. no positioning issue was reported during the case run. as clinical details mention, the patient was on crrt and bivad support, which had some difficulty in balancing flows, with the left support currently higher than the right and possible suction alarms noted. the suction events were most likely related to patient condition and the complexity of managing concurrent bivad support, based on clinical details and data log review. the cause of the tachycardia, sepsis and renal failure were unable to be determined due to insufficient clinical details.
5399 | 1220648-2025-46631 | 23097954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097954&pc=OZD | 08/20/2025 | 09/18/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026615115 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 74 KG | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient admitted in with acute myocardial infarction with cardiogenic shock. this is the second cp and when it was placed there was optical signal loss that they tried to troubleshoot with aic replacement, but when this did not remedy the issue they replaced with a third cp. | [5 empty] | The investigation of the reported failure mode has been completed. optical sensor issue: clinical review reports placement signal not reliable (psnr) alarms during priming of the pump on two separate consoles. it was replaced with a new impella cp that functioned without issue. product was not returned for investigation. since data logs were inconclusive and product wasn't returned for investigation, the cause of the optical sensor issue could not be determined. device history review: the device passed all the post sterile inspection checks. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5400 | 1220648-2025-46632 | 23097960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097960&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692122 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Use of Device Problem | [5 empty] | Nerve Damage | Pain | Embolism/Embolus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted to support a patient with cardiac history and new acute myocardial infarction/cardiogenic shock. suction events occurred and a clot was visualized at the inlet. the thrombus embolized and caused nerve numbness and pain. the pump was explanted and the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5401 | 1220648-2025-46634 | 23097962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23097962&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026698670 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support the patient went into sustained ventricular tachycardia and required defibrillation and several rounds of cardiopulmonary resuscitation. an echocardiogram showed the pump was deep so it was pulled back but was difficult to position. a right arterial distal perfusion catheter and a left femoral arterial sheath were placed to address loss of pulses. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Product complaint #: (b)(4). investigation summary: no product was returned. tachycardia/ ischemia: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device in wrong position (positioning issue): no product was returned. pump was positioned deep but then pulled back resolved under echo. the cause of the positioning issue cannot be determined since the insufficient clinical details were provided. device history lot: device lot: 1946925. device history batch: subcomponent lot: n/a. device history review: device (sn: (b)(6)) passed all post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5402 | 1220648-2025-46633 | 23098052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23098052&pc=OZD | 08/21/2025 |  |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707508 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 152 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had ventricular tachycardia and atrial fibrillation with rapid ventricular rate. the patient underwent coronary artery bypass graft and mitral valve repair. the plan was to remove the impella 5.5 prior to cross clamp and evaluate how the patient did post-operatively after coming off the bypass. the patient did not tolerate coming off the bypass and had higher pressor and worsening lung function. a new impella 5.5 was placed but three days later the patient¿s care was withdrawn and the patient expired. | [5 empty] | No product was returned for investigation. the root cause of the tachycardia/paroxysmal atrial fibrillation was unable to be determined due to insufficient clinical details. the device passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5403 | 1220648-2025-46620 | 23098190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23098190&pc=OZD | 08/21/2025 |  |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707148 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 82 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was placed on impella 5.5 support after presenting with decompensating heart failure and required inotropic support yesterday. the patient had a bloody bowel movement and hemoglobin was low. a unit of packed red blood cells were given and anticoagulation was stopped. the patient received an additional blood transfusion. the patient remained stable and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5404 | 1220648-2025-46635 | 23099750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23099750&pc=OZD | 08/21/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702818 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 121 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an echocardiogram showed the impella cp was measuring over 5.5 centimeters from the annulus. the patient was having more ventricular tachycardia runs overnight. the impella cp was adjusted with echocardiogram. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. device in wrong position: clinical details state that echo showed the impella greater than 5.5 cm from annulus on the second day of support. the cause of the positioning issues was not determined since insufficient clinical details were provided and no product was returned. tachycardia: the cause of the injuries was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5405 | 1220648-2025-46636 | 23099774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23099774&pc=OZD | 08/21/2025 | 10/27/2025 |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026713144 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 72 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support of the impella 5.5, placement signal not reliable with absent ao and left ventricle placement signals appeared. the representative called the account to help with persistent peak signal-to-noise ratio alarm with absent pulmonary stenosis waveforms. the registered nurse reported an order to verify the position. he denied any kinks and confirmed the white plug was secured in the automated impella controller. alarm was reviewed and explained that it does not impact performance of the device. instructions were provided to disable the alarm audio. | [5 empty] | The investigation for the reported optical signal issue has been completed. clinical report missing ao and left ventricle placement signals. the logs show placement signal not reliable alarms and the 5s parameters are of the pattern identified as an oscillating contrast issue. the root cause of the optical signal issue has been characterized as a console-related issue. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5406 | 1220648-2025-31590 | 23099823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23099823&pc=PYX | 10/25/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 10/25/2024 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025551059 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp flex support and during support, placement signal was lost. there was a placement signal unreliable alarm right after hd catheter placed. decision made to withdraw care, patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product returned for investigation. data logs show sudden loss of cvp metrics and drop in signal not reliable (snr) to 0. there is also a mc spike and speed deviation at this time indicating an interaction with the impeller. mc recovers but cvp and snr do not. placement snr alarm and suction are triggered. clinical mentions placement snr alarm after insertion of hd catheter. the cause of the optical signal issue was most likely damaged sensor due to hd catheter interaction. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5407 | 1220648-2025-46637 | 23099884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23099884&pc=OZD | 08/21/2025 |  |  | 09/19/2025 |  |  | 08/21/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026703474 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp and had episodes of ventricular tachycardia requiring cardioversion. a nurse reported the patient had a hematoma that was resolved by the night shift nurse. heparin drop titrated per hospital protocol was noted. later, the patient was successfully weaned from the pump and survived. | [5 empty] | Then investigation has been completed. tachycardia : in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematoma : the cause of the injury was not determined since the product was not returned and insufficient clinical details were provided. device history review: the pump passed all the post sterile inspection checks. | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5408 | 1220648-2025-31591 | 23100096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100096&pc=PYX | 11/12/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 11/12/2024 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025564020 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that post decannulation the patient lost pa waveform. placement signal not reliable alarm occurred. it was reported that the impella still had flows. patient expired on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. log review showed a motor current spike followed by a drop in signal not reliable and papi, cvp, and placement signal going out of range. no dp parameters were affected. ecmo was being decannulated at this time. the root cause of the optical signal issue was most likely optical sensor damage since the motor current spike and affected parameters in the data logs occurred when ecmo was being decannulated. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5409 | 1220648-2025-31592 | 23100124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100124&pc=OZD | 10/05/2023 | 09/19/2025 |  | 09/19/2025 |  |  | 07/27/2023 | 09/19/2025 |  |  | 07/31/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381136 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal for the impella cp does not work. the physician removed the pump for inspection, and there were no anomalies noted. the same pump was used but the signal issue was not resolved. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. data logs were not retrievable. as per clinical placement signal did not work since case start, but pump continued to be in support. the cause of the optical signal issue was not determined since product was not returned and data logs being unable to be reviewed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5410 | 1220648-2025-31593 | 23100142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100142&pc=PYX | 06/27/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 06/27/2024 | 09/19/2025 |  |  | 12/31/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025442197 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella rp placement signal not reliable alarm occurred. x-ray completed, which confirmed position. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for investigation. data logs show that three hours into the case the central venous pressure (cvp) goes out. the placement signal not reliable alarm is triggered. the signal-to-noise ratio drops to 0, cvp goes out of spec, contrasts decreases, lamp increases, and gain and light decrease slightly but are overall stable. there is also a motor current and placement signal spike at this time, indicating something hitting the impeller. the optical signal never recovers, and the case continues for 48 more hours. clinical states that continuous renal replacement therapy (crrt) was being used while on support. the root cause of the optical signal issue was most likely a damaged sensor due to interaction with crrt lines. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5411 | 1220648-2025-31594 | 23100232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100232&pc=OZD | 02/02/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 03/06/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1111982 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp for the support of a patient. the day after the implant the pump alarmed, and the team disabled the placement signal. the patient expired while on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data analysis: the case started on (b)(6) 2025, and within just 40 seconds of activating the pump, several alarms were triggered, including the "impella in ventricle" alarm along with error messages stating, "invalid sample due to signal-to-noise (snr) ratio below minimum" and "algs suspended opm signal invalid." after 18 hours of pump operation, another alarm indicated a "placement signal low," accompanied by multiple notifications of "impella position unknown," "impella in aorta," and "impella in ventricle." notably, the monitoring for the placement signal was disabled after one day of pump operation, which led to no further alarms being triggered. the pump was ultimately stopped after 1 day and 9 hours of continuous use. the lt log and rt log data exhibited low snr values, and the placement signal was also found to be low. however, the amplitude for the placement signal appeared to be satisfactory. further investigating found that the same low snr was seen throughout the next cases. device analysis: the console was returned to field service for further investigation, as the reported complaint could not be reproduced. during the inspection, visible light was observed coming out from the external optical pump connector. additionally, an analog output from the charge-coupled device array of the optical bench did not show any distortions. the voltage supplied to both the optical bench and the bulb was confirmed to be functioning correctly. however, upon measuring the "5s" parameter using the optical cavity, it was found that the snr (signal-to-noise ratio) values were lower than the specified thresholds outlined in the pm procedure. root cause: the root cause for the placement signal being dampened and positioning alarms was due to the defective optical bench with low snr. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5412 | 1220648-2025-31595 | 23100300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100300&pc=OZD | 03/20/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 03/20/2024 | 09/19/2025 |  |  | 09/30/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438864 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp for support of a patient. the cp was placed as care escalation from an intra-aortic balloon pump. the team observed placement signal loss after admission to the intensive care unit from the cardiac catheterization lab. the optical loss did not cause any harm or injury. the pump was not replaced. the pump support ran for a total of 61 days before the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no pump or data logs returned. the cause of the optical signal issue was not determined since product and data logs were not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5413 | 1220648-2025-31596 | 23100324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100324&pc=PYX | 05/30/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 05/30/2025 | 09/19/2025 |  |  | 12/31/2024 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661978 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp flex placed to support the patient who was failing after fistula surgery. the rp lost the optical signal, but the pump remained on until the family made decision to withdraw care. the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the optical signal issue was updated to placement signal issue since product evaluation revealed an issue with the dp sensor. clinical details state that there was a placement signal not reliable (psnr) alarm on the day of implant. they planned to leave in the pump despite the alarms. data log review showed stable metrics until about 8 hours into the case when the placement signal (ps) drifts down slightly before going out of range with the pump flow. optical parameters were stable during this time. psnr alarm 1051 and 1052 were triggered, indicating a dp sensor issue. ps returns briefly shortly after but then is out of specification for the rest of the case. the pump was returned and evaluated. the electrical resistance readings of the dp sensor lines were below spec. there was no biomaterial or damage found on the sensor. there was no evidence of blood in the blue handle, and no holes were found in the catheter. the root cause of the placement signal issue was most likely an electrical short due to the low sensor resistance readings but the cause of the short is unknown. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5414 | 1220648-2025-31597 | 23100394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100394&pc=OZD | 06/25/2024 | 09/19/2025 |  | 09/19/2025 |  |  | 06/25/2024 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512678 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Unknown |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and there was a placement signal issue. the sensor worked at the start of the case, but midway through it stopped working. the pump continued to support the patient. the alarms were disabled. the decision was made to withdraw care, and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the optical sensor stopped working after the pump was removed and reinserted. the product was not returned for investigation. the data logs show a decrease in the signal-to-noise ratio (snr) and confirm loss of placement signal with raw optical signal values railing negative. the data logs from the pumps used before the complaint pump had snr values within normal range. the root cause of the optical signal issue is not determined as no product was returned for investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5415 | 1220648-2025-46638 | 23100487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100487&pc=OZD | 08/20/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026706374 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 130 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Appropriate Term/Code Not Available | [5 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a 5.5 pump placed to support he patient admitted in with a cardiac history and a prior pump explanted 10 days prior. the pump was placed as the team made decision to perform the percutaneous coronary intervention(pci). the pump allowed for high-risk percutaneous coronary intervention and was supporting when it was explanted and replaced on day 6 due to concerns of pump not unloading. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction: e4; and h6 medical device problem code 140508 was inadvertently omitted on the initial manufacturer device report.
5416 | 1220648-2025-31598 | 23100495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100495&pc=OZD | 02/07/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 02/07/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596517 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Death | Yes | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a cp pump placed. the team performed a percutaneous coronary intervention and placed a temporary pacer. the team had the support with the cp ongoing when after they rolled the patient, the placement signal was lost. the pump remained on until the patient's family made choice to withdraw care after a failed wean. the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned for investigation. data log review confirmed the placement signal not reliable (psnr) alarm triggered on (b)(6). at this time, the signal-to-noise ratio (snr) dropped to zero, contrast increased in amplitude, lamp and gain increased, and light dropped. none of the optical signal metrics recover for the remainder of the case. these changes in the optical signal metrics align with a damaged optical fiber and with the clinical narrative provided that the patient rolled over in their sleep at the time of the alarm. based on clinical details provided and data log review, the root cause of the optical signal issue was determined to most likely be a damaged optical fiber. since product was not returned, the location of the damage could not be determined. psnr occurred upon patient movement. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5417 | 1220648-2025-46639 | 23100507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100507&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026633577 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Backflow | Restricted Flow rate | Increase in Pressure | Patient Device Interaction Problem |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo) and was supported by an intro-aortic balloon pump prior to the decision to implant the impella 5.5 for additional support. on the 9th day of support the patient was successfully weaned and decannulated from the va ecmo. on the 36th day of support, it was reported that the purge pressure was increasing, and the purge flow was lower than expected. the purge cassette was exchanged, and the purge line was inspected for kinks. no kinks were observed, and the purge cassette change did not resolve the purge issues. the medical team then elected to add tissue plasminogen activator (tpa) to the purge solution. after the tpa was completed, the purge pressures came down into expected ranges and purge flows increased to expected ranges. additionally, it was recommended several times to the medical team to replace the impella 5.5 as the device use has exceeded 30 days of use. the patient continues to remain on impella 5.5 support awaiting a heart transplant as of the date of writing this report. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was cannulated for venoarterial extracorporeal membrane oxygenation (va ecmo) and was supported by an intro-aortic balloon pump prior to the decision to implant the impella 5.5 for additional support. on the 9th day of support the patient was successfully weaned and decannulated from the va ecmo. on the 36th day of support, it was reported that the purge pressure was increasing, and the purge flow was lower than expected. the purge cassette was exchanged, and the purge line was inspected for kinks. no kinks were observed, and the purge cassette change did not resolve the purge issues. the medical team then elected to add tissue plasminogen activator (tpa) to the purge solution. after the tpa was completed, the purge pressures came down into expected ranges and purge flows increased to expected ranges. additionally, it was recommended several times to the medical team to replace the impella 5.5 as the device use has exceeded 30 days of use. the patient continues to remain on impella 5.5 support awaiting a heart transplant as of the date of writing this report. new information was reported noted that a thrombus was noted to be forming around the impella cannula at the aortic valve and that the medical team was handling the patient¿s hemolysis. the medical team was resistant to exchange the pump, which had exceeded 30 days of use, as the patient did not have an exit strategy. the medical team considered increasing the patient¿s anticoagulation therapy or potentially adding tissue plasminogen activator in the purge solution, but neither were expected to resolve the complication. additionally, there were retrograde flow alarms and low flow values observed at performance level 7. the patient¿s lactate dehydrogenase had been increasing over 24 hours, and urine production was decreased. the patient was noted to be very lethargic, with an ejection fraction of 10%. rather than risk a catastrophic event, the decision was made to place the patient in comfort care and withdraw impella support. the patient expired after support was discontinued. further information noted treatment was withdrawn as heart failure deteriorated "significantly" and there was no exit strategy (patient was ineligible for ventricular assist device or heart transplant). |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Fields b1 and b2: corrected reportability type. new information was reported that the patient expired following support discontinuation. the exact date of death is unknown but has been captured as the date of explant. if specific information regarding the date of death becomes known, a supplemental report will be submitted accordingly. field b5 is being updated to include new information that was reported to the manufacture and was not included in the initial report. the revised b5 provides a complete event narrative. d6b. explant date was populated based on additional information. field h1: type of reportable event corrected to death. new information was reported that the patient expired following support discontinuation. field h6: complaint coding was updated to better describe the reported event based on additional information. consequently, h6 health effect: clinical/impact and medical device problem coding were repopulated/updated, accordingly. investigation: the device was not returned as per the customer it was no longer available. therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. purge pressure high: clinical reported high purge pressure which resolved when tissue plasminogen activator (tpa) protocol was initiated. no product or logs were returned. the cause of the issue was not established as no product or logs were returned. bleed and hemodynamic instability: the cause for these events was not determined due to insufficient clinical details. mechanical interaction with blood/hemolysis: clinical reported hemolysis; however, no product or logs were returned. the cause of the issue was not established as no product or logs were returned. regarding thrombus, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. and retgrograde flow/low pump flow: clinical reported intermittent retrograde flow alarms and low values seen at p-7. the cause of the issue was not established as no product or logs were returned. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation.
5418 | 1220648-2025-46640 | 23100549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100549&pc=OZD | 08/20/2025 | 11/06/2025 |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124473 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Network Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) failed to boot up successfully after patient support concluded. it was reported after a patient was removed from support, the medical team attempted to reboot the aic to connect to impella connect services. upon reboot, it was observed that the amber airplane light and the blue cloud light were blinking simultaneously and the boot up sequence would not complete. the aic was power cycled again, it was attempted to be put into flight mode, and factory reset with no change to the startup failure. there was no patient involvement. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: product ic2110 was returned to field service for investigation. analysis: rl00252 was unable to complete boot up due to partition switching as each boot up would fail in one partition and then attempt to load into the other partition unsuccessfully reproducing the complaint. replacing the usd card with a known-good allowed the remote lan module (rlm) to boot up as expected. root cause: the rlm was unable to boot up due to a corrupted usd card but the cause of the corruption could not be determined. device history batch: subcomponent name: usd card. material number: 2000-0017. lot number: 2024286437. summary: there were no issues related to this complaint discovered when reviewing the product history of console ic2110.
5419 | 1220648-2025-46641 | 23100956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100956&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026707153 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction (stemi)/cardiogenic shock was being implanted with an impella 5.5 device for mechanical circulatory support. the patient presented to the hospital with complaints of dyspnea, orthopnea, and edema, was found to be hypotensive and hypoxic, and required rapid sequence intubation and intravenous pressor therapy. an electrocardiogram revealed a normal sinus rhythm that was influenced by a pacemaker. a computed tomography scan revealed pulmonary edema but was negative for pulmonary embolism. the patient was determined to be in stemi, and the patient was taken emergently to the cardiac catheterization laboratory for right and left heart catheterization. after a percutaneous coronary intervention was performed to place three stents, the patient began to experience cardiogenic shock, and the decision was made to pace an intra-aortic balloon pump (iabp) through the right femoral artery. it was noted that the patient had poor prefusion in the right lower extremity of the iabp site and the patient began showing signs of cardiogenic shock again. the decision was made to surgically place an impella 5.5 for increased support in the right axillary artery. the left ventricle was successfully accessed with a guidewire, however during delivery of impella 5.5 difficulties were met when attempting to access the left ventricle. a lunderquist wire was used as a buddy wire without success. viperslide and positioning of shoulder were attempted without success. a contrast injection revealed stenosis of axillary artery so a mustang 6.0x8 balloon dilatation catheter was used to dilate the axillary with appropriate axillary expansion. left ventricle access was again obtained with a guidewire and the impella 5.5 was again attempted to be delivered but was unsuccessful. the impella 5.5 was removed, and the iabp support was reinitiated to evaluate hemodynamics. the decision was then made to place an impella cp through the axillary graft placed for the impella 5.5. the impella cp was placed without difficulties and impella support was initiated with iabp support. the decision was then made to remove the iabp and monitor the patient. on the second day of impella cp support, the patient¿s lactate dehydrogenase and alanine aminotransferase blood levels were noted to be elevated, and the patient¿s urine was pink-tinged. there was noted to be no suction issues with impella, and a plan was made to explant the cp later that day. the following day the impella cp was explanted, and the patient was noted to have survived at the outcome of impella support, however overall patient outcome is unknown. this complaint involves two impella devices: pump 1: impella 5.5 with serial number (b)(6). pump 2: impella cp with serial number (b)(6). this mdr specifically addresses impella 5.5 with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary unable to deliver or advance : the cause of the deliver issue was determined to be patient condition as the patient had stenosis in the axillary artery preventing successful delivery of impella. device history lot: device lot : 1945135 device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5420 | 1220648-2025-46569 | 23100960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100960&pc=OZD | 08/16/2025 |  |  | 09/19/2025 |  |  | 08/16/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628478 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that suction events occurred during impella cp support so the p level was adjusted to p4. oozing was present at the access site and the patient had fever overnight. the pump was explanted. no blood loss was observed from the access site during explant and the patient lost pulses to the leg. the patient was taken to surgery for an embolectomy. | [5 empty] | No product was returned for investigation. data logs are not available. the cause of the suction was unable to be determined since product/logs were not returned and insufficient clinical details were provided. the cause of the access site bleeding was unable to be determined since insufficient clinical details were provided. the cause of the ischemia was unable to be determined due to insufficient clinical details. the cause of the fever cannot be determined as insufficient clinical details were provided. the device passed all post-sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5421 | 1220648-2025-46642 | 23100972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23100972&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635338 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Backflow | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Cardiac Arrest | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hematuria | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and during the 25+ day support on the impella 5.5 pump there were harms observed. the patient was admitted to the hospital on (b)(6) 2025 in anterior wall st elevation myocardial infarction, received two stents placed into coronary arteries, and was discharged a couple of days later. the patient reported having ongoing chest pains and shortness of breath and returned to the emergency room on (b)(6) 2025. the patient subsequently had 3 episodes of cardiac arrest and was brought emergently to the cardiac catheterization laboratory where an angiogram revealed patent stents. the advanced heart failure medical team recommended impella cp placement. right heart catheterization was performed, and right ventricular dysfunction was observed, and the patient¿s mixed venous oxygenation was 45%. the medical team then decided to implant an impella cp and canulate the patient with extracorporeal membrane oxygenation (ecmo) therapy. the patient remained on emco and impella support for 4 days, and the medical team decided to escalate the patient to an impella 5.5 with the intention of weaning the patient from ecmo. the patient was brought to the operating room for impella 5.5 implant, impella cp explant and possible emco wean. there were no issues reported with the implant of impella 5.5 and explant of impella cp, however during the wean of ecmo, the patient¿s pressures would drop despite the p level of the impella being at p8. the decision was then made to have the patient remain on emco with impella 5.5 support and attempt to wean ecmo after more rest and recovery. on the second day of support, it was reported that there was a retrograde flow alarm to the impella support center. the support team provided options to the medical team, however there is no information provided on any actions the medical team took to resolve the issue. on the 6th day of impella 5.5 support, it was reported that the patient had pink-tinged urine the previous day, and that the lactate dehydrogenase blood result increased from 1100 u/l to 1783 u/l. an echocardiogram performed (b)(6) 2025 was reviewed and it was determined that the impella was close to the septum. a plan was made to reimage the impella placement and reposition if needed. a plan was made to wean and decannulate ecmo later that evening. the following day it was reported that the patient experienced pulseless electrical activity arrest during the attempt of weaning ecmo. ecmo was recannulated and the patient was cooled, then rewarmed overnight. the patient¿s kidney function decreased, and the patient was placed on continuous renal replacement therapy. a formal echocardiogram performed on (b)(6) 2025 revealed the impella to be in good position. on the 15th day of impella support, the patient experienced bleeding of the liver and spleen. heparin was discontinued and the patient was transfused with a unit of packed red blood cells as a result of the bleeding. it was also noted that there were ischemic areas of the patient¿s brain and upper vasculature. on the 19th day of support, the patient was successfully weaned and decannulated from ecmo. heparin was restarted, and the patient¿s plan was to continue to rest and recover. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complications. | [5 empty] | Product complaint # (b)(4). investigation summary: clinical symptoms (cardiac arrest, ischemia):in order to make a cause determination, all of the clinical details outlined in (b)(6) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. clinical symptoms (device repositioning, hemorrhage/bleeding): the cause of the injury was not determined due to insufficient clinical details. mechanical interaction with blood(hemolysis, hematuria): the pump was not returned for investigation. the data log was only available up to 16-aug, while hemolysis was reported on 20-aug. the cause of the hemolysis issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. retrograde pump flow: there was some motor speed and motor current ramp-up noted during the time of the reported retrograde flow notification while running on p2. the cause of the retrograde pump flow issue could not be determined without the returned product and sufficient clinical information. device history lot: device batch: 1900755. device history batch: subcomponent lot: na. device history review: the pump sn passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5422 | 1220648-2025-46643 | 23101077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23101077&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | No Clinical Signs, Symptoms or Conditions | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that patients treated with impella pumps experienced ischemia and strokes during support. the medical staff reported that they had two or three strokes and three ischemic arms from axillary impellas this year. two of the strokes occurred on patients with heartmate 3 devices, and one was a catastrophic stroke. the complainant¿s laboratory found the partial thromboplastin time and anti xa results to be sub-therapeutic, and this has been corrected. patient information was not provided. no further information was provided this complaint involves three patient/device interactions. separate mdrs will be submitted for each of the events associated with this complaint. | [5 empty] | Actual date of events were not reported and therefore date of awareness has been documented as event date. patient information was not provided. therefore, section a was left blank. section d4: device information was not provided. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint #(b)(4).
5423 | 1220648-2025-46644 | 23101100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23101100&pc=OZD | 08/20/2025 |  |  | 09/19/2025 |  |  | 08/20/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | IMPELLA |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had ischemia and a stroke during support. the medical staff reported that they had 2 or 3 strokes and 3 ischemic arms from axillary impella's this year. apparently banner tucson's lab stated that their lab had shown partial thromboplastin time and anti xa's being sub-therapeutic. this has been corrected. they stated that 2 of the strokes were also hm3 patients and one was a catastrophic stroke. the account expressed their hesitation to use impella in recent months because of this. there are 3 suspected patients having had stroke and ischemia, however, the specific cases linked to them is unknown. | [5 empty] | Information was received that found medwatch report 1220648-2025-46644 was a duplication of medwatch report 1220648-2025-30314. all information regarding this patient will be reported under medwatch report 1220648-2025-30314. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. patient information unknown.
5424 | 1220648-2025-31599 | 23101890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23101890&pc=PYX | 07/25/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 07/25/2025 | 09/19/2025 |  |  | 10/31/2024 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274765 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiac Arrest | Unspecified Vascular Problem | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with cardiogenic shock and heart failure. while on support, there was a placement signal malfunction. cardiopulmonary resuscitation was performed prior to the signal issue. there was no further information provided. the patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. data logs show suction conditions throughout the case with suction alarms and negative cvp. there are placement signal not reliable (psnr) op out of range alarms when cvp passes the threshold of -100. snr, contrast, ps, and mc follow the same trends throughout the case. gain and light are stable and within normal ranges. flows also show instances of potential lack of volume. patient was likely deteriorating as cpr had to be performed. the cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5425 | 1220648-2025-31600 | 23102599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23102599&pc=PYX | 07/25/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 07/25/2025 | 09/19/2025 |  |  | 07/31/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024385433 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient presented with heart failure for circulatory support. during support the pump triggered intermittent 'placement signal not reliable' alarms. at the time of the noted call, the displayed placement signal was within normal limits and positioning was confirmed via chest x-ray. the staff was advised that the signal was not required for support and the alarm was silenced. the family made the decision to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was not returned for investigation. data logs were investigated. when the pump is started, the cvp/ optical sensor readings drops and goes into negative. the snr dropped at the same time. this suggests that the pump's strut (optical sensor area) was against the patient's anatomy which creates a local suction. this cannot be confirmed sine the product has not been returned for investigation. therefore, the root cause of the optical signal issue was not determined since product was not returned to conclude anything. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5426 | 1220648-2025-31601 | 23102795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23102795&pc=PYX | 07/25/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 07/25/2025 | 09/19/2025 |  |  |  |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025560596 |  | 2000341 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. it was reported that during support; there were placement signal not reliable alarms. the alarms were disabled and support continued. no further information was provided. the family of the patient decided to withdraw care and the patient expired. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for analysis. no data logs were provided for review. the root cause of the optical signal issue was not determined as no product was returned for analysis and no data logs were provided for review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5427 | 1220648-2025-31602 | 23102887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23102887&pc=OZD | 07/25/2025 | 09/19/2025 |  | 09/19/2025 |  |  | 07/25/2025 | 09/19/2025 |  |  | 09/30/2025 |  | 09/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399246 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient needing mitral valve reconstruction. the complainant reported the aortic placement signal curve and the impeller flow failed; however, the pump continued to run reliably. the patient had to go into surgery two hours later, but it is not known whether the pump was replaced or pulled completely. no further information was provided. the patient expired while on support. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. no data logs were returned for evaluation. clinical details noted that both the ao placement signal and impeller flow graph failed, but pump continued to run reliably. the root cause of the optical signal issue cannot be determined as there is limited clinical details and no product or data logs to review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5428 | 1220648-2025-46560 | 23082576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23082576&pc=OZD | 08/08/2025 |  |  | 09/18/2025 |  |  | 08/08/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676843 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 42 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered sluggish flow down the distal extremity during insertion of an impella cp. as a result, the team elected to place an anti-grade/retrograde fem to fem bypass. later, the patient was successfully weaned from the impella cp and brought to the operating room for planned impella 5.5 escalation. this report represents the impella cp. a separate report was submitted to represent an issue that occurred with the impella 5.5. refer to section h.10 for the related report number. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the ischemia was unable to be determined due to insufficient clinical details. the pump passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5429 | 1220648-2025-46559 | 23082804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23082804&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026710672 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 42 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had previous been on impella cp support and was escalated to an impella 5.5 for support. while on impella 5.5 support, the patient developed hematuria. it was noted that the patient's pink urine had cleared up after giving the patient one unit of albumin and one unit of blood. no further issues were noted. this report represent the impella 5.5. a separate report was submitted for the impella cp. refer to section h.10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. hematuria: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5430 | 1220648-2025-46562 | 23082941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23082941&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026708824 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed to support a percutaneous coronary intervention procedure. after using shockwave, the placement signal disappeared. there was no harm to the patient. | [5 empty] | The investigation of the reported failure mode has been completed. the pump and purge cassette were returned. the data logs show that the 5s parameters and placement signal (ps) is stable until the ps spikes to 3000 and signal not reliable (snr) and contrast drop. it returns and then does this again multiple times until at the end of the case the ps is completely out. the returned product revealed that the sensor membrane appeared damaged with the 5s parameters reflective of a sensor break. the pump passed laser continuity. based on the clinical details, data log analysis and returned product, the root cause of the optical signal issue is most likely due to a damaged sensor due to shockwave. the distance between impella and shockwave is unknown. device history review: the device passed all the post sterile inspection checks. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5431 | 1220648-2025-46564 | 23083054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23083054&pc=OZD | 08/20/2025 |  |  | 09/18/2025 |  |  | 08/20/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1763601 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 64 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | False Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that as soon as the automated impella controller (aic) was turned on, an ¿impella defective¿ alarm was seen. the alarm resolved as soon as it appeared. after the alarm resolved it was connected to a pump and no issues were noted to have occurred throughout the procedure. the procedure was successful, and the patient was weaned from the pump and survived. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5432 | 1220648-2025-46566 | 23083481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23083481&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696988 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 117 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Inaccurate Flow Rate | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 generated concerns for flow saturation given increased left ventricle signal pressures and dampened motor current, as well as unexpectedly high flow on p-7. there was also a concern of hemolysis as lactate dehydrogenase increased from the 400's to 900's. the impella 5.5 was exchanged for a new one with improved waveforms and appropriate flows. this report represents the first impella pump. another report was submitted to represent the issues that occurred on the second impella pump. refer to section h.10 for the related report number. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The impella cp was returned for evaluation. data logs were reviewed and flows were confirmed to be saturated for the majority of the case. there was no damage noted to the impeller or cages of the returned pump. there was a small amount of biomaterial around the impeller, which remained after soaking the pump. there were no signs of biomaterial in the purge gap. since the pump was returned cut, the pump was unable to be run to determine if the complications reproduced. based on the limited clinical details, inconclusive data logs, and incomplete returned product, the cause of the hemolysis could not be determined. since data logs were inconclusive and product was returned incomplete, the cause of the saturated flows could not be determined. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The device was returned for investigation. data logs were reviewed and the complaint condition was confirmed. there was no damage noted to the impeller or cages of the returned pump. there was a small amount of biomaterial around the impeller, which remained after soaking the pump. there were no signs of biomaterial in the purge gap. since the pump was returned cut, the pump was unable to be run to determine if the complications reproduced. based on the limited clinical details, inconclusive data logs, and incomplete returned product, the cause of the hemolysis could not be determined. since data logs were inconclusive and the product was returned incomplete, the cause of the saturated flows could not be determined. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the device passed all post-sterile inspection checks.
5433 | 1220648-2025-46567 | 23083694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23083694&pc=OZD | 08/23/2025 |  |  | 09/18/2025 |  |  | 08/23/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696989 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 117 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with existing protek duo, right ventricular assist device with an oxygenator had an impella 5.5 placed. on day two of impella support, there was concern for flow saturation given increased left ventricle signal pressures and dampened motor current, as well as unexpectedly high flow on p-7. there was also a concern of hemolysis as lactate dehydrogenase increased from the 400's to 900's. the impella 5.5 was exchanged for a new one with improved waveforms and appropriate flows. on day four of impella 5.5 support, the patient was septic and cultures were positive. intervention is unknown. on day eight of impella 5.5 support, ventricular tachycardia (vt) occurred with one shock. formal echocardiogram showed the impella measuring deep at 7.3 centimeters (cm). the device pulled back 2 cm and the vt resolved. it was further reported that the family decided to withdraw care and the patient expired. this report represents the second impella pump. another report was submitted to represent the issues that occurred on the first impella pump. refer to section h.10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. no product was returned. tachycardia, sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. positioning issue: no product was returned. echo showed impella measuring deep at 7.3cm. the cause of the positioning issue was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
5434 | 1220648-2025-31541 | 23083701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23083701&pc=OZD | 04/23/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 04/23/2024 | 09/18/2025 |  |  | 12/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440353 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a discrepancy between the placement signal and the invasive ao pressure. this occurred day 1 of impella cp placement. it appears that due to the inaccurate placement signal, when attempting to wean pressors, the subsequent drop in the patient blood pressure triggers a ¿placement signal low¿ alarm. upon confirming correct placement via echocardiogram, customer turned off the alert. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation. no physical abnormalities were observed on the returned pump. the pump passed the laser continuity test, and all 5s optical parameters were found to be within specification. further investigation of the optical sensor pressure test revealed no issues with the sensor. data logs show the placement signal trending down with low alarm throughout the case run without noting any positioning-related alarms. placement signal low alarms were observed on both consoles ic10253 and ic9778 used with no major impact to optical 5s parameters. on second console ic9778 signal not reliable was observed to be low range below 1500 throughout the case. as per the given clinical data, there is a discrepancy between the placement signal and invasive ao pressure. the cause of the optical signal issue was patient condition related, based on returned product and log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5435 | 1220648-2025-46568 | 23083859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23083859&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  | 08/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384764 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge pressure of an implanted impella 5.5 dropped while the impella flow increased. the purge flow rated remained consistent. no patient harm was reported. | [5 empty] | Only event logs were returned for evaluation. with no device and incomplete event logs, no root cause could be determined. the pump passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5436 | 1220648-2025-31542 | 23083950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23083950&pc=OZD | 07/14/2025 | 11/25/2025 |  | 09/18/2025 |  |  | 09/09/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155470 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) issue with the placement signal. the impella cp was used in the procedure with no issues with insertion. impella was turned on in auto and within a few minutes "placement signal not reliable" alarm was presented. flows were kept within expected range, patient remained stable throughout and impella was successfully weaned and explanted without issue. doctor was aware that the impella was still flowing, the waveforms were just unreliable. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation has been completed. data analysis: imclog & ltlog data confirmed the reported failure mode given there was a placement signal not reliable alarm (#1036, opm invalid signal) triggered during the clinical procedure on 14/07/25 . real time (rt) logs indicated not only invalid values (-32768) of placement signal, but also erroneous values of the raw optical data presenting as upward spikes by approx. 200 mmhg, or 350 mmhg respectively to average signal level (approx. 70mmhg), most likely result of incorrect peak detection on the fringe pattern across ccd detector, due to low snr and/or noisy signal. device analysis: console (ic3034) was sent back to fs for further investigation, and 5s test results indicated that optical bench had contrast below spec. measurement of the analog output of ccd detector revealed localized signal distortion which differently affected optical parameters (snr, contrast) depending on cavity length. during console inspection some minor but persistent contamination of optical pump connector had been observed (unlikely to have contributed to this complaint). further log analysis and device testing revealed corruption of microsd card in the rlm (unrelated to complaint). root cause: the root cause for placement signal not reliable alarms was caused by failure of optical bench. d10 concomitant medical products and therapy dates updated.
5437 | 1220648-2025-31543 | 23084448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23084448&pc=OZD | 04/21/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 07/30/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1505149 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, the pump experienced a placement signal not reliable alarm. the investigating engineer attributed the issue to the automated impella controller (aic). no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs confirms that all signals received from the optical bench were confirmed. console was returned for investigation. the device issues could not be reproduced. during the inspection, visible light was observed coming out from the external optical pump connector. additionally, measured "5s" parameter using the optical cavity were within the specification and the analog output from the ccd array of the optical bench did not show any distortion. the root cause for the placement signal not reliable alarm was most likely due to the intermittent failure of the optical bench (defective lamp or electronic components). a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5438 | 1220648-2025-31544 | 23084608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23084608&pc=OZD | 06/09/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 06/09/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 | [5 empty] | US |  | 2025578556 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 110 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an impella 5.5 support and during support, there were placement signal not reliable alarms. no further information has been provided. support continued. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. the data logs show that the placement signal (ps) was relatively stable until about 80 hours into support. during this time though, there was some variability in signal-to-noise ratio (snr). at the 80 hour mark, the ps increases to unrealistic values and then returns to values within normal expectations for the ps while the snr drops to 0, contrast drops but gain, light, and lamp are stable. this continues for about an 1 hour and 43 minutes and triggers placement signal not reliable (psnr) alarms. the placement signal returns to being stable and then there are other increases seen while on support on the automated impella controller, but not sustained or high enough to trigger psnr. the pump is transferred to another console on 6/4. throughout support there are multiple instances of the same repeated issue, except impella position unknown alarms are also being triggered. the alarm was accurately triggered due to the increases in placement signal but was a false alarm as positioning was confirmed. due to lack of clinical information and no product returned, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5439 | 1220648-2025-46571 | 23084814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23084814&pc=PYX | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025531493 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received that the impella high purge pressure that was believed to be a kink was resolved by straightened the tubing. additionally, it was reported that after the impella was explanted the new hemostatic suture that was placed was not due to bleeding. it was following best practices for impella site closure. | The complainant reported that an impella rp flex had a brief high purge pressure alarm reviewed via impella connect. it appeared most likely to be a kink. it was advised that if it happened again and the purge flow remained below two milliliters per hour then a change in purge solution may be warranted. it was further reported that the impella rp flex was explanted as planned and after explant, a new hemostatic suture was placed followed by manual compression until adequate hemostasis was achieved. the patient survived. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for evaluation. high purge pressure: clinical details noted a brief high purge pressure alarm that clinical team thought was a kink. data logs were reviewed and the logs showed a sudden purge pressure spike with drop in purge flow that resolved within 5 minutes. purge flow ticks were also stalled at that time. the cause of the high purge pressure was due to a kink per clinical details and data log analysis. product damage (purge tubing kink): clinical details noted that pump had a brief high purge pressure alarm occurred and clinical team suspected a kink. no additional clinical details were provided. data logs were reviewed and showed a sudden purge pressure spike that resolved within 5 minutes consistent with a kink. the cause of the product damage (purge tubing kink) was due to a purge tubing kink based on returned data logs. device history review: the device passed all the post sterile inspection checks.
5440 | 1220648-2025-31545 | 23084885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23084885&pc=OZD | 05/26/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 05/26/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653604 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 89 KG |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support, and during support there were placement signal alarms with absent placement signal. staff monitored and support continued. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. clinical details reported placement signal not reliable (psnr) alarms with absent placement signal on (b)(6) 2025. data log review shows a rise in purge pressure (pp) past 300 mmhg with intermittent spiking (ps going in and out of range) on (b)(6) 2025, which triggered psnr alarms. contrast, gain, and signal-to-noise ratio (snr) are variable during this time with no true pattern. the variable ps seemed to resolve within a day. there was another instance of psnr alarms about 11 days later, but then the ps issues seemed to resolve for the rest of the case. product was not returned. there is a lack of clinical details around the time of the psnr alarms, and it is unclear if there were other devices being used or specific clinical events happening. there is no evidence to suggest that the console is the cause of the issue. the root cause of the optical signal issue was not determined since insufficient clinical details were provided and product was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5441 | 1220648-2025-31546 | 23084956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23084956&pc=OZD | 05/07/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 05/07/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624660 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support, and, during support, all torque was removed from the cable and started at p2, increasing to p6. placement signal not reliable alarms occurred, left ventricle was reading 3000/10, and aortic placement signal off the screen. all other waveforms and flow were within normal limits. whilst reviewing impella connect, it appears it may not have zeroed appropriately without staff¿s knowledge. it appears however, that the white plug was not fully seated into the automated impella controller, and, after pushing it in, waveforms normalized. no patient harm or injury occurred. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. according to the clinical notes, the placement signal waveform normalized after pushing the white plug. data logs show that the 5s optical parameters and placement signal were as expected before insertion and pump start. there was a first impella pump plug disconnect before the pump was started, and it ran for about one minute without issue, followed by a loss in signal not reliable alarm and placement signal alarm after the pump was turned to p0. a second impella plug connection notification was issued. the pump started again, and the placement signal issue persisted for about 20 minutes before resolving for the remainder of the case run. the imc log was reviewed, and the following observations were made: the calibrated sensor value was off by 21 during the initial and first impella disconnect. the optical sensor calibration values matched the factory calibration during the third impella connection. the cause of the optical signal failure was most likely air gap from between the automated impella controller and the patient cable plug based on given clinical details and data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5442 | 1220648-2025-31547 | 23084987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23084987&pc=OZD | 05/07/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 05/07/2024 | 09/18/2025 |  |  | 01/31/2026 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025457823 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and there were placement signal issues. for an unknown reason, the placement signal waveforms and hemodynamics then went blank around the same time that shockwave was initiated. shortly after, placement signal not reliable alarm occurred. patient hemodynamics were monitored and impella position was verified with imaging. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. data logs were reviewed and rt logs from case show 5s parameters change right before "placement signal not reliable" alarms is triggered consistent with damage to the sensor face. some 5s parameters were able to recover slightly before final "placement signal not reliable" alarm was triggered and again 5s parameters show change consistent with sensor face damage. motor current and pump flow were not affected. clinical details noted that placement signal waveforms and hemodynamic went blank on the automated impella controller (aic) at same time that shockwave device was initiated. impella position confirmed to be in correct position with imaging and no kinks in the catheter were noted. the root cause of the optical signal issue was due to sensor damage as a result of shockwave use. a trend chart was pulled from (b)(6) 2023 to (b)(6) 2024 for all cp optical pumps with optical signal issues with root cause damaged sensor and no trend was observed. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5443 | 1220648-2025-31548 | 23085011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085011&pc=OZD | 04/10/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 04/10/2024 | 09/18/2025 |  |  | 12/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448757 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. after insertion, the placement signal was lost. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. the logs from beginning of case show a sudden drop in signal not reliable and a spike in placement signal to 3000 for the rest of the case. clinical details noted that motor current was unaffected after placement signal was lost. the root cause of optical signal issue is most likely due to a broken fiber. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5444 | 1220648-2025-46572 | 23085015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085015&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026625843 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 encountered bleeding during support. additional information indicated that the bleeding was a mild intracerebral hemorrhage. the bleeding was treated conservatively with close monitoring of platelet counts and anticoagulation. the bleeding may be related to impella, but there is no direct evidence. the team also switched to bicarb purge fluid. the patient had history of ischemic insult in the brain and there were some pathological vessels visible on ct which was a probable risk factor. impella support continues. | [5 empty] | A4, a5, and a6 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D.6b date of explant was added. no product was returned for investigation. investigation summary: no product was returned for investigation. major bleed: the bleeding was bleeding was a mild intracerebral hemorrhage. there are no issues with the pump; motor current is good, activated clotting time/ aptv in range. the root cause of the bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all post sterile inspection checks.
5445 | 1220648-2025-46573 | 23085036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085036&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026696998 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 88 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during the support the anticoagulation was held due to a gastrointestinal bleed and there was persistent mispositioning of the pump and ventricle tachycardia. the pump was repeatedly repositioned and blood given for treatment. after 18 days care was withdrawn and patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5446 | 1220648-2025-46574 | 23085046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085046&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026709835 |  | 000000000001000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was tried to be implanted with an impella 5.5 device for mechanical circulatory support. the native anatomy precluded the placement of the 5.5 and so the team made decision to place a cp. no patient harm was reported due to the issue. | [5 empty] | Correction: g1 manufacturing site name and address. the investigation of the reported failure mode has been completed. unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had an anatomical narrowing directly on the anastomosis, preventing successful delivery of impella. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5447 | 1220648-2025-31549 | 23085060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085060&pc=OZD | 04/26/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 04/26/2024 | 09/18/2025 |  |  | 11/30/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal was lost after shockwave therapy. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and 5s parameters are consistent with a broken sensor face. motor current remained stable throughout case. the root cause of the optical signal issue was due to a broken sensor from shockwave use. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5448 | 1220648-2025-31550 | 23085197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085197&pc=OZD | 05/09/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 05/09/2024 | 09/18/2025 |  |  | 01/31/2026 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454461 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, the pump experienced a placement signal alarm occurred. the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs show the temporary placement signal was not reliable for about 12 hours, with the air gap trend on the 5s optical parameter. the root cause of the optical signal failure was most likely an air gap from between the console and the patient cable plug or within the middle of the red handle ferrule. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5449 | 1220648-2025-31551 | 23085292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085292&pc=OZD | 06/05/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 06/05/2024 | 09/18/2025 |  |  | 02/28/2026 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025479019 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal issues were noted. the pump¿s position was verified via echocardiography. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product or data logs were returned. the root cause of the optical signal issue was not determined as no product or data logs were returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5450 | 1220648-2025-46576 | 23085317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085317&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2026672308 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp failed to deliver due to the patient's anatomic limitations. the pump was replaced with another impella cp. | [5 empty] | B.7 added information that was inadvertently omitted from the initial report that was submitted. investigation summary: unable to deliver or advance: the cause of the deliver issue was determined to be patient condition as the patient had an anatomical narrowing directly on the anastomosis preventing successful delivery of impella. device history review: the pump passed all the post sterile inspection checks. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5451 | 1220648-2025-31552 | 23085328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085328&pc=OZD | 07/25/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 07/25/2024 | 09/18/2025 |  |  | 02/28/2026 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025466195 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal issues were observed. the audio alarms were silenced, and the pump¿s position was verified via echocardiography. despite these issues, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. the pump was not returned for investigation. the data logs show that the placement signal is not reliable with values of 3000, which depends on the optical signal parameters showing trend of fiber damage on all 5s parameters. the reported failure mode of the placement signal issue was changed to an optical signal issue based on the clinical and data log review. the cause of the optical signal issue was most likely damaged optical fiber line, based on data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5452 | 1220648-2025-31553 | 23085414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085414&pc=OZD | 08/02/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 08/02/2024 | 09/18/2025 |  |  | 08/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385374 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal unreliable were observed. despite this issue, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. data logs were investigated, and the alarm was confirmed. returned product was investigated and the 5s parameters were out of specification. visual inspection revealed that there was a kink in the catheter roughly 1 inch away from the kink protector. the pump failed the laser continuity test and the light leak was found to be at the mentioned kink. the failure mode was changed from placement signal issue to optical signal issue because the issue was with the optical sensor, not the dp sensor, as it is a cp pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5453 | 1220648-2025-31554 | 23085468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085468&pc=OZD | 07/31/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 08/05/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 180730 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. while on support, the automated impella controller (aic) displayed a ¿placement signal not reliable¿ alarm, and no light was visible on the aic. an aic replacement was not performed during treatment due to the impella connect connection. the patient was not harmed, and there were no issues with pump function. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. data logs were investigated, and are consistent with complaint and optical placement signal suddenly became unreliable, and placement signal not reliable alarm presented. console was tested and the issue was not reproduced - there was visible light from optical pump connector in the console, and ¿5s¿ parameter testing showed that optical system was within specification. despite not being able to reproduce, data analysis indicated pattern consistent with intermittent failure of the optical bench due to either defect of lamp or electronic components. the root cause of the automated impella controller issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5454 | 1220648-2025-31555 | 23085512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085512&pc=OZD | 08/12/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 08/12/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525256 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal was unreliable. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. the failure mode was changed to optical signal issue because placement signal issue is used for issues on pumps with a differential pressure sensor, which the cp does not have. the pump was also run in the lab and the device functioned to specification. the data logs do not show any abnormalities and the 5s parameters from the field use were stable. the data logs of the 2 cases prior to complaint case, using the same console, were checked and there were no similar issues. the root cause of the optical signal issue was most likely patient condition because the issue resolved with no troubleshooting and the returned pump functioned to specification. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5455 | 1220648-2025-46575 | 23085550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085550&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704087 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hematoma | Ischemia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in with cardiac history and a code in the field. the medical doctor obtained access to the right femoral artery, prepared and placed the impella cp in normal fashion. it was reported that there was leaks from the purge disk during set up and had to replace the purge cassette. no issues with priming of the impella afterwards. furthermore, a hematoma was noted at the site and femstop was placed above the impella. the hematoma was managed. distal pulses were not dopplerable to either foot and the patient is going to be monitored in the intensive care unit. levophed was running, and heparin elected to be paused for now and bicarbonate was used in the purge. it was also reported that albumin was given due to suction alarms after the impella cp was dropped from p8 to p6. the patient's urine was noted to be tea colored. bruising was noted but no bleeding or hematoma was reported. pulses were doppler. levophed and amiodarone were weaned off. dobutamine remained. the patient received blood products. later, the patient was weaned and survived. additional information was received. the team later concluded that most of the hematoma was caused from the venous sheath placed on that side in the left ventricular failure. it was noted to be unrelated to the impella repositioning sheath. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H6: added codes based on the results of the investigation. investigation summary no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed suction alarms. purge system blocked and high purge pressure alarms were seen, with an isolated instance of purge system blocked alarm which resolved within a minute, and persistent alarms on the same day. purge pressure spikes with drop in purge flows were seen during these events, for the same p-level. no motor current spikes or trends were seen during this time. a placement signal alarm was seen at the beginning of the case, but resolved within a minute. an impella position in ventricle alarm coincided with the placement signal low alarm an resolved within the same timeframe. placement signal and motor current showed no spikes or trends outside of p-level changes. pump suction: clinical details mention albumin being administered to the patient for blood volume during impella support, with suction alarms resolving, as confirmed by data log analysis. the patient was making 400 ml/hr of urine output, and left ventricular end-diastolic pressure (lvedp) of 31. the pump flow had to be reduced due to suction alarms, and was only able to be restored to p-8 after blood volume increase after albumin administration. the patient's hemodynamic parameters were out of regular ranges with cardiac output (co) 3.7, cardiac index (ci) 1.9, mixed venous oxygen saturation (svo2) 42%, and lactate in mmol/l (la 3.3). patient was administered blood products before explanting the pump. the root cause of the pump suction was most likely patient condition related based on provided clinical details and data log analysis. access site adverse event: clinical details mention bruising at the impella access site. activated clotting time (acts) were 109, with the doctor not going forward with anticoagulation management at the time. the patient received tenecteplase (tnkase) later and had bleeding from the venous sheath. the root cause of the access site bleeding was not determined due to lack of sufficient clinical details. ischemia: clinical details mention that the patient's feet were cold to touch per the nurse. knee immobilizer was requested for the patient and the patient was presumed to be neuro intact per the doctor. in order to make a cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: clinical details mention the patient having a urine output above 30 ml/hr and tea colored. the cause of the injury was not determined due to insufficient clinical details. device history review pump passed all post-sterile inspection checks.
5456 | 1220648-2025-46577 | 23085899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23085899&pc=OZD | 08/19/2025 | 10/22/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680821 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female | 112 KG | Non Hispanic | Black Or African American | Death | Yes | No | Patient Device Interaction Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on bipella support had lactate dehydrogenase (ldh) spiked. the impella cp was verified with echocardiogram. the patient was on dialysis. the team continued to trend ldh for possible hemolysis. it was further reported that care was withdrawn and the patient expired while on impella cp support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and thrombosis has been completed. data logs were reviewed for the impella cp, and there were no abnormalities around the time of the reported onset of hemolysis symptoms. there were no motor current spikes during the case either. pump flows were variable during the case, however, they were generally within specification. review of the rp flex data logs confirmed an ingestion event. up to that point, pump flows were within specification and other metrics were generally stable. therefore, it couldn't be confirmed based on data logs that the biomaterial that was ingested was responsible for hemolysis that was reported the day before. additionally, the patient was also on dialysis at the time of hemolysis and the rp flex was repositioned. product was not returned for investigation. since data logs were inconclusive, clinical details were limited, and product wasn't returned for investigation, the root cause of the hemolysis could not be determined. the root cause of the thrombosis was unable to be determined due to insufficient clinical details
5457 | 1220648-2025-46578 | 23086182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086182&pc=OZD | 08/18/2025 | 10/30/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2026700883 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Decrease in Pressure | Priming Problem |  |  |  |  | Arrhythmia | Hemolysis | Tachycardia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient encountered suction, ventricular tachycardia (vt), ventricular fibrillation (vf), low purge pressure and air in purge during impella cp support. the patient had continuous runs of vt and vf requiring defibrillation intra-procedurally. the surgeon repositioned the impella catheter under echocardiogram guidance to 3.82 centimeters. in total, the patient had experienced five episodes of sustained ventricular tachycardia intra-procedurally requiring five defibrillations. the patient also encountered low purge pressure alarms. however, no leaks were noted and recommendation was given to change the purge cassette should an issue arise again, which it did not. air in purge was also encountered for which steps were walked through on how to deal with which resolved the alarm. pump suction was also encountered throughout support and hemolysis was also found and confirmed via labs. however, after five days of support, care was withdrawn and the patient passed away. | [5 empty] | A.4 and a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported tachycardia, hemolysis, purge pressure, and pump suction has been completed. data logs were reviewed, and there were no clear abnormalities. there was a concentration of suction and position unknown, however, they resolved. there were no improper position alarms or mc spikes during the entire case. pump flows were variable throughout support; however, this cannot be definitively considered a cause for the hemolysis. since insufficient clinical details were provided, data logs were inconclusive, and product wasn't returned for investigation, the cause of the hemolysis could not be determined. data logs were reviewed for the purge pressure and the alarms were confirmed. at the time that bag/cassette change procedures were being done. the purge pressure low alarms were the first to trigger, then air in purge. based on the clinical details provided supported by the issue resolving after 20 minutes in data logs, the cause of the purge pressure low and air in purge alarms were determined to be a use issue. data logs were reviewed, and there was a concentration of diastolic and systolic suction alarms on the final day of support. based on the clinical details provided that the patient's condition was worsening and ultimately care was withdrawn in combination with the mc/ps trends seen at the end of the case in data logs, the cause of the pump suction was determined to be patient condition related.
5458 | 1220648-2025-46580 | 23086271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086271&pc=OZD | 08/19/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026720538 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed in the left femoral artery and was placed during percutaneous coronary intervention (pci) as bailout. the impella was removed due to lower flows than desired. replaced with new impella cp intra - procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). the investigation for the reported low pump flow has been completed. data log review showed that in the 25 minutes the pump was on, it was in continuous suction as the motor current had a minimum of 210 ma. the placement signal was also seen trending down throughout the case, which was likely due to the patient being on extracorporeal membrane oxygenation support. motor current spiked to >1500 ma around 2 minutes after the pump was started. towards the end of the case, motor current and low increased, but the motor speed remained at the same remote patient monitoring, which indicates ingestion, as it required more work to maintain the constant motor speed. the root cause of the low pump flow was most likely biomaterial ingestion, as data logs show a motor current spike at the time of suction, as well as an increase in motor current and flow without any change to motor speed. this is one of two related reports. this file represents the first impella pump and another report was submitted for the second impella pump. refer to section h.10 for the related report number.
5459 | 1220648-2025-46581 | 23086316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086316&pc=OZD | 04/09/2025 | 10/30/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601593 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella 5.5 support. after explant the patient had a ischemic arm and the clot moved down the patient's arm. the patient remained stable and survived. | [5 empty] | The investigation for the reported limb ischemia, thrombosis, and stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia, thrombosis, and stroke could not be determined. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
5460 | 1220648-2025-46579 | 23086364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086364&pc=OZD | 08/19/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599975 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 87 KG | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered delivery issues with an impella 5.5. the decision was made for impella placement. a 10mm graft was prepped and due to lack of space on the anterior aorta, the physician elected for right axillary insertion. cutdown was made and the graft was anastomosed successfully. the impella 5.5 was prepped and purged. a 23f peel away sheath was inserted into the graft and back-bled, intending to treat the patient with the impella, but the physician made one more attempt to come off bypass and the patient's ejection fraction recovered dramatically to 50-55%. therefore, the decision was made to abort the implant and to close the chest. the patient was sent to the intensive care unit. | [5 empty] | The investigation of the reported failure mode has been completed. unable to deliver or advance : the cause of the delivery issue was unable to be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks. | A.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5461 | 1220648-2025-46582 | 23086367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086367&pc=OZD | 08/18/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711537 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Appropriate Term/Code Not Available | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted to support a patient with ami/cgs. the patient had a left main coronary 100% occluded. the pump generated low purge pressure so the purge cassette was replaced. additionally, the guidewire was replaced as it appeared defective. after eight days of support the patient expired from multi organ failure. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
5462 | 1220648-2025-46583 | 23086419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086419&pc=OZD | 08/19/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  | 08/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391763 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Malposition of Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the device was explanted, the patient outcome was not reported. | The user facility had a cp pump placed for the patient admitted in with cgs and cardiomyopathy. the pump was placed as the team planned for heart transplant. during the support the team added ecmo. the pump was alarming for high pp and suction as the pump was in poor position at the chordae. the pump was repositioned and for the high pp they were monitoring alone. the pump remains on support, currently on day 32 while awaiting a heart. |  |  |  |  | Additional information received, the device was explanted, b5 and d6b updated. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. patient information unknown.
5463 | 1220648-2025-31556 | 23086422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086422&pc=OZD | 09/01/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 09/01/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532092 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, placement signal unreliable were observed. despite this issue, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. the failure mode was changed from placement signal issue to optical signal issue because the complaint pump is a cp. there is no dp sensor on a cp, only an optical sensor. data logs were reviewed, and roughly 30 minutes into the case, the "placement signal not reliable" alarm triggered. none of these optical parameters resolved for the remainder of the case. based on clinical details, this loss of optical signal occurred immediately after an interaction with shockwave lithotripsy. the root cause of the optical signal issues was determined to be a damaged optical sensor due to shockwave interaction. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5464 | 1220648-2025-31557 | 23086455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086455&pc=OZD | 09/06/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 09/06/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531369 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a cp pump was used to support a patient undergoing high-risk percutaneous coronary intervention. during support, the placement signal was unreliable. despite this, the pump remained operational until weaning and explantation. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. impella cp returned for evaluation. as the impella cp does not have a differential pressure sensor, the failure mode investigated was changed from "placement signal issue" to "optical signal issue" due to the failure mode of "placement signal issue" only being used for issues related to the differential pressure sensor. upon visual inspection no abnormalities were noted. pump passed laser continuity test and no biomaterial was present. data logs from field were returned and showed slowly drop to near 0 over approximately 50 minutes before rising to a normal range. the root cause of the optical signal issue cannot be determined as no issues were noted with returned pump and limited clinical details were provided on patient's condition. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5465 | 1220648-2025-31558 | 23086497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086497&pc=OZD | 11/14/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 11/14/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. while on support, the automated impella controller (aic) displayed a ¿placement signal not reliable¿ alarm. support was continued, and the patient was successfully weaned off the device. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. the data logs confirm placement signal was maxed out. the console was received for investigation and fringe pattern were analyzed but showed no abnormalities. the optical connector had significant detritus which was able to have some of it cleaned away. the failure mode was unable to be reproduced after cleaning the optical connector. the root cause of the consistently unreliable placement signal was most likely due to the contaminated optical connector. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5466 | 1220648-2025-46584 | 23086824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086824&pc=OZD | 08/19/2025 | 10/22/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026732055 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 89 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The first impella was removed due to low flows and the second impella cp was successfully inserted. the patient exhibited prolonged polymorphic ventricular tachycardia but remained stable with a mean arterial pressure of 65 with extracorporeal membrane oxygenation and impella. in the intensive care unit, the plan was to address the ventricular tachycardia with hopes to improve impella flows and patient hemodynamics, but the patient expired on support. this report represents the second impella pump. another report was submitted for the issues with the first impella pump. refer to section h.10 for the related report number. | [5 empty] | The investigation of the reported failure mode has been completed. no product was received for evaluation. tachycardia: the root cause of the injury was unable to be determined due to insufficient clinical details. low pump flow: no product was returned. pump was explanted due to low flows. after review of logs, suction alarms along impella position unknown alarms were observed and motor spike. the cause of low pump flow was most likely biomaterial ingestion per log review. the investigation of the reported failure mode has been completed. no product was received for evaluation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation for the reported tachycardia and low pump flow has been completed. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. no product was returned. pump was explanted due to low flows. after review of logs, suction alarms along impella position unknown alarms were observed. the root cause of low pump flow was most likely patient condition related per log review.
5467 | 1220648-2025-46585 | 23086829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086829&pc=OZD | 08/27/2025 |  |  | 09/18/2025 |  |  | 08/11/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  | 2025608644 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male | 126 KG | Non Hispanic | Black Or African American | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Necrosis | Sepsis | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 developed sepsis. the patient¿s left foot and second toes are necrotic, and the patient has positive blood cultures. impella support coontinues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product complaint # (b)(4). d.6b explant date was added. the device was not returned for investigation. e.1 added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. h.11 added instructions for use as they were inadvertently omitted from the initial report that was submitted. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the complaint was not confirmed. the device passed all post sterile inspection checks. | B2: date of death is unknown. b3: corrected the event date. investigation summary: no product was returned for investigation. sepsis: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. necrosis, hypotension: the cause of the injury was not determined due to insufficient clinical details. major bleed: the cause of the bleed was not determined due to insufficient clinical details. pump suction: no logs are available for the date of the event. the cause of the pump suction was not determined due to insufficient clinical details. clinical details: on (b)(6), 2025: impella 5.5 s2 was implanted at 7:04 pm. on (b)(6) 2025: no impella issues. no changes in plan of care. sepsis. cultures pending. on (b)(6) 2025: l foot first and second toes necrotic. pt has positive blood cultures. on (b)(6) 2025: volume given overnight for low flow events on impella/extracorporeal membrane oxygenation (ecmo). = 1 red blood cell (rbc). on (b)(6) 2025: plan for volume due to hypotension. on (b)(6) 2025: arterial bleed from sacral wound noticed last night. 2 packed red blood cells (prbc)s given. on (b)(6) 2025: impella explanted. device history review: device with passed all post sterile inspection checks.
5468 | 1220648-2025-46586 | 23086865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086865&pc=OZD | 08/22/2024 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500684 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The study reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient had impella removed, and a hematoma formed at the impella sheath site after removal. patient received 1 u red blood cells and this hematoma was resolved after 45 minutes of manual compression. next day, patient developed new groin hematomas (l & r) and a new fever, for which the patient received vancomycin and cefepime. patient's hemoglobin dropped so another unit of red blood cells was given. cultures came back negative and fever subsided. repeat imaging of groin showed stable hematomas without concern of further extravasation. patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. hematoma: the cause of the injury was not determined since product was not returned and insufficient clinical details provided. fever: the cause of the clinical symptoms cannot be determined as insufficient clinical details were provided. device history review: the device passed all the post sterile inspection checks.
5469 | 1220648-2025-46587 | 23086875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086875&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701588 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after just 46 hours the pump was explanted due to clot concerns in the left ventricle. the patient was transferred to the tertiary center on extracorporeal membrane oxygenation (ecmo) support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. thrombosis : the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5470 | 1220648-2025-46589 | 23086994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23086994&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696409 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Inaccurate Flow Rate | Increase in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp had purge system blocked alarm. troubleshooting was done but unsuccessful. the impella started exhibiting signs of pump saturation and increased motor current. shortly after, pump briefly stopped and was not able to restart. a new impella cp was placed. no patient harm was reported due to the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. the ifu states: impella cp with smartassist. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: the device was returned for evaluation. pump stop - following uneventful implant procedure, pump stop occurred while the registered nurse (rn) was investigating purge issues. data logs recovered did not show a pump stop. the returned pump had biomaterial wrapped around the impellor blades. the pump was unable to be started during investigation due to the biomaterial. biomaterial could not be fully removed from the purge gap to allow for pump start. the cause of the pump stop was biomaterial build-up. saturated flow - during purge flow investigation by the rn, saturated flows occurred. recovered data logs did not show saturated flows, and not all data logs were recovered. the returned pump had biomaterial wrapped around the impellor blades. the pump was unable to be started during investigation due to the biomaterial. the cause of the saturated flow was not determined as the data logs showed no saturated flow in the field. purge blocked - following uneventful impella implant, purge system blocked alarms occurred. data log review showing a gradual build up of purge pressure with decreasing purge flows. the cause of the purge system blocked was biomaterial build up. device history review: pump passed all post sterile inspection checks.
5471 | 1220648-2025-46591 | 23087582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23087582&pc=OZD | 08/12/2025 |  |  | 09/18/2025 |  |  | 08/12/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1378679 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Smoking | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during preventive maintenance the automated impella controller generated a bootup/self-check alarm and gave off a burning odor. there was no patient use at the time of the event. | [5 empty] | The reported event did not occur during patient use. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. data analysis: the complaint was reported on aug 12th, 2025; however, according to the logs, the issue first occurred on aug 8th, 2025. during initial boot up, the console restarted automatically due to a pmd communication issue. due to a persistent pmd communication issue, the console rebooted several times and displayed ¿system self check failed.¿ when the console was rebooted again on august 27th, 2025, the issue persisted, and the console continued to display "system self check failed" due to the ongoing pmd communication issue. this confirms the reported failure mode. device analysis: this console was tested after arriving and reproduced the reported failure mode. during the initial boot-up, the console restarted automatically due to the pmd communication issue, resulting in the ¿system self check failed¿ screen. the fwcheck via telnet confirmed the problem was due to the impellatronic pca board, as the hw revision was n.a. for the sau_motor_1. the issue was resolved by replacing the old impellatronic pca with a known good one. after the replacement, no further issues were observed with the known good pump and purge cassette attached. further no burning smell or burned component was noticed during investigation. root cause: the root cause of the "system self check failure" was determined to be a defective impellatronic pca board. no issues related to burning smell or damaged components were observed within the automated impella controller.
5472 | 1220648-2025-46593 | 23088076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23088076&pc=PYX | 08/19/2025 | 11/11/2025 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600762 |  | 000000000001000323 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Female | 41 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem |  |  |  |  | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that during support of the impella rp flex, that hemolysis noted with decreased renal function and urine output. also pump flow was at p-8 with suction alarms. impella flows decreased the following morning. there was some bleeding noted at the rp flex site and surgery added stitch with appropriate hemostasis. the patient was later weaned successfully and survived. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the pump was returned. pump suction - data log review confirmed suction events at p-8 which ceased when pump was turned below p-7 and did not occur for remainder of case. the returned pump had catheter damage, impacted the motor lines but not the 5s parameters. when attempting to start pump, motor current (mc) high alarm triggered and impella stopped. the cause of the pump suction was patient condition related due to low volume. hemolysis and suction resolved with lowered p-levels. mechanical interaction with blood - it was reported that the hemolysis and contributing hemolysis resolved by giving blood volume. data log review confirmed suction events at p-8 which ceased when pump was turned below p-7 and did not occur for remainder of case. placement signal (ps) was low at the time of the hemolysis/suction and upon suction resolving, ps gradually increased over the remainder of the case. no abnormalities related to hemolysis were noted on the returned pump. the cause of the mechanical interaction with blood was patient condition related due to suction events occurring above p-8, when hemolysis was noted. hemolysis and suction resolved with lowered p-levels. access site adverse event - the returned pump showed no abnormalities related to pt bleeding. the cause of the access site bleeding was use related as the bleeding ceased upon proper stitch placement leading to appropriate hemostasis. renal failure: the cause of the renal failure could not be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks. | The following fields were corrected: b1, d2a, d2b, h6. upon investigation closure, a supplemental manufacturer device report will be filed.
5473 | 1220648-2025-46590 | 23088550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23088550&pc=OZD | 08/18/2025 |  |  | 09/18/2025 |  |  | 08/18/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026711984 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp was placed to support a patient in non-st elevated myocardial infarction. while on support, the patient experienced oozing. the access site was re-sutured with forward tension sutures in place and the site was angle matched. pressure was held to dissolve the hematoma. the perclose suture was secured and the hematoma was mostly resolved. after transfer to the intensive care unit, it was noted the hematoma worsened and critical care placed a fem stop for a few hours. bruising was noted from the hematoma. one unit of packed red blood cells and fresh frozen plasma was transfused. it was reported that hemostasis was achieved through the femstop. additionally, intermittent suction alarms were observed. the case proceeded as the patient was transferred from the impella cp to an impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | No product was returned for investigation. suction alarms were observed in the data logs. the cause of pump suction cannot be determined due to insufficient clinical details. the device passed all post sterile inspection checks.
5474 | 1220648-2025-46594 | 23089025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23089025&pc=OZD | 08/20/2025 |  |  | 09/18/2025 |  |  | 08/20/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687549 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 67 KG |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Low Blood Pressure/ Hypotension | Inflammation | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while awaiting a heart transplant. on the second day of support, it was reported that the patient was experiencing a new onset of ectopy. an echocardiogram and chest x-ray were ordered to evaluate the position of the impella 5.5 and the swan. it was also noted that the patient¿s arm distal to the impella insertion was swollen and the impella access site was redressed with pressure dressing. the following day, it was reported that there were suction events which were orthostatic in nature and it was noted that patient had a very low central venous pressure. as a result of the hypotension issues, the patient¿s lasix intravenous drip was discontinued. due to the patient¿s renal function and likelihood for dialysis, the patient was deemed to not be a candidate for a heart transplant but was approved for a durable left ventricle assist device (lvad). the patient successfully underwent surgery for an lvad at which time the impella 5.5 was successfully weaned and explanted. | [5 empty] | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary ==> no product was returned. arrhythmia: in order to make a cause determination, additional clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. pump suction: no product was returned. orthostatic suction events were present in the morning. after reviewing logs, suction alarms and position unknown alarms were observed that resolve. no motor current (mc) spikes outside of p-level changes. the cause of the suction event was likely patient condition related per clinical since the suction was orthostatic in nature, the patient had very low central venous pressure (cvp), and log review showed no abnormalities confirming the suction resolved. device history lot ==> device lot: 1932174. device history review ==> device (sn: (b)(6)) passed all post sterile inspection checks.
5475 | 1220648-2025-46595 | 23090199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23090199&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026698836 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 140 KG | Non Hispanic |  | Injury | Yes | No | Defective Component | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [28 empty] | Initial submission | Followup | [8 empty] | An impella cp was implanted in a patient undergoing an emergency procedure for coronary artery complication in percutaneous intervention. a groin bleed occurred as the tissue track was too long for the reposheath due to body habitus. the team explanted the pump and performed a contralateral leg closure and covered stenting. no patient harm was reported. | [5 empty] | A6 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | E.1. added initial reporter phone number as it was inadvertently omitted from the initial report that was submitted. the investigation has been completed. no product was returned for investigation. major bleed: the patient started bleeding from the groin. the cause of the bleeding was determined to be patient condition because due to the patient's body habitus, the repo sheath did not totally reach the artery after inserting all the way to the skin and was just not long enough for the patient to create hemostasis. device history review: the device passed all post sterile inspection checks.
5476 | 1220648-2025-46596 | 23090321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23090321&pc=OZD | 08/05/2025 |  |  | 09/18/2025 |  |  | 08/20/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1470310 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while supporting an implanted impella cp, an automated impella controller generated low purge flow alarms. the purge cassette was replaced and the amount of heparin added to the purge fluid was doubled. additionally, an intraventricular alarm occurred. the pump position was adjusted but the alarm did not resolve. later, the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
5477 | 1220648-2025-46597 | 23091685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23091685&pc=OZD | 08/20/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 08/20/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661770 |  | 000000000001000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered purge system blocked, and the patient having a bleeding, atrial fibrillation (afib) and tamponade during impella 5.5 support. during support, the patient had a tamponade event due to bleeding and oozing. the following day, heparin was stopped due to bleeding and increased chest tube output. additionally, the patient encountered atrial fibrillation and the team managed the rhythm. finally, a "purge system blocked" alarm was encountered and the team made the decision to remove the device that same day. device was successfully removed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Correction e4 the investigation of the reported failure mode has been completed. the impella device was returned for investigation major bleed/cardiac tamponade: the cause of the injury was not determined due to insufficient clinical details. arrhythmia: the root cause of the injury was unable to be determined due to insufficient clinical details. purge system blocked: the pump was returned cut therefore product testing was limited. biomaterial was observed in the purge gap. the data logs show a gradual rise in purge pressure for about an hour before alarms were triggered. there were no motor current spikes outside p-level changes before purge pressure rise. the high purge pressure was sustained till explant. the cause of the purge system blocked was due to biomaterial build up. device history review: the device passed all the post sterile inspection checks.
5478 | 1220648-2025-31559 | 23092566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092566&pc=PYX | 11/16/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 11/16/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025483298 |  | 1000323 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp flex support. while on support, placement signal not reliable (psnr) alarm with absent pulmonary artery pressure (paps) and impella flow numbers was present. the automated impella controller (aic) showed increase paps and low flows with stable motor current (mc) prior to alarm. alarm audio was turned off. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. log review showed placement signal drifting up with a simultaneous decrease in pump flow starting on 11/16. motor current remained stable and reacted to the p-level changes. cvp also remained stable while the placement signal drifted upward. a dp placement signal not reliable alarm triggered as well. the root cause of the placement signal issue was mostly likely sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5479 | 1220648-2025-31560 | 23092567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092567&pc=OZD | 10/05/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 10/05/2024 | 09/18/2025 |  |  | 01/31/2026 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524042 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, the impella cp pump triggered a optical signal loss. hemodynamics were confirmed and the functionality of the pump remained unaffected. the device was successfully weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. neither product nor data logs were returned for investigation. as product was not returned for investigation and clinical information was limited, the root cause of the optical signal issues could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5480 | 1220648-2025-31561 | 23092569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092569&pc=OZD | 11/18/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 11/18/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525047 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during use, a ¿placement signal not reliable¿ alarm suddenly occurred, and the placement signal disappeared. the motor waveform displayed a pulsed pattern. the placement signal subsequently reappeared and disappeared intermittently. despite this, the impella continued to provide effective hemodynamic support, and its operation was not affected. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the returned pump was inspected and there were no physical abnormalities. the 5s parameters were within expected range. the data logs were reviewed and there were 'controller error' alarms. no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5481 | 1220648-2025-31562 | 23092584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092584&pc=OZD | 03/21/2023 | 09/18/2025 |  | 09/18/2025 |  |  | 03/21/2023 | 09/18/2025 |  |  | 06/30/2024 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235322 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during use, an optical signal failure. despite this, the impella continued to provide effective hemodynamic support, and its operation was not affected. there was no reported injury or harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was returned and a damaged fiber line was found in the red handle. a loose bond was observed between the kink protector and the catheter allowing the fibers and cables to rotate. the data logs were reviewed and indicate a fiber break. the root cause of the optical signal issue was determined to be a damaged optical fiber line in the red handle resulting from a failed adhesive bond between the catheter and the kink protector. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5482 | 1220648-2025-31563 | 23092592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092592&pc=OZD | 02/23/2023 | 09/18/2025 |  | 09/18/2025 |  |  | 02/23/2023 | 09/18/2025 |  |  | 12/31/2023 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 380 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024299465 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. it was reported that the impella cp lost placement signal after the use of the shockwave. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data log review indicates the 5s parameters as expected at the start of the case. upon initiation of shockwave, 5s parameters immediately become consistent with a sensor break. the pump was investigated and the sensor was removed from the pump, revealing a broken sensor face. based on the clinical description that ps was lost coinciding with the initiation of shockwave, the root cause of the optical signal issue is due to a damaged sensor as a result of interaction with shockwave. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5483 | 1220648-2025-31564 | 23092597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092597&pc=OZD | 10/10/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 10/10/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025483535 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was implanted in a patient for circulatory support. the complainant reported that the impella 5.5 had a 'placement signal not reliable' alarm. the alarm was disabled and motor currents were monitored. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. there are placement signal not reliable (psnr) alarms and controller error alarms about 81 hours into the case but then they resolve. during this time the snr increases, contrast increases, lamp increase and gain and light remain stable. the placement signal is noisy throughout the case and snr is low. there was no product returned. due to the inconclusiveness of the data logs and lack of product returned, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5484 | 1220648-2025-31565 | 23092609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092609&pc=OZD | 12/01/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 12/01/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569420 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. the complainant reported that there was a placement signal issue. the impella red alarm indicating that the pump was in the aorta appeared, however, the doctor stated that the position was checked several times and the impella is approx. 5 cm below the aortic valve. it was also noted that if the doctor pulled the pump slightly more (around 5mm) into the ventricle, the "impella in the ventricle" alarm was triggered immediately. however, whenever the pump was correctly positioned at 5 cm, the impella in the aorta alarm was displayed. patient was stable even under p2. it was discussed that there was a possible sensor error and recommendation was given to deactivate the placement signal. the placement signal monitoring was switched off. the pump remained on until the pump was successfully weaned and explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella 5.5 was returned for evaluation. upon visual inspection, biomaterial was present on the impeller blade. optical parameters were checked and all within a normal range. pump also passed laser light continuity test. data logs were reviewed and lt logs from case showed ¿impella position in aorta¿ alarms with concurrent ¿impella position in ventricle¿ alarms noted at time of reported issue and troubleshooting. continuous suction alarms also noted throughout case. no hard failures of the optical sensor were seen in the logs during the case. a few hours before "impella in ventricle/aorta" alarms, a drop in motor current pusatility and a drop in pump flows with no change in p level were seen. additionally, placement signal decreases at the same time as pump flow drop. clinical details noted that when pump was correctly positioned at 5cm, "impella in aorta" alarms were being triggered. the failure mode investigated was changed from "placement signal issue" which only encompass dp sensor issues to "optical signal issue" as the impella 5.5 which only has an optical sensor. the root cause of the optical signal issue was most likely due to biomaterial. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5485 | 1220648-2025-31566 | 23092628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092628&pc=OZD | 12/16/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 12/16/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569414 |  | 1000482 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted in a patient for circulatory support. while on support, there was a placement signal and left ventricle curve not displayed alarm. the motor current pulsatile verified the impella was in the correct position. the device was successfully weaned and explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was not returned for investigation. data logs were not returned as well. without the product and data logs, the failure cannot be investigated further. therefore, the root cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5486 | 1220648-2025-31568 | 23092643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092643&pc=PYX | 05/08/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 05/08/2024 | 09/18/2025 |  |  | 10/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025448699 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during planned explant of an impella rp the device generated a placement signal alarm. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data logs show that about 4 hours into support the placement signal begins to drift. it continues to drift upwards for 33 hours and then shifts and drifts downwards for another 33 hours before stabilizing for the remainder of the case. on the last day of support at the very end of the case the cvp spikes to 300 mmhg triggering placement signal not reliable (snr) (op) alarm. at the same time, snr drops to 0, contrast drops, lamp maxes at 100, gain increases, and light decreases. the optical parameters do not recover. the returned 5s parameters are not within specification and the cavity length is maxed out. the pump passed the laser continuity test and electrical testing. psnr alarm was reproduced in the lab. imaging the op sensor showed damage to the sensor face. the root cause of the optical signal issue was most likely a damaged sensor. running the pump in the pressurized flow loop reproduced the dp sensor drift issue. the root cause of the placement signal issues was most likely dp sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5487 | 1220648-2025-31569 | 23092651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092651&pc=PYX | 10/13/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 10/13/2024 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025508363 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported that the placement signal not reliable alarm occurred. the staff reported that the patient coughed and the placement signal numbers became skewed. the audio was then disabled. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issues was most likely patient condition related as no issues with the optical sensor were noted and the central venous pressure was able to return to normal. the root cause of the placement signal issues was most likely due to sensor drift of the dp sensor. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5488 | 1220648-2025-46588 | 23092660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092660&pc=OZD | 08/19/2025 | 01/06/2026 |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1490418 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they suspected an automated impella controller (aic) failure. when connecting the purge set and pump for impella use, the aic screen skipped the priming process and changed to the screen immediately before starting. after disconnecting the purge set and pump, resetting them, and restarting the aic, the aic process proceeded normally. additional information on august 21, 2025, indicated that the aic operated without any problems from the start of support until the end of support. no patient harm was reported due to the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental/final report will be filed.
5489 | 1220648-2025-31570 | 23092661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092661&pc=PYX | 03/07/2023 | 10/23/2025 |  | 09/18/2025 |  |  | 03/07/2023 | 09/18/2025 |  |  | 12/31/2024 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024295647 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. the cp was placed along with an impella rp for bipella support. the rp had lost placement signal after two days of support. the audio was disabled. the pump remained on for support until weaning and explantation. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was not determined since neither data logs nor product were returned. all pertinent information available to abiomed has been submitted at this time. | D6b "if explanted, give date" corrected.
5490 | 1220648-2025-46600 | 23092662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092662&pc=OZD | 08/19/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1115051 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during preparation of the impella cp, the priming of the system failed due to the automated impella controller (aic) having a broken purge disc bracket that was no longer in place. a back up aic unit was used and the impella cp was successfully primed, and support initiated. the patient remained on impella cp support, and it was noted the impella cp was explanted; however, patient outcome at time of explant is unknown. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental report will be filed.
5491 | 1220648-2025-46599 | 23092665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092665&pc=OZD | 08/01/2025 |  |  | 09/18/2025 |  |  | 08/19/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601224 |  | 000000000001000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was being implanted with an impella 5.5 for increased mechanical circulatory support. the patient had already been implanted with impella cp and veno-arterial extracorporeal membrane oxygenation to manage myocarditis. the medical team decided to escalate the patient from the impella cp to an impella 5.5 and it was reported that the guidewire was kinked when removed from the box. a new guidewire was used. additionally, it was reported that the impella 5.5 was ultimately aborted due to patient anatomy and instead an intra-aortic balloon pump was successfully placed. it was noted that no patient harm was reported due to the reported issues. this complaint involves one guide wire 0.018" ptfe coated and one impella 5.5 pump: guide wire 0.018" ptfe coated with lot number 8853993 pump 2: impella 5.5, with serial number (b)(6). this mdr specifically addresses impella 5.5, with serial number (b)(6), which is the subject of this report. separate mdrs will be submitted for the other devices associated with this complaint. | [5 empty] | Product complaint # (b)(4). no device received, investigation complete: the investigation into the reported event has been completed. no device was received for evaluation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5492 | 1220648-2025-31571 | 23092667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092667&pc=PYX | 11/15/2023 | 09/18/2025 |  | 09/18/2025 |  |  | 11/15/2023 | 09/18/2025 |  |  | 08/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024392247 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. the complainant reported there was a placement signal unreliable alarm. no other issues were noted. the patient was successfully weaned from the pump and the device explanted. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal failure was most likely a small air gap from between the automated impella controller and the patient cable plug based on the data log and clinical review. all pertinent information available to abiomed has been submitted at this time.
5493 | 1220648-2025-31572 | 23092681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092681&pc=PYX | 09/12/2023 | 09/18/2025 |  | 09/18/2025 |  |  | 09/12/2023 | 09/18/2025 |  |  | 09/30/2023 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199492 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 and impella rp pump were implanted as bipella support in a patient for circulatory support. while on bipella support, after a coughing event and a turn; a patient signal not reliable alarm on the impella rp occurred. flow stayed the same with no change in patient or pump performance. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was not determined since the product was not returned and sufficient clinical information was not provided. all pertinent information available to abiomed has been submitted at this time.
5494 | 1220648-2025-31573 | 23092684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092684&pc=OZD | 11/27/2023 | 09/18/2025 |  | 09/18/2025 |  |  | 11/27/2023 | 09/18/2025 |  |  | 09/30/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403399 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a "placement signal not reliable" alarm was displayed on the automated impella controller (aic). a transthoracic echocardiogram (tte) was performed, and the impella 5.5 was noted to be in good position with no thrombus noted. the alarm was then silenced. there was no reported patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the placement signal issue was not determined since the product was not returned and the data logs were not returned as well. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5495 | 1220648-2025-31574 | 23092688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092688&pc=OZD | 12/16/2023 | 09/18/2025 |  | 09/18/2025 |  |  | 12/16/2023 | 09/18/2025 |  |  | 05/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366664 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal on the impella cp was excessively reading lower than the patient¿s arterial line. proper positioning was confirmed via echocardiography, and, since the pump was functioning well, the menu was changed in the placement signal value to match the patient¿s arterial line. after the change all alarms cleared. there was no harm to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue was not determined since product was not returned for investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5496 | 1220648-2025-46601 | 23092692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092692&pc=OZD | 08/18/2025 |  |  | 09/18/2025 |  |  | 08/20/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026713046 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. it was reported that following percutaneous coronary intervention, in intensive care unit the patient lost pulse to left leg. the patient was taken back to the cath lab. angio showed occluded superficial femoral artery and dissected left iliac. common femoral ballooned and iliac stented to restore flow. the patient remained stable and was weaned successfully. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | Product complaint #(b)(4). investigation summary: no product was returned for investigation. pain: the cause of the clinical symptom was not determined due to insufficient clinical details. ischemia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the ischemia was unable to be determined due to insufficient clinical details. vascular dissection: angio showed occluded sfa and dissected left iliac. the cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history lot: device lot: 1937885. device history batch: subcomponent lot: n/a. device history review: device with sn: (b)(6) passed all post sterile inspection checks.
5497 | 1220648-2025-31575 | 23092693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092693&pc=OZD | 02/01/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 02/01/2024 | 09/18/2025 |  |  | 10/31/2025 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414764 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted in a patient for circulatory support. while on support, there was a placement signal not reliable alarm. the audio was disabled and the patient remained stable. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue cannot be definitively determined without inspection of the pump. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5498 | 1220648-2025-31576 | 23092706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092706&pc=PYX | 09/13/2024 | 09/18/2025 |  | 09/18/2025 |  |  | 09/13/2024 | 09/18/2025 |  |  | 02/28/2026 |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478584 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impela rp after a coronary artery bypass graft and repeat percutaneous coronary intervention in the catheterization lab. the rp was supporting the patient despite a loss of the placement signal. the optical loss caused no harm or injury and did not prompt intervention. the pump supported until wean and explant after seventy three hours. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. based on clinical details and data logs, the root cause of the placement signal issue was determined to most likely be dp sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5499 | 1220648-2025-31577 | 23092712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23092712&pc=PYX | 01/03/2025 | 09/18/2025 |  | 09/18/2025 |  |  | 01/03/2025 | 09/18/2025 |  |  |  |  | 09/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548241 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the central venous pressure and pulmonary artery pulsatility index readings were inaccurate following an impella rp flex implant. support with the implanted pump was uninterrupted despite the noted issue. the patient was successfully weaned from the pump and the device explanted. there was no reported harm or injury to the patient. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue was not determined since product/logs were not returned and insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5500 | 1220648-2025-46540 | 23071604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23071604&pc=OZD | 08/18/2025 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665770 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction/cardiogenic shock implanted with an impella cp for mechanical circulatory support and experienced loss of pulsatility requiring intervention and would subsequently expire. the patient was admitted with non-st-segment elevation myocardial infarction and sent to the unit. the patient presented to the catheterization lab and during a left heart catheterization, upon initial engagement of the left main with a diagnostic catheter, rhythm change was noted, and the patient became hypotensive and unresponsive. angiogram showed ostial left main lesion with possible clot of the distal left main and no flow into the circumflex or left anterior descending arteries. the decision was made to place an impella cp device. the patient required continuous cardiopulmonary resuscitation during insertion and an epinephrine drip was started. the left femoral artery access was obtained. the impella cp was implanted and started on auto. the physician was able to wire the circumflex and left anterior descending arteries. angioplasty of both vessels, along with the left main was completed. penumbra was used in the left main and left anterior descending arteries. flow was restored in the vessels. however, it was noted pulsatility was absent when chest compressions stopped. eventually, the decision was made to attempt placement of an impella rp flex via the right femoral vein. however, the attempt was unsuccessful as while attempting to get good wire placement into pulmonary artery for impella rp flex placement, the pulmonary artery catheter popped out into right ventricle. at this time, the patient¿s family decided to cease resuscitation efforts and the impella rp flex implant was aborted, and the patient expired. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ related medical device reports: this impella cp report is one of two devices associated with the event. there is not enough information to dissociate the impella cp from the demise outcome. however, it should be noted, the patient was under active cardiopulmonary resuscitation at the time of the impella cp implant and likely expired from the cardiac event. refer to the related manufacturer report number noted in section h.10 for the report on the impella rp flex. | The device was not returned; therefore, an evaluation/analysis of the device was not possible. the investigation was completed and noted a device history record review confirmed the device passed all the post sterile inspection checks. regarding the ischemia, in order to make a cause determination, all of the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. h6 investigation: type, findings and conclusion codes and h6 component code were updated accordingly based on the completed investigation. related medical device reports: this impella cp device report is one of two devices associated with the event. refer to the related manufacturer report number noted in section h.10 for the report on the impella rp flex.
5501 | 1220648-2025-46542 | 23071855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23071855&pc=OZD | 08/15/2025 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685622 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient was admitted to the hospital for heart failure treatment when the patient experienced ventricular tachycardia and cardiopulmonary arrest. the patient then was cannulated for extracorporeal membrane oxygenation, and an intra-aortic balloon pump was placed. the patient then underwent a percutaneous cardiac intervention, and the decision was made after to implant an impella cp pump for escalated mechanical support. it was reported that after the impella cp support was initiated, the patient was returned to the cardiac care unit. upon arrival at the room, an ultrasound was performed and confirmed that the impella had fallen out of the left ventricle and was in the aorta. the impella cp was immediately removed and replaced with a new impella cp. three days later, the patient expired on support with the second impella cp pump. no adverse events or issues were reported with the second impella cp pump. | [5 empty] | H6 code b18 was reported incorrectly on the initial report that was submitted. b17 is the code that should of been on the initial report. h6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary: the investigation has been completed. no product was returned. device in wrong position (positioning issue): the cause of positioning issue cannot be determined since insufficient clinical details were provided. device history review: the device passed all post sterile inspection checks. | A.2 and a.4 are unknown. based on the information available, the first impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5502 | 1220648-2025-31515 | 23071949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23071949&pc=OZD | 07/09/2025 | 11/24/2025 |  | 09/17/2025 |  |  | 09/10/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1537012 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after inserting the impella cp and performing pci, the position waveform disappeared before leaving the catheter room, and the placement signal not reliable alarm sounded. the position waveform disappeared for several hours after that, but the waveform reappeared several hours later and was displayed without any problems. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that shortly after start, for 3 hours placement signal was not reliable (with snr close to 0), which presented to the operators as placement signal absent from the aic screen. after the signal returned, there was a large number of suction alarms, and the average snr of optical signal was below 1500 throughout the case. console was tested after arriving at fs, and its optical system was found to be failing ¿5s¿ specification for test cavities with different lengths: test cavity ¿a¿, 16560nm test cavity ¿b¿, 16280nm, lamp: 100.00%, light: 2.70v, cont.: 55.18%, gain: 1.41, snr: 699.2, mano: 757.1 lamp: 100.00%, light: 2.72v, cont.: 60.93%, gain: 1.32, snr: 811.4, mano: 757.1. inspection of the console revealed minor contamination of the optical pump connector fiber, but it did not contribute to the optical system malfunction. the reported disappearance of position waveform and psnr condition were caused by failure of optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h3 updated. d10 concomitant medical products and therapy dates updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
5503 | 1220648-2025-31516 | 23072090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072090&pc=OZD | 02/04/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 02/04/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585331 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during therapy, the pump generated a false in-aorta alarm and experienced loss of signal. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned for investigation. data logs were not provided for this case run. the root cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5504 | 1220648-2025-31517 | 23072579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072579&pc=OZD | 01/20/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 01/20/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568879 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, a discrepancy was observed between the aortic pressures displayed on the console and peripheral pressure measurements. left ventricular systolic and diastolic values were also inaccurate. the console was replaced for troubleshooting; however, the signal discrepancy persisted. support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. there were no data logs returned. the product was returned and there were no damages or abnormalities seen. the 5s parameters were within specification and no placement signal not reliable alarms were reproduced. however, since field logs were not returned, the initial calibration of the sensor can not be seen and it can not be confirmed if the user applied an offset, influencing the calibrated g0 from the lab run. due to lack of data logs returned, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5505 | 1220648-2025-46543 | 23072680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072680&pc=OZD | 08/04/2025 |  |  | 09/17/2025 |  |  | 08/04/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676221 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 69 KG | Non Hispanic |  | Death | Yes | No | Difficult or Delayed Positioning | Restricted Flow rate | Patient-Device Incompatibility |  |  |  |  | Arrhythmia | Fever | Tachycardia | Ventricular Fibrillation | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient presented to the emergency department with shortness of breath, chest pain, tachycardia, and in non- st elevation myocardial infarction. the patient was intubated and taken to the cardiovascular laboratory where the patient was placed on an intra-aortic balloon pump to stabilize, but the patient continued to decompensate, and the decision was made to the place the patient on an impella 5.5 for mechanical circulatory support until a percutaneous or surgical intervention could be done. it was reported that during impella 5.5 implantation, the physician had issues navigating through the patient¿s calcified arteries. the physician utilized balloon angioplasty and a 20f dry seal flex sheath to successfully pass the calcification. during the second day of impella support, it was noted that while suctioning the patient, the patient experienced ventricular tachycardia (vt) which self-resolved. later the same day, the patient had additional runs of vt. it was noted there were no impella alarms at that time. the following morning, echocardiography was performed and found the pump to be too shallow. the physician repositioned the impella, but the pump possibly went too deep during positioning and the patient then went into another episode of vt that required defibrillation to resolve, and the patient was given amiodarone. the patient was then transferred to another hospital for advanced support on (b)(6) 2025. after arriving at the new hospital, the patient went into ventricular fibrillation (vf). the patient was cardioverted back into sinus tachycardia. echocardiography was performed after arrival and found the impella slightly shallow at 4.2 centimeters (cm) and was advanced to 4.6 cm. the following day, it was reported that the patient had another episode of vf while sitting in a chair which self-resolved. it was noted that the patient had a percutaneous cardiac intervention (pci) on (b)(6) 2025, and the following day it was reported that the patient was experiencing lidocaine toxicity, and the patient was weaned from lidocaine and precedex. during the wean, the patient began experiencing ectopy. during this time, there was an impella alarm for suction and to resolve the p level was decreased and the patient was given albumin. on (b)(6) 2025 it was reported that the patient had agitations during the previous night and was started on precedex and given a unit of packed red blood cells (prbcs). it was noted that the patient¿s blood loss may have been due to blood loss during the pci. it was also noted that the patient had been febrile over the last few days but was improving. on (b)(6) 2025 it was noted that the patient¿s hematocrit was continuing to drop, however, there was no root cause established for the blood loss. the patient coded during the morning of (b)(6), and the patient was placed on levophed and vasopressin. it was noted that the impella again migrated into a shallow position and was repositioned. the following day, the patient¿s hematocrit continued to drop, and the patient was transfused with two additional units of prbcs. it was noted there were no signs of bleeding or hematoma at the access site, and the cause of the drop in hematocrit was unknown. the patient began weaning from the impella, and it was noted that on (b)(6) 2025 the patients status was updated to do not resuscitate. the following day, the decision was made by the family to withdraw care. care was withdrawn and the patient expired. | [5 empty] | A.6 is unknown. based on the information available, the impella 5.5 cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not collected at the time of explant by the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary : pump was not returned. clinical symptoms:- vt/vf/arrhythmia: in order to make a cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. fever: the cause of the injury was not determined due to insufficient clinical details. pec:- patienty anatomy or condition prior to therapy: during implant it was reported that the physician had some trouble navigating through the calcified artery during impella implant. a balloon angioplasty and 20 fr dry seal flex sheath were used to pass through the calcification. the cause of patient anatomy or condition prior to therapy is due to patient condition as patient had pre-existing calcified vessels. difficult to place or position: as per clinical, the impella was repositioned as it continued to move into a shallow position. the patient coded in the morning, intubated, and placed on levophed and vasopressors. the medical team was very uncertain about what was causing all of the issues at this point as the cardiac numbers were good. the cause of position issues was not determined due to insufficient clinical information. pump suction: as per clinical one suction alarm occurred and resolved after a p-level decrease. the patient was given a concentrated albumin. data logs reviewed and suction alarm were observed and p-level decrease was done to immediately resolve suction. the cause of pump suction was patient condition related based on clinical and log review. device history lot : device lot: 1924527. device history batch : subcomponent lot: n/a. device history review : the pump sn (b)(6) passed all post sterile inspection checks.
5506 | 1220648-2025-46544 | 23072739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072739&pc=OZD | 08/18/2025 |  |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026701135 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Thrombosis/Thrombus | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support was bronched twice due to bleeding from the endotracheal tube. as a result, anticoagulation was withheld. on day six of impella support, the patient was bradycardic and then hypotensive with maps in low 40s. cardiopulmonary resuscitation was initiated and return of spontaneous circulation was achieved after approximately five minutes. the physician feels that this is not impella related but ultimately caused by a very weak heart on presser which is making the heart extra irritable. additionally, the patient had a bronchoscopy, and the lungs were found to have clots. the patient received inhaled tissue plasminogen activator. the patient hemodynamics worsened and the expired-on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | No product was returned for investigation. the cause of the hypotension was not determined due to insufficient clinical details. the root cause of the bradycardia/thrombosis was unable to be determined due to insufficient clinical details. the cause of the bleeding is determined to be patient condition because the device passed all post sterile inspection checks.
5507 | 1220648-2025-46541 | 23072759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072759&pc=OZD | 08/18/2025 | 10/31/2025 |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 85 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a patient with a history of cardiomyopathy awaiting a left ventricular assist device. the pump was supporting but cardioversion was needed. the patient also had heparin withheld due to gastrointestinal bleeding. after six days of support, the family made decision to withdraw care. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the cause of the bleed is most likely patient condition related since bleeding from multiple sites. the pump passed all the post sterile inspection checks.
5508 | 1220648-2025-31518 | 23072893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072893&pc=OZD | 02/01/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 02/01/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604018 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during use, placement signal issues were observed. device positioning was verified, and support was continued despite the alarms. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was not returned. data logs confirmed even on new console, impella position unknown alarms persisted. the root cause of the optical signal issue was not determined since product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5509 | 1220648-2025-46531 | 23072983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23072983&pc=OZD | 08/13/2025 | 01/29/2026 |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026709958 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Death | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hematoma | Nerve Damage | Pain | Thrombosis/Thrombus | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | It was reported that a patient was implanted with an impella 5.5 and developed a hematoma which prompted withholding of heparin. the patient also had nerve numbness and pain in their arm. the patient was treated with tylenol. impella support continues. | Additional information was received. the hematoma was old meaning occurred post operatively per the clinical consultant. the numbness has gone away. the patient is awaiting a heart transplant and is still on support. a return of the device can not be guaranteed. | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient presented to an outside hospital where a right heart catheterization revealed high filling pressures. the patient was transferred to the complainant hospital for advanced therapy evaluation. the patient was then brought to the operating room for an impella 5.5 implant as a bridge to heart transplant. the impella 5.5 was implanted via the right axillary artery without complications. on support day 7, it was noted that the patient experienced shortness of breath when laying flat and sometimes when sitting. an echocardiogram was performed and confirmed the pump was in good position. the patient¿s mitral regurgitation was severe and would require increased medical management. the patient remained on impella support awaiting heart transplant. on support day 12, it was noted that the patient¿s heparin was withheld due to a hematoma in the right armpit of the patient. the medical staff reported that the patient¿s hematoma was stable. the patient experienced discomfort to the right shoulder and arm, and had occasional numbness to the right hand. on support day 14, the patient continued to remain off anticoagulation therapy because of the hematoma that presented two days previously. on support day 23, the impella had alarms for suction. the patient received fluids and the alarms were resolved. during the suction alarms, the patient reported to feel ¿not well¿ and diaphoretic, however the patient improved after fluids were infused and suction alarms resolved. on support day 28, it was noted that there were some suction alarms which occurred when the patient coughed and self-resolved. on support day 46, it was noted that the nurse requested troubleshooting review for intermittent suctions alarms. the patient was noted to be stable and was still awaiting a heart transplant. on support day 49, it was reported that the previous day the patient was found slumped in the bathroom, unconscious, with a left-sided weakness. the patient had a head computed tomography (ct) scan which revealed a clot in the right middle cerebral artery. the patient was determined to not be a candidate for a thrombectomy. the day following the stroke, the patient continued to have weakness in the left upper extremity, however the left lower extremity weakness had resolved. on support day 50, the patient was alert and able to sit out of bed in a chair. there were no neurological deficits noted from the stroke, only weakness noted to the right arm. the patient maintained a status 2 for high medical urgency for heart transplant. on support day 51, it was noted that there were suction alarms on the impella when the patient¿s mean arterial pressure dropped below 70 mmhg. the suction alarms continued intermittently throughout the following night and the patient was weaned to performance (p) level 4. on support day 54, the patient was noted to be status 7, temporarily inactive, for transplant until cleared by neurology. the patient¿s left arm and hand function were slowly returning. it was additionally reported that over the weekend, the impella had alarms for placement signal not reliable, and eventually the placement signal sensor failed. the medical team was advised to monitor motor current and pump flows on the device, and to monitor pump position with chest x-rays and transthoracic echocardiogram. there were no plans to replace the impella 5.5 due to the patient not being cleared by neurology for surgery. on support day 62, the patient experienced a hemorrhagic conversion of the previous ischemia stroke which was noted on ct scan. the patient experienced a loss of function on the left side and was minimally responsive. on support day 63, the patient regained some left upper extremity strength. it was noted that anticoagulation therapy was on hold due to the hemorrhagic stroke. the patient was delisted from the transplant list until cleared by neurology. on support day 65, the patient experienced generalized pain which the patient stated he believed it was caused by being bedridden since the stroke last week. the patient¿s left side strength was noted to be increasing. a recent ct scan revealed that there was no worsening changes observed. systemic anticoagulation therapy was restarted. placement signal remains absent on impella after sensor failure. on support day 78, the patient because experiencing new weakness in the left arm and hand. ct scan was performed which did not reveal additional changes. the patient was on anticoagulation therapy, and the patient was taken emergently to radiology to evaluate for a thrombotic event. on support day 79, the patient could not make sense of words, was mumbling, and was in and out of delirium. the patient did not have movement in the left upper extremity. two days later the patient was transferred to palliative care, and care was withdrawn. |  |  |  | Investigation summary: no product was returned for investigation. hematoma: hematoma noted on right armpit. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. nerve damage/pain): the cause of the injury was not determined due to insufficient clinical details. thrombosis/stroke: in order to make a cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. optical sensor issue: clinical reported placement signal drift noted prior to placement signal not reliable alarm on (b)(6). the data logs returned were from on (b)(6). the logs show a sudden drop in the placement signal by ~50 mmhg followed by the downward trend of placement signal with an increase in pump flows (without a p-level change) for over 7 hours before placement signal went out of range and #1048 placement signal not reliable alarm was triggered. before going out of range, there was a step back up in the placement signal which correlates with user applied adjustment. this occurred on (b)(6) which was day 53 of support. the 5s parameters before this event were stable. the pattern identified is characteristic of sensor wear. the cause of the optical sensor issue was determined to be sensor wear, based on data log review. device history lot device lot: 1946932. device history batch: subcomponent (0550-2501, optical sensor) lot: 2025534837. device history review: device with sn: (b)(6) passed all post sterile inspection checks. device traveler review: q1) any qns/mrrs/reworks associated with the complaint device: no. q2) were there any failures of the device to meet the requirements for manufacturing testing, qualification, and inspection? no. q3) were there any other product malfunction complaints in this device lot related to this failure mode? no. q4) subcomponent inspection review: did any subcomponents fail incoming lot inspection? no. q5) subcomponent lot review: were there any other product malfunction complaints in this subcomponent lot related to this failure mode? yes, there was one other pm in the subcomponent lot#: 2025534837. | A.4 and a.5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.5 additional information was received and added. no product was returned for investigation. hematoma: a hematoma was noted on the right armpit. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. nerve damage/ nerve injury: the cause of the injury was not determined due to insufficient clinical details. the device history review was completed and the device passed all post sterile inspection checks. | B2. is death has been updated to reflect patient outcome. the exact date of death is unknown but has been captured as the date of explant. if additional information becomes known, a supplemental report will be submitted. b3. date of event, has been updated. b5. is being updated to include related device information not included in the initial report. the revised b5 provides a complete event narrative. h1: type of reportable event updated to death based on new information provided. h6 health effect - clinical code, health effect - impact code, and medical device problem code have been updated to reflect the complete correct coding.
5510 | 1220648-2025-31519 | 23073059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23073059&pc=OZD | 03/18/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 03/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025606961 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella rp pump was implanted in a patient for circulatory support. while on support there were placement signal issues. placement was verified and audio was silenced. support from the pump was continued. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. data logs show intermittent placement signal not reliable (snr) alarms starting on (b)(6) 2025 and ending on (b)(6) 2025. at the time of these alarms, the placement signal goes out of range with snr dropping to zero, gain decreasing, lamp increasing, and contrast barcoding. this issue was intermittent meaning that the placement signal was able to recover temporarily until the next alarm. upon this data log review, it is apparent that the console is the potential cause of the issue. please refer to (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5511 | 1220648-2025-31520 | 23073214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23073214&pc=OZD | 07/25/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 09/08/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1658143 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 support had placement signal not reliable alarm. the aic is being requested to be returned for evaluation. | [5 empty] | The impella device has been received from the customer. upon completion of the investigation, a supplemental mdr will be filed.
5512 | 1220648-2025-31521 | 23073354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23073354&pc=OZD | 03/22/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 03/22/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580303 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that impella 5.5 was in good position, but it looked like the impella was getting false impella in aorta alarms. the rn was advised to disable the placement signal monitoring. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. clinical details reported that on the first day of support, position in aorta alarms were triggering even though echo confirmed proper pump placement. data log review confirmed many positioning alarms during the first 30 hours of the case. there were also 42 placement signal not reliable (snr) alarms, mostly in the first half of the case. when these alarms were triggering, snr was quite variable, dropping when there was a rapid and relatively random oscillation between a larger and smaller amplitude raw/calculated optical signal waveform. this was causing for the calculated placement signal to be out of synch with the motor current, then causing false position in aorta alarms. on 29/mar/25, most all psnr and positioning alarms resolved when the ps stepped down in pressure measurements, though there was not a clear trend noted through the case that snr was dropping/improving at specific pressure measurement values. clinical details did not report anything that would explain the abnormal waveform/rhythm seen or anything to align with psnr resolving on 29/mar/25. returned product was investigated and there were no visual abnormalities. the pump's 5s parameters were in specification when connected to a console, and it passed both laser continuity and uson testing. since limited clinical detail was provided, returned product did not exhibit any optical signal issues, and there were no definitive causes seen in data logs, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5513 | 1220648-2025-46547 | 23073619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23073619&pc=OZD | 08/18/2025 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578998 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient was brought to post-operative recovery following a primary percutaneous coronary intervention with impella cp support. after arriving to the recovery area, the patient became hypoxic, pressures dropped, the patient coded, and the decision was made to bring the patient to the cardiac catheterization laboratory. it was noted that when the impella cp was first placed, initial flows were 3.3 liters per minute (l/min). however, the flows decreased to 0.5 l/min. the patient was then cannulated for extracorporeal membrane oxygenation and a swan was placed. the patient was prepared for transport to another hospital. it was then noted there was a hematoma and profuse bleeding around the impella cp repositioning sheath. the physician tried to readjust the angle with gauze, reapplied forward tension sutures, pushed the blue suture hub further in, and cut wings off the blue suture hub to advance in skin tract. the patient continued to bleed and the hematoma was unresolved. the decision was made to explant the impella cp and replace it with a new impella cp. during insertion of the new impella cp, the impella sheath kinked from the hematoma created from the first implant. a 14f edward¿s introducer placed, but the impella cp was unable to get past the kink. a 14frx25cm sheath placed and the sheath kinked. the decision was then made to abort the impella cp. the patient outcome is unknown. this complaint involves one impella cp device and two introducers: pump 1, impella cp with serial number (b)(6). 14fr introducer with lot number s9153522, used with pump 1 impella cp with serial number (b)(6). 14fr introducer with lot number s9467412, used with attempted implant of pump 2, impella cp. this report specifically addresses impella cp with serial number (b)(6). separate reports will be submitted for the other devices associated with this complaint. refer to section h.10 for the related report numbers. | [5 empty] | Product complaint #(b)(4). investigation summary: impella cp sn (b)(6) returned. hypotension: clinical details report that the patient's pressures tanked so they were put back on impella cp support. the cause of the clinical symptom cannot be determined as insufficient clinical details were provided. low pump flow: clinical details report that pump flows suddenly dropped from 3.3 l/min of flow to 0.5 l/min, so ecmo was cannulated and a swan was placed. later that same day, there was profuse bleeding from the access site. the pump was explanted due to the bleed. data logs were reviewed, and the low pump flows were confirmed. while attempting to run at p-9, there was a significant drop in mc and flows. flows were also abnormal while running at p-2 the rest of the case. during the time where pump flows dropped significantly at p-9, suction #73 triggered, indicating continuous suction conditions with low mc values. the ps was also abnormal throughout the case. at case start, it appears that the optical sensor had a 30 mmhg pressure applied during the calibration procedure, resulting in a shift down from that point forward, though the timing couldn't be confirmed as imc logs weren't available. there was also periods of very low pulsatility and dramatic shifts in the pressure readings. this could be explained by the report that the patient's pressures were unstable, however, it's unclear. returned product was investigated and there were no visual abnormalities that would explain the low flows: no biomaterial, cannula kink, or impeller blade damage. the pump was connected to a console and run in a bucket, and pump flows were within target range for p-9. while running in a bucket, the ps readings were -20 mmhg, likely confirming the optical sensor calibration error. that being said, it was determined to be unrelated to the reported low flows, considering that suction alarm #73 is triggered purely based on mc values. though it was confirmed that there was no deficiency of the product, insufficient clinical details were provided to determine the cause of the low flows, as no troubleshooting methods were provided beyond starting the patient on ecmo. for these reasons, the cause of the low pump flows could not be determined. access site adverse event/major bleed/hematoma: clinical details report bleeding and a hematoma at the access site 1 hour after the repositioning unit was advanced. several different methods were used to try and troubleshoot but none worked. this included adjusting insertion angle, applying new forward tension sutures, and shifting the placement of the repositioning unit in the access site. acts at the time were unknown. returned product was investigated and it was confirmed the butterflies were cut off of the blue suture hub. it was deemed to be unrelated to the cause of the access site bleed, as it was reported this was done to troubleshoot. no other abnormalities were noted with the repo unit. since insufficient clinical details were provided, the cause of the access site adverse event could not be determined. device history lot: device lot: 1863061. device history batch: subcomponent: na. device history review: pump sn (b)(6) passed all post-sterile inspection checks. | D.9 additional information was received that the pump was returned but the exact date is unknown. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
5514 | 1220648-2025-46548 | 23073883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23073883&pc=OZD | 09/06/2024 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076153 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Literature | Health Professional | User facility | Company representation |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp encountered "placement signal not reliable" alarms. placement monitoring was disabled along with the audio, and the patient was monitored via motor current and echocardiogram for placement. | [5 empty] | A.1, a.4, a.5, and a.6 are unknown. the automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. literature citation: khattak, s., shamkhani, w., maqableh, g. m., al-shatanawi, a., & khan, s. q. (2025). tissue plasminogen activator to resolve elevated purge system pressure in a catheter-based ventricular assist device. jacc: case reports, 30(18), 103890. https://doi.org/10.1016/j.jaccas.2025.103890.
5515 | 1220648-2025-31522 | 23074167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074167&pc=OZD | 04/01/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 04/01/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612029 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. after 6 days of support, there were optical signal issues. the team troubleshot effectively and did not need intervention. the pump remained on until the pump was successfully weaned and explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella 5.5 returned for evaluation. upon visual inspection of the pump, no abnormalities were noted. optical 5s parameters of returned pump were measured and no abnormalities were noted. pump passed laser continuity testing and no biomaterial was present on returned pump. some dried dextrose was present on returned pump. data logs were reviewed and low signal not reliable at beginning of case with some variability in contrast and all other optical parameters remaining stable. clinical details noted that "placement signal low" alarms occurring during the case and ao offset was adjusted. no problem was found with the returned pump. upon review of data logs and returned pump, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5516 | 1220648-2025-46549 | 23074223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074223&pc=OZD | 08/18/2025 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026704092 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 76 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that during impella cp support, the patient lost pulses down the leg where the impella cp was inserted. vascular surgery was consulted. the vascular surgery team performed left femoral artery cut down and repair. the decision was made to remove the impella cp and replace with an impella 5.5 inserted in the axillary artery. the impella cp was successfully explanted and replaced by the impella 5.5. the patient remained on the impella 5.5 until successfully weaned and explanted. additional information was received. computed tomography was negative. the patient did have some weakness but cerebrovascular accident was not confirmed on testing. the patient recovered and was discharged home. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
5517 | 1220648-2025-31523 | 23074284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074284&pc=OZD | 04/05/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 04/05/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619570 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient with acute myocardial infarction and cardiogenic shock. during support, there was an optical signal issue. the nurse reported that there was a "placement signal unreliable" alarm and that no kinks were found. intermittent placement signal and left ventricle signal failures were also encountered. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: after 3 days on support, placement signal not reliable (psnr) alarm reported with intermittent placement signal failure. no kinks were noted by care team. pump remained implanted for 3 additional days. data logs were not recovered in their entirety as only rt logs through (b)(6) 2025 at 0505 were recovered; however, psnr complaint occurred on (b)(6) 2025. the pump was returned and inspected. no issues noted with 5s parameters nor electrical issues. psnr did not reproduce. pump was run in the lab in a bucket of water where the ps ran at 37mmhg upon pump start. uson testing showed a consistent 36-38mmhg elevated reading by the pump relative to the pressure environment, indicating the sensor is appropriately responding to the change in uson pressures. no offsets were applied during lab run. the cause of the optical signal issue was not determined since complete data logs were not returned and psnr could not be reproduced during investigation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5518 | 1220648-2025-46550 | 23074516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074516&pc=OZD | 05/31/2025 | 09/15/2025 |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. - 1220648 |  |  | US |  |  |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Injury | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had impella 5.5 support published in june of 2025. the patient had a coronary artery bypass graft (cabg) and was on impella and extracorporeal membrane oxygenation support. the patient sat upright in bed on day one post operative for the cabg, and bleeding was observed from the chest tubes. the patient was transfused and found to have a left ventricle (lv) rupture/perforation. the patient survived the surgical intervention. this was at the site of the prior infarcted lv tissue. | [5 empty] | Section a - patient information is unknown. d.4 - serial number, lot number, expiration date and primary udi number are unknown. section e - initial reporter information is unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. literature citation: hall, j., raza, h., lee, s., bryce, n., & abrol, s. (2025). successful treatment of left ventricle inferior wall perforation and rupture associated with an impella 5.5: a case report. reports, 8(2), 98. https://doi.org/10.3390/reports8020098. | E.1 added initial reporter facility name, address and phone number since the initial report was submitted. investigation summary: no product was returned for investigation. heart injury: the cause of the injury was not determined due to insufficient clinical details. device history review: the device serial number was not returned in the clinical details.
5519 | 1220648-2025-31524 | 23074524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074524&pc=OZD | 04/08/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 04/08/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619160 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support in escalation from an impella cp. the 5.5 pump was found to have optical signal loss; however, did not prompt any intervention or explant. the pump remained on to support the patient. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the clinical stated there was a drift in the placement signal about 35 days into support. however, due to lack of data logs and product returned, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5520 | 1220648-2025-31525 | 23074731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074731&pc=OZD | 04/11/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 04/11/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616580 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support, shockwaves were used and the placement signal went away. thirty seconds later there was a placement signal 'not reliable' alarm. the case continued with adequate flows from the impella and with the motor current monitored. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. clinical details report that while using shockwave in the lad, the impella placement signal went out for the remainder of the procedure. the team felt that shockwave was operated sufficiently far from the pump, but no measurement was provided. the pump was not removed due to the complication. neither the pump nor data logs were returned for investigation. since neither data logs nor the pump were available for investigation and a measurement of the distance between shockwave and impella at the time the placement signal went out was not provided, the root cause of the optical signal could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5521 | 1220648-2025-31526 | 23074821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074821&pc=OZD | 04/17/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 04/17/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609435 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. the pump lost placement signal during shockwave use; however, did not require intervention or explantation. the pump remained on until the pump was successfully weaned and explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: clinical details state that the placement signal went out after using shockwave during a pci. the pump was not removed due to this issue. data log review shows the placement signal going out of range about an hour into the case. signal not reliable (snr) drops to zero, contrast decreases, and gain and lamp increase. the pump was returned and evaluated. a 5s parameter check showed snr near zero and contrast below spec. the pump passed laser continuity testing. this evidence supports a sensor break. per cp ifu (0048-9007 rs), cp use with shockwave intravascular lithotripsy (ivl) catheter at a distance of less than 20 mm between the optical sensor and ivl device may interfere with or damage the impella optical sensor. the distance between shockwave and the optical sensor in this case is unknown. the root cause of the optical signal issue was most likely a damaged sensor since the loss of placement signal seen in the data logs went out after using shockwave according to clinical details and testing of the returned product showed evidence of a damaged sensor. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5522 | 1220648-2025-46551 | 23074924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074924&pc=OZD | 09/06/2024 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658693 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Literature | Health Professional | User facility | Company representation |  | 54 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella cp developed a large parenchymal hemorrhage. the patient underwent an emergency craniotomy, was transfused two units of blood, had heparin withheld, and was treated with protamine. the patient did not recover well and the patient's family chose to withdraw care. the patient ultimately expired. | [5 empty] | A.1, a.4, a.5, and a.6 are unknown. b.2 exact date of death is unknown. e.1 initial reporter is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. literature citation: khattak, s., shamkhani, w., maqableh, g. m., al-shatanawi, a., & khan, s. q. (2025a). tissue plasminogen activator to resolve elevated purge system pressure in a catheter-based ventricular assist device. jacc: case reports, 30(18), 103890. https://doi.org/10.1016/j.jaccas.2025.103890.
5523 | 1220648-2025-31527 | 23074987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23074987&pc=OZD | 04/25/2025 | 09/17/2025 |  | 09/17/2025 |  |  | 04/25/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623937 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. during support, there were false 'impella in aorta' alarms. position of the pump was verified via an echocardiogram allowing the placement monitoring to be disabled. the patient was successfully weaned from the pump and the device was explanted. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: clinical details state false impella in aorta alarms after ~37 days of 5.5 support. no product/logs were returned. the root cause of the optical signal issue was not determined since no product/logs were returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5524 | 1220648-2025-46552 | 23075297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23075297&pc=OZD | 08/19/2025 |  |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627640 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Ischemia | Ventricular Fibrillation | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented and was found to be hypotensive and had witnessed cardiac arrest with immediate cardiopulmonary resuscitation including one round with subsequent return of spontaneous circulation. the decision for an impella cp was made. the impella cp was independently placed via the right femoral artery. the patient was noted to be having refractory ventricular fibrillation during the procedure requiring defibrillation x 15. lidocaine and amiodarone drops were infusing. mean arterial pressure was stable with the impella on auto and levophed was infusing. rhythm was stabilized. it was noted that this was a difficult case. the patient remained on significant pressor and inotrope with impella support through the day. attempts were made to de-escalate the impella due to limb ischemia and hemolysis. there was critical limb ischemia in the right lower extremity as confirmed by doppler. there was gross hemolysis of lab specimens to the extent that the lab values would not result and the patient had also stopped making urine. neuro exam was questionable. the family was presented with the option to escalate to an impella 5.5. however, they elected to withdraw care. the patient was made do not resuscitate. the family was at the patient's bedside for withdraw of care and discontinuation of impella. the patient expired. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. the pump was not returned for investigation. clinical symptoms (ischemia/ventricular fibrillation): the root cause of the injury could not be determined due to insufficient clinical details. mechanical interaction with blood (hemolysis): according to the clinical details, hemolysis was reported on 20-aug. the data log does not show any major suction alarms or positioning issues during the case run. the placement signal showed some deterioration trend before the reported hemolysis and later some improvement. no motor current spikes were reported during the case run. the cause of the hemolysis issue was not determined, as the product was not returned for investigation and sufficient clinical information was not provided. device history review: the device passed all the post sterile inspection checks.
5525 | 1220648-2025-31528 | 23075441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23075441&pc=OZD | 11/06/2023 | 10/24/2025 |  | 09/17/2025 |  |  | 11/07/2023 | 09/17/2025 |  |  | 05/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024364955 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the placement signal of an implanted impella cp pump became unreliable during patient support. it was noted that the aortic placement signal was significantly lower than the aortic line by 30 points. despite the observation, support with the pump was maintained without hinderance. no patient harm has been reported. | [5 empty] | D6a: corrected the explant date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for investigation. data logs show the 5s parameters to be stable throughout the case until a spike in placement signal, contrast, and gain 8.5 days into support, but they immediately return to normal values. there are placement signal low and suction alarms at the same time at which the pump's position is unknown, however, there are no accompanying positioning alarms. the root cause of the optical signal issue cannot determined since product was not returned. all pertinent information available to abiomed has been submitted at this time. | G3: corrected date.
5526 | 1220648-2025-46553 | 23075464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23075464&pc=OZD | 08/19/2025 |  |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580943 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support of the impella 5.5, there was a hematoma at the axillary site after sneezing and now complains of pain. five units of packed red blood cells were given over the weekend. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for investigation. hematoma : the cause of the injury is most likely use related since hematoma at axillary site after sneezing. device history review: the device passed all the post sterile inspection checks.
5527 | 1220648-2025-46554 | 23075607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23075607&pc=OZD | 08/19/2025 |  |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660016 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella 5.5 support, a hematoma was found on computed tomography scan. the patient was switched to argatroban and one unit of packed red blood cells were given. | [5 empty] | A.4 and a.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Investigation summary: no product was returned for investigation. hematoma: hematoma was noted in computed tomography scan. the location of hematoma and the treatment were unclear from the provided clinical details. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history review: the device passed all the post sterile inspection checks.
5528 | 1220648-2025-46555 | 23075662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23075662&pc=OZD | 08/19/2025 |  |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691960 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and while on support, the team was concerned that the patient was septic. it was also reported that the patient had poor distal perfusion on the impella side. a decision was made that the patient does not hemodynamically need impella. the impella was removed to restore distal flow. additionally, the vascular surgery team did a cut-down for impella removal and patched the vessel. the team also proactively attempted to de-clot, but no clot was observed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was returned for investigation. ischemia/ sepsis: the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5529 | 1220648-2025-46556 | 23076297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23076297&pc=OZD | 08/19/2025 | 11/10/2025 |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702633 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | Pericardial Effusion | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient presented to the emergency department and was diagnosed with st-segment elevation myocardial infarction. percutaneous coronary intervention was completed with an intra-aortic balloon pump (iabp). the patient remained hemodynamically unstable, and a decision was made to place an impella cp. the iabp was removed from the right femoral artery (rfa) and impella peel-away sheath was inserted into the rfa. the patient went into pulseless electrical activity and advanced cardiac life support was initiated. the impella cp was prepped and primed and inserted into left ventricle and started in auto. return of spontaneous circulation was achieved. the physician had an echocardiogram completed in the cardiac catheterization lab that revealed pericardial effusion and pericardiocentesis was performed. the patient was taken to intensive care unit. the nurse was unable to doppler the distal pulses to the right lower extremity and the medical doctor did not want any further intervention. the family decided to make the patient do not resuscitate and the patient expired. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the reported failure mode has been completed. no product was returned for investigation. cardiac arrest/ ischemia: the root cause of the injury was unable to be determined due to insufficient clinical details. pericardial effusion: the cause of the clinical symptom was not determined due to insufficient clinical details. device history review: the device passed all the post sterile inspection checks.
5530 | 1220648-2025-31529 | 23076577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23076577&pc=OZD | 12/01/2023 | 09/17/2025 |  | 09/17/2025 |  |  | 12/01/2023 | 09/17/2025 |  |  | 09/30/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410756 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during support, the optical sensor signal was lost. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was an air gap from between the automated impella controller and the patient cable plug. all pertinent information available to abiomed has been submitted at this time.
5531 | 1220648-2025-31530 | 23076616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23076616&pc=OZD | 12/17/2023 | 09/17/2025 |  | 09/17/2025 |  |  | 12/17/2023 | 09/17/2025 |  |  | 12/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421432 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted for circulatory support. during use, the optical sensor signal was lost. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5532 | 1220648-2025-46557 | 23076742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23076742&pc=OZD | 08/17/2025 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026702827 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | Tachycardia | Renal Impairment | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support encountered possibly fever and infection. the patient had elevated temperatures so blood cultures were collected and antibiotics were started. additionally, the patient experienced several runs of supraventricular tachycardia over the night and morning which required being shocked three times. also, the patient's kidneys were of concern so continuous renal replacement therapy was started. however, after five days of support, care was withdrawn and the patient ultimately expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary: no product was returned for investigation. acute kidney failure/ fever: the cause of the injury was not determined due to insufficient clinical details. tachycardia/ infection: in order to make a root cause determination, the clinical details would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: device passed all post sterile inspection checks.
5533 | 1220648-2025-46558 | 23076832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23076832&pc=PYX | 08/19/2025 |  |  | 09/17/2025 |  |  | 08/19/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026683288 |  | 000000000001000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 66 KG |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support and was to be implanted with an impella rp flex for bipella support. the bipella case was aborted due to patient anatomy. there was no reported patient harm. | [5 empty] | D.2b revised as it was previously entered incorrectly on the initial report that was submitted. g.4 revised as it was previously entered incorrectly on the initial report that was submitted. investigation summary: the investigation has been completed. no product was returned for investigation. delivery issue: the cause of the issue was determined to be patient condition because due to patient anatomy and the doctor was unable to get any catheter into pulmonary artery as extracorporeal membrane oxygenation catheters are in the way and unable to be moved. the impella rp flex insertion was aborted. device history review: the device passed all the post sterile inspection checks. | A.5 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
5534 | 1220648-2025-46514 | 23077046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23077046&pc=OZD | 08/18/2025 |  |  | 09/17/2025 |  |  | 08/18/2025 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688065 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp pump was implanted to support a patient admitted with st elevated myocardial infarction/acute myocardial infarction/cardiogenic shock. the patient developed a fever and was treated with antibiotics. the patient also suffered from ventricular tachycardia for which the patient was defibrillated. the non-impella site had a bleed which was managed with transfusions of red blood cells. after four days of support, the patient was successfully weaned from the pump and survived. additional information was received. the clinical consultant reported that they did not allege that the pump cause these harms but rather the myocardial infarction that the patient had. they did not save the pump for investigation as they were not concerned with the impella. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental report will be filed. | No product was returned for investigation. the cause of the fever was not determined due to insufficient clinical details. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of the bleeding is determined to be patient condition because the bleeding is from other locations unrelated to impella. the device passed all post sterile inspection checks.
5535 | 1220648-2025-31531 | 23077134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23077134&pc=OZD | 03/12/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 03/12/2024 | 09/17/2025 |  |  | 11/30/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435475 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a ¿placement signal not reliable¿ alarm triggered during impella cp support. it was noted that the placement signal readings were outside of normal parameter. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was not determined as the failure did not reproduce with the returned product and since data logs were not returned for investigation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5536 | 1220648-2025-31532 | 23077167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23077167&pc=OZD | 03/21/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 03/21/2024 | 09/17/2025 |  |  | 05/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365491 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, there was a placement signal alarm. the positioning was confirmed with echocardiography. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were investigated. the root cause of the placement signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5537 | 1220648-2025-31533 | 23077283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23077283&pc=OZD | 03/27/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 03/27/2024 | 09/17/2025 |  |  | 07/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377087 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, the placement signal on the automated impella controller (aic) disappeared. there was no measurable pressure in the aorta and no waveform was visible. the positioning was confirmed with echocardiography. neither the console nor the pump were replaced as a result of the failure. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were not returned. the root cause of the optical signal issue was determined to be a damaged sensor due to interaction with shockwave device. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5538 | 1220648-2025-31534 | 23077487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23077487&pc=OZD | 04/11/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 04/11/2024 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469289 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, there were placement signal issues. during device initiation there were multiple alarms, and low placement signal issues were observed. multiple attempts to reposition the device were unsuccessful. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was patient condition related, based on returned product review. all pertinent information available to abiomed has been submitted at this time.
5539 | 1220648-2025-31535 | 23077918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23077918&pc=OZD | 04/18/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 04/18/2024 | 09/17/2025 |  |  | 10/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420788 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and there were placement signal issues. when the shockwave was used, the placement signal went out, and monitoring disabled was showing. the motor current was still pulsatile, and flows were 3.4-3.6. support continued. the pump was eventually weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. product and data logs were not returned. as per clinical, placement signal was lost with placement signal not reliable alarm for cp pump during shockwave use. the cause of the optical signal issue was most likely a damaged sensor due to shockwave interaction. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5540 | 1220648-2025-31536 | 23078154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23078154&pc=OZD | 04/27/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 04/27/2024 | 09/17/2025 |  |  | 11/30/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429380 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a false position in aorta alarm appeared during impella cp support. it was noted that a percutaneous coronary intervention was being performed at the time. positioning was verified and the device was repositioned, but the alarm persisted. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the field data logs were reviewed and a motor current spike was found at the time that the position in aorta alarm triggered. all pertinent information available to abiomed has been submitted at this time.
5541 | 1220648-2025-31538 | 23078381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23078381&pc=OZD | 07/05/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 07/05/2024 | 09/17/2025 |  |  | 01/31/2026 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509026 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, the optical signal became unreliable. the pump was successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the priming issue and optical signal issue has been completed. the root cause of the optical signal issue was most likely patient condition related. all pertinent information available to abiomed has been submitted at this time.
5542 | 1220648-2025-31537 | 23078390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23078390&pc=OZD | 05/19/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 05/19/2024 | 09/17/2025 |  |  | 12/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456430 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal was lost during impella cp support. the lost signal was thought to be due to priming and insertion issues, but support continued. the device was successfully weaned and explanted. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the priming issue and optical signal issue has been completed. the root cause of the priming issue was determined to be use-related due to improper technique when inserting the purge disk. the root cause of the optical signal failure was determined to be a corrupted eeprom as evidenced by the calibrated g0 value being zero. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5543 | 1220648-2025-31539 | 23078565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23078565&pc=OZD | 07/26/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 07/26/2024 | 09/17/2025 |  |  |  |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524219 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, there were placement signal issues. the placement signal was absent with dash marks. the impella support continued and the placement signal returned. the pump was successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was not returned. upon review of data logs, it is apparent that the automated impella controller is the potential cause of the issue. all pertinent information available to abiomed has been submitted at this time.
5544 | 1220648-2025-31540 | 23078672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23078672&pc=OZD | 09/03/2024 | 09/17/2025 |  | 09/17/2025 |  |  | 09/03/2024 | 09/17/2025 |  |  | 12/31/2025 |  | 09/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450336 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, there were placement signal issues. the pump was successfully weaned and explanted. no patient harm has been reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of the optical signal issue was most likely patient condition related as no issues with the returned pump were noted. all pertinent information available to abiomed has been submitted at this time.
5545 | 1220648-2025-46509 | 23061637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23061637&pc=OZD | 09/03/2025 | 10/10/2025 |  | 09/16/2025 |  |  | 09/03/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1435251 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that they encountered automated impella controller (aic) issues. it was reported that the aic was powered on three times, with a hard reset two times. each time the console failed to boot from the hard drive. as a result, the complainant contacted the field service to request a loaner and the original console was left with the biomedical department. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | H6: component code added. h6: h6: updated codes for type of investigation, investigation findings, and investigation conclusions. h11: updated based on the results of the investigation. investigation summary data analysis: a review of the console logs confirmed the reported failure mode. device analysis: the device was returned for evaluation. visual inspection confirmed the contamination on the pbm, which impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5546 | 1220648-2025-46510 | 23061890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23061890&pc=OZD | 08/18/2025 | 09/15/2025 |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682012 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support, the patient was moved onto the stretcher and went into sustained ventricular tachycardia (vt) and was shocked out of it into normal sinus rhythm. the impella positioned was checked with point-of-care ultrasound and position in the mid ventricle still at 4.7 centimeters. no alarms on the impella were noted. the plan was to rest and optimize until high-risk percutaneous coronary intervention. in the morning the (b)(6), drained 300 milliliters from the jp at the site. bival was on hold and bleeding has since stopped. the patient had vt arrest and was intubated. lidocaine drops were started and one unit of packed red blood cells were given. the patient was taken emergently for chloride and drug-eluting stent x 2 to the right coronary artery. the patient was stable during the procedure and the plan is to do percutaneous coronary intervention of the left coronary artery. the impella was explanted on the fourth day of support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | The investigation of the reported failure mode has been completed. no product was received for investigation. tachycardia : the root cause of the injury was unable to be determined due to insufficient clinical details. major bleed : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks.
5547 | 1220648-2025-46511 | 23061927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23061927&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645656 |  | 000000000001000080 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male | 77 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella cp had a purge system open alarm. a leaking purge sidearm was found. the impella cp was replaced with a new impella cp and there was no patient harm. | [5 empty] | The investigation of the reported failure mode has been completed. the pump and purge cassette were returned for investigation. purge leak-pump: no visual damage was noted on the returned product. the pump was then primed, and a purge leak was observed from the sidearm filter hole. the filter was further examined, no damage was seen at air hole and weld line. data log review was not required for this case. as per the clinical details, purge side arm filter leak was observed during pump use. the cause of the purge leak was sidearm filter damage where purge fluid leaked out for the filter vent hole. device history review: the device passed all the post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed.
5548 | 1220648-2025-31472 | 23061932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23061932&pc=OZD | 04/21/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 05/29/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1292759 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was implanted for circulatory support. review of the data logs showed that the console experienced an optical signal issue. it is unknown how the issue was resolved; however, no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs confirm at the time of the contrast oscillating. console was not returned for investigation. the root cause of the intermittent unreliable placement signal and oscillating contrast could not be determined as console was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5549 | 1220648-2025-46512 | 23062037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23062037&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684261 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced gastrointestinal bleeding. in response, heparin was withheld. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | Product complaint # (b)(4). investigation summary- no product was returned for investigation. major bleed: gi bleed was reported and heparin was stopped as per clinical details. the act values or cause of the bleeding were provided in the clinical details. the root cause of the injury was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. device history lot- device lot: 1939798. device history batch- subcomponent lot: n/a. device history review- device with sn: (b)(6) passed all post sterile inspection checks.
5550 | 1220648-2025-46513 | 23062127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23062127&pc=OZD | 08/18/2025 |  |  | 09/16/2025 |  |  | 08/18/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675729 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male | 60 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a pediatric patient on impella 5.5 support for the hospitalization and plan for care. while on support, the patient continued to have elevated high lactate dehydrogenase which was a concern for hemolysis. the decision was made to explant the impella 5.5 and place the patient on a different left ventricular assist device for more durable support. the patient survived the surgery. | [5 empty] | A.5 - a.6 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the reported failure mode has been completed. hemolysis: the product was not returned for investigation, and data logs were not provided. the cause of the hemolysis could not be determined due to insufficient clinical details, no product returned, and no data logs provided. device history review: the device passed all the post sterile inspection checks.
5551 | 1220648-2025-46515 | 23062440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23062440&pc=OZD | 08/17/2025 | 09/16/2025 |  | 09/16/2025 |  |  | 08/17/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683253 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 57 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented feeling hot and impending unwell. the patient was taken for left heart catheterization and became acutely hypotensive. moderate to severe mitral regurgitation was noted on echocardiogram. the decision was made to place an impella cp for cardiogenic shock. the patient was on support for a total of 60.18 hours. it was noted that the impella cp was weaned and explanted, and the patient looked great hemodynamically. however, groin run offs showed dissection in both the common femoral artery and the right femoral artery from groin interventions at the previous hospital and left femoral artery probably related to this admission. the patient was sent to the operating room for patch angioplasty. the patient survived. | [5 empty] | A.5 - a.6 are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of the reported failure mode has been completed. vascular dissection : the cause of the injury was not determined since product was not returned and insufficient clinical details provided. device history review: the device passed all the post sterile inspection checks.
5552 | 1220648-2025-31473 | 23062441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23062441&pc=OZD | 10/17/2023 | 09/16/2025 |  | 09/16/2025 |  |  | 10/17/2023 | 09/16/2025 |  |  | 06/30/2025 |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377360 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was implanted to support a patient undergoing a high-risk percutaneous coronary intervention. during support there were placement signaling issues while doing shockwave (3x12mm) of the proximal of the left anterior descending artery. there was no reported harm or injury to the patient. | [5 empty] | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the cause of the issue was a most likely a damaged sensor due to interaction with shockwave device. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
5553 | 1220648-2025-46320 | 23062637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23062637&pc=OZD | 08/16/2025 |  |  | 09/16/2025 |  |  | 08/17/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653890 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that a patient presented to the hospital with an anterior in st elevation myocardial infarction after feeling uncomfortable for several days following an unrelated medical procedure. the patient was found to have a flush occlusion of their left anterior descending artery (lad) and ejection fraction of 20%. the physician then elected to place an impella cp to treat the lad. it was reported that while the patient was in the operating room there was an alarm for air in purge and low purge pressure. it was noted that the medical team attempted to de-air the purge without success. the purge cassette was changed and the low purge pressure and air in purge alarms were resolved. it was also reported that on (b)(6) 2025, the patient experienced an episode of atrial fibrillation with no information on interventions that were performed to resolve the issues. additionally, the day before the patient had bleeding from the groin site and as a result of the bleeding heparin was withheld and manual compression was held. the bleeding complication persisted, and the patient was transfused with two units of red cells and a purse string closure was placed at the impella access site. the patient was then escalated later to an impella 5.5 for increased hemodynamic support. the patient experienced a hematoma at the axillary graph anastomosis site that was tunneled with no further complications and bleeding at the impella access site that was sandbagged to resolve. additionally, there were false alarms for impella in the aorta when the patient coughed. the patient¿s care was withdrawn, and the patient expired on impella 5.5 support. this complaint involves one impella cp, one purge cassette used during impella cp support, one impella 5.5, and one automated impella controller which alarmed during impella 5.5 support. purge cassette lot number is unknown. impella cp with serial number: (b)(6), impella 5.5 with serial number: (b)(6), automated impella controller with serial number: (b)(6). this report specifically addresses the impella cp. separate reports were submitted for the other devices associated with this complaint. refer to section h.10 for related report numbers. | [5 empty] | A.4 and a.5 are unknown. based on the information available, the impella cp cannot be excluded as a possible contributing factor to the patient¿s outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Medical safety review was completed and noted the following: the event may have cascaded from the impella cp. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the outcome of the bleed for the patient requiring two units of prbcs and is a serious injury (as well as the arrhythmic event) that should be reported. the patient was escalated to an impella 5.5 and experienced a hematoma at the access site that was sandbagged. the patient would eventually expire on the impella 5.5. regarding the automated impella controller (aic) sensing issue there was no report or indication of aic-to-aic exchange contributing to a disruption in or lack of support. the patient expired for reasons unrelated to the aic (this type of event is a product malfunction). impella 5.5 pump positioning was confirmed, and support had continued. there is no reason to believe that this hematoma is associated with the demise outcome given the information at hand. however, the impella 5.5 device was present at the time of the demise outcome and dissociation cannot be made. however, the patient's underlying condition (cardiogenic shock, inability to wean from ecmo and not a transplant candidate) most likely contributed to the outcome and the care was withdrawn which appears to have led to the demise outcome. medical safety officer reviewed the event and noted the same; the impella 5.5 device cannot be dissociated from the outcome and impella cp is a serious injury and aic is a product malfunction type of reportable events. change in reportability: after review of the case by the medical safety officer, it was determined the impella cp is a serious injury type of reportable event. sections b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes) have been updated, corrected accordingly. the event story also involves two other devices. refer to the related manufacturer report number as noted in section h10 for the medical device report on the automated impella controller and impella 5.5 which is associated with the demise outcome. h6 codes updated for investigation findings. investigation summary access site adverse event (major bleed): the bleeding at multiple access sites occurred simultaneously during the case. the cause of the access site adverse event issue was most likely related to the patient condition, as bleeding was observed at multiple access sites at the same time. clinical symptoms(arrhythmia): in order to make a cause determination, additional clinical details outlined would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. device history review: the pump (b)(6) passed all post sterile inspection checks.
5554 | 1220648-2025-46508 | 23063056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063056&pc=OZD | 08/17/2025 | 11/13/2025 |  | 09/16/2025 |  |  | 08/17/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657460 |  | 000000000001000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 74 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy awaiting heart transplant was implanted with an impella 5.5 for mechanical circulatory support. on the 17th day of support, the medical team reported that the impella 5.5 had alarms for high purge pressure. there were no visible kinks noted, and the medical team decided to run tissue plasminogen activator (tpa) through the purge solution of the impella. the tpa resolved the high purge alarms. on the 28th day of support, the patient experienced an embolic stroke. systemic anticoagulation was withheld as a result and the patient was placed on a temporary inactive status on the heart transplant waitlist. impella 5.5 support was continued for the patient to rest and recover as much as possible. of note, no impella issues were observed. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary- pump was not returned. purge pressure high - after 16 days on support, high purge pressure (hpp) alarms trigger. no kinks were noted and the purge cassette was not replaced. tissue plasminogen activator (tpa) use is not known. data log review showed purge pressure ramp up and purge flows trend down over the course of 2 hours leading up to hpp alarm. issue resolves within 12 hours after a bag change. the cause of the purge pressure high was an adverse event due to patient condition due to the gradual buildup of biomaterial. stroke - after 27 days on support, patient had embolic stroke. the cause of the stroke / embolism was not determined due to insufficient clinical details. device history lot - device lot: 1911456. device history review- pump s/n (b)(6) passed all post sterile inspection checks. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
5555 | 1220648-2025-46517 | 23063192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=23063192&pc=OZD | 08/17/2025 |  |  | 09/16/2025 |  |  | 08/17/2025 | 09/16/2025 |  |  |  |  | 09/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661851 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 72 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | Hematuria | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp pump following an episode of ventricular fibrillation (vf) arrest. the patient had been admitted to the hospital in non- st-segment elevation myocardial infarction and on (b)(6) 2025 had a high-risk percutaneous coronary intervention that was supported with a different impella cp pump. the first impella cp pump was successfully weaned and explanted on (b)(6) 2025, and two days later following the vf arrest the decision was made to replace the impella cp with a new pump. the patient was noted to have a low threshold for vf and ventricular tachycardia (vt) and the same day impella cp support was initiated, the patient required four rounds of defibrillation to resolve vf and vt episodes. the following day, the patient was noted to have red urine. the next day, it was noted that the patient¿s urine had still not cleared and was pink/red in color. additionally, the patient had more episodes of vt during continuous renal replacement therapy (crrt). crrt was aborted as a result of the arrhythmia. the patient¿s urine cleared to yellow the following day. the device was subsequently weaned and explanted successfully, and the explant was not performed earlier than planned as a result of the complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental report will be filed. | H6 component code, type of investigation, investigation findings, and investigation conclusions codes have been updated to reflect investigation closure. investigation summary. ventricular fibrillation/tachycardia: in order to make a cause determination, additional clinical would have to be provided. the root cause of the injury was unable to be determined due to insufficient clinical details. hematuria: the cause of the injury was not determined due to insufficient clinical details. device history lot: device lot: 1920338 device history review: the pump s/n: (b)(6) passed all the post sterile inspection checks.