8371 | 1220648-2025-29488 | 22286719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286719&pc=OZD | 05/21/2025 | 12/03/2025 |  | 06/18/2025 |  |  | 05/21/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1202894 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon turning the pump on to p2, flows displayed 0.0/0.0, and the pump reverted to p0. several attempts were made, however the p-level consistently defaulted back to p0 before the automated impella controller (aic) turned off. the aic was restarted prompting impella defective" alarm and alert to "replace impella." a new aic was brought in but did not resolve the issue. after several attempts the decision was made to explant the defective pump and implant a new 5.5 impella. the patient later expired and that is being reported under manufacture report number 1220648-2025-29322. | [5 empty] | The investigation has been completed. the cause of the issue was not determined since issue could not be consistently reproduced. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the patient later expired and that is being reported under manufacture report number 1220648-2025-29322.
8372 | 1220648-2025-29491 | 22286741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286741&pc=OZD | 05/29/2025 | 07/16/2025 |  | 06/18/2025 |  |  | 05/29/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669983 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was implanted and during the percutaneous coronary intervention procedure, the cp experienced high purge pressure alarm with no outwardly visible cause. the impella was repositioned, but the alarm would resolve intermittently throughout the case. the cp was maintained in p-4 throughout and weaned successfully. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed | Thrombus fm was added based on the investigation. high purge pressure: the data logs show that adequate priming was done of the pump before insertion. after the pump was inserted, purge pressure increased slightly and then shortly after the pump was started there was a small spike in motor current and purge pressure began to rise even more. it rose for about 12 minutes before it triggered a purge pressure high alarm. it remained high despite troubleshooting with changing the cassette, checking lines, etc. flows on the pump were variable. pump was explanted after 1 hour and 18 minutes. the returned product had a thrombus on the impeller and a slight film in the gap. high purge pressure reproduced when the pump was connected to the lab console with the returned cassette. a window was cut in the catheter and there was no blood seen in the purge line. the catheter was then cut near the motor housing and the high purge pressure resolved and fluid expelled from the red handle side. the blockage was isolated to the motor side of the catheter. the pump was torn down and biomaterial was found under the sleeve bearing. there was no biomaterial seen on the ball bearings or magnet. based on the data logs and returned product, the root cause of the high purge pressure is most likely due to biomaterial. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
8373 | 1220648-2025-29492 | 22286751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286751&pc=OZD | 06/02/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 06/02/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026663235 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 insertion was aborted due to the patient's anatomy. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella unable to pass through the vasculature due to patient anatomy.
8374 | 1220648-2025-29494 | 22286862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286862&pc=OZD | 05/29/2025 | 07/21/2025 |  | 06/18/2025 |  |  | 05/29/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636191 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the pump was placed to the left ventricle (lv) but there was poor positioning that prompted the pump to be explanted. the pump had been placed directly into the aorta while in the or with transesophageal echocardiography (tee) guidance. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of positioning issues most likely was use related due improper technique since the physician disliked impella position after troubleshooting. physician made decision to abort impella.
8375 | 1220648-2025-29493 | 22286863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286863&pc=OZD | 05/27/2025 | 09/09/2025 |  | 06/18/2025 |  |  | 05/30/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885741 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO, IMPELLA 5.5 SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while supporting a patient implanted with an impella 5.5 a purge block alarm occurred. the purge cassette was replaced but the issue did not resolve. replacement of the automated impella controller resolved the issue and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge issue - other has been completed. the device and data logs were received for investigation. the root cause of the ¿purge pressure high¿ and ¿purge system blocked¿ alarms, as well as the inability to complete the cassette change procedure, was due to purge pressure sensor malfunction. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8376 | 1220648-2025-29495 | 22286871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286871&pc=OZD | 05/29/2025 | 07/21/2025 |  | 06/18/2025 |  |  | 05/29/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661919 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support placed for the patient with known vascular and cardiac disease. while on impella support the patient¿s labs were trending poorly showing signs of hemolysis. lactic acid increasing, liver function tests increasing, creatinine increasing, urine output decreasing. vasopressin and dobutamine added overnight. the patient was escalated to an impella 5.5 in result of these findings. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. data logs show pump was ran at p-9 through the case with good flows and no spikes in motor current. on (b)(6)the day of the hemolysis, a trend down in ps was observed. the impella device was not returned for evaluation. the root cause of the hemolysis was most likely patient condition related since the patient had severe bi-atrial enlargement, severe right and left ventricle enlargement, severe mitral and tricuspid regurgitation and left atrial appendage (laa) thrombus.
8377 | 1220648-2025-29496 | 22286914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286914&pc=OZD | 06/02/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/02/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657618 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 anticontamination sleeve ripped from impella hub. the catheter was cleaned and the sleeve was reattached, but sterility of the product was broken. no patient harm was reported and the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8378 | 1220648-2025-29497 | 22286922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286922&pc=OZD | 05/27/2025 | 06/25/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1088058 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during yearly preventive maintenance the automated impella controller displayed a battery communication error. new batteries were installed however that did not resolve the issue. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) has been completed. the console logs from the reported day of event are consistent with complaint and show battery communication failure alarm #353 and battery level low alarm 23. issue did not resolve upon console power-cycling. there were no such alarms prior to, or at the very beginning of preventative maintenance procedure. inspection of the console¿s batteries revealed both led indicators showing that batteries were fully charged. after swapping the batteries (battery 1 to power battery manager (pbm)2, battery 2 to pbm 1), batttest tool indicated that information from pbm channel 2 was not received, this time from battery 1: this proves that smbus communication failure was caused by the pbm, as both batteries were able to communicate on pbm channel 1. the cause of the aic issue was a failure of pbm board.
8379 | 1220648-2025-29498 | 22286951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286951&pc=PYX | 05/30/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/30/2025 | 06/18/2025 |  |  | 07/31/2025 |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024399879 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced notable bleeding from their access site following the impella rp implant. an additional suture was placed, and manual pressure was held to manage the condition. the following day the patient was put on extracorporeal membrane oxygenation and the rp was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8380 | 1220648-2025-29499 | 22287008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287008&pc=OZD | 05/23/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/02/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1727322 |  | 0042-0040-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the purge disc was not detected by the automated impella controller (aic). a second aic was tried unsuccessfully. the purge system was exchanged and immediately resolved the issue. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8381 | 1220648-2025-29500 | 22287072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287072&pc=OZD | 05/30/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/30/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584289 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following cardiopulmonary resuscitation (cpr), a clot was found in the cannula of an implanted impella cp device. the pump was replaced due to the noted clot. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8382 | 1220648-2025-29501 | 22287138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287138&pc=OZD | 05/31/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/31/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662785 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support had a head computed tomography (ct) scan, and a large stroke was observed. additionally, anticoagulant was held due to oozing at the surgical site. the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8383 | 1220648-2025-29502 | 22287215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287215&pc=OZD | 06/02/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/02/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669138 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and developed right femoral artery hematoma. the pump was explanted and an impella 5.5 was implanted. the patients outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8384 | 1220648-2025-29503 | 22287266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287266&pc=OZD | 05/30/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/30/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669871 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed for the patient admitted in with known ejection fraction of 5. while on cp support, the impella was turned to p7 due to hemolysis. the pump was explanted however the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8385 | 1220648-2025-29505 | 22287435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287435&pc=OZD | 05/29/2025 | 07/21/2025 |  | 06/18/2025 |  |  | 05/29/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657892 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was placed emergently overnight due to cardiogenic shock. early the next morning, darker red urine appeared. echocardiogram obtained and showed cp in good positioning. the impella was decreased to p5 and a decision was made to escalate to 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The pump was returned for investigation. no physical damage was noted on the returned product. no biomaterial was observed on the pump. the pump was run per specifications in a bucket of water. the pump was sent for hemolysis testing, and it passed with an mih result of 5.32. according to the clinical notes, hemolysis was suspected due to red urine on the morning of 29-may. the doctor reduced the p level to p5, and urine output (uop) returned to normal, reported around 4 pm. data logs were provided until noon on 29-may. the available data log shows no signs related to the patient condition, improper positioning, or interaction with biomaterial. the cause of the hemolysis issue was not determined since issue did not reproduce on returned product and data log was not provided to confirm the field run activity during time of hemolysis.
8386 | 1220648-2025-29506 | 22287464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22287464&pc=OZD | 05/30/2025 | 07/26/2025 |  | 06/18/2025 |  |  | 05/30/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650454 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp displayed thrashing behavior that resulted in access site bleeding. hemostasis was achieved with steep angle matching and the transfusion of four units of packed red blood cells. the patient¿s labs appeared hemolyzed, so the impella was repositioned. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Hemolysis: data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. access site bleeding: product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to patient movement since the patient thrashing in bed causing impella site to bleed profusely.
8387 | 1220648-2025-29507 | 22288072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22288072&pc=OZD | 06/06/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/10/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The abiomed team in japan located a publication inclusive of impella use and extracorporeal membrane oxygenation (ecmo) in japan at a (b)(6) hospital. the version of the impella is not known, nor the actual support. the patient developed a subcutaneous hematoma was observed at the puncture site of the left femoral artery on the ecmo but, the contribution of the impella is not known. the hemodynamics remained stable, and the patient was able to be weaned from ecmo on the fourth day and from impella on the seventh day. the patient's condition continued to improve, and he was discharged on his own on the 35th day without any neurological sequelae. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
8388 | 1220648-2025-29508 | 22288120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22288120&pc=OZD | 06/06/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/06/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A publication titled, "a case of fulminant myocarditis in which impella cp was useful for acute treatment, but ventricular arrhythmia was observed during the course of the disease and a wearable automatic cardioverter defibrillator was introduced", described a patient who had an impella implant. during support, the patient developed sustained ventricular tachycardia and required defibrillation. considering the possibility that mechanical stimulation of the impella was the cause, the impella was removed on the same day. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi)# are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
8389 | 1220648-2025-29357 | 22222941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22222941&pc=OZD | 05/21/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/21/2025 | 06/17/2025 |  |  | 05/31/2025 |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352385 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support for 18 days, there was a gastrointestinal bleed that prompted the pain to be held and one unit of packed red blood cells was given. after 25 days of impella 5.5 support, the patient was successfully weaned and the pump was explanted. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8390 | 1220648-2025-29358 | 22224783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22224783&pc=OZD | 05/31/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/31/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674015 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support (mcs) and developed a hematoma requiring intervention with device removal. the patient was scheduled as an inpatient left and right heart catheterization. the patient became hypotensive. catheterization revealed proximal chronic total occlusion of the right coronary artery with collateralized to the left system. the distal main had an 80% lesion and diffused disease, severe mitral valve regurgitation, mild aorta regurgitation and atrial septal defect (asd). the decision was made to place an impella cp device, which was successfully completed and consult with cardiothoracic surgery for coronary artery bypass graft, mitral valve replacement and potential asd repair. the patient was transferred to the unit in stable condition. the patient developed a hematoma in the groin and manual pressure was applied. the next day, it was noted the access site hematoma was still being managed and the issue was unable to be controlled even after several dressing changes and angle managing by the physician. consequently, the impella cp was explanted. the status of the patient following the event was indicated as unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8391 | 1220648-2025-29359 | 22225778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22225778&pc=OZD | 05/31/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/31/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676879 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp pre percutaneous coronary intervention for mechanical circulatory support and developed an access site hematoma and subsequently expired. the access site was the right femoral artery using micro puncture and ultrasound for access. the impella cp was implanted. after heparin, activated clotting time (act) was 290 seconds. the patient had 99% occluded left main and ostial left anterior descending arteries that were direct stented. upon completion, the 14 french peel away was removed and the device was secured in the right groin. prior to transport to the unit, a small hematoma was noted at the access site. the patient arrived at the unit and femstop was placed. hematoma and oozing around site were noted. the leg was cool and discolored but doppler popliteal and posterior tibial pulses were present. femstop was removed the next morning and oozing had stopped with the leg looking pinker and warmer. foley patent with yellow urine. act was 129 seconds. there was no anticoagulation due to the hematoma. hemoglobin was at 7.8, hematocrit 23.2. and packed red blood cells were ordered. the impella cp was at performance level p-8 flowing 3.6 liters. this same day per the family, care was withdrawn and the patient subsequently expired. | [5 empty] | Given the impella-related access site complication with transfusion, there is not enough information to exclude the impella as an associated factor with the demise outcome. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8392 | 1220648-2025-29360 | 22226865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22226865&pc=PYX | 05/31/2025 | 07/17/2025 |  | 06/17/2025 |  |  | 05/31/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661879 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp and impella rp flex for mechanical circulatory support. following the impella cp insertion, the impella rp flex was implanted, and post implant, imaging revealed outlet of the impella rp flex migrated over into the left pulmonary artery. mattress suture was tightened and additional suture was applied to secure eyelet. manual pressure over site and then dressing was applied. final repositioning to take slack out of impella rp flex optionally alongside the impella cp was completed. the patient was sent to the unit on bipella support in stable condition. the patient remained on support for approximately four days before the patient was successfully weaned, and device was explanted with a survived outcome. the patient was on dobutamine: 7.5 m/k/m, levophed: three microgram per minute with a plan to re-assess labs/hemodynamics one hour post explant, wean and extubate later that day as the patients progress allowed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | No product was returned for analysis. clinical mention after successful pump insertion, peel-away removed and repositioning sheath advanced, imaging revealed outlet of rp flex migrated over into the left pa. the dislodgement occurred before fixation step, after advancing repo sheath. the root cause of the positioning issue was most likely use related as evidenced by clinical details of dislodgement occurring before fixation step after advancement of repo sheath
8393 | 1220648-2025-29361 | 22227584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22227584&pc=OZD | 05/22/2025 | 08/13/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593560 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was being prepped for an impella cp implant and during the setup, the device was not recognized. it kept asking for grey to grey connection like with the old impella cp. it was an impella cp smart assist. it eventually then seem to rectify itself so the staff continued with set up but on insertion, it would not ramp up past p3 when turned on and gave a suction alarm. position was checked and there were no issues at all. another device was used which worked fine and the case was completed as planned. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of pump not detected and low pump flow has been completed. the impella device was returned for evaluation. pump not detected: clinical details state that the impella cp could not be recognized at first and the aic kept asking to connect the grey-to-grey connectors. the issue eventually resolved. the data log review that the pump was not recognized during the first attempt at case start wizard. case start wizard was started again and the pump was recognized. the pump was returned and evaluated. the pump was run and had no detection issues. the pump was then reprogramed to run through case start as a new pump and it was able to be detected. there was no problem found with the returned pump. upon review of data logs and returned pump, it is apparent that the console is the potential cause of the issue. low pump flow: data log review showed sustained low pump flow and a suction alarm. upon startup, there were two distinct motor current spikes unlike the singular motor current spike typically seen upon pump startup. the pump was returned and evaluated. there was a kink found in the mid-section of the cannula and biomaterial found wrapped around the tip of the impeller. low pump flow was reproduced in the lab. the root cause of the low pump flow was most likely biomaterial as the returned product had biomaterial wrapped around the tip of the impeller, low pump flow reproduced, and data log review indicated ingestion. product damage (cannula kink): the failure mode product damage (cannula kink) was added since a cannula kink was found on the returned pump. it is unknown how/when the kink occurred. the root cause of the product damage (cannula kink) was not determined since insufficient clinical details were provided about the source and timing of the kink. thrombus: the failure mode thrombus was added since data log review for low pump flow showed a motor current spike indicating ingestion. there was biomaterial found on the tip of the impeller on the returned pump. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added.
8394 | 1220648-2025-29362 | 22227976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22227976&pc=OZD | 05/21/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/21/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650342 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, after just under 2 hours the pump was removed and replaced due to low flows and the team successfully placed a new impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8395 | 1220648-2025-29363 | 22228450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22228450&pc=OZD | 05/16/2025 | 08/14/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a customer was requesting a pump reorder due to not being able to implant the pump due to a very complex anatomy. additional information was provided that the implantation had to be aborted due to a peripheral artery disease. no other information is available. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the patient had a very complex anatomy. the implantation had to be aborted due to a peripheral artery disease. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 model number, catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
8396 | 1220648-2025-29364 | 22228765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22228765&pc=OZD | 05/22/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661287 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support had ventricular fibrillation and intermittent non-sustained ventricular tachycardia (vt) while on impella support. procainamide was added. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8397 | 1220648-2025-29366 | 22229253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22229253&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640438 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella cp implant, and during the implant, the pump was unable to navigate aortic arch because of struts. the implant was aborted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk pci: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of delivery issue was most likely patient condition since the implant of pump was aborted as it was not able to pass through the aortic arch because of struts.
8398 | 1220648-2025-29367 | 22229288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22229288&pc=PYX | 05/22/2025 | 06/27/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026639567 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient endured an episode of ventricular tachycardia during impella rp flex support. the patient was cardioverted and amiodarone was pushed to treat the condition. it was noted that the measured flow rate around this time was below the expected range for the set p level. the patient decompensated into liver failure later that day and passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of low pump flow and vf/vt/af/at has been completed. the impella device was not returned for evaluation. low pump flow: the data logs show on (b)(6) a motor current spike which caused a shift in pump parameters: there was a sharp decrease in purge flows with a rise in purge pressure. the pump flows also decreased at that instant and there was a placement signal spike. the flows remain lower after the motor current spike. placement signal trends up for the remainder of support while pump flows trend down. the root cause of the low pump flow was most likely biomaterial as evidenced by the motor current spike and its effects on pump parameters in the data logs. vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. thrombus failure mode added due to log analysis revealing ingestion. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
8399 | 1220648-2025-29365 | 22229297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22229297&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  | 01/31/2026 |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454449 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision was made for impella 5.5 placement on a patient with ef 10-15% post cabg. during 5.5 insertion, it was noted that there was difficulty getting impella 5.5 passed the innominate bend. there were multiple wires used as buddy wires and also pulling right arm out straight. rota glide was also used and an angio was done finding tortuous anatomy. it was noted that after over 6hrs of attempting to cross impella 5.5 beyond the innominate bend, utilizing multiple techniques and wires, it was decided to abort further 5.5 placement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The root cause of delivery issue is determined to be patient condition because the pump was unable to pass through tortuous patient anatomy despite using multiple techniques and wires. aborted the implantation after trying for 6 hours.
8400 | 1220648-2025-29368 | 22229378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22229378&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026664367 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was being transferred to bed post impella 5.5 procedure and the red impella plug started oozing. attempts made to stop the bleeding were unsuccessful. decision made to explant 5.5 and increase inotropes. the impella 5.5 was removed. | [5 empty] | The investigation of product damage (hole in catheter) has been completed. after transferring the patient to a hospital bed post implant procedure, oozing noted at red impella plug. attempts made to stop the bleeding were unsuccessful. decision made to explant 5.5 and increase inotropes. the pump was returned and inspected. the returned pump plug was found to have dried blood leaking out. multiple holes punctured in the catheter were found at the 30cm mark. opening the red handle showed blood inside. the cause of the product damage (hole in catheter) was use related due to holes being punctured in the catheter at the 30cm mark causing bleeding into the red pump plug. | The device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
8401 | 1220648-2025-29370 | 22229481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22229481&pc=OZD | 05/22/2025 | 07/14/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659429 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an ureteral stent placed and experienced complications requiring higher level of care and was taken to the catheterization lab for impella cp placement. the impella cp was successfully placed. during support, it was noted that on echocardiogram the impella cp was measuring at 5.8-6 centimeters. the doctor expressed that red urine could be due to hemolysis from the impella cp being deep. however, the doctor was happy with the position and did not adjust. the patient was successfully weaned off of all inotropes/pressors. additionally, the impella cp was also weaned without issue and the patient tolerated well. impella cp was explanted and the patient survived the procedure. | [5 empty] | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. positioning issues: the cause of reported issues most likely due to patient movement since good position confirmed before transferring the patient and after patient transferred echo showed impella was deep. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8402 | 1220648-2025-29369 | 22229621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22229621&pc=OZD | 05/23/2025 | 10/13/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147292 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) alarmed for failure. the aic was replaced and impella support continued. no patient harm was reported. | [5 empty] | Patient information is unknown. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) ¿ controller failure (red alarm) has been completed. the impella device was returned for evaluation. there was no log evidence to confirm the ¿controller failure¿ alarm. it is likely that the ¿controller error¿ alarm was misreported as a ¿controller failure¿ alarm. loss of optical placement signal data (placement, signal not reliable, contrast, lamp, gain) and triggering of a controller error.
8403 | 1220648-2025-29371 | 22230299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22230299&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  | 01/31/2026 |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453633 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, labs hemolyzed but the impella was found too deep. the pump was repositioned and staff monitored. signs of hemolysis were still present so volume was given and the need for lowering the p-level was reassessed due to hemolysis. the pump was eventually escalated to an impella 5.5 and that patients' outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The pump was not returned for investigation. the data log shows the trending placement signal and the pump flow while running on a steady p-level. there was no conclusive evidence of improper positioning or interaction with the impella, on data log. according to the clinical notes, the pump was repositioned during hemolysis, but there was no improvement. later, the cp pump was replaced with a 5.5, and the same issues of suction, positioning, and red urine were reported. the cause of the hemolysis issue was most likely patient condition related based on data log trend. the cause of the positioning issue was not determined without sufficient clinical information.
8404 | 1220648-2025-29372 | 22230899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22230899&pc=OZD | 05/23/2025 | 12/04/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186693 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) exhibited an unexpected shutdown during preparation for support. the device was removed and a replacement was requested. no patient harm was reported. | [5 empty] | Data analysis: the aic logs show a controller error alarm (opm comm stopped ¿ event set # 1043). triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are lastly reported as +16384 values and then no more samples were recorded. the absence of snr samples impacted on the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. the logs show the presence of an ¿unexpected controller shutdown¿ alarm on the next boot, confirming the reported complaint. device analysis: this console was tested after arriving at fs. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. root cause: the root cause of the controller error and loss of placement signal was due to an aic software issue. | Patient information is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8405 | 1220648-2025-29373 | 22231573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22231573&pc=OZD | 05/22/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653998 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella cp support experienced a slightly livid and cool leg in the past and already had a successful extracorporeal fem-fem bypass. the next day it was noted that the fem-fem bypass was no longer needed and the patient was escalated to an impella 5.5. | [5 empty] | A4 and a5, ethnicity, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8406 | 1220648-2025-29374 | 22232157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22232157&pc=OZD | 05/22/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570268 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support during a high risk cardiac procedure. it was later discovered that the patient developed sepsis symptoms and therefore, the pump was explanted early. the patient survived the procedure and was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8407 | 1220648-2025-29375 | 22233098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22233098&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the wheel arm of a cart had snapped off when the package was opened. there was no patient involvement. | [5 empty] | No patient was involved in the reported event. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8408 | 1220648-2025-29376 | 22233875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22233875&pc=OZD | 05/22/2025 | 08/13/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636899 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted. no issue with the device was reported. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because insertion was aborted due to patient anatomy. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8409 | 1220648-2025-29377 | 22233889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22233889&pc=OZD | 05/25/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/25/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654964 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced atrial fibrillation (af). amiodarone was administered and support continued. additionally, the patient had a gastrointestinal bleed. heparin was withheld and two units of packed red blood cells were transfused. the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8410 | 1220648-2025-29378 | 22234083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22234083&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639461 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 85 KG | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had hemolyzed labs and was producing red urine. because the patient had no neuro function, care was discontinued and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes; may also play a role in patient susceptibility to hemolysis."
8411 | 1220648-2025-29379 | 22234671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22234671&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1223031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preventive maintenance that the automated impella controller (aic) was no longer charging. the fuses were found to be blown and were replaced. however, when connected to power, the new fuses also blew. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. the console logs did not contain any data relevant to this complaint. the console was tested after arriving from field service. reported inability to charge and blown fuse issue was confirmed. console would not charge on an engineering lab bench when connected to a/c power and the fuses were found to be blown. the root cause of the reported issues was able to be isolated to the power supply after troubleshooting and was resolved after testing the aic using new fuses and a known good power supply unit.
8412 | 1220648-2025-29380 | 22234834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22234834&pc=OZD | 05/23/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Perforation of Vessels | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, intervention was completed to chronic total occlusion (cto) stent to left anterior descending (lad) and diag. when going to pull the wires after the intervention the diag wire curled up and broke off causing a perforation in the diag and shutting down the lad. echo in the cardiac care lab (ccl) showing an effusion. shortly after the patient had a cardiac arrest requiring cpr. a pericardial drain was placed. impella cp was placed emergently during cpr. once it started, the patient obtained rosc. balloon tamponade was still up on diag. the physician was concerned for rv and didn't think cp was enough support. patient cannulated and impella dropped to p4. intervention to lad utilizing penumbra from prox to distal lad. ptca to om. bleeding had slowed down from pericardial drain. swan placed but had difficulty to get a crossed, so ecmo flows were decreased and then impella suction alarms present blood was given and then increased flows. dr stated he had concerns that impella perf through left ventricle (lv). he said he would open the patient in the icu to access. it was noted that the diag was still bleeding. there was no lv perf. it was also reported that the patient had a stroke. the care was ultimately withdrawn and the patient expired. | [5 empty] | The investigation of vascular damage - peripheral vessel, stroke/cva, ventricle perforation, and low pump flow has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: the cause of the vascular damage - peripheral vessel was not determined since product was not returned and insufficient clinical details. ventricle perforation: there was no issue found during the case. clinical details state although left ventricle (lv) perforation was feared, lv perforation did not occur. stroke/cva: as clinical information was not provided, the true root cause of the stroke/cva could not be determined. low pump flow: data log review showed suction events and low flows following implant procedure until impella was turned down to p-1, then suction resolved. the cause of the low pump flow was most likely patient condition related due to pt on ecmo, suction events, and low volume of the patient. flows were increased to p-2 after clinical details stated blood volume was given and the suction alarms ceased and flow returned to nominal. health effect - impact code 4643 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29380. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella cp with smartassist system and impella 5.5 with smartassist for use during. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8413 | 1220648-2025-29382 | 22234996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22234996&pc=PYX | 05/23/2025 | 07/22/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625254 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, there were placement signal issues. the rp flex appeared to still be operating within normal limits and support continued. additionally, the patient had bleeding of unknown etiology. eight (8) units of blood products were given to the patient. heparin was used in the purge prior to the patient bleeding. care was withdrawn and the patient expired. | [5 empty] | Vascular damage - peripheral vessel: the patient had bleeding of unknown etiology. 8 units of blood products were given to the patient. the root cause of vascular damage - peripheral vessel was not determined as insufficient clinical information was provided optical signal issue: no product was returned for analysis. the data logs confirm psnr alarm and show upon insertion, a gradual decrease in raw optical values while snr and contrast remained stable. cvp is below -100, triggering alarms. motor current also trends down as raw optical does. the pump may have been experiencing suction condition as low motor current indicates limited volume. since the optical sensor is trending more negative with the motor current also trending down, and optical improves (goes positive) as motor current increases, it is possible that it is a suction condition causing psnr. clinical mention that issue resolved after device was slightly repositioned. the root cause of the optical signal issue was unable to be determined as no optical failure pattern was identified in the logs and no product was returned to confirm possible root cause with limited clinical information. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp smartassist system with automated impella controller. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | B1 product problem was erroneously entered on the initial manufacturer device report number 1220648-2025-29382. h6 medical device problem code 2917 was erroneously entered on the initial manufacturer device report number 1220648-2025-29382. h10 investigation results for optical signal issue was erroneously entered on manufacturer device report number 1220648-2025-29382-1.
8414 | 1220648-2025-29384 | 22235131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235131&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643344 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the team placed the cp for the support with difficulty due to femoral calcification. there was rising pp that prompted the team to add tissue plasminogen activator (tpa) to the purge. the patient also had signs of hemolysis. the pump remained on despite this for a total of five (5) days until the family made the choice to withdraw care and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8415 | 1220648-2025-29385 | 22235257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235257&pc=OZD | 05/22/2025 | 06/27/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661780 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, purge system blocked was noted. purge flow was 2.7 on ateplase purge. the plan was to continue ateplase, check the position, wean p level, and check baseline partial thromboplastin time. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of high purge pressure has been completed. no product was returned for analysis. the data logs confirm 'purge system blocked' and 'purge pressure high' alarms that started minutes after the pump was started. the logs show a gradual rise in purge pressure with a corresponding decrease in purge flow before threshold was met and alarms triggered. the gradual rise was over about 4 minutes. there was no motor current spike before high purge pressure event. the purge pressure values remain elevated for about 30 hours before returning to normal values. clinical mention tissue plasminogen activator (tpa) was added to the purge and support continued. the root cause of the high purge pressure was most likely biomaterial as evidenced by the gradual rise in purge pressure in the data logs and the resolution of issue after tpa was added. instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29385.
8416 | 1220648-2025-29387 | 22235353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235353&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657606 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the impella 5.5 pump triggered placement signal unknown alarms. it was noted that the waveforms showed a partially ventricular placement signal with diastolic in the low teens. it was additionally noted that the patient experienced some bleeding around this time from an unknown source. there were concerns of a hemothorax and five units of blood products were given. the patient's hemodynamics were stable despite the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
8417 | 1220648-2025-29381 | 22235354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235354&pc=OZD | 05/22/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624404 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and had a computed tomography scan that revealed possible retroperitoneal bleeding. heparin was paused and one unit of red blood cells was transfused. additionally, the patient had atrial fibrillation with rapid ventricular response due to dobutamine. amiodarone was administered. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8418 | 1220648-2025-29386 | 22235434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235434&pc=OZD | 12/02/2023 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  | 10/31/2025 |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415945 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 49 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received an impact study with a patient that had neurological impairment with a possible relationship between the impella device and the impella procedure. the study stated, "pat has mild loss of strength in the arm and mild hanging mouth corner, no other signs of neurologic deficit, no diagnostic, no therapy. the patient was examined neurologically by the ward physician during the daily ward round. from day to day there was always an improvement in symptoms. only minimally drooping corner of the mouth on the left. no ct performed and neurological consultation performed." | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8419 | 1220648-2025-29383 | 22235505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235505&pc=OZD | 12/28/2024 | 07/15/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533153 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella cp. labs were hemolyzed. it was further reported that the patient coded requiring compressions. return of spontaneous circulation was achieved and the impella cp was repositioned. the patient had aortic stenosis making proper positioning challenging. the patients prognosis was poor, and care was withdrawn. the patient expired. there is not enough evidence to exclude the impella cp device as an associated factor in the patient's outcome. | [5 empty] | The investigation of hemolysis, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. hemolysis: the cause of the hemolysis was not determined since product was not returned and insufficient clinical information for if the hemolysis resolved. positioning issues: the cause of the positioning issue was patient condition related due to the patients noted aortic stenosis (as) and small lv that made proper positioning challenging. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8420 | 1220648-2025-29388 | 22235833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235833&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657722 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Stroke/CVA | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanting an impella 5.5 a suction event occurred and the patient coded requiring cardiopulmonary resuscitation, epinephrine and bicarbonate. the following day, the patient was not able to move their right arm. pulses were present and no swelling of the extremity was noted leading the physician to believe there was possible nerve compression. two days later, a computed tomography scan showed evidence of a neurology deficit and stroke leading to right arm weakness. the patient slowly began to regain strength in their right arm, however labs came back positive for heparin induced thrombocytopenia. the doctor opted to switch from lovenox to argatroban in response to the condition. support continued uninterrupted with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Access site bleeding: the root cause of the access site bleeding is determined to be use issue because the ptt values are higher than the therapeutic range of about 400 as per the clinical details at the time of bleeding. vf/vt/af/at: "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.""
8421 | 1220648-2025-29389 | 22235911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22235911&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521484 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support. after the impella was explanted, some oozing was noted at impella site, 40 mins manual pressure held, protamine was given, hemostasis achieved. the patient tolerated the procedure well. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8422 | 1220648-2025-29390 | 22236058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236058&pc=OZD | 05/22/2025 | 07/16/2025 |  | 06/17/2025 |  |  | 05/22/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635252 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia (vt) and was defibrillated five times. the following day, vt occurred again. echocardiogram at bedside showed the impella 5.5 had not moved at 4.8 centimeters but was up against the ventricle wall. the impella 5.5 was repositioned to 4.2 centimeters to be moved from the ventricle wall. it was further reported that patient care was withdrawn and the patient expired. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. ¿positioning: pump repositioned to 4.2cm to hopefully help get off the wall some. myocardium extremely sensitive and extremely small left ventricle (lv) cavity. the cause of the positioning issues is most likely patient condition related due to patient's extremely small lv. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8423 | 1220648-2025-29391 | 22236203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236203&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660966 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient came into the hospital as a non-st-segment elevation myocardial infarction and with chest pain. the doctor brought the patient to the cardiac catheterization lab. a 6fr sheath was placed in radial and the patient was found to have 99% lymphatic malformation (lm) / left anterior descending (lad) artery / circumflex artery (cx)/ right coronary artery (rca) / with a end diastolic pressure of 30 with a ejection fraction of 20%. cardiothoracic surgery was consulted and dr. (b)(6) was going to take him to the operating room tomorrow for coronary artery bypass grafting. intra-aortic balloon pump was going to be placed but the patient decomposed in the chronic lymphocytic leukemia. decision by the doctor to place impella cp for support and then to fix all the vessels with stents. return of spontaneous circulation was achieved very quickly by staff. the doctors all assisted with the insertion of impella and stents. patients lm / lad / cx / rca was all stented while on support with a good result. patient was intubated. sheath was removed from the groin and the repo was placed. no bleeding was noted at the site. plan is to rest the patients heart for a few days and hope for heart recovery. the following day the patient developed fixed and dilated pupils last night. a computed tomography scan showed right cerebral hemorrhage with intraventricular and subarachnoid extension. anticoagulation is on hold due to brain bleed. right, groin access site remains clean, dry and intact. it was further reported that the patient expired while on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8424 | 1220648-2025-29392 | 22236224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236224&pc=OZD | 04/23/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594767 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Unspecified Infection | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through abiomed's long-term outcome and quality indicator impella registry that a patient implanted with an impella 5.5 and became febrile and experienced an infection that was treated with antibiotics. ventricular tachycardia was noted which was treated with cardioversion, and polymorphic ventricular tachycardia which also treated by cardioversion. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock and cp with smartassist. section: table 3.2 impella catheter components: "the infusion filter prevents bacterial contamination and air from entering the purge lumen." section: warnings and cautions: warnings: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers." | The pump was returned cut at the catheter. no other abnormalities. infection/sepsis - reported infection at access site. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician. -physical examination findings. -results of special investigations : laboratory test results & imaging studies. -microbiological culture and sensitivity results. -response to previous treatments and therapies. -any relevant epidemiological information or exposure history. -concomitant devices in use. -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined. vt/vf - reported vt and polymorphic vt. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
8425 | 1220648-2025-29393 | 22236249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236249&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656678 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support was taken to operating room and a clot was removed from the mediastinum. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8426 | 1220648-2025-29394 | 22236310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236310&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648806 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, there was a hematoma at the access site. a femstop and sandbag were placed. 2 l of lactated ringer¿s solution, 2 plasma, 1 platelet were given to the patient. support continued. additionally, the patient started to show signs of sepsis, urine and blood culture work up along with addition of iv antibiotic treatment were implemented. later weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8427 | 1220648-2025-29395 | 22236408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236408&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic had a purge disk/flag issue. the purge disc could not be recognized. however, the aic was replaced and it was recognized on another console. there was no reported patient harm. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the console was tested and the issue of purge disc not detected was reproduced by connecting the purge cassette. it was confirmed that the purge flag was missing/broken. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the root cause for the purge disc not detected alarm was broken purge flag.
8428 | 1220648-2025-29396 | 22236443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236443&pc=OZD | 05/23/2025 | 07/21/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657853 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there were purge blocked alarms. tissue plasma activator protocol was initiated. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | High purge pressure: clinical details state that a purge system blocked alarm appeared on day of implant. the pump was not replaced, and tissue plasminogen activator (tpa) purge protocol was initiated but clinical details do not specify if that resolved the issue. data log review shows a 12-minute rise in purge pressure (pp) from 500 to greater than 1000 mmhg, which begins about 6 minutes after starting the pump. there was a motor current spike with a stable motor speed prior to the rise in pp. the purge pressure was sustained above 1000 mmhg for approximately 7 hours before gradually decreasing for another 10 hours before being back to 500 mmhg. about 3 hours into the high purge pressure alarms, it appears that purge fluid was changed. no product was returned. the root cause of the high purge pressure was most likely biomaterial ingestion due to the motor current spike prior to the rise in pp seen in data logs. thrombus: the failure mode thrombus was added since the root cause of high purge pressure was most likely biomaterial ingestion. as the necessary information was not provided, the root cause of the thrombus was not determined.
8429 | 1220648-2025-29397 | 22236497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236497&pc=OZD | 04/22/2025 | 08/13/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient had developed hemolysis with a relationship to the pump that was treated by blood products. | Abiomed received a loqi study complaint that reported an adverse event for worsening ventricular tachycardia with relationship to impella cp listed as possibly. "loqi states: vtach happened after impella placement. levophed started. patient was defibrillated with 200j. patient returned back to v paced rhythm." later the patient arrested and passed away. |  |  |  |  | The investigation of hemolysis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information received from the clinical site h6 updated with the additional information received from the clinical site instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29397. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8430 | 1220648-2025-29398 | 22236672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236672&pc=OZD | 05/24/2025 | 07/18/2025 |  | 06/17/2025 |  |  | 05/24/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665191 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp to support a patient admitted in with acute myocardial infarction and stemi. the pump was placed but removed/replaced and escalated to the 5.5 pump after just over 5 hours. the patient had an access site bleed and signs of hemolysis. the cp site's bleed/hematoma was treated by manual hold prior to explant. | [5 empty] | The pump was not returned for investigation. the data log shows a motor current spike, along with some temporary upward trends in the purge pressure parameter. additionally, there were fluctuating trends in pump flow and placement signal during the test run. no conclusive evidence of positioning issue during the case run. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was most likely biomaterial, based on given clinical details. | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8431 | 1220648-2025-29399 | 22236673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22236673&pc=OZD | 05/23/2025 | 08/20/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611558 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the motor current curve of the impella cp suddenly became flatter. an echocardiogram was performed to check the impella cp position, and the result was that the pump is well and freely positioned in the left ventricle. the medical team suspects that this phenomenon could be a thrombotic event. it is known that the patient has an left atrial appendage thrombus. the pump was successfully weaned and explanted five days later. the patient survived the procedure. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. returned complaint device visually inspected. biomaterial observed around impeller and purge gap. no broken blade or cannula kink observed. clinical stated, motor current curve of the impella cp suddenly became flatter. an echocardiogram was performed to check its position, the pump was well and freely positioned in the left ventricle. the distance to the aortic valve is 3.5 cm. the vena cava was also checked. filling is adequate and good. suspected that this phenomenon could be a thrombotic event since the patient had an left atrial appendage thrombus. data log reviewed, many suction alarms occurred. suction alarms observed at the date of reported low flow issue followed by sudden drop in flow just right before the pump is turned down from p-8, flow remains low despite p-level changes up and down. suddenly while at p-2 there is a motor current (mc) spike, speed deviation, increase in purge pressure, and increase in mc with dampened mc and saturated flow for the remainder of the case with suction alarms. the clot was likely in the cannula and was then ingested. low pump flow: the cause of low pump flow issue most likely was biomaterial ingestion based on returned product and log review. saturated flow: saturated flow added per data log review. the cause of saturated flow was most likely biomaterial ingestion based on returned product and log review. thrombus: added per data log review and device evaluation. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
8432 | 1220648-2025-29400 | 22237087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237087&pc=OZD | 05/20/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651262 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient had primary percutaneous coronary intervention (ppci) and a gastroinstestinal bleed, that started after percutaneous coronary intervention (pci). however, heparin was stopped. the impella cp was successfully weaned and explanted. the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8433 | 1220648-2025-29401 | 22237243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237243&pc=OZD | 05/24/2025 | 07/09/2025 |  | 06/17/2025 |  |  | 05/24/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667466 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 64 YR |  | Male |  |  |  | Death | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during removal of the peel-away sheath, the physician inadvertently began to crack and peel the component away despite the distal end still being partly within the patient's body. upon noticing that the sheath was not fully outside of the body, the physician proceeded with their attempt. a hematoma developed immediately and both manual pressure, sutures, and femostop were placed to manage the condition. blood was given to counteract the condition. support continued as planned with the newly implanted impella cp pump. the patient was eventually transitioned to comfort care and passed away. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding and difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. access site bleeding - major: the cause of the access site bleeding - major was use related due to the physician breaking the peel away sheath with still inside the patient, causing a hematoma. difficulty inserting/removing introducer: the cause of the difficulty inserting/removing introducer was not determined since product was not returned and insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29401. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29401. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8434 | 1220648-2025-29402 | 22237378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237378&pc=OZD | 05/24/2025 | 08/12/2025 |  | 06/17/2025 |  |  | 05/24/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608612 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp had a purge leak while on patient support. changing the purge cassette did not solve the problem. a purge filter bypass was done successfully. additionally, the patient was in septic condition. the patient escalated to an impella 5.5. | [5 empty] | The investigation of purge leak ¿ pump and infection/sepsis has been completed. the impella device was not returned for evaluation. purge leak - no product was returned. data logs were not recovered. the cause of the purge leak - pump was most likely a leak from the sidearm but the exact location of the leak was unknown. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
8435 | 1220648-2025-29403 | 22237398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237398&pc=OZD | 05/27/2025 | 07/03/2025 |  | 06/17/2025 |  |  | 05/27/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640266 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump's access site was seen to have a bleed that prompted the team to treat with femostop, manual hold, additional suture, and lidocaine with epinephrine. later, it was reported that there was light pink urine. p-level reduced to p-2 and the pump was pulled back two centimeters. support was continued until weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. access site bleeding: clinical mention bleeding occurred after insertion of guidewire repositioning unit. activated clotting time (act) was noted to be 280. the root cause of access site bleeding - major was most likely due to anticoagulation management because the act values reported during the time of bleed are higher than the suggested therapeutic range during the tome of bleed. hemolysis: no product was returned for analysis. the data logs from the case do not show any motor current spikes outside p-level changes and there were no suction or pump positioning abnormalities. clinical mention scant light pink uop began overnight with imaging used for pump position which revealed pump position was not good. echo shows inlet 5.3cm from av. the pump was pulled back 2cm. complication was managed by position adjustment and -level reduction. subsequent patient updates after repositioning revealed normal urine output not pink or yellow indicating hemolysis resolution. the root cause of the hemolysis was most likely due to improper positioning as evidenced by the imaging revealing pump not in good position and the issue resolution after pump repositioning.
8436 | 1220648-2025-29404 | 22237466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237466&pc=OZD | 05/26/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/27/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient arrived by emergency medical services and was seen to be in need of impella cp and extracorporeal membrane oxygenation support. while on support, the impella cp was alarming impella in the aorta and suction. an echocardiogram by the physician assistant was questionable at best due to the patient's previous chest compressions and chronic obstructive pulmonary disease. the impella cp was positioned just across the valve and advanced which improved motor current and flows. the following day there was a short episode of ventricular tachycardia with mouth care. four days later, care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8437 | 1220648-2025-29405 | 22237608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237608&pc=OZD | 05/24/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/24/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561462 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella 5.5 support and after a difficult delivery, the pump migrated back into the aorta. upon removal of the device, the catheter was visibly damaged. the pump was replaced. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8438 | 1220648-2025-29406 | 22237679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237679&pc=OZD | 05/24/2025 | 07/16/2025 |  | 06/17/2025 |  |  | 05/24/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657742 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp placed for the support of a patient admitted in with a cardiac history and stemi. the patient had been shocked repeatedly prior to the pump placement and was intubated and sedated in poor condition. the cp was placed, the pump was seen to be in poor position at the valve and there was signs of hemolysis. the team did not reposition but instead when patient made a dnr the patient expired after just 39 hours of support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of hemolysis, access site bleeding-minor, and positioning issues has been completed. data review: data logs showed pump ran without issue during support. there were multiple suction alarms triggered during the case. failure mode changed from access site bleeding - minor to access site bleeding - major because blood product were given in addition to manual pressure to help managed the oozing. the cause of the bleeding was most likely patient condition related since patient bleeding from all orifices. the root cause of hemolysis was most likely due to pump was not in the proper position since the physician confirmed pump might be interfering with the mitral valve (mv). the root cause of positioning issue was most likely patient condition related because pump was noted to be possibly interacting with mv but when repositioned, triggered ectopy. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8439 | 1220648-2025-29407 | 22237812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237812&pc=OZD | 05/20/2025 | 08/13/2025 |  | 06/17/2025 |  |  | 05/20/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658651 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Arrhythmia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella 5.5 support, there was suction on p8 with hypotension which was resolved by lowering to p7. the impella was showing the inlet close to the septum and repositioning was attempted, but echo kept showing the inlet still close to the septal wall. the patient was shocked four times and was re-intubated. it was further reported that the patient was made dnr and expired on support. | [5 empty] | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow and positioning issues was not determined since there was a lack of clinical details, no product returned, and insufficient data logs. the patient had a history of ventricular tachycardia (vt) with an automatic implantable cardioverter defibrillator prior to impella support. the patient went into vt storm. the medical team did not attribute vt to the impella, thus h6 was revised to remove 1721 and changed to 4445. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8440 | 1220648-2025-29408 | 22237873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237873&pc=OZD | 05/30/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/30/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660184 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hernia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 support developed oozing in the pocket. heparin was held. it was further reported that the patient experienced brain bleed and brain herniation. the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8441 | 1220648-2025-29409 | 22237983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22237983&pc=OZD | 05/24/2025 | 07/21/2025 |  | 06/17/2025 |  |  | 05/24/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while transferring the patient to their bed, the implanted impella cp pump was inadvertently pulled back across the valve and into the aorta. attempts to reposition the pump were unsuccessful and the decision was made to explant the device. patient survived the procedure. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the positioning issue is most likely due to patient movement as the clinical details state that when the patient was transferred to his bed the impella was pulled across the valve and into the aorta. | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8442 | 1220648-2025-29410 | 22238210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22238210&pc=OZD | 05/23/2025 | 06/17/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1332595 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a self-check failure and was removed from service. there was no reported patient involvement. | [5 empty] | The impella console was received from the customer and an investigation regarding the returned device has been completed. the console logs are consistent with the complaint and showed communication issues with the power battery manager (pbm) during the initial bootup. due to the communication issue, the automated impella controller console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
8443 | 1220648-2025-29411 | 22238764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22238764&pc=OZD | 05/23/2025 | 07/15/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644378 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp for mechanical circulatory support. the impella cp was inserted via the left femoral artery through the peel away sheath. during insertion, the doctor was unable to advance the impella after multiple attempts to advance. the pump was removed and a kink was noted in the catheter. a new device was inserted successfully. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings, ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. a catheter and cannula kink were noted after the pump was removed. a new pump was delivered successfully. impella cp was returned and a cannula kink was found. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient had tortuous vasculature at the femoral artery. cannula kink - the root cause of the issue was a kinked cannula due to patient condition since the patient had tortuous vasculature at the femoral artery.
8444 | 1220648-2025-29412 | 22239584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22239584&pc=OZD | 05/23/2025 | 09/22/2025 |  | 06/17/2025 |  |  | 05/23/2025 | 06/17/2025 |  |  |  |  | 06/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659985 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient endured multiple episodes of ventricular tachycardia during impella 5.5 support. the patient was shocked in response to the condition and started on lidocaine. it was also noted that an intermittent impella position unknown message kept appearing during support. no intervention was required and support continued uninterrupted. the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned. the data logs show the placement signal was relatively stable. during these increases, "impella position unknown" alarms would trigger correctly, but are a not accurate as positioning was confirmed. due to lack of clinical details and no product returned, the root cause of the optical signal issue cannot be determined. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-29412.
8445 | 1220648-2025-29299 | 22210116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210116&pc=OZD | 05/30/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/30/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670601 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support for about 20 minutes and during support, the patient developed a hematoma with bleeding at the access site and the pump was explanted. the patient's outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8446 | 1220648-2025-29300 | 22210189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210189&pc=OZD | 05/30/2025 | 07/26/2025 |  | 06/16/2025 |  |  | 05/30/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675693 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the physician attempted to reposition the device and noted it would not move freely. during manipulation, the catheter was inadvertently over torqued which caused a pump stop. further examination of the exposed graft noted a hardened clot which had prevented catheter movement. the pump was explanted and replaced. the patient's outcome was critical. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Fm optical signal issues added based on returned product and data log investigation. positioning issues the clinical stated that the physician attempted to reposition the device and further examination of the exposed graft noted there was a hardened clot preventing movement of the catheter. based on the clinical details provided, the root cause of the positioning issues is most likely due to the patients condition as there was thrombus formation in the graft, however the cause of the thrombus is unknown. thrombus in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. pump stop & optical signal issue the clinical stated that during manipulation the catheter was inadvertently over torqued causing a pump stop. the data logs show that the pump run was stable with stable pp, mc and ps for about no as put into surgical mode about 239 hours in for about 2 hours. during this time there were some increases in the ps to 3000. the pump was then turned back on for 20 minutes and the ps spiked to 3000, gain decreased, light decreased, snr went to 0, contrast barcoded and lamp went to 100. the mc also spiked and the pump stopped, triggering psnr alarms, 115 and 119 pump stop alarms. the returned product was cut, but there was no biomaterial seen in the inflow or outflow cage. there was significant stretching of the catheter where the catheter jacket is in between the coils between the 60-70cm mark. based on the clinical details, product and data log analysis the root cause of the pump stop is most likely due to open electrical lines from excessive force. the root cause of the optical signal issue is most likely due to a damaged fiber line that could be see in the case logs to be torqued and then fully damaged when the pump stop occurred, due to excessive force.
8447 | 1220648-2025-29301 | 22210260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210260&pc=OZD | 05/30/2025 | 07/18/2025 |  | 06/16/2025 |  |  | 05/30/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656671 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be advanced through the patient's vasculature during attempted delivery. it was noted that despite multiple attempts, the pump would not advance beyond the anastomosis resulting in a kinked catheter. the implant was aborted as a result of the noted resistance and there was no patient harm. | [5 empty] | Delivery issue : impella 5.5 pump could not be advanced through the patient's vasculature. the root cause of the delivery issue was most likely patient condition related, impella 5.5 pump could not be advanced through the patient's vasculature during attempted delivery. product damage ( canula kink ) : the pump would not advance beyond the anastomosis resulting in a kinked catheter. the root cause of the product damage issue was most likely a catheter kink, the pump would not advance beyond the anastomosis resulting in a kinked catheter. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8448 | 1220648-2025-29302 | 22210368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210368&pc=OZD | 05/23/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/23/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513958 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and had large clots and was bleeding from the chest cavity requiring washouts multiple times. the bival was held and blood products were administered. after impella 5.5 was explanted, it was noted there was a thrombus attached and hanging off the outlet. the thrombus was not seen on echocardiogram, only after the impella was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8449 | 1220648-2025-29303 | 22210456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210456&pc=OZD | 05/20/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/20/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654184 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient developed heparin induced thrombocytopenia during impella 5.5 support. treatment was switched to argatroban in response to the condition. the patient's underlying condition failed to improve five days later and the decision was made to withdraw care. the patient passed away shortly after. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8450 | 1220648-2025-29304 | 22210459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210459&pc=OZD | 05/30/2025 | 07/26/2025 |  | 06/16/2025 |  |  | 05/30/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660257 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to traverse into the ascending aorta during attempted delivery due to the patient's anatomy. after multiple failed attempts, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of delivery issue was most likely use issue since non-abiomed product (10 mm hemashield platinum graft) was used which is not recommended per ifu.
8451 | 1220648-2025-29305 | 22210566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210566&pc=OZD | 05/20/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/20/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658740 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Other |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced multiple complications during impella 5.5 support. roughly one day into support, the patient experienced persistent bleeding. the patient was taken to the operating room for a washout and support continued with the current setup. the following day, the patient endured a stroke. the patient remained sedated on ecpella support. on that same day, the pump was repositioned due to several runs of ventricular tachycardia and ectopy. the p-level was limited due to the patient's cardiac function. the decision was ultimately made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8452 | 1220648-2025-29306 | 22210634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210634&pc=OZD | 05/30/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/30/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645713 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a ventricular tachycardia/ventricular fibrillation arrest patient admitted in with acute myocardial infarction/cardiogenic shock. after three days, the pump was weaned and explanted. upon explant, it was observed that the limb developed signs of ischemia. the team had vascular surgery intervene and the decision was made to take the patient to the operating room emergently for exploration and endarterectomy. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8453 | 1220648-2025-29307 | 22210727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210727&pc=OZD | 05/29/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/29/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671733 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump escalated to an impella 5.5 for the patient admitted in with a non-st-segment elevation myocardial infarction. the impella 5.5 was supporting when the patient was turned and encountered self limiting runs of ventricular tachycardia. the registered nurse said they charted it and would discuss if the intensivist was around. pads were acquired but not applied to the patient. a code cart was outside of room but no report of it being used. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8454 | 1220648-2025-29308 | 22210797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210797&pc=PYX | 05/20/2025 | 07/15/2025 |  | 06/16/2025 |  |  | 05/20/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625507 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Break | Fracture | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while implanted and providing support the impella rp's clear purge side arm broke off. the patient remained on support. | [5 empty] | The investigation of purge leak - pump has been completed. the clinical details state that the clear purge sidearm broke off at the tubing of an rp flex and an emergent purge bypass had to be performed. no excessive force had been applied. the rp flex had been used for about 16 days at this point. data log review shows a purge pressure rise above 1000 mmhg followed by a drop to 0 mmhg, which recovers within 30 minutes. product was returned but only included the cannula and part of the catheter, which had been cut from the rest of the pump. this returned product was insufficient for evaluation. the root cause of the pump purge leak was most likely a damaged sidearm but the exact location of the leak is unknown since product/images were not returned. d9 device returned to manufacturer date added g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2025-29308. h6 medical device problem code 1260 was erroneously entered on the initial manufacturer device report number 1220648-2025-29308. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿
8455 | 1220648-2025-29309 | 22210807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210807&pc=OZD | 05/27/2025 | 07/14/2025 |  | 06/16/2025 |  |  | 05/27/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance beyond the sheath due to the patient's vascular anatomy. as a result of the noted resistance, the decision was made to abort the implant. the patient outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through introducer sheath and aborted insertion due to patients vascular anatomy. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8456 | 1220648-2025-29310 | 22210845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210845&pc=PYX | 05/21/2025 | 07/08/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561539 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 68 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Failure to Sense | [5 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex and during support, there were placement signal issues. additionally, bulk air was noted in both the left ventricle and right ventricle. the patient was de-aired and support continued. the patient also had issues with volume status and there were low flows. blood products were given and support continued. | [5 empty] | The investigation of embolism and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data logs were reviewed and the logs showed pump flows below intended p-level range at beginning of case. motor current and placement signal were also variable at that time with cvp going out of range at times of low pump flow. clinical details noted that the patient was having volume issues on the first day of support and required multiple blood products and was experiencing low flows on both impella and left ventricular assist device (lvad) devices. the root cause of the low pump flows was most likely patient condition related based on data log analysis and clinical details which noted that both lvad device and impella rp flex was experiencing low flows as patient had a volume issue. embolism: clinical details noted that chest was open during lvad and impella rp flex placement. when closing chest after procedure, a bulk of air was noted in both left ventricle (lv) and right ventricle (rv) with diminished flows on both devices. patient wet back on bypass and de-air of chest was performed again. the root cause of the embolism could not be definitively determined as no product was returned and limited clinical details were provided. h6 health effect - clinical code 4582 and medical device problem code 1559 were erroneously entered on the initial manufacturer device report number 1220648-2025-29310. | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8457 | 1220648-2025-29311 | 22210985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22210985&pc=OZD | 02/08/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/27/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED. INC. |  |  | US |  | 2025573388 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via impact study noting that the patient experienced ventricular tachycardia and device suction issues during impella support. the patient was administered an amiodarone bolus. additionally, the impella performance was reduced for the ventricular tachycardia. the support was completed successfully and the device was weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8458 | 1220648-2025-29312 | 22211202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211202&pc=OZD | 05/27/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/27/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026652142 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. on day five of support, the team noted a "rupture on the reposition sheath". the product damage did not result in patient harm or any further action. the patient continues on support. | [5 empty] | The impella device was not received from the customer as it is still supporting the patient. therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8459 | 1220648-2025-29313 | 22211230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211230&pc=OZD | 05/21/2025 | 08/15/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026640854 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be advanced through the patient's vasculature during attempted delivery due to vascular calcification. the implant was aborted and an impella cp pump was subsequently placed for patient support. | [5 empty] | The investigation of delivery issues has been completed. the impella 5.5 and purge cassette were returned. no damages or abnormalities were found. the root cause of the delivery issue was most likely patient condition related since the patient had vascular calcification. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿
8460 | 1220648-2025-29314 | 22211275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211275&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664009 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support experienced limb ischemia. consequently, the device was explanted. the impella cp device was implanted in the setting of acute-onset stress cardiomyopathy. the patient arrived in the catheterization lab on epinephrine and levophed with a blood pressure of 40/20 with no appreciable peripheral pulses. the right femoral artery was accessed with ultrasound and micro puncture. short peel-away was placed. impella was advanced into position without difficulty and placed in auto mode but taken shortly thereafter to performance level p-9. however, after sustaining suction alarm, the performance level was reduced to p-8. there were some continued issues with hemodynamic instability; however, after correcting issues with the intravenous lines, the patient stabilized with no further impella alarms. the peel away sheath as exchanged, and site sutured, and angle matched. touhy borst valve was locked at 82 cm. an echocardiogram revealed device slightly deep at 4.5 cm; however, it was noted to be in "excellent" orientation with pigtail free and inlet in mid-ventricular space. the patient was sent to the unit on support. two days after the start of the impella mcs, the team was unable to find pulses per doppler. attempt was made to warm the extremity and wean pressors as tolerated; however, this was unsuccessful. the decision was made to explant the impella cp device and once removed, doppler pulses were obtainable. the patient survived the explant; however, the patient's condition following the event was indicated as unknown. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
8461 | 1220648-2025-29315 | 22211367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211367&pc=OZD | 05/27/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/27/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026663310 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. once positioned into place, the team felt the inlet was too close to the mitral valve despite attempts to reposition and starting at p2 with wire removed. the impella 5.5 was removed from the body and was placed in surgical mode while ventricular access was reobtained. the impella 5.5 with wire was re-directed towards left ventricle apex. the patient continued on support until the device was successfully weaned. the patients outcome was reported as stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8462 | 1220648-2025-29316 | 22211404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211404&pc=OZD | 05/21/2025 | 07/17/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1679798 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during a case the automated impella controller was during on an error message displayed indicating to do not use; replace impella. it was noted the procedure outcome was unknown; it is also unknown if the device was replaced as a result of the reported event. there was no harm to the patient reported, and no further information was provided. | [5 empty] | Data analysis: the case before the complaint date, there is consistent pbm1 communication issues through 5/14. the console was next booted up at 9:21 on 5/21 with continued communication issues but the console was automatically shut down due to no user interaction 30 minutes later ((b)(6)). the console was booted up again at 10:14 but shutdown by the user five minutes later. neither of these boot ups had any alarms to the user outside of ac power disconnected despite the pbm1 communication issues. device analysis: the pbm1 communication issues were reproduced but like in the field, there was no alarms ((b)(6)). the pbm was taken out and examined but there was no evidence of what could have been causing the pbm1 comm issues. a known-good pbm was inserted and the pbm comm issues went away. root cause: there was no evidence of a controller error in the logs. repair: please precautionary replace the pbm (0042-1125) due to the pbm communication issues, then perform sections 18, 19, and 20 of the pm procedure (0042-7309) and functional check (0042-7316) on (b)(6). | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
8463 | 1220648-2025-29317 | 22211423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211423&pc=OZD | 06/06/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 06/06/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information reported from literature review stated that the impella cp was removed without resistance while confirming the site of stenosis in the common iliac artery under fluoroscopy. however, immediately after removal, there was no detectable pulse in the femoral artery, raising suspicion of thrombotic occlusion based on the findings. angiography showed complete thrombotic occlusion of the right external iliac artery, necessitating surgical thrombectomy. postoperatively, the dorsalis pedis artery was improved. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 first name and last name are unknown. h.4 device manufacture date is unknown.
8464 | 1220648-2025-29318 | 22211439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211439&pc=OZD | 05/20/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637058 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed a hematoma at the access site. a dissection was noted. the doctor felt like it was due to advanced patient age and the impella was removed. the artery was repaired. the doctor then decided that the aorta looked too fragile to place on the impella 5.5 without dissecting it also. the doctor then abandoned the pump bypass and completed two vessel off pump bypass. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8465 | 1220648-2025-29319 | 22211551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211551&pc=OZD | 05/18/2025 | 08/14/2025 |  | 06/16/2025 |  |  | 05/18/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584891 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision for impella 5.5 was made on a patient with cardiogenetic shock. however, it was noted that during implantation of the impella 5.5, it was not possible to advance the device through the graft. several attempts were made; however, they were all unsuccessful. instead, an impella cp was successfully implanted. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was returned for evaluation. product damage(cannula kink) was added after product evaluation since a cannula kink was found. reported impella 5.5 unable to advance pass the graft. impella 5.5 delivery was aborted and impella cp was successfully implanted. impella 5.5 was returned and cannula kinks were found. delivery issue - the root cause of the delivery issue was most likely patient condition related since the pump was unable to pass the graft. cannula kink - the root cause of the product damage(cannula kink) was most likely patient condition related since the pump was unable to pass the graft. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. there are two manufacture report # for the same patient. this report captures the delivery issue with the impella 5.5 (serial number (b)(6) and the other one will capture the patient's outcome (death) on the impella cp (serial number (b)(6). as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8466 | 1220648-2025-29320 | 22211640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211640&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601140 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had limb ischemia. a bilateral distal perfusion catheter was implanted and eventually, the pump was explanted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8467 | 1220648-2025-29321 | 22211648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211648&pc=PYX | 05/28/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/28/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659341 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted, as elective placement, with an impella rp flex for mechanical circulatory support. following the implant, the motor current spiked momentarily at performance level p-8 and flows dropped to 2.2 liters per minute. flows did not recover and the patient's urine began to turn red. the pump was weaned and explanted. upon explant, a fibrous yellow clot was noted at the inflow cage. the patient was noted to be in stable condition following the event. | [5 empty] | The complainant shared the impella rp flex device was discarded as it was thought there was no issue with the device, and also stated the ingestion of particulate was likely chronic clot that came from an old pacemaker pacing lead. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8468 | 1220648-2025-29322 | 22211713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211713&pc=OZD | 05/21/2025 | 11/17/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658361 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient struggled to wean from the bypass, necessitating revision of some grafts and a bentall procedure. despite these efforts, the patient was unable to come off the pump and required five shocks and was placed on central va ecmo for support. the decision was made to place an impella 5.5 as a vent. the team anticipated difficulties crossing the ao valve due to lack of pulsatility. after multiple attempts with various catheters, the al1 successfully crossed the valve, and the impella was placed without further issues. however, upon turning the pump on to p2, flows displayed 0.0/0.0, and the pump reverted to p0. several attempts were made; however, the p-level consistently defaulted back to p0 before the automated impella controller (aic) turned off. the aic was restarted prompting impella defective" alarm and alert to "replace impella." a new aic was used, however, that did not resolve the issue. after several attempts the decision was made to explant the defective pump and implant a new 5.5 impella. again, crossing the valve proved to be difficult. light chest compressions were performed to help open the aortic valve; however, efforts were unsuccessful. ecmo flows were decreased, and a slight increase in epi was administered to assist with crossing the valve, however this also failed. the chest dressing was removed, and manual heart massage was performed to facilitate opening of the ao valve, yet crossing remained unsuccessful. bleeding was observed during the procedure, and the patient received 20 units of ffp, 19 units of rbc, 2 units of platelets, and 3 units of cryoprecipitate. the patient remained non-pulsatile with compromised hemodynamics and continued to require ongoing transfusions of blood products. ultimately the decision was made by the family to withdraw care and the patient expired. | [5 empty] | The pump was returned for investigation. vascular damage: patient was bleeding from chest cavity across the open chest procedure and required multiple blood products. the bleeding occurred across multiple impellas. the cause of the vascular damage was not determined due to lack of clinical details. pump not detected: the pump was inserted with difficulty and started when it was noted torque was built up in the white cable. the pump was un-plugged and the torque removed and plugged back in. upon reconnection with the console, upon turning the pump on to p2, flows displayed 0.0/0.0, and the pump reverted to p0. several attempts were made, but the p-level consistently defaulted back to p0 before the aic turned off. the aic was restarted prompting impella defective" alarm and alert to "replace impella." replacing the console did not resolve the issue so the pump was replaced. data log review showed the pump was manually disconnected while running and threw impella defective alarms upon reconnection. imc data logs show a ¿pumpcalibdata unknown type 0.' imc logs also show the user attempting to increase the p-level prior to the eeprom being fully read. the pump was returned and inspected. when plugging the pump into the aic showed no abnormalities when the eeprom was allowed to be read. when the pump was connected and the p-level was changed prior to the eeprom being read, impella defective alarms occurred. ct imaging of the pump showed no electrical or circuit abnormalities. the cause of the pump not detected issue was an eeprom issue however the cause of the eeprom issue could not be determined since the escalated investigation did not determine a true root cause. the issue was able to be recreated in the lab using both complaint aic and known good aic's. complaint aic was unable to recreate the issue on known good pumps. h3: added information regarding the status of the device investigation. h6: added codes for the result of the investigation. h10: added investigation information. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
8469 | 1220648-2025-29323 | 22211773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211773&pc=PYX | 05/28/2025 | 09/16/2025 |  | 06/16/2025 |  |  | 05/28/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025528878 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Low Blood Pressure/ Hypotension | Hypoxia | Perforation | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella 5.5 for mechanical circulatory support and escalated to bipella support with implant of an impella rp flex for right ventricular disorder. the following night, the patient experienced pulmonary bleeding after the impella rp flex placement. medication was administered via an endotracheal tube to control the bleeding. partial thromboplastin time remained subtherapeutic all night in the 200s. the patient experienced an acute hypotensive episode resulting in cardiac arrest. return of spontaneous circulation was achieved. the patient was hypoxic. the left lung was unable to ventilate and was noted to be obstructed by a clot. bedside bronchoscopy was being performed to evacuate the blood clot. the patient was transfused with blood products. the following day, it was noted that the patient's condition stabilize. the bronchoscopy was successful in evacuation of a clot from the left lung. a couple of days later, it was noted, however, that the patient continued to be hypoxic and a repeat bronchoscopy with old thrombus deep in the bronchioles was noted and was unable to be reached. the decision was made to explant the impella rp flex and cannulate for veno-arterial extracorporeal membrane oxygenation (va ecmo) with the impella 5.5 device which was successfully completed. latest update noted the patient remains on impella 5.5 and va ecmo support. | [5 empty] | In supplemental manufacturer device report vt was captured however this complaint is not related to vt. | Additional information was received and noted a perforation occurred in the lung and was believed to be from the abiomed 0.27 wire. the explanted impella rp flex pump was discarded. the impella device was discarded, and it is unknown if the guidewire is available. therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the investigation of vf/vt/af/at has not been completed. should new information be received, a supplemental manufacturer device report will be filed.
8470 | 1220648-2025-29324 | 22211964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22211964&pc=OZD | 05/28/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/28/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612736 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced multiple episodes of ventricular tachycardia with defibrillation one day after start of support. the patient remained on mcs and isuprel was started. the following morning, it was noted there were no premature ventricular contractions. the next day thereafter, it was noted the patient was successfully weaned and the impella cp was explanted with a survived outcome. the patient remained on isurpril. it was noted the patient's home doses of methadone daily was likely the cause of the long qt and premature ventricular contractions. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8471 | 1220648-2025-29325 | 22212056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212056&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620061 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had torsades rhythm overnight requiring two (2) shocks to go back into normal sinus. the patient had frequent rhythm issues prior to impella placement and correct positioning evaluated on transesophageal echocardiography (tee) and the patient successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8472 | 1220648-2025-29326 | 22212229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212229&pc=OZD | 05/21/2025 | 07/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025566126 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the aortic sensor numbers were not correlating. the impella was repositioned, but the sensor issue was not resolved. during the ecmo clamp trial the impella at p4 was hitting the septum wall. the impella was repositioned and able to go up to p8 without alarms. | [5 empty] | Post 5.5 implant, patients ao on impella 5.5 has been multiple point difference, echo confirmed positioning multiple times. no alarms on pump. even after repositioning no changes. the pt had a starting sbp of 107mmhg and a map of 100mmhg. the pump was returned with the catheter cut distal to the red handle so the pump could not be run or identified. optical signal issue: data log review showed no abnormalities outside of lower than the reported ps pressure values for the patient, around 50mmhg. the cause of the optical signal issue was not determined due to inconclusive data log review and returned product was cut and the sensor could not be tested. positioning issues: after 2 days on support, during ecmo clamp trial, impella hitting septum wall. repositioned to 38cm from 40cm. the cause of the positioning issue was most likely use related due to pump being placed too deep during ecmo trials, requiring it to be pulled back. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8473 | 1220648-2025-29327 | 22212301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212301&pc=OZD | 05/28/2025 | 08/13/2025 |  | 06/16/2025 |  |  | 05/28/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026651971 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful due to the patient's anatomy. the device was unable to be advanced. the impella 5.5 implant was aborted. subsequent treatment is unnoted as well as status of the patient following the event. however, there was no report or indication of adverse effects as a result of this event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of delivery is determined to be patient condition because the pump was unable to pass through the vasculature due to resistance in axillary artery. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8474 | 1220648-2025-29328 | 22212502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212502&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560160 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had hematoma at the insertion site. aggrastat was turned off, and pressure dressing was applied. support continued. additionally, the patient's cmp/lactate labs were continuing to hemolyze and support continued. the patient was eventually escalated to impella 5.5 with removal of impella cp from left axillary. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8475 | 1220648-2025-29329 | 22212589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212589&pc=OZD | 05/19/2025 | 07/18/2025 |  | 06/16/2025 |  |  | 05/19/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614907 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision for impella 5.5 was made on a patient with cardiogenetic shock. however, it was noted that during implantation of the impella 5.5, it was not possible to advance the device through the graft. several attempts were made; however, they were all unsuccessful. instead, an impella cp (b)(6) was successfully implanted. during cp support, there was intermittent suction which they increased the volume. position was verified by echo. the patient was scheduled for a bypass and ventricular septal defect closure surgery with a 5.5 escalation. it was noted that the impella 5.5 was in direct aortic and extracorporeal life support central implanted after successful heart surgery. the patient was with open thorax on ward. also noted was that after surgery that lactate was dropping and parameters were getting better. however, the patient ultimately expired on support. this case is for the cp (b)(6). the medical safety officer reviewed and determined there is not enough information to exclude the impella cp pump 2 device used as an associated factor in the patient's outcome given there was intermittent suction with support. the patient expired the following day after escalation in therapy to impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the medical safety officer reviewed and determined there is not enough information to exclude the impella cp pump 2 device used as an associated factor in the patient's outcome given there was intermittent suction with support. the patient expired the following day after escalation in therapy to impella 5.5. the impella 5.5 for delivery issue is reported on manufacture report number 1220648-2025-29319. | The investigation of low pump flow has been completed. the pump was not returned. data logs were not recovered. the cause of the low pump flow was most likely patient condition related due to issue resolving around the time additional volume was given and the p-level was reduced.
8476 | 1220648-2025-29330 | 22212794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212794&pc=OZD | 05/21/2025 | 07/15/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638710 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump came out of position to the aorta and was readvanced, later the pump was seen to be in the valve. the patient also had signs of hemolysis but not confirmed. the patient had bilateral limb access site bleeding and was treated by removal of anticoagulation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high-risk pci. instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position. ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | Clinical stated: patient reported bleeding from bilateral groin sites, left abdominal surgical site, and urine via foley. cangrelor and heparin gtt stopped until generalized bleeding is under control. while preparing the patient for icu the impella was pulled into the aorta and was re-advanced back into the lv without a wire. access site bleeding: major: the cause of the access site bleeding: major issue patient condition related since bleeding was in multi sites (bilateral groin sites, left abdominal surgical site, and urine via foley). positioning issues: the cause of the positioning issues most likely was use related due improper technique since while preparing the patient for icu the impella was pulled into the aorta and was readvanced back into the lv without a wire.
8477 | 1220648-2025-29331 | 22212903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212903&pc=OZD | 05/28/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/28/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658406 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant placed an impella 5.5 pump to support a patient with cardiac history and in need of care escalation from an intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation. the impella 5.5 was placed axillary and was found to have a hematoma at the impella insertion site and oozing at old iabp insertion site to the right groin. heparin currently was on hold and red blood cells and platelets infusion were given to the patient. later, the family decided to withdraw care and the patient expired. the relationship between the impella 5.5 and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8478 | 1220648-2025-29332 | 22212928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22212928&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  | 06/30/2025 |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364884 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient was returned to the operating room for a pocket clean out. one unit of blood products was given to the patient and support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8479 | 1220648-2025-29333 | 22213136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213136&pc=OZD | 05/21/2025 | 07/17/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662641 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had atrial fibrillation (af) and intrinsic rhythm competing with epicardial pacing wires. epicardial pacing was turned off to allow intrinsic rhythm and support continued. additionally, the patient's access site was oozing. the site was redressed, and blood products were ordered and support continued. the patient expired on support. | [5 empty] | Access site bleeding: the root cause of the access site bleeding is determined to be use issue because the ptt values are higher than the therapeutic range of about 400 as per the clinical details at the time of bleeding. vf/vt/af/at: "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined."" | Access site bleeding: the root cause of the access site bleeding is determined to be use issue because the ptt values are higher than the therapeutic range of about 400 as per the clinical details at the time of bleeding. vf/vt/af/at: "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
8480 | 1220648-2025-29334 | 22213210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213210&pc=OZD | 05/20/2025 | 07/09/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639564 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient endured non-sustained ventricular tachycardia and an access site bleed during impella cp support. the ventricular tachycardia (vt) was believed to be possibly related to impella support, while the bleed was noted to have a definite relationship. the bleed was managed after a bandage change and fresh frozen plasma (ffp) infusion. the patient was escalated to an impella 5.5. | A later notification from abiomed¿s long-term outcome and quality indicator (loqi) impella registry additionally reported the patient endured recurring thrombocytopenia and procedural mediastinal hemorrhaging during support. the hemorrhage was believed to have a definite relationship to the impella device while the thrombocytopenia was said to have a possible relationship to the impella device. the bleed was noted during a scheduled mitral valve repair. the conditions were treated with blood products and platelets. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B5 updated with additional information received from the clinical site. h6 health effect - clinical codes updated to reflect additional failure modes. instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29334.
8481 | 1220648-2025-29335 | 22213308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213308&pc=OZD | 05/22/2025 | 09/17/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658384 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented for a surgical mitral valve replacement/tricuspid valve replacement with impella 5.5 placement. the patient was successfully supported for a total of 188.83 hours. however, it was noted that after successfully weaning and during explanting the 5.5 the cannula broke at the junction of the outlet flow. there was no harm to the patient reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported product damage (detached inflow/outflow - great vessel) has been completed. there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. the product had a clamp mark but it was stated that that was from retrieving the cannula. additionally, the bond was ripped as the cannula bond was jagged remaining on the cannula and not the motor housing. on the inside of the cannula there was also residual glue marks where it would be attached to the motor housing. there was a slight bend in the catheter but no kinks and blood seen around the impeller. due to a lack of sufficient evidence during the removal of the 5.5, the root cause of the product damage (detached inflow/outflow) cannot be determined. corrections to the initial mfr report: h3 has been updated to device eval completed. h6 type of investigation code 10 added.
8482 | 1220648-2025-29336 | 22213387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213387&pc=OZD | 05/21/2025 | 07/08/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658635 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella 5.5 for mechanical circulatory support. during the implant, the staff attempted to pass the impella unsuccessfully. rotoglide was utilized and a catheter was placed over amplatz to give more stability and unable to pass the impella. the staff attempted to torque the impella 5.5 around the curvature unsuccessfully. calcium in the innominate and on the curvature of the axillary was preventing the impella 5.5 from passing. the implant was aborted and an impella cp was placed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to calcium noted in the innominate and on the curvature of the axillary preventing the impella from passing.
8483 | 1220648-2025-29338 | 22213485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213485&pc=OZD | 05/22/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661754 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support had runs of atrial fibrillation. also, computed tomography scan result showed a large subacute right middle cerebral artery infarct. the family elected to withdraw care and patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8484 | 1220648-2025-29339 | 22213526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213526&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431278 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facilities engineer reported that the impella automated impella controller (aic) had a battery failure and was removed from service. there was no reported patient involvement. | [5 empty] | The impella device was received from the customer and an investigation regarding the returned device has been completed. console logs were consistent with what was reported in the complaint description. an instance of battery failure (transport connection blown/missing) alarm was observed in the logs on the reported complaint date. the console was tested and the reported battery failure was confirmed and able to be reproduced on an engineering lab bench. during inspection of the batteries on the automated impella controller, it was observed that the liquid-crystal display indicator on battery 1 exhibited a pattern that corresponds to cell imbalance. troubleshooting was performed by swapping battery 1 with a known good battery which resolved the issue. the reported battery failure was isolated to battery 1. the reported battery failure issue was determined to be caused by cell imbalance in battery 1 due to a defective battery.
8485 | 1220648-2025-29341 | 22213652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213652&pc=OZD | 05/20/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/20/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631899 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that purge solution was leaking from the red impella plug during impella 5.5 support. the decision was made to continue with support while monitoring the leak. four days later, the bedside nurse reported that the patient had a ventricular tachycardia event early in the morning requiring cardioversion at bedside. four more days later, integrated purge side arm was leaking purge solution. however, after 16 days of support, care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
8486 | 1220648-2025-29342 | 22213814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213814&pc=PYX | 05/22/2025 | 07/21/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025547321 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella rp flex being prepped to support a patient a right heart failure post-right ventricular assist device patient. the rp flex had an optical signal loss and so troubleshooting was performed. even automated impella controller swapping did not resolve issue and so the rp itself was exchanged. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue has been completed. the data logs were reviewed and right after plugging the pump in, placement signal not reliable (psnr) alarms are triggered. the snr was 0, gain was 2.5, light was around 2, and contrast was 10. the logs were not returned so calibration of the sensor could not be checked. the product was returned and the laser continuity test showed light exiting the sensor face. the light was very faint. the 5s parameters were taken and they were not within specification. the sensor was imaged under the keyence and there was material on the sensor face. the pump was soaked in water and then the sensor was imaged again but the material remained. this material on the sensor face, likely damaged it. the root cause of the optical signal issue is due to a damaged sensor due to material on the sensor face, but the material/cause of the material in unknown.
8487 | 1220648-2025-29343 | 22213846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22213846&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the patient had oozing at the access site. purse sting sutures were placed cinching the sheath to the skin. heparin was withheld. 5 units of blood products were given, 2 units of fresh frozen plasma and hemoglobin increased to 8.7. patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of access site bleeding is determined to be use issue because doctor resutured suture hub and added a purse string suture cinching the sheath to the skin and bleeding at site was stopped.
8488 | 1220648-2025-29345 | 22214134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214134&pc=PYX | 05/22/2025 | 06/27/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625308 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Failure to Pump | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp-rp flex support ongoing for the bipella needs of cardiogenic shock patient in right heart failure. the pump was in position when there was need to reposition and a that time the pump was removed and then redelivered. the pump may have been damaged in the redelivery and shortly afterwards had a pump stop. the team decided to remove the rp flex and medically manage while awaiting transfer to the tertiary center. | [5 empty] | The investigation of delivery issues, pump stop, and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: clinical mention the pump was in position when there was need to reposition. no details were given as to why there was a need to reposition. the root cause of the positioning issue was not determined as limited clinical information related to failure was provided and no product was returned. delivery issues: clinical noted challenges when attempting redelivery and that pump met resistance in femoral vein and was unable to advance past vasculature. the root cause of the delivery issue was most likely patient condition related as evidenced by the clinical details of being unable to pass pump through vasculature. pump stop: the data logs show several motor current spikes outside p-level changes with placement signal shifts and effects on pump flow while the pump was at p9. the pump was turned down to p0, turned up to p2 and back down to p0. when the pump was not running at p0, purge pressure began to gradually rise and purge flow started to decrease suggesting potential biomaterial build up. clinical mention pump damage and compression to the inlet at re-delivery attempt. it is likely that the inlet struts were damaged and caused interaction with the impeller. at restart attempts, there were motor current spikes with 'impella stopped - motor current high' alarms. there were 6 failed attempts after which pump was explanted. the root cause of the pump stop was unable to be determined as no product was returned for analysis to confirm suspected root causes of biomaterial or impeller and inlet damage and data log analysis is inconclusive h6 health effect - impact code 4627 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29345. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8489 | 1220648-2025-29337 | 22214220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214220&pc=OZD | 05/22/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647242 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed symptoms of right-hand paresthesia during impella cp support. the pump was subsequently explanted later that same day. the patient survived the operation. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8490 | 1220648-2025-29340 | 22214289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214289&pc=OZD | 05/22/2025 | 07/10/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025527612 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an attempt to place the impella 5.5 in an escalation was unsuccessful due to not being able to pass the 5.5 into the ascending aorta. it was noted that the more the physician tried to place the pump into the left ventricle, the catheter began buckling close to the motor. the decision was made to open a new pump and try with a new wire with success. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The root cause of delivery issue was most likely use issue since there was difficulty implanting the first pump due to the catheter buckling and aborted but a new pump was implanted successfully.
8491 | 1220648-2025-29346 | 22214487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214487&pc=OZD | 05/20/2025 | 07/15/2025 |  | 06/16/2025 |  |  | 05/20/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643219 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patients urine was red. the pump was deep this morning and was pulled back under echocardiogram. fluid bolus was given. additionally, there was oozing at the impella site and sutures were added. it was further reported that the patient¿s family decided to withdraw care and the patient expired. there is not enough evidence to exclude the impella used as an associated factor. | [5 empty] | The investigation of access site bleeding, hemolysis, and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding - major: after 1 day on support, oozing was noted at impella site requiring the physician to place sutures. the pump was not returned. the cause of the access site bleeding was most likely use related due to sutures required that resolved the issue. hemolysis: the pump was not returned. data log review showed adequate flows and placements signal values during day of compliant, however, 8 hours of the logs were unaccounted for. the cause of the hemolysis was most likely improper positioning of the device due to clinical details stating the pump was too deep in the left ventricle (lv). positioning issues: data log review showed adequate flows and placements signal values during day of compliant, however, 8 hours of the logs were unaccounted for. the cause of the positioning issue was most likely use related due to the pump being initially placed too deep in the lv. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8492 | 1220648-2025-29344 | 22214493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214493&pc=OZD | 05/20/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/20/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665978 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 73 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an extensive cardiac history/comorbidities was on impella rp flex and impella cp support. the patient had been found down and unresponsive in the community. cardiopulmonary resuscitation was given and the patient was intubated for percutaneous coronary intervention. it was noted that the heparin drip was to be started when left femoral oozing was controlled. ultimately, the patient expired on support. | [5 empty] | Abiomed's medical safety officer reviewed and there is not enough information to exclude the impella cp device used as an associated factor to the patient's demise outcome given the access site bleeding event. heparin in purge may have impacted/exacerbated the outcome of the bleed for the patient. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8493 | 1220648-2025-29347 | 22214607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214607&pc=OZD | 05/22/2025 | 07/07/2025 |  | 06/16/2025 |  |  | 05/22/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658626 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery for patient support. due to the tortuosity of the patient's artery, the decision was made to abort the implant. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature.
8494 | 1220648-2025-29348 | 22214651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214651&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654220 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.5 device for mechanical circulatory support. during the implantation, the physician was unable to safely insert the impella 5.5 due to the patient¿s anatomy and an impella cp was placed via the axillary instead. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings, ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8495 | 1220648-2025-29349 | 22214845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214845&pc=OZD | 05/21/2025 | 07/10/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643387 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, upon startup, the impella cp pump was unable to achieve flows above one liter per minute. using fluoroscopy, it was revealed that the cannula kinked distal to the impeller motor. repositioning was performed but the kink remained. the decision was made to replace the pump with no issues. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The pump was returned for investigation. a data log was not provided, and it was not retrieved from the remote server. no physical damage was observed on returned product, except the kinked cannula. the impeller was free of biomaterial. the pump was tested in a bucket of water without any abnormalities. the cause of the low pump flow issue was most likely due to a kinked cannula; however, the cause of the cannula kink remains unknown. product damage (cannula kink) failure mode was added as an investigated failure mode. the cause of the product damage issue was kinked cannula; however the cause of the cannula kink was unknown without sufficient clinical information.
8496 | 1220648-2025-29350 | 22214951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22214951&pc=OZD | 05/21/2025 | 07/18/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659977 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during the impella insertion, the aortic cannula was pulled out of the aorta. insertion of impella was paused and focus was put on crashing back on bypass and reclamping the aorta. it was noted that the patient was losing a significant amount of blood. once the patient was back on bypass and the new cannula was secure, the insertion of the impella was continued. during this time, there was verbal confirmation in the room that the wire was out. when the pump was started, there was a ¿pump stopped¿ alarm and asked if wire was out. the response was it is now. the pump was restarted and the alarm was resolved. the patient continued on support and was slowly weaned and explanted. the patients outcome is stable. | [5 empty] | H6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29350. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. entering cardiac output (impella cp with smart assist). section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The pump was not returned for investigation. the data log shows that the pump stopped at the start of the case but was able to restart and ran as expected. according to the clinical notes, the wire was not pulled out during the temporary pump stop. the cause of the temporary pump stop issue was most likely use related to removal of guidewire before start the pump, based on given clinical details. the cause of the access site issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the positioning issue was most likely use related since it was it was accidently pulled in aorta. | The pump was not returned for investigation. the data log shows that the pump stopped at the start of the case but was able to restart and ran as expected. according to the clinical notes, the wire was not pulled out during the temporary pump stop. the cause of the temporary pump stop issue was most likely use related to removal of guidewire before start the pump, based on given clinical details. the cause of the access site issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the positioning issue was most likely use related since it was it was accidently pulled in aorta.
8497 | 1220648-2025-29351 | 22215050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22215050&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626231 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the access site developed a hematoma that was treated by manual hold, femostop, two units of red blood cells, and stopping of the systemic heparin. the pump remained on support for six days. the impella cp was successfully weaned and removed. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8498 | 1220648-2025-29352 | 22215175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22215175&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613502 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient had draining from the access site. the patient had some swelling on the side. angiomax was held and support was continued. | [5 empty] | The impella device was not returned, and the investigation has been completed. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
8499 | 1220648-2025-29353 | 22215274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22215274&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653442 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had some ectopy. the impella was removed the following day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8500 | 1220648-2025-29354 | 22215364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22215364&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662653 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient who was implanted with an impella 5.5 device for mechanical circulatory support experienced ventricular tachycardia and was treated by cardioversion and p level adjustment. additionally, there was a bleed that prompted the team to remove systemic heparin. the pump was on sodium bicarbonate in the purge rather than heparin. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8501 | 1220648-2025-29355 | 22215496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22215496&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661952 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. three days into support, the patient endured several episodes of ventricular tachycardia storm. the patient was shocked 23 times and received 20 seconds of chest compressions to counteract the condition. five days later, the patient continued to decompensate, and care was withdrawn. the patient expired upon removal of support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8502 | 1220648-2025-29356 | 22215643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22215643&pc=OZD | 05/21/2025 | 06/16/2025 |  | 06/16/2025 |  |  | 05/21/2025 | 06/16/2025 |  |  |  |  | 06/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680562 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Pericardial Effusion | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the physician was able to induce ventricular tachycardia (vt) with successful mapping. during the end of the ablation, a pericardial effusion was noted. the physician placed a drain, and the remainder of the procedure was aborted due to the patients decline. the decision was made to leave the impella cp in and send to the unit to be placed on multiple anti arrhythmic and continue impella cp support. the plan was to stabilize the patient and correct ventricular tachycardia, labs, and hypotension during ventricular tachycardia episodes. it was noted that during impella cp support, the patient's neurology status had improved. however, vasopressor requirements were still very high. the decision was made to place the patient on comfort measures only and withdraw care . the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8503 | 1220648-2025-29275 | 22201612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22201612&pc=OZD | 05/20/2025 | 07/03/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631302 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella 5.5 implant and during the implant, the axillary artery was visualized to be large enough for the 5.5. the area of stenosis was about 5 cm proximal to graft encountered and the pump could not be advanced. staff attempted to dilate with 22fr dilator. the 5.5 could still not pass. catheter ended up bending into a spiral with amount of axial pressure exerted. the implant was aborted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The root cause of the delivery issue was most likely patient condition related, unable to advance 5.5 past subclavian artery.
8504 | 1220648-2025-29276 | 22201780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22201780&pc=OZD | 05/20/2025 | 07/14/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Delivered as Unsterile Product | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while cutting off the repositioning unit of an impella cp pump in preparation for axillary access, the catheter jacket was inadvertently perforated. the surgeon was advised of the bleedback should they proceed with the implant. the pump was swapped for a new impella cp device for planned support. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Product damage (hole in catheter): the failure mode was updated from product damage ( sleeve damage) to product damage (hole in catheter) since upon investigation, it appears that the catheter was perforated. clinical details state that the catheter was perforated while prepping the cp for axillary insertion. the repo unit was cut off to avoid interference with ability to get vascular control during insertion or occlude the vessel. another cp was opened and prepped. no product/images were returned and logs are not applicable. the root cause of the product damage (hole in catheter) was most likely use-related since clinical details state that the catheter was perforated while prepping the cp for axillary insertion. h6 medical device problem code 1421 was erroneously entered on the initial manufacturer device report number.
8505 | 1220648-2025-29277 | 22202203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22202203&pc=OZD | 05/20/2025 | 07/07/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612739 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Extravasation | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that shortly after arrival at icu from impella cp placement oozing was observed for the right groin site. heparin was paused and two units of prbcs were administered. additionally, the rn was unable to hear previously positive doppler right pedal pulses. a reperfusion catheter was placed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: access site bleeding - section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ limb ischemia - impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of limb ischemia could not be determined since the product was not returned and insufficient clinical details were provided. the root cause of access site bleeding was most likely patient movement since patient started oozing after arriving to the post procedural area. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29277. h6 health effect - clinical code 1842 was erroneously entered on the initial manufacturer device report number 1220648-2025-29277. correct code added h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29277.
8506 | 1220648-2025-29278 | 22202343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22202343&pc=OZD | 05/20/2025 | 08/13/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653951 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in germany had a 5.5 pump support ongoing for the patient admitted in with cgs and cardiomyopathy. the pump was supporting when there was poor positioning and signs of hemolysis. the team could not remedy either and so decisions/conversations ongoing for wean/explant and or lvad placement instead of the 5.5 pump. support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: hemolysis - section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. the data log shows several suction alarms on the reported day of hemolysis. no trends in motor current or placement signals justify the cause of hemolysis during the case run. additionally, no positioning-related alarms were reported. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was most likely improper positioning of the device, based on returned product investigation. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 health effect - impact code 4641 was erroneously entered on the initial manufacturer device report number.
8507 | 1220648-2025-29279 | 22202376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22202376&pc=OZD | 05/20/2025 | 07/07/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659996 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a right upper extremity deep vein thrombosis in the same side as the impella. the patient was brought to the operating room due to swelling of the right arm and weakness requiring a washout. a small thrombus under the muscle closure could have caused the pressure on the brachial plexus. the patient was stable post axillary washout. a right brachial vein venous thromboembolism was found which was likely the cause of the arm swelling and not from the axillary site. the patient survived and support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Thrombus: "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined." stroke/cva: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. -timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient. -neurological examination findings and any focal deficits observed. -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). -laboratory test results, including coagulation profiles and cardiac markers. -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿
8508 | 1220648-2025-29280 | 22202469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22202469&pc=OZD | 05/20/2025 | 06/26/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Thrombosis/Thrombus | [10 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient endured atrial fibrillation rvr and ventricular tachycardia during impella cp support. the pump was repositioned in response to the condition. it was additionally noted that hemolysis had developed. plans were made to initiate continuous renal replacement therapy. however, care was withdrawn the following day and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of hemolysis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. hemolysis: hemolysis was noted with no additional information. the data logs show a motor current with speed deviation about 2 hours after support started. there was a motor current spike with speed deviation and a pump auto-reduction in p-level to p1. the pump was turned to p6 at that instant there was also a suction alarm which resolved later. clinical reported hemolysis signs about 3 hours after this probable ingestion event. the root cause of the hemolysis was most likely biomaterial as evidenced by the motor current spike with speed deviation occurring before clinical report of hemolysis signs vt/vf/at/af: as the necessary information was not provided, the root cause of the vt/vf was not determined. thrombus: based on the data logs and clinical details thrombus was added as a failure mode. as the necessary information was not provided, the root cause of the thrombus was not determined b2 date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29280. b2 life threatening was erroneously selected on the initial manufacturer device report number 1220648-2025-29280. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-29280. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: hemolysis - section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8509 | 1220648-2025-29281 | 22202580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22202580&pc=OZD | 05/20/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593515 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Life Threatening | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella 5.5 was implanted on a patient that presented with cp and sob. during support it was noted that heparin was being stopped due to a cranial bleed. additionally, the patient experienced episodes of suction and was given fluid bolus. it was also noted that out of the blue the patient was experiencing hematuria and that the current metrics was showing limited hemolysis. the patient continues on heart recovery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8510 | 1220648-2025-29282 | 22202970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22202970&pc=OZD | 05/20/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660213 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed a 5.5 pump for the support of a patient with cardiac and renal history. the patient has dialysis and prior mi. the pump was placed along with ecmo for organ support. after 8 days the team observed lv thrombosis. the clot burden has not prompted intervention or explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: thrombus- section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8511 | 1220648-2025-29283 | 22203012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22203012&pc=OZD | 05/21/2025 | 07/21/2025 |  | 06/13/2025 |  |  | 05/21/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815750 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Other | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported during preventive maintenance, no patient involvement, the automated impella controller (aic) display flickered when powered on. the aic was sent for data and device analysis. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. data analysis: console logs from the reported day of event do not contain any information relevant to this failure mode. device analysis: the console was sent back to field service (fs) for further investigation and was tested after arriving in fs. the reported issue was reproduced, and intermittent lines were observed on lcd screen. further testing aimed at determining whether the issue was caused by the 4-in-1 assembly or the lcd were inconclusive. root cause: the root cause of the flickering screen (horizontal intermittent line) was mostly due to the 4-in-1 assembly or the lcd. | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
8512 | 1220648-2025-29284 | 22203120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22203120&pc=OZD | 05/21/2025 | 07/08/2025 |  | 06/13/2025 |  |  | 05/21/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1705772 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support (mcs). during preparation, the automated impella controller (aic) failed to recognize an installed purge disc. the purge disc was replaced; however, the issue persisted which led to an aic-to-aic exchanged to resolve the failure. there was no harm to the patient per the report. the patient underwent pci, coronary angioplasty and stent placement, to the ramus artery. post procedure the pump was successfully weaned, the device explanted with a survived outcome. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. data analysis: the complaint date revealed that the console started case start and got to step 2 of the case wizard. the purge cassette was inserted and the user was asked to insert the purge disc before issuing a purge disc not detected alarm (event set 402) for two and a half minutes as a purge disc was never recognized as inserted. the purge disc not detected alarm resolves when the purge cassette is removed. device analysis: the console was examined and the purge unit was clean and the purge cassette and disc were able to be recognized during testing. conclusion: the cause of the purge disc detection issue was not determined as the failure mode was not able to be reproduced. | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
8513 | 1220648-2025-29285 | 22203167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22203167&pc=OZD | 05/20/2025 | 08/12/2025 |  | 06/13/2025 |  |  | 05/21/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648154 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and had to be explanted soon after it was placed. while placing the impella cp device, the cannula kinked just below the outlet cage. upon placement, the impella was unable to build up flow and suction alarms triggered. the impella cp device was pulled back into the descending aorta to straighten the kink; however, upon readvancement, the cannula re-kinked. consequently, the impella cp device was explanted and replaced. there were no adverse effects as a result of the event. latest updates noted the patient remains on impella cp mcs. | [5 empty] | The investigation of low pump flow and product damage (cannula kink) has been completed. the impella device was returned for evaluation. low pump flow: returned complaint cp pump visually inspected. cannula kinks near outlet and bending area observed. no biomaterial or broken blade seen. clinical stated, while placing impella cp in the left ventricle, a kinking just below the outlet cage was observed. impella was not able to build up flow, suction alarm observed, auto mode with 0.5l flow. the root cause of low pump flow issue was due to cannula kink, but the cause for cannula kink is unknown product damage (cannula kink): while placing impella cp in the left ventricle, a kinking just below the outlet cage was observed. the cause for how cannula kink occurred was unknown due to unknown information on patient condition or use issue. the root cause of product damage(cannula kink) was not determined since the cause for cannula kink is unknown due to insufficient clinical details. | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
8514 | 1220648-2025-29286 | 22204314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22204314&pc=OZD | 05/21/2025 | 08/15/2025 |  | 06/13/2025 |  |  | 05/21/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666362 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and ethnicity) undergoing high-risk percutaneous coronary intervention (pci) was preoperatively implanted with an impella cp device for mechanical circulatory support (mcs); however, it was noted the pump did not "work correctly; wrong flow." the impella cp device was exchanged to complete the protected pci. the pci was completed, replacement impella cp device was successfully weaned and explanted with a survived outcome. | [5 empty] | The investigation of low pump flow has been completed. low pump flow: returned complaint cp pump visually inspected. no cannula kink observed. no biomaterial observed. no broken blade found. impella cp connected to console to reproduce reported low pump flow issue at it run for more than 10 minutes. flow was with specified limit at p-9. no low flow can be reproduced and no other issue observed. clinical stated: after impella placement, pump did not work correctly. wrong flow, wrong alarms. unsuccessful repositioning attempted. pump exchange. the cause of low pump flow issue cannot be determined due to lack of clinical details, no abnormalities on returned product, low flow not reproduced on returned pump, and data log not available to review. | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
8515 | 1220648-2025-29287 | 22204377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22204377&pc=OZD | 05/17/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/17/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675346 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump to support a patient admitted in with st-elevated myocardial infarction and high risk cardiac procedure. the patient had ventricular tachycardia (vt) runs during the support which required defibrillation, which was successful. the physician reported events were not due to impella. patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8516 | 1220648-2025-29288 | 22204521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22204521&pc=OZD | 05/20/2025 | 07/15/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657866 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella cp implant and the impella cp kinked at upper part of cannula near motor when advancing through new tavr valve within an old bioprosthetic tissue valve. when in auto mode, the aic (console) displayed a yellow suction alert continuously. staff couldn¿t break suction when lowering p level or repositioning. staff decided to replace impella cp with a new one. the patient outcome was unknown. | [5 empty] | Reported pulling the pump out of position when transferring the patient from the or table to the bed. pump was removed and replaced with ecmo. the root cause of the positioning issue was most likely patient movement since the pump was pulled out when transferring the patient. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
8517 | 1220648-2025-29289 | 22204675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22204675&pc=OZD | 05/20/2025 | 07/29/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5-EU |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026659099 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella 5.5 implant and during the implant, the pump got stuck at junction of axillary artery and truncus brach. the implant was aborted and an impella cp was implanted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since impella 5.5 got stuck at junction of axillary artery and truncus branch and since it was not possible to advance impella 5.5 over axillary artery into the aorta along with calcification/ stenosis observed. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
8518 | 1220648-2025-29290 | 22204766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22204766&pc=OZD | 05/20/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664790 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 65 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemothorax | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced ventricular fibrillation requiring defibrillation. the patient then developed a hemothorax requiring a chest tube and transfusion of two units of red blood cells. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8519 | 1220648-2025-29291 | 22204884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22204884&pc=OZD | 05/20/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/20/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620668 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed a fever. cultures were taken and antibiotics were administered. additionally, the patient had heparin-induced thrombocytopenia. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & the instructions for use for the cp with smartassist states the following: section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8520 | 1220648-2025-29292 | 22205053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22205053&pc=OZD | 05/14/2025 | 07/15/2025 |  | 06/13/2025 |  |  | 05/14/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662861 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | Failure to Read Input Signal | Device Sensing Problem |  |  |  |  | Hemodynamic instability | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and decompensated right after implant. vasopressin, epinephrine, and amiodarone drips were started. the patient had suction alarms and performance level p-6 was the highest level able to be achieved. echocardiogram had poor windows; however, the physician was content with the impella cp position. the patient continued to have poor oxygenation and potential plan for extracorporeal membrane oxygenation (ecmo). the patient continued to be labile. the patient was planned for transfer and during transfer the following day, upon automated impella controller (aic) to aic transfer, ¿impella defective¿ alarm appeared. the new aic used and ¿placement signal not reliable¿ alarmed. per transport personnel, the impella shut off three times and restarted. it was noted the impella cord not completely plugged in as small gap was visible between console plug in and impella cord. the patient's mean arterial pressure was in the 70s and ecmo at 3l. impella cord was unplugged to assess prongs, which were all straight then replugged in and still not fitting flush into plug in. the impella cp device was at p-6/ with flows of 2.4l and placement signal was not visible on aic. the physician was okay with muting the placement signal unreliable alarm which was completed with guidance. an echocardiogram was completed and the impella cp measured 3.46cm within mid left ventricle. the patient was noted to also has severe mitral regurgitation and the physician felt, however, the mitral regurgitation was causing a false normal ejection fraction. the patient remains intubated and sedated on support. however, the patient would eventually expire. care was withdrawn. | [5 empty] | The investigation for the optical signal issue, temporary pump stop, and the pump not detected issue has been completed. the data logs show that the impella was running without issue 5/13 - 5/14 till 11:49 am on console ic9283. there were no impella defective alarms, pump stops or placement signal issues. there is a gap in time in which the pump was plugged into ic2043 in which they received an impella defective alarm. there are no data logs returned for ic2043. the root cause of the pump not detected cannot be determined due to lack of data logs or product returned. 5/14 at 1:39pm the pump is plugged into ic2045, the pump is beginning to ramp up p-level until it has alarm 115 and 119 pump stops with ps not reliable alarms. the snr is 0, gain and light are increased, lamp is maxed out and contrast is variable. there are no spikes in motor current (mc) or increases in purge pressure, but the purge flow increases from 5 to about 17ml/hr. 5/14 at 1:52 the pump is running again, motor current was very stable, ps was maxed out at 3000 triggering psnr alarms and there were controller error alarms. the snr is 0, gain and light are increased, lamp is maxed out and contrast is variable. the pump runs until about 5/26 where is it ultimately explanted. due to lack of product returned, the root cause of the temporary pump stop cannot be determined. due to lack of product returned the root cause of the optical signal issue cannot be determined. mso review: the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the patient's death is assigned to both the pump and aic (console). there is not enough evidence to exclude the impella cp pump and automated impella controller used as an associated factor in the patient's outcome given event details on both devices. added h6 med dev prob codes 1503, 2917 | Medical safety reviewed the event and determined that the demise outcome is assigned to both the impella cp pump and automated impella controller. there is not enough evidence to exclude the impella cp pump and automated impella controller used as an associated factor in the patient's outcome given event details on both devices. the impella cp is one of two devices associated with the event. refer to initial medical device report for the automated impella controller serial number (b)(6) event date 14-may-2025 and patient identifier end in (b)(6). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8521 | 1220648-2025-29293 | 22205130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22205130&pc=OZD | 07/12/2023 | 06/13/2025 |  | 06/13/2025 |  |  | 06/06/2025 | 06/13/2025 |  |  | 03/31/2024 |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-JP |  |  | OZD |  |  | IMPELLA 2.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023199499 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 10 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A publication titled "two cases of combined therapy with extracorporeal membrane oxygenation (ecmo) and intracardiac pump catheter for circulatory support (impella) in pediatric fulminant myocarditis", was reviewed and during support, a hematoma spread from the insertion area into the retroperitoneal cavity, ultimately leading to lower limb weakness, likely due to hematoma compression. the intervention for the hematoma is unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8522 | 1220648-2025-29294 | 22205523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22205523&pc=OZD | 05/09/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/09/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657751 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp placed to support a patient with possible heart transplant rejection. while one support the patient had runs of ventricular tachycardia and no intervention noted. the patient remained on support for 3 more days until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8523 | 1220648-2025-29295 | 22205706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22205706&pc=OZD | 05/14/2025 | 07/15/2025 |  | 06/13/2025 |  |  | 05/14/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122628 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and decompensated right after implant. vasopressin, epinephrine, and amiodarone drips were started. the patient had suction alarms and performance level p-6 was the highest level able to be achieved. echocardiogram had poor windows; however, the physician was content with the impella cp position. the patient continued to have poor oxygenation and potential plan for extracorporeal membrane oxygenation (ecmo). the patient continued to be labile. the patient was planned for transfer and during transfer the following day, upon automated impella controller (aic) to aic transfer, ¿impella defective¿ alarm appeared. the new aic used and ¿placement signal not reliable¿ alarmed. per transport personnel, the impella shut off three times and restarted. it was noted the impella cord not completely plugged in as small gap was visible between console plug in and impella cord. the patient's mean arterial pressure was in the 70s and ecmo at 3l. impella cord was unplugged to assess prongs, which were all straight then replugged in and still not fitting flush into plug in. the impella cp device was at p-6/ with flows of 2.4l and placement signal was not visible on aic. the physician was okay with muting the placement signal unreliable alarm which was completed with guidance. an echocardiogram was completed and the impella cp measured 3.46cm within mid left ventricle. the patient was noted to also has severe mitral regurgitation and the physician felt, however, the mitral regurgitation was causing a false normal ejection fraction. the patient remains intubated and sedated on support. however, the patient would eventually expire. care was withdrawn. | [5 empty] | The investigation of pump not detected has been completed. the cause of the impella defective alarm was not determined since product was not returned and logs were not available. | Medical safety reviewed the event and determined that the demise outcome is assigned to both the impella cp pump and automated impella controller. there is not enough evidence to exclude the impella cp pump and automated impella controller used as an associated factor in the patient's outcome given event details on both devices. the automated impella controller is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-29292 for the report on the impella cp. the automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8524 | 1220648-2025-29296 | 22205787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22205787&pc=OZD | 05/09/2025 | 07/08/2025 |  | 06/13/2025 |  |  | 05/09/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618538 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications during impella 5.5 support. throughout the course of support, the patient experienced bleeding from their exlap site and abdominal region. blood products were given, heparin was withheld, washouts were performed, wound vac was placed and dressings were changed to treat the bleeds. in addition, the patient developed an infection for which blood cultures came back positive. iv antibiotics were given to the patient to treat the condition. support continued uninterrupted following the noted interventions. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of vascular damage - peripheral vessel, infection/sepsis, and thrombocytopenia has been completed. the impella 5.5 was returned cut at the catheter. no other abnormalities. vascular damage - the root cause of the vascular damage was not determined due to lack of clinical details. infection/sepsis - the root cause of the infection cannot be determined. thrombocytopenia - reported thrombocytopenia on (b)(6) 2025. the cause of the thrombocytopenia was not determined due to insufficient clinical details. h6 health effect - clinical code 4431 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29296. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8525 | 1220648-2025-29297 | 22205826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22205826&pc=OZD | 05/14/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/14/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643309 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia requiring surgical intervention and thereafter would expire. after impella cp device implant, the patient underwent emergent percutaneous coronary intervention with one drug-eluting stent (des) to the left main and three des to the left anterior descending artery with impella cp device implanted for cardiogenic shock management. the patient was transferred to the unit on support. later that day, the patient developed limb ischemia and underwent a fem-fem bypass. the patient continued on impella cp mcs until hours later before being escalated to an impella 5.5 device. the following day after escalation, approximately 16 hours later, the patient would expire. care was withdrawn. | [5 empty] | Medical safety reviewed the event and determined there is not enough information to exclude the impella cp used as an associated factor to the patient's outcome given the device-related ischemia event requiring surgical intervention and patient expiring shortly thereafter. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
8526 | 1220648-2025-29298 | 22206197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22206197&pc=OZD | 05/15/2025 | 06/13/2025 |  | 06/13/2025 |  |  | 05/16/2025 | 06/13/2025 |  |  |  |  | 06/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597071 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a high-risk percutaneous coronary intervention successfully supported by the impella cp. there were multiple balloon inflations and shockwave were needed to deliver 3 des. the patient remained stable throughout the case. after 2 hours of support the was impella weaned and explanted. the access site closed via two perclose and a 6fr angioseal was added due to track ooze. later it was reported that about 21 days the patient was found unresponsive following a motor vehicle accident. a large amount of blood was noted under the patient that was coming from right femoral access site (impella cp access site). paramedics were unable to revive the patient and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8527 | 1220648-2025-29236 | 22192593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22192593&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583041 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support had a ventricular tachycardia (vt) episode. lidocaine drops were initiated. impella flow increased to p6 during the event and no further vt was noted per then nurse. the patient was appropriately weaned down on the impella device without issues. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8528 | 1220648-2025-29237 | 22192694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22192694&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648836 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp placed to support a patient admitted in with acute myocardial/cardiogenic shock. the patient was placed via single access and was supporting when there was limb ischemia bilaterally. both bilateral lower extremity pulses were absent via doppler. the doctor was made aware of the patient needing more support and needing to mitigate complication. the doctor ordered an arterial us study for bilateral lower extremities. the study showed decreased flow to bilateral lower extremities via collateral flow. another doctor was consulted for impella 5.5 transition, but the decision was made to transfer the patient to another facility for left ventricular assist device/transplant workup. two days later, the pump was explanted and the patient was escalated to an impella 5.5. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8529 | 1220648-2025-29238 | 22192753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22192753&pc=OZD | 05/18/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/18/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653736 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, an ischemic stroke was noted. the patient was not waking up from sedation after surgery. care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8530 | 1220648-2025-29239 | 22192826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22192826&pc=PYX | 05/18/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/18/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590529 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an implanted impella rp flex was inserted via the right ij without issue. it was noted that there was no placement signal or waveform and alert that signal was unreliable. they were unable to calculate flows due to this. the physician was not comfortable leaving the device. the device was removed and a new pump was placed. no patient harm was noted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8531 | 1220648-2025-29240 | 22193065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193065&pc=OZD | 03/10/2025 | 10/24/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324023 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The us complainant had a cp pump support placed for a hrpci. the pci was successfully supported by the pump despite the optical signal being at fault/intermittently malfunctioning. the event is not reportable. later the investigating team wished to further evaluate the automated impell controller (aic). aic was not exchanged during the patient use, but there was low nsr across pumps. during aic investigation a self check error/malfunction was found within the data logs. | [5 empty] | H6: investigation finding was inadvertently not included in the previous report. | H4 device manufacture date updated. | The investigation of automated imeplla controller (aic) - system self check failure has been completed. based on the data and device analysis the root cause of the "system self check failure" was determined to be power battery manager (pbm) corrosion due to leakage of the electrolyte from the battery failure alarm super caps h6 type of investigation, investigation findings, and investigation conclusions were erroneously entered on the initial manufacturer device report number 1220648-2025-29240, as the investigation results were available at the time of the initial report. h10 investigation results were was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29240. | Device was returned on 25-apr-2025 and evaluation/investigation is under way. upon completion of the investigation a supplemental will be filed.
8532 | 1220648-2025-29241 | 22193095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193095&pc=OZD | 05/18/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/18/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592359 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and pre procedure had chronic gastrointestinal (gi) bleeds. a day after impella 5.5 implant, the patient still had low hemoglobin and shock liver. the patient received a total of four untis of packed red blood cells, one unit of platelets, two units of cryoprecipitate, and two units of fresh frozen plasma. anticoagulation was on hold until computed tomography output slowed. the care team continued to replace volume loss. the gi team was consulted as the patient had a history of chronic gi bleeds. it was also noted that the patient had rhythm swaps between normal sinus rhythm and atrial fibrillation. the following day it was confirmed that the patient had a gi bleed. the patient continued to receive units of packed red blood cells over the next few weeks. the patient was not a candidate for advanced heart faliure therapies and consulted interventional cardiology for possible high risk percutaneous coronary intervention. the patient and family decided to go on hospice and impella 5.5 successfully weaned. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8533 | 1220648-2025-29242 | 22193308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193308&pc=OZD | 05/09/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486952 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a 5.5 pump placed to support a patient with known cardiac and systemic comorbidities and nstemi. the pump was placed and supporting when after 4 days there was a need for cardioversion due to afib rvr confirmed to be not due to the use of the pump. the pump remained on and on day 15 of the support there was a cva. the cva caused limb weakness but thrombectomy was not performed and patient remained on the anticoagulation regimen systemically. the pump purge is sodium bicarb and not heparin. the patient remained on support for another day til wean and explant. patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
8534 | 1220648-2025-29243 | 22193385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193385&pc=OZD | 05/18/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/18/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644268 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon turning the support level of an implanted impella cp pump to p-3, the patient endured an immediate responsive run of ventricular tachycardia and junctional rhythm. the pump was returned to the previous setting and the device continued to support uninterrupted. the patient outcome was stable. | [5 empty] | The impella device was discarded, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the venticular tachycardia was not determined.
8535 | 1220648-2025-29244 | 22193591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193591&pc=OZD | 05/20/2025 | 08/25/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639803 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump stopped unexpectedly during patient support. attempts to restart the pump were unsuccessful and the decision was made to remove the device. | [5 empty] | Device was returned on 04-jun 2025 and an investigation is underway. upon completion of the investigation, a supplemental will be filed. instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Corrections to the initial mfr report: g1 mdr reporting contact email. h6 type of investigation code added.
8536 | 1220648-2025-29245 | 22193653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193653&pc=OZD | 05/20/2025 | 10/24/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653590 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. the implanted pump began to trigger suction alarms roughly nineteen days into support. volume was given but the issue persisted. attempts were then made to reposition the device, but the staff was unable to rotate the pump away from the mitral valve. a placement signal unknown alarm subsequently sounded, for which the team was instructed on how to silence the alarm. support was continued with the setup and the pump was successfully repositioned under echocardiogram the following day. no patient harm was reported. | [5 empty] | H4 device manufacture date updated. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for analysis. the data logs show a loss of placement signal. placement signal was seen trending down for about 5 hours before it went out of range. the pattern identified matches that of a durability sensor wear. the placement signal not reliable alarm was confirmed. placement signal did not recover during support. the root cause of the optical signal issue was most likely due to a sensor wear as evidenced by durability sensor wear pattern observed in the data logs. the root cause of the positioning issue was not determined as limited clinical information related to how pump became mispositioned was provided. the root cause of the low pump flow was not determined as no product was returned, no motor current spike was observed in the data logs and limited clinical information was provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-29245.
8537 | 1220648-2025-29246 | 22193837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193837&pc=OZD | 05/16/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/16/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619061 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump became stuck under the papillary muscle and could not be repositioned bedside. the patient was subsequently taken to the cath lab, for proper repositioning of the device and the support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8538 | 1220648-2025-29247 | 22193904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193904&pc=OZD | 05/20/2025 | 07/03/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649047 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp placement, the pump was unable to pass through the vasculature. therefore, the procedure was aborted due to the patient¿s anatomy. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to the patient's anatomy.
8539 | 1220648-2025-29248 | 22193957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22193957&pc=PYX | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025611152 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Unspecified Infection | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was on va ecmo for pulmonary htn, however was referred to heart team for impella rp placement and ecmo removal. rp was successfully placed and va ecmo was removed. it was noted that during rp support the patient was hemodynamically stable and weaning down. however, there was a bump in lactate and fever was reported. the patient was started on antibiotics. additionally, it was noted that there was an active purge system blocked alarm. they checked for kinks and changed out the pf bag/cassette prior to calling to no available resolution. also noticed a drop in flows. currently on p7 with flows of 2.2. tpa protocol was initiated. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8540 | 1220648-2025-29249 | 22194011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22194011&pc=OZD | 05/20/2025 | 07/02/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604706 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an attempted impella 5.5 implant and during the implant, the staff was unable to advance the impella 5.5 past subclavian artery. an angiogram was performed revealing significant stenosis. interventional cardiology was attempted and could not advance pump either. the implant was aborted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of the delivery issue was most likely patient condition related, unable to advance 5.5 past subclavian artery.
8541 | 1220648-2025-29250 | 22194068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22194068&pc=OZD | 05/17/2025 | 07/14/2025 |  | 06/12/2025 |  |  | 05/17/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648158 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was being implanted with an impella cp device for mechanical circulatory support and the attempt to place the impella cp was aborted due to a femoral angiogram that was taken, and the vessel was deemed too small for impella delivery. no surgical backup was available for an axillary approach, so the pump insertion attempt had to be aborted due to patient anatomy. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related to patient's small vessels as stated in the clinical. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8542 | 1220648-2025-29251 | 22194170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22194170&pc=OZD | 05/20/2025 | 07/07/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  | 08/31/2025 |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392910 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that hematoma was noted at the impella cp insertion site upon arrival at icu. fem stop was placed. it was further reported that the patient was taken to the or to repair the access site. no other information was provided. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided.
8543 | 1220648-2025-29252 | 22194259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22194259&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626889 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was being placed to support a patient during a post-op bentall procedure. the chest was left open and washed out. the team attempted to deliver the 5.5 pump with multiple wires and diagnostic catheters but, the pump could not be fully delivered and was aborted. no harm or injury from any delivery attempts was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8544 | 1220648-2025-29221 | 22194390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22194390&pc=OZD | 05/19/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/19/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145434 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 105 KG |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a purge disc not detected alarm. troubleshooting was unsuccessful and the aic was exchanged. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. the imc logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined after arriving in fs. a broken purge disc flag was identified. the root cause of this issue was a broken purge disc flag.
8545 | 1220648-2025-29253 | 22194565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22194565&pc=OZD | 05/20/2025 | 07/03/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655586 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and upon explant, the graft was exposed and the impella was removed through the graft. when the impella was removed, the cannula was detached from the outflow on motor housing catheter. the cannula was at the proximal end of the graft and was removed. graft was stappled and incision was closed. | [5 empty] | Device was returned on 28-may-2025 and an evaluation/investigation is underway. upon completion of the investigation, a supplemental mdr will be filed. | Product was returned. there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. this issue has previously been seen on direct cases during removal and can be referenced in the previous escalation pia-0000148. the design testing has stated that the bond is tested to a 30 n specification per dir550-1.1.7.2. the product had a clamp mark but it was stated that that was from retrieving the cannula. additionally, the bond can be seen attached to the motor housing with no rips or any other indications of damage. on the inside of the cannula there was also residual glue marks seen, indicating there was adhesion of the cannula to the motor housing. there was a slight bend in the catheter near the motor housing but no kinks. there was no clot found. due to a lack of sufficient evidence during the removal of the 5.5, the root cause of the product damage (detached inflow/outflow) cannot be determined.
8546 | 1220648-2025-29254 | 22195003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22195003&pc=OZD | 05/20/2025 | 07/07/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143006 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella automated impella controller (aic) had a controller error alarm. the aic was exchanged. there was no reported patient harm. | [5 empty] | Device was returned on 21- may - 2025 and evaluation/investigation is underway. upon completion of the investigation, a supplemental mdr will be filed. | Data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. see the attached imc_ic2698.pdf. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. see the attached image rt_log.png. device analysis: the console was sent back to field service for further investigation and was tested after arriving in fs. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. further, optical bench was checked and no distortion in the analog output or issues were found. root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue (optical bench fw anomaly). corrective actions: before returning the console to the customer, the following actions are instructed to perform: perform section 13 ¿ 15 from preventive maintenance procedure (0042-7309). perform a full functional check on the console and optical system (0042-7316).
8547 | 1220648-2025-29255 | 22195258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22195258&pc=OZD | 05/19/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/19/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635390 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient had two runs of ventricular tachycardia that required defibrillation. amiodarone was running. the patient continued on support until the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8548 | 1220648-2025-29256 | 22195371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22195371&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577892 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 68 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient presented implanted with an impella cp and with a mottled right leg. the patient was taken to the cath lab to be escalated to an impella 5.5. after an external fem-fem bypass distal pulses were present. the patient expressed they did not want extreme measures taken and care was withdrawn. the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8549 | 1220648-2025-29257 | 22195437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22195437&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664078 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support and had a run of ventricular tachycardia requiring cardioversion. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8550 | 1220648-2025-29258 | 22195807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22195807&pc=OZD | 05/20/2025 | 07/02/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645851 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on day 20 of support ,computed tomography (ct) head showed a small subarachnoid bleed. no neurology intervention was planned. on day 21, the motor current was high and the impella stopped. multiple attempts to restart the impella were unsuccessful and the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of pump stop and stroke/cva has been completed. the impella device was not returned for evaluation. pump stop: data logs were reviewed and the real time (rt) log at the time of the pump stop showed a spike in motor current with a speed deviation and complete saturation of pump flow. a small deviation in placement signal was also noted and purge pressure increased slightly with a decrease in purge flow. immediately after the motor current spike, motor current stepped up without a change in p-level. another motor current spike with speed deviation and momentary drop in pump flow occurred 5 minutes before the pump stop occurred. pump was attempted to be restarted without success. clinical details noted that pump had been running for 21 days and patient was not being anticoagulation due to bleeding. the root cause of the pump stop could not be definitively determined as insufficient evidence was provided without the returned product. stroke/cva: noted that patient had subarachnoid bleeding on ct. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. instructions for use: impella cp with smart assist system for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h6 health effect - clinical code 1888 was erroneously entered on the initial manufacturer device report number 1220648-2025-29258. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29258.
8551 | 1220648-2025-29259 | 22195898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22195898&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662844 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during implantation of an impella cp the pump folded over itself on the catheter portion behind the motor. the pump was stuck in the patient and after several maneuvers the pump was unfolded and removed. a new impella cp was successfully implanted to support the patient. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8552 | 1220648-2025-29260 | 22196097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196097&pc=OZD | 05/20/2025 | 07/17/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661970 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Defective Device | Device-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | 14 FRENCH INTRODUCER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was attempted. the physician was unable to advance the companion sheath over the wire past the hemostatic valve. the companion sheath was replaced to successfully complete the procedure. additionally, the physician had difficulty removing the wire from the peel away sheath. the physician suspected the hemostatic valve was defective. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Single access issue: clinical details reported that when trying to do a single-access pci procedure, the 7fr companion sheath failed to advance through the hemostatic valve at the 10:00 position. it was replaced with a new companion sheath, which was stuck at the 2:00 position, and no further issues were noted. product was not returned for investigation. since product was not returned for investigation and limited clinical details were provided, the root cause of the single access issue could not be determined. introducer damage - mechanical: during a single-access procedure, the 7fr companion sheath was unable to stick the hemostatic valve at the 10:00 position. a second companion sheath was able to stick at the 2:00 position. when removing the guidewire from the peel away sheath, there was difficulty noted by the physician. for these reasons, it was believed that the hemostatic valve may be defective. product was not returned for investigation. since insufficient clinical details were provided and product wasn't returned for investigation, the root cause of the introducer damage could not be determined.
8553 | 1220648-2025-29261 | 22196135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196135&pc=OZD | 05/20/2025 | 07/14/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp exhibited bleeding around the impella sheath as well as the endotracheal tube and the nose. in response, the pump was explanted. the patient survived. | [5 empty] | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleed: product was not returned for investigation. though there were several contributing factors reported in clinical details, the root cause of the access site bleed was determined to most likely be a use issue due to not placing forward tension sutures. vascular damage: since limited clinical details were provided and the pump wasn't returned for investigation, the root cause of the vascular damage could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high-risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8554 | 1220648-2025-29263 | 22196317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196317&pc=OZD | 05/20/2025 | 07/03/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661020 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient began to experience pain in their right leg during impella cp support. minimal flow was shown in the leg and the patient was brought to the operating room for either an external fem-fem bypass or explant. upon weaning the pump down, the patient went into ventricular tachycardia and a shock was required. a temporary pacer was placed and the procedure was aborted. support continued with the current setup with both legs being monitored. | [5 empty] | The investigation of vf/vt/af/at and limb ischemia has been completed. impella cp was returned and no damages or abnormalities were found. limb ischemia: as all the necessary information was not provided, the root cause of the limb ischemia was not determined. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8555 | 1220648-2025-29262 | 22196323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196323&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659385 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and exhibited red urine. the pump was found to be shallow in the left ventricle and was repositioned to resolve the hemolysis. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8556 | 1220648-2025-29264 | 22196417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196417&pc=OZD | 09/16/2024 | 07/08/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544689 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received from the clinical site "hb results: (b)(6) 2024: 11.4 mg/dl, (b)(6) 2024: 8.5 mg/dl, (b)(6) 2024: 9 mg/dl, (b)(6) 2024 8.6 mg/dl. 16se 2025: the right groin; new sutures and stitching. pi relatedness: the adverse event is ¿possible¿ related to the impella device." | It was reported through abiomed's protect iv study that a patient on impella cp support had bleeding from the impella access site. one unit of packed red blood cells was transfused. bleeding was resolved on the same day. the patient was successfully weaned from the pump and survived. |  |  |  |  | B.5 additional information was received regarding further treatment received by the patient. h.6 added code to reflect treatment the patient received. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-29264. e.1 revised facility name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-29264. g.1 added mdr reporting address as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-29264. | A4 and a5, ethnicity, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8557 | 1220648-2025-29265 | 22196728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196728&pc=PYX | 05/20/2025 | 07/02/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578250 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 67 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex exhibited signs of hemolysis. attempts to reposition the pump were unsuccessful. the patient was successfully weaned from the pump and survived. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The pump was returned for investigation. according to the clinical notes, hemolysis was reported on 20-may in the afternoon, with some resolution noted. later in the evening, the md was unable to reposition due to hemodynamics, but lab values were improving despite the same-colored uop. later that morning, the uop turned pink-amber, and some resolution was observed after giving volume. by 21-may in the afternoon, the uop was red but slightly improving, and the md decided to keep the pump running at p5. the data log suggests a motor current spike with several suction alarms and low pump flow on the morning of 20-may, with trends returning to normal in a short time. there were fluctuations in pump flow while running at a steady p-level until the night of 20-may. later, higher pump flow was noted while running at p5, with some resolution and a return to high pump flow over time, accompanied by a negative placement signal. no issues were noted regarding purge parameters. the cannula was found kinked on the returned product. no other physical damage was noted. biomaterial was found on the impeller, which was flushed during the test run. the pump was able to run as expected in a bucket of water. the cause of the hemolysis issue was most likely biomaterial, based on returned product investigation and data log review. the cause of the positioning issue was not determined since sufficient clinical information was provided. thrombus failure mode was added as an investigated failure mode since it was found on the returned product and could have contributed to hemolysis during the case run. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
8558 | 1220648-2025-29266 | 22196735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196735&pc=OZD | 05/26/2025 | 09/17/2025 |  | 06/12/2025 |  |  | 05/26/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094546 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an automated impella controller (aic) had a controller failure alarm. the aic was exchanged to continue patient support. no patient harm was reported. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | Data analysis: on 5/19 5.5 pump (b)(4) was connected. starting at 16:36:32 on 5/20 there was an invalid placement signal that two minutes later triggered a placement signal not reliable alarm to the user ((b)(6)). the imc logs show on 5/25 there was intermittent #(b)(4) event alarms for crc error that would intermittently trigger and resolve until the pump is unplugged on 5/26 ((b)(6)). the lt logs show that the placement signal went to 3000 mmhg and remained for the duration of the case ((b)(6)). the rt logs show that on 5/20, in addition to the placement signal, the raw optical sensor went to -3276.8 and snr went to nearly 0 ((b)(6)). this pattern of optical signals is indicative of an air gap. starting on 5/25 at 22:43:20, all of the 5s signals would intermittently go to -16384 before returning to the raw optical previous 5s signals where raw optical sensor is -3276.8 and snr is near 0. device analysis: neither of the failure modes naturally reproduced. the initial failure could be reproduced by slightly disconnecting the pump where there is electrical connection but an optical gap. root cause: there was no optical signal issue found with the console. the cause of the optical signal failure was most likely a damaged sensor during the shockwave procedure see 25-44648-2. there was no evidence of a battery failure. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. copy and paste "device subcomponent" here: [0042-96254]. repair: please clean the optical connector (0042-2736) then perform a functional check (0042-7316) on (b)(6). | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-29266. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-29266. h10: additional information regarding the dhr review.
8559 | 1220648-2025-29267 | 22196815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196815&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661966 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placement occurring for the cardiogenic shock patient with coronary artery disease. the pump was unable to be delivered due to iliac stenting. the patient went to or for the coronary artery bypass graft when the percutaneous cardiac intervention failed. patient continued on ecmo support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8560 | 1220648-2025-29268 | 22196824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22196824&pc=OZD | 05/20/2025 | 07/07/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632457 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5, the surgeon inadvertently advanced the impella 5.5 through the aortic valve leaflet. support was continued until the patient was successfully weaned from the pump. | [5 empty] | The returned product was observed and there were no damages noted with the pump. based on the clinical details, the root cause of the heart valve damage is most likely due to use as the pump was pushed through the av leaflet during attempted direct placement of the pump. | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter. ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
8561 | 1220648-2025-29269 | 22197000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22197000&pc=OZD | 05/27/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/27/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient lost distal pulses due to the pump occluding distal flow. the pump was removed and pulses returned to baseline. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8562 | 1220648-2025-29270 | 22197050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22197050&pc=OZD | 05/20/2025 | 06/12/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659222 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp was prepped and inserted in a 83 year old male patient. a percutaneous transluminal coronary angioplasty (ptca) was performed and then a balloon post-dilatation was performed on to confirm results. impella was weaned and explanted post procedure. however, vascular surgeon was called to removed 14 fr. sheath and an attempt at perclose was made without achieving hemostasis. the patient was then transported to or for vascular repair. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8563 | 1220648-2025-29271 | 22197182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22197182&pc=OZD | 05/27/2025 | 07/16/2025 |  | 06/12/2025 |  |  | 05/27/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662016 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp attempted to be placed for a patient admitted in with a non-st-segment elevation myocardial infarction and undergoing high-risk percutaneous coronary intervention. it was reported that the patient has heavily calcified illiacs. a long impella peel away sheath crossed, however, the impella pump could not cross. impella placement and percutaneous coronary intervention were aborted. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of the delivery issue was most likely patient condition related, impella unable to pass through as patient has heavily calcified iliacs.
8564 | 1220648-2025-29272 | 22197251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22197251&pc=OZD | 05/27/2025 | 07/11/2025 |  | 06/12/2025 |  |  | 05/27/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539110 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was not able to pass through the graft. several attempts were made but unsuccessful. the impella 5.5 insertion was aborted and the impella cp was placed. the patients outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely use issue since the doctor used vascutek gelweave 8mm along with steel core wire both of which are non-abiomed product which are not recommended to use per instructions for use. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8565 | 1220648-2025-29273 | 22197278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22197278&pc=OZD | 05/20/2025 | 07/08/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657788 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, however the impella was wrongly positioned in the in the papillary muscle. decision made to pull out impella and place intra-aortic balloon pump. | [5 empty] | The investigation of positioning issues has been completed. pump was not returned for investigation. data log review was not required for this case. the cause of the positioning issue was most likely use-related since it was mistakenly placed in the papillary muscle. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8566 | 1220648-2025-29274 | 22197614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22197614&pc=OZD | 05/20/2025 | 07/14/2025 |  | 06/12/2025 |  |  | 05/20/2025 | 06/12/2025 |  |  |  |  | 06/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed without difficulty. after the pci procedure while the 14fr short sheath was being removed, impella cp came back across the aortic valve. when trying to re-advance into the lv, the impella prolapsed, and the pigtail was hooked on the cannula. we got radial access and advanced a snare to catch the pigtail of the impella. simultaneously, we pulled back with the snare attached to the pigtail and pulled back on the 9 fr impella catheter at the right groin site and were able to straighten the impella out. the cp was explanted and a new cp was implanted successfully. | [5 empty] | The investigation of positioning issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. positioning issues: the cause of positioning issues most likely was use related due to improper technique as impella cp came back across the aortic valve while the sheath was being removed. product damage (cannula kink): the cause of product damage (cannula kink) issue most likely was cannula kink due to improper technique as it occurred when trying to re-advance impella into the left ventricle (lv), the impella prolapsed, and pigtail was hooked on cannula. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8567 | 1220648-2025-29182 | 22182750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22182750&pc=OZD | 05/23/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/23/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was inserted to support a patient during percutaneous coronary intervention procedure. during support, the patient had ischemia. the patients superficial femoral artery appeared less than 5 millimeters. after the surgeon successfully placed the impella device, it was noted that placing the companion sheath was very difficult. the side port was used to take an angiogram of the vessel. distal flow was compromised. the decision was made to remove the 14 french peel away sheath and place a repositioning sheath. a second angiogram showed flow was restored. the peel away sheath was removed and the repositioning sheath was sutured in place for the remaining case. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8568 | 1220648-2025-29183 | 22182800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22182800&pc=OZD | 05/24/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/24/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671047 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the following day, the patient continued to have chest pain overnight and developed arrhythmias, going in and out of normal sinus rhythm. the patient went into second and third degree blocks. nitroglycerin was started for the chest pain and a temporary pacer was placed. the patient continued on support and three days later ejection fraction was noted to have recovered from 17% on admission to 48%. neurological status was intact. the next day, the patient was successfully weaned and the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8569 | 1220648-2025-29184 | 22182859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22182859&pc=OZD | 05/24/2025 | 07/15/2025 |  | 06/11/2025 |  |  | 05/24/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient experienced access site bleeding the following day. it was noted that te impella site remained "oozy" along with all puncture sites (extracorporeal membrane oxygenation, central lines). a total of four units of blood and one unit of platelets was administered. heparin was briefly stopped and restarted. left lower extremity pulses were dopplerable. urine remained bluish/green (from methylene blue) and adequate. the patient continued on support and work up for transplant. the patient was reported to be in stable condition following the event. four days later, it was noted, the patient was successfully weaned and the impella cp was explanted with a survived outcome. care was escalated to an impella 5.5. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient condition because the of the bleeding cause from all the puncture site along with all the access site. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8570 | 1220648-2025-29185 | 22182998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22182998&pc=OZD | 05/16/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/16/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553841 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 66 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impell 5.5 support had intermittent ventricular tachycardia. milrinone was stopped and the patient was started on amiodarone. the patient survived. | [5 empty] | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8571 | 1220648-2025-29186 | 22183024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183024&pc=OZD | 05/17/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/18/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653825 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was placed on a patient and would be worked up for ventricular assist device. initially during impella 5.5 support it was noted that the patient was doing well and hemodynamically stable. however, several days after support started the patient began to have a gastrointestinal bleed and heparin was being held. additionally, the patient was clostridioides difficile positive. it was noted that after that, the patient continue to have bleeds with dark stools and the patient received several units of blood. ultimately, it was noted that the patient began having abdomen pain resulting in a computed tomography that revealed free air in abdomen. general surgery was consulted. however, the patient began to decompensate quickly. it was believed that the patient had ischemic bowel and surgery was deemed not to be an option. the patient continued to decompensate and expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8572 | 1220648-2025-29187 | 22183039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183039&pc=OZD | 05/16/2025 | 07/31/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646988 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male | 105 KG | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a ventricular tachycardia ablation was being performed on a patient implanted with an impella cp. while mapping with an optrell catheter, the catheter got tangled with the pigtail of the implanted impella cp pump. with some difficulty, the physician was able to free the catheter and remove the optrell device. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation of device interaction or damage by another device has been completed. the impella device was not returned for evaluation. based on the clinical description, the root cause of the device interaction is most likely due to use as the optrell catheter became caught in the pigtail of the impella during mapping.
8573 | 1220648-2025-29188 | 22183096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183096&pc=OZD | 05/17/2025 | 07/14/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597153 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed, pink-tinged urine. additionally, the patient had a placement signal low alarm for about 10 min in the cath lab. the impella was repositioned to resolve the issues. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis and positioning issues has been completed. reported the pump was placed too deep when implanted and placement signal low alarms occurred. pump was pulled back to good position after 10-15 minutes and alarms resolved. patient had hemolysis a day after. no clot was noted and no blood product were required. fluid management cleared hemolysis. no product was returned. data logs show low placement signal placement signal low alarms when the pump is started. there was a step up in placement signal to normal range after 15 minutes and alarms resolve. positioning issue - the root cause of the positioning issue was most likely use related since the pump was placed too deep by the physician. hemolysis - the root cause of the hemolysis was most likely patient condition related since fluid management resolved the hemolysis.
8574 | 1220648-2025-29189 | 22183106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183106&pc=OZD | 05/17/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657623 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella 5.5 support experienced non-sustained ventricular tachycardia and atrial fibrillation with rapid ventricular response. the patient received a unit of blood and hemodialysis was initiated to remove fluid. a few days later it was noted that the patient was going in and out of atrial fibrillation with rapid ventricular response again. rhythm was tried to be controlled prior to removing the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8575 | 1220648-2025-29190 | 22183189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183189&pc=OZD | 05/17/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642516 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 137 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited bleeding and a hematoma. heparin was withheld and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8576 | 1220648-2025-29191 | 22183329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183329&pc=OZD | 05/17/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656317 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Infusion or Flow Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support had ventricular tachycardia/fibrillation requiring one shock. an echocardiogram was done for placement after shock and no adjustments was required. flow saturation was also noted that same day. tissue plasminogen activator was initiated. additional information was provided that the patient continued to deteriorate and required more pressor/inotrope support. as a result, the family decided to withdraw care. the patient expired shortly after. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8577 | 1220648-2025-29192 | 22183518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183518&pc=OZD | 06/03/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 06/03/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information was obtained from a literature report (b)(4) reporting that a patient presented with acute myocardial infarction. the patient was managed with ecpella because of hemodynamic compromise. however, a week later, there was a sudden increase in venous oxygen saturation. transthoracic echocardiography revealed ventricular septal rupture, and free wall rupture. intra-ventricular thrombus was also observed despite standard anticoagulation therapy. even with double cardiac rupture, ecpella could facilitate left ventricular unloading and sustain hemodynamics. however, because of the risk of device failure due to thrombus aspiration into the lmpella cp, the patient underwent repair surgery. postoperatively, the patient was temporarily weaned off of ecpella, and his hemodynamics deteriorated again, and expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | A.1, a.4, and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system: section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8578 | 1220648-2025-29193 | 22183692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183692&pc=OZD | 05/27/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 06/03/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647425 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a percutaneous coronary intervention (pci) was performed with intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation. after pci was completed, the impella cp was inserted in place of the iabp, but the femoral artery dissected, so insertion was discontinued. per the doctor it is unclear when the dissection occurred (when the iabp was removed or when the impella sheath was inserted. the patient recovered and survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8579 | 1220648-2025-29194 | 22183790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183790&pc=OZD | 06/04/2025 | 08/29/2025 |  | 06/11/2025 |  |  | 06/04/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, a "pump stopped_aic failure" alarm occurred and the pump stopped. the automated impella controller was replaced and the pump was restarted. support continued and the patient remained stable. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. the impella pump was returned for investigation. the console displayed ¿impella stopped. controller failure,¿ event #1047, due to a pmd communication issue for more than 17 seconds. rt logs confirm the pump stop (rt log pump stop.png). this aligns with the reported failure mode. this console was tested and unable to reproduce the reported failure mode after running the pump for 1 hour. measured 20.12 v on connector j001.2 of the impellatronic board, confirming no issue with the pbm. the ribbon cable is appropriately connected between the impellatronic and the 4-in-1. the root cause of the intermittent pmd communication issue was most likely the impellatronic board malfunction.
8580 | 1220648-2025-29195 | 22183909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22183909&pc=OZD | 06/09/2025 | 09/16/2025 |  | 06/11/2025 |  |  | 06/09/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316937 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on the impella 5.5 support and during support, there was a purge system failure alarm that occurred when replacing the purge set. even after replacing the purge set with a new one, there was no improvement and priming did not proceed. the automated impella controller (aic) was replaced and the issue resolved. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. data analysis: the case began on (b)(6) 2025, with the first purge cassette change initiated approximately four days later. shortly after the purge cassette was replaced with a new one, a purge system failure (piston block #417) alarm along with complete the procedure ((user interaction stopped) complete the purge wizard #420) alarm was triggered. later, the purge cassette was removed and reinserted, but the alarm persisted. the initially used purge cassette was then reused, yet the alarm continued to remain. finally, a third new purge cassette was utilized, but despite this, the alarm still existed without resolution. ultimately, the console was shut down and replaced. real time log data showed around zero purge pressure and erratic purge flow ticks. device analysis: the console was returned for further investigation. while investigating the aic, a thin layer of glucose deposit was detected underneath the purge stepper motor gasket and piston. no issues were identified with the aic's purge disc or the purge flag. the error was successfully reproduced by connecting and disconnecting the purge cassette. however, even after cleaning the residue around the motor, the issue persisted, indicating it was most likely a mechanical or electrical malfunction within the purge system. the reported issue of purge system failure (piston blocked #417) was not determined, as the exact cause for the issue was not identified. d9 device returned to manufacturer date added d6a implant date should have been left blank on the initial manufacturer device report number 1220648-2025-29195.
8581 | 1220648-2025-29196 | 22184084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184084&pc=OZD | 05/17/2025 | 07/15/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636181 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was placed on a patient that had placement suction occur when the pump was decreased to p-6 level. the patient continued to have intermittent suction events. the impella 5.5 was repositioned, but suction was not resolved. it is believed that the right side of the heart was failing. it was further reported that care was withdrew and the patient expired. there is not enough information to exclude the impella device used as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Complaint device not returned for analyzing. data log reviewed: no mc spike outside of p-level change observed. many suctions occurred during entire impella support with impella position unknown alarms. placement signal slightly trending down at end of support. clinical stated: patient had rv dysfunction and vsd issues, continues to have intermittent suction events, most likely related to right-sided failure and negative volume status. repeated events after a few days of support due to right sided dysfunction. pump repositioned did not resolve issue. blood volume not given. no cannula kinks. no lv thrombus observed. low pump flow: the cause of low pump flow issue most likely was patient condition related due to rv dysfunction right sided failure.
8582 | 1220648-2025-29197 | 22184298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184298&pc=OZD | 05/18/2025 | 07/10/2025 |  | 06/11/2025 |  |  | 05/18/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675399 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that attempts to reposition an implanted impella cp were unsuccessful, as the pump was pulled into the aorta. the pump was explanted and replaced with a new pump. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The cause of the positioning issue most likely was use related due to improper technique since the impella was pulled into the aorta during repositioning by the physician and wasn't able to reach optimal position.
8583 | 1220648-2025-29199 | 22184620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184620&pc=OZD | 05/17/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609377 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Remove | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during implantation of an impella cp the doctor was not able to remove the 0.0018 wire out of the impella. the doctor tried to remove the wire and the impella flipped over and dislocated out of left ventricle. the doctor decided to pull back the impella, including the impella sheath. a new impella was implanted successfully. the patient survived. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
8584 | 1220648-2025-29198 | 22184622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184622&pc=OZD | 05/18/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/18/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513940 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and exhibited access site bleeding. systemic heparin was withheld and approximately six units of blood were transfused. vascular was brought in and did not intervene as they had oversewn the graft prior to closing. they advised to apply a pressurized dressing to the impella 5.5 site. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8585 | 1220648-2025-29200 | 22184671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184671&pc=OZD | 05/30/2025 | 08/19/2025 |  | 06/11/2025 |  |  | 06/02/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026652103 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complainant in japan reported the patient was on impella cp support and during support, a purge pressure low alarm occurred about 7 hours after insertion. the purge set was replaced, but the situation did not improve, and the automated impella controller was replaced, but there was no improvement. while pump replacement was considered, the purge fluid was changed from 5% glucose to 10%, and the purge pressure returned to normal. support continued and there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak pump has been completed. the data logs show that around 7 hours purge pressure low alarms were triggered and the purge flow increased to about 30ml/h. the returned product was purged with a syringe and fluid was noted to be built up and leaking around the reservoir. based on the data log and returned product analysis, the root cause of the purge leak pump is most likely to a damaged sidearm reservoir.
8586 | 1220648-2025-29201 | 22184812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184812&pc=OZD | 06/02/2025 | 08/29/2025 |  | 06/11/2025 |  |  | 06/02/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645673 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that extracorporeal membrane oxygenation (ecmo) and impella cp were placed. after that, impella was set to p-2-4, and the suction alarm was occasionally sounding. the pump flow rate showed 0.0 liters. volume was added and the position was adjusted, but there was no improvement. the impella was replaced due to concerns that the flow rate would not be secured when weaning from ecmo in the future. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the low pump flow has been completed. the returned cp pump was visually inspected. biomaterial observed wrapped around impeller and purge gap. no broken blade or no cannula kink observed. the biomaterial was soaked and removed to run test the pump where low flows were not reproduced. clinical details state that repositioning pump and giving blood volume did not resolve flow issue, presence of lv thrombus was confirmed before insertion. additionally, after removal, when water was turned on outside the body, a small amount of what looked like fine tissue fragments was found. also, even outside the body, the suction alarm sounded when the p-5 level was reached or higher. the cause of the low pump flow issue was most likely biomaterial as observed on the returned product clinical details correlate with finding root cause. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
8587 | 1220648-2025-29202 | 22184908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184908&pc=OZD | 05/17/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660992 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Tachycardia | Ventricular Fibrillation | Ectopic Heartbeat | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella 5.5 support and was having a lot of ectopy believed to be catheter induced and the plan was to reposition bedside under enchocardiogram. four days later, the patient coded and went into ventricular tachycardia (vt)/ventricular fibrillation. the patient was shocked back to normal sinus rhythm. an hour later the patient went into supraventricular tachycardia (svt) and was treated with metoprolol. the following day, the patient was still experiencing intermittent runs of vt and svt. the swan was pulled yesterday in hopes that it was culprit. the plan was to wean and possibly discontinue milrinone today. the discussion was underway to slowly wean and explant the impella 5.5 soon. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8588 | 1220648-2025-29203 | 22184926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22184926&pc=OZD | 06/03/2025 | 08/29/2025 |  | 06/11/2025 |  |  | 06/03/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2025588028 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while attempting to remove a .018 guidewire following impella cp placement, resistance was met and the wire came out frayed. a new wire was subsequently inserted without issue and the pump was successfully implanted. the patient remains on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. the 0.018" guidewire was returned for investigation. a data log review was not required for this investigation. a kink was reported at the guidewire tip. additionally, the tip core was found to be fractured below the observed kink on tip, and the coil was frayed. the cause of the guidewire issue was not determined since sufficient clinical information was not provided. a2 age at the time of event was updated as it was left blank on the initial manufacturer device report number.
8589 | 1220648-2025-29204 | 22185007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185007&pc=OZD | 05/19/2025 | 06/18/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1494422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Use of Device Problem | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automatic impella controller (aic) had a purge failure. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. data review: the console logs were partially consistent with what was reported in the complaint description. a purge error was reported but analysis of the logs show that there was no purge errors but instead there was evidence of a pump detection issue during the purge case start setup. during step 3 of the purge case start, the user was instructed to insert the pump but the console did not detect one for 20 seconds which resulted in a epm reset. after the epm reset there was a connect grey connectors screen which is indicative of an epm crystal issue. device analysis: the console was tested after arriving in field service. the pump detection issue was unable to be reproduced on an engineering lab bench. console had no issues detecting the pump. conclusion: the pump detection issue was due to the rarely occurring epm crystal issue. aic - purge issue - other there was no purge error found during the case. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-29204. h6 medical device problem code 1670, investigation findings 3233 and investigation conclusions 11 were erroneously entered on the initial manufacturer device report number 1220648-2025-29204. h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-29204. h10 narrative was erroneously as the investigation results were complete and the investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29204. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8590 | 1220648-2025-29205 | 22185022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185022&pc=OZD | 06/04/2025 | 09/02/2025 |  | 06/11/2025 |  |  | 06/04/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5-JP |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026656049 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, a "pump stopped_aic failure" alarm occurred and the pump stopped. the automated impella controller (aic) was replaced and the pump was restarted. support continued. | [5 empty] | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. the clinical details state that "a pump stopped- controller failure" alarm occurred and the pump stopped. after replacing the aic, the alarm was cleared and the pump restarted. the pump ran with no issues for the remainder of the case. the data logs show that the pump stopped with a controller failure alarm about 280 hours into support. there were no changes in placement signal (ps), motor current, flow or purge pressure during/before this pump stop occurred. the pump was stated to have been stable and ran for 21 more days. upon review of data logs, it is apparent that the console may be a potential cause of this pump stop, as there was no problem found with the pump. please refer to manufacturer device report 1220648-2025-29194 for investigation results of the console. d6b explant date was added h6 component code 765 was erroneously entered on the initial manufacturer device report number 1220648-2025-29205. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu state: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist for use during cardiogenic shock and high risk cpi section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8591 | 1220648-2025-29206 | 22185048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185048&pc=OZD | 05/19/2025 | 07/15/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661752 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. after about 30 minutes of support, the patient began to have multiple suction alarms and decreased motor current pulsatility and decoupling of waveforms. positioning was optimized under transesophageal echocardiogram, and the patient's volume status was acceptable. an attempt to increase support was made, however, they were unable to flow higher than p4. the decision was made to exchange for a new impella 5.5 due to the high likelihood of motor ingestion of some sort. | [5 empty] | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the pump was returned for investigation. no physical damage was observed on the returned product. pump flow was noted to be saturated for some time and then gradually decreased to low flow (below 5 l/min on p9) during the test run in the bucket of water. the pump was tested overnight, but the flow remained below 5 l/min, accompanied by suction alarms. the data log shows an hour-long case run where pump flow dropped mid-case with suction alarms. suction alarms and low flow were observed while running above p4. no motor current spikes, placement signal trends, or purge-related abnormalities were reported. the cause of the low pump flow issue was not determined without sufficient clinical information during case run.
8592 | 1220648-2025-29207 | 22185101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185101&pc=OZD | 05/19/2025 | 07/15/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560067 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp the cannula kinked. attempts to reposition the pump were unsuccessful. provider repositioned a couple times and kink remained at same place. the pump was explanted and a new pump was implanted successfully. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (cannula kink) and positioning issues has been completed. the returned complaint device visually inspected. cannula kink observed near outlet. no other abnormality observed. clinical details indicate that impella cp was placed over 0.18 impella wire, after achieving optimal position, 0.018 wire removed, and impella cannula kinked, provider repositioned couple time and kink remained at same place, and from another side mentioned insertion aborted due patient anatomy. product damage (cannula kink): the cause of product damage cannula kink issue was most likely patient condition related due to patient's short aorta. positioning issues: the cause of positioning issues was most likely patient condition related due to patient's short aorta. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29207.
8593 | 1220648-2025-29208 | 22185304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185304&pc=OZD | 05/19/2025 | 12/02/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1824431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) was in for preventative maintenance (pm). upon boot up, the aic wan nota able to enable the network. the aic log could not be downloaded. the pm could not be completed as the console failed to communicate. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. device analysis: console was tested after arriving from field service. the reported mobaxterm communication issue was able to be reproduced. an incorrect configuration of the consoles network settings was observed while troubleshooting the issue. root cause: the root cause of the mobaxterm communication issue was identified as a misconfiguration of the consoles network settings. h6 investigation type, findings and conclusions h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8594 | 1220648-2025-29209 | 22185331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185331&pc=OZD | 05/18/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681938 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 54 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Nerve Damage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp the patient experienced access site bleeding. two units of blood products were transfused to the patient. additionally, the patient was not responding to pain in the impella leg and experienced two episodes of ventricular tachycardia. support continued until the patient was successfully weaned off the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
8595 | 1220648-2025-29210 | 22185384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185384&pc=PYX | 05/19/2025 | 07/14/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1198720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a yellow alarm, "placement signal not reliable," and was unable to display blood flow, "flow calculation disabled." the staff switched the catheter back to the hospital's aic and all the waveforms and blood flow displayed correctly. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: rp pump (b)(4) was plugged into (b)(6) on 5/2 having run for 1685 minutes and immediately triggered an event #1052 for placement signal not reliable with an rp dual sensor ((b)(6) imclog psnr .png). the rt logs confirm for the duration of the pump connected to (b)(6), the optical placement signal was fine but the electrical (differential) placement signal was around -6000 mmhg which is inaccurate ((b)(6) rtlog psnr dual sensor.png). device analysis: the failure mode was reproduced as upon running a known good rp pump, the placement signal not reliable alarm immediately triggered and remained until the pump was unplugged. the rt log from this test showed electrical placement signal was -5350 mmhg which again was drastically inaccurate ((b)(6) rtlog test pump psnr.png). the impellatronic board was replaced with a known good and the console performed as expected with both electric and optical placement signal reasonable. root cause: the cause of the malfunctioning differential placement signal was a defective impellatronics board. repair: please replace the impellatronic board (0042-1125 or equivalent) then run an rp pump for 10 minutes confirming no placement signal not reliable alarms and perform the scheduled pm (0042-7309)
8596 | 1220648-2025-29211 | 22185416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185416&pc=OZD | 05/30/2025 | 12/04/2025 |  | 06/11/2025 |  |  | 06/02/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1336513 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, a purge pressure low alarm occurred about 7 hours after insertion. the purge set was replaced, but the situation did not improve, and the automated impella controller was replaced, but there was no improvement. while pump replacement was considered, the purge fluid was changed from 5% glucose to 10%, and the purge pressure returned to normal. support continued and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data analysis: a review of the imc logs reveals the occurrence of 5 "purge pressure low" alarms, which is consistent with the complaint. additionally, the rt and lt logs show that before these alarms began to occur, the purge pressure had slowly decreased from between ~400 and 500 mmhg to ~300 mmhg. device analysis: the console was not returned for investigation. root cause: the cause of the issue was not determined since product was not returned. s d10 concomitant medical products and therapy dates updated.
8597 | 1220648-2025-29212 | 22185818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185818&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658585 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Bruise/Contusion | Tachycardia | Swelling/ Edema | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. three days into support, the patient endured an episode of ventricular tachycardia while in the computed tomography. the patient was cardioverted back to sinus rhythm. in addition, the cutdown site was seen to have some swelling and bruising which was not acutely worsening. three days later, the anticontamination sleeve was found to be damaged at the distal end. the sleeve was cleaned and wrapped with tegaderm to manage the damage. support was continued uninterrupted and the procedure outcome was successful. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8598 | 1220648-2025-29213 | 22185858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22185858&pc=OZD | 05/31/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 06/02/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656053 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Septic Shock | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support the patient developed septic shock but the timing of onset and treatment are unknown. in addition, since many catheters were inserted in addition to impella, such as extracorporeal membrane oxygenation, soluble guanylate cyclase, and controlled ventilation, it was unclear whether there was a direct causal relationship between sepsis shock and impella. the doctor reported that there were no problems with the impella sterilized sleeve or skin problems at the impella insertion site. it was also reported that multiple organ failure progressed due to sepsis shock and the patient died. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8599 | 1220648-2025-29214 | 22186066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186066&pc=OZD | 06/02/2025 | 08/28/2025 |  | 06/11/2025 |  |  | 06/05/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426098 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that at the end of the case, the aic (console) could not be turned off or forced to shut down. after trying several times, the screen went black and an abnormal noise was heard. when it was restarted, a alarm message about an aic abnormality appeared. the abnormal sound occurred in (b)(6), impella 5.5 was replaced. there was no reported patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. e1- initial reporter contact name unknown. | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. the aic and data logs were received for investigation. based on the data log analysis and device evaluation the reported issue was unable to be confirmed or reproduced. the console functioned as designed.
8600 | 1220648-2025-29215 | 22186204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186204&pc=OZD | 05/19/2025 | 07/14/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620845 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the surgeon attempted to reposition the pump and when repositioning, the surgeon pulled the pump back across the aortic valve and was unable to readvance back into the ventricle. the decision was made to remove the impella 5.5 and reinsert the wire and catheter across the valve. once the pump was removed, there was significant tissue noted on the inlet cage making the surgeon concerned about reinserting the same pump. the decision was made to abort and use a new device. | [5 empty] | The investigation of positioning issues and thrombus has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump accidentally was pulled out of the lv during reposition. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8601 | 1220648-2025-29216 | 22186519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186519&pc=OZD | 05/17/2025 | 07/14/2025 |  | 06/11/2025 |  |  | 05/17/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664796 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the patient had hemolysis. some blood was noted in the urine. the pump appeared deep on echocardiogram. the pump was pulled back and the p level was decreased. the hemolysis improved and the urine cleared. plasma free continued to be elevated at 120, which was down from 230. support continued. the device was successfully weaned and the patient survived the explant. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the pump was not returned for investigation. the data log shows no conclusive trends on the placement signal, motor current, or any reported positioning alarms during time of reported hemolysis. according to the clinical notes, hemolysis improved after the pump was repositioned when it was discovered to be deep. the cause of the hemolysis issue was most likely improper positioning of the device, based on provided clinical details. the cause of the positioning issue was not determined since sufficient clinical information was not provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8602 | 1220648-2025-29217 | 22186522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186522&pc=PYX | 04/02/2025 | 09/25/2025 |  | 06/11/2025 |  |  | 04/02/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025570989 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Pump | Device Sensing Problem | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the staff was unable to adequately calibrate / zero the signal of an implanted impella rp pump. roughly six days into support, positioning issues were encountered following swan ganz catheter removal and the pump stopped unexpectedly multiple times. the pump was replaced in response to the noted failure. further details state that the patient was unstable and care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | H6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-29217. d10 concomitant medical products and therapy dates updated. h6 medical device problem code 3009 positioning problem has been updated. | The pump was returned for investigation. no physical damage was observed on the returned product. the pump passed all electrical resistance tests on the motor cable and the dp sensor. the pump was further connected to the aic, and the impella stopped; a controller failure alarm was observed during the start of the pump. a significant amount of biomaterial was observed on the impeller. high purge pressure was reproduced during the test. further teardown of the inflow cage was performed to clear the biomaterial. after clearing the biomaterial, the pump was able to run in a bucket of water with no purge-related issues and also confirmed that the purge gap appeared normal on the pump. pump stop: the data log shows several pump stop alarms with motor current spikes at the end of the case run on day 7. purge pressure was slightly elevated during the pump stop event. the cause of the pump stop issue was biomaterial ingestion during the case run, based on returned product investigation and data log review. placement signal issue: the lt log shows notifications of completing the zeroing sensor event. additionally, the imc log was reviewed and noted that the zeroing of the dp sensor was per specification and that the correct offset was maintained until the end of the case run. the imc log confirms that the doctor pressed the p-0 pump start input three times after zeroing the sensor. the cause of the placement signal issue was not determined, as the data log suggests no issues in calculating the offset during the zeroing process, and sufficient clinical information was not provided to confirm any user-related aspects. the cause of the positioning issue was not determined since sufficient clinical information was not provided. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8603 | 1220648-2025-29218 | 22186604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186604&pc=OZD | 05/18/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/18/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638644 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured episodes of ventricular tachycardia during impella cp support. support continued with the implanted pump following the event. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8604 | 1220648-2025-29219 | 22186613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186613&pc=OZD | 04/03/2025 | 10/01/2025 |  | 06/11/2025 |  |  | 04/03/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589749 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported the patient was on impella 5.5 support and there was placement signal issues. the optical sensor was damaged and not working appropriately intermittently. support continued. furthermore, there were suction alarms that occurred during support. the pump was repositioned. this issue reoccurred with patient coughing. further details state that care was withdrawn and the patient expired. | [5 empty] | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details provided and the data log analysis, the root cause of the low pump flow is most likely due to improper positioning. due to lack of clinical details of repositioning of the pump, the root cause of the positioning issues cannot be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the data logs show that there are non-cyclical increases in the raw optical signal. during this time the signal-to-noise ratio (snr0 drops and contrast is variable while lamp is maxed out. there's no significant changes in gain or light. there are positioning alarms during this whole time period as well. due to lack of clinical details regarding additional devices or the patient¿s condition as well as no product returned, the root cause of the optical signal issue cannot be definitely determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-29219. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. although the patient expired on the impella 5.5 device and there was low pump flow and positioning issues, the impella cp related events cascaded and contributed more to the patient's outcome. the demise outcome is associated with the impella cp under manufacturer mdr report #1220648-2025-27411. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8605 | 1220648-2025-29220 | 22186779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22186779&pc=OZD | 05/21/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/21/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632459 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 placed to support and replace a prior impella 5.5 that malfunctioned on the first console. a second pump failed to deliver despite all efforts and was removed. additionally, bleeding was observed during the procedure, and the patient received 20 units of fresh frozen plasma, 19 units of red blood cells, 2 units of platelets, and 3 units of cryoprecipitate. the patient remained non-pulsatile with compromised hemodynamics and continued to require ongoing transfusions of blood products. ultimately, the decision was made by the family to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8606 | 1220648-2025-29222 | 22187051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187051&pc=OZD | 05/23/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/23/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp was inserted to support percutaneous coronary intervention (pci). after the pci, a decision was made to remove the impella due to cold mottled leg. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8607 | 1220648-2025-29223 | 22187099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187099&pc=OZD | 05/23/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/23/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667231 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp for a st-segment elevation myocardial infarction patient. the pump was supporting when there was a gastrointestinal bleed. the bleed prompted the team to remove heparin. the purge had no heparin in the fluid. the pump remained on for a total of 27 hours til wearned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8608 | 1220648-2025-29224 | 22187224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187224&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645875 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, and after the explant, the post closure did not work. a balloon tamponade was attempted, and it did not work. a dissection was noted. new access was obtained, and a covered stent was utilized, and angioseal was used to close access. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8609 | 1220648-2025-29225 | 22187293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187293&pc=OZD | 05/23/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/23/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645754 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient was planning on having an impella 5.5 implant. however, a left ventricle thrombus was noted. the impella 5.5 implant was cancelled and the impella cp was explanted. there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8610 | 1220648-2025-29226 | 22187354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187354&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657689 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient endured diffuse bleeding from their axillary pocket immediately following the impella implant. anesthesia was unable to oxygenate the patient with frank red blood out of the et tube and the diffuse bleeding. while attempting to cannulate, the patient coded and passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8611 | 1220648-2025-29227 | 22187461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187461&pc=OZD | 05/23/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/23/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186692 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | A complainant reported that there was an automated impella controller failure during prep. a backup controller was successfully used. the impella cp was running without any problems. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8612 | 1220648-2025-29228 | 22187542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187542&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651668 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, after explant, an angioseal failed and a dissection was noted to the femoral artery. a covered stent was placed, and hemostasis was achieved. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8613 | 1220648-2025-29229 | 22187582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187582&pc=OZD | 05/23/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/23/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641410 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the staff was unable to obtain pulses after the removal of the peel away sheath. the patient coded and the pump was turned to p2. however, the patient later expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8614 | 1220648-2025-29230 | 22187832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187832&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622522 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. due to the patient's worsening underlying condition, the decision was made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8615 | 1220648-2025-29231 | 22187862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22187862&pc=OZD | 05/19/2025 | 07/16/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568137 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Embolism/Embolus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, air was noted in the left ventricle (lv) with a possible air embolus to the coronaries at the time of the noted ventricular tachycardia (vt). the air was removed via a lv vent. high purge pressure alarms appeared on the automated impella controller (aic). heparin was run through the purge line to manage the issue. however, a chest wall bleed was noted leading to blood being given and an additional stitch being placed at the graft. support was continued following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of vf/vt/af/at, high purge pressure, and embolism has been completed. the pump was not returned for investigation. the data log shows an approximate 1-hour gradual rise in purge pressure. high purge pressure was resolved with a gradual decrease trend; however, the purge pressure remained on the higher end but within the specification range for the remainder of the case run. the data log also confirms that priming was successfully completed before pump insertion. high purge pressure: the cause of the high purge pressure issue was biomaterial buildup based on data log review and provided clinical details. embolism: the cause of the embolism issue was not determined, as sufficient clinical information was not provided. the relationship between the impella and the embolism remains unknown. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. investigation results for vascular damage - peripheral vessel were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29231. should any new information be received, a supplemental manufacturer device report will be filed. h6 health effect - clinical code 1888 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29231.
8616 | 1220648-2025-29232 | 22188054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22188054&pc=OZD | 05/19/2025 | 07/02/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, when attempting to reposition the impella cp in the catheterization lab, the anticontamination sleeve ripped exposing the catheter. sterile tegaderm was placed. also, echocardiogram was performed, and the impella pigtail was caught on the mitral valve chordae and inlet near the pap muscle. additionally, a small hematoma developed on arrival to the cardiovascular intensive care unit from the catheterization lab on the right femoral site. pressure was held and they are continuous monitoring. the patient was eventually escalated to an impella 5.5. | [5 empty] | The investigation of product damage (sleeve damage) and positioning issues has been completed. the impella device was not returned for evaluation. product damage (sleeve damage): when attempting to reposition in the catheterization lab and while doing so sterile sleeve around catheter ripped exposing catheter. sterile tegaderm placed. the pump was not returned. the cause of the product damage (sterile sleeve damage) was not determined since product was not returned and insufficient clinical details. positioning issues: following implant procedure, it was noted the pigtail was caught on mv chordae and inlet near pap muscle. the pump was not returned. the cause of the positioning issue was most likely use related due to the incorrect initial placement causing the pigtail to be caught on mitral valve (mv) chordae and inlet sitting near the papillary muscle which required repositioning. h6 component code 4723 was inadvertently omitted on the initial manufacturer device report. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8617 | 1220648-2025-29233 | 22188064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22188064&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609684 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had experienced a new onset atrial fibrillation (afib) with rapid ventricular response (rvr). amiodarone gtt initiated for arrhythmia, and after 45 hours of support, the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8618 | 1220648-2025-29234 | 22188144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22188144&pc=OZD | 05/21/2025 | 07/17/2025 |  | 06/11/2025 |  |  | 05/21/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657761 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the impella 5.5 has a purge fluid leaking out of the sidearm at the site of the hole on the membrane of the purge reservoir. the purge flows and purge pressures have remained stable. the local team advised to continue to monitor purge flows/pressure closely at this time. the patient was successfully weaned and the patient outcome was stable. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. data log review showed a very slight, gradual rise in purge pressure and decrease in purge flow beginning on (b)(6) 2025. however, purge pressure and purge flow were fairly stable throughout the case. there were no signs of low purge pressure or high purge flow that would indicate a purge leak. no product was returned. the root cause of the pump purge leak was most likely a damaged sidearm but the exact location of the leak is unknown since product was not returned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | Purge leak - pump: clinical details state that purge fluid was leaking out of the sidearm at the site of the hole on the membrane of the purge reservoir, but purge flow and pressure remained stable. data log review showed a very slight, gradual rise in purge pressure and decrease in purge flow beginning on (b)(6) 2025. however, purge pressure and purge flow were fairly stable throughout the case. there were no signs of low purge pressure or high purge flow that would indicate a purge leak. no product was returned. the root cause of the pump purge leak was most likely a damaged sidearm but the exact location of the leak is unknown since product was not returned.
8619 | 1220648-2025-29235 | 22188497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22188497&pc=OZD | 05/19/2025 | 06/11/2025 |  | 06/11/2025 |  |  | 05/19/2025 | 06/11/2025 |  |  |  |  | 06/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 18 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant published a case support in which there was impella use and hemolysis. the patient with known nonischemic idiopathic dilated cardiomyopathy and biventricular failure. the patient was being medically managed but when the patient still needed escalation the team placed an impella cp, but the hemoglobin levels dropped and there was anemia and hemolysis which was treated by blood transfusions, p-level adjustments, and pump explant. after explanting the impella, the hemolysis rapidly resolved. | [5 empty] | D4: model number, catalog number, serial number and lot number were not provided. d6a / d6b: implant and explant dates not provided. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8620 | 1220648-2025-29139 | 22174514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22174514&pc=OZD | 05/24/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/24/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655089 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and the following day experienced positioning issues and hemolysis which eventually lead to the device being explanted. an impella cp was placed via the right femoral artery and plan of care noted probable escalation to an impella 5.5 at some point soon due to the need for long-term support. after the start of mcs, the following day, early morning/overnight, impella in the aorta alarm triggered. heart failure and cardiothoracic surgery attempted to recross the aortic valve after noticing the pigtail in the aorta via echocardiogram. the decision was made to bring the patient back to the catheterization lab and attempt to reposition. once in the catheterization lab, a stiff 0.035 guidewire was used through the reposition port to attempt to enter the left ventricle and prolapse the aortic valve. however, this was unsuccessful. the decision was made to remove the impella cp, maintain wire access, place a new peel away sheath and place the same impella cp back into the ventricle. the peel away was then removed and the repositioning sheath was re-inserted. touhy borst was secured, and the patient returned to the unit continuing on impella mcs. later in the day, late morning/early afternoon, the patient's labs were hemolyzing (actions taken were unnoted). the next day, however, the decision was made to place the patient on an impella 5.5 in exchange for the impella cp given the plans for long-term care, bridge to left ventricular assist device or transplant. the patient remained on impella 5.5 support for approximately 11 days before being removed without any problems. the patient was noted to look and feel well and in great spirits. the plan of care was unclear for the patient at such time as the patient declined any durable left ventricular assist device or transplant. ejection fraction was 25% on last echocardiogram. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8621 | 1220648-2025-29140 | 22174578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22174578&pc=OZD | 05/24/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/24/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658076 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs). during support, the following day, the patient was noted to have more chest pain with st changes. the patient was taken to the catheterization lab. catheterization showed previously placed stents were patent. thus, there was no in-stent restenosis noted. left circumflex artery was stented during this time. low flows on the impella cp were noted, after closure of the left femoral artery, which was accessed for percutaneous coronary intervention. the decision was made to reposition the impella cp device. however, at this time, it was noted the anticontamination sleeve cover was completely severed from the hub. therefore, repositioning could not be performed without contamination. consequently, the impella cp device was explanted and replaced. the patient continued on mcs with the replacement impella cp and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8622 | 1220648-2025-29141 | 22174635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22174635&pc=OZD | 05/23/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/23/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669141 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support (mcs) and experienced limb ischemia which resulted in the device being explanted. the patient presented to the emergency department in complete heart block, coded and was emergently transported to the catheterization lab after return of spontaneous circulation and intubation. upon arrival to the catheterization lab, a temporary pacer was placed followed by impella cp insertion via the right femoral artery without incident. left heart catheterization revealed left anterior descending and diagonal arteries disease. after multiple attempts to intervene the left anterior descending artery, the physician decided to abort the intervention and consult for surgery. the patient was taken to the unit on impella mcs and after arriving to the unit, it was noted there was no pulse present with doppler. an arterial doppler order was placed and an echocardiogram was completed with position verified at 3.6 centimeters across the aortic valve and 94 centimeters externally. the patient remained on support and the following day; it was noted the performance level was set to p-4 and flowing 2.5 liters per minute during this time. the access site was intact; however, pedal pulse was not present. the plan of care was to take the patient back to the catheterization lab for impella cp explant, perform a valvuloplasty and if needed at the end of the procedure place the patient on an impella 5.5 device. the impella cp device was successfully weaned and explanted. hemodynamically, the patient tolerated the procedure well and was transported back to the unit without impella mcs and reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8623 | 1220648-2025-29142 | 22174720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22174720&pc=OZD | 05/24/2025 | 07/08/2025 |  | 06/10/2025 |  |  | 05/24/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665957 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. the physician performed a cutdown approach to right axillary. an 8 millimeter vascutek graft was placed. after using a c-arm to place the impella .018 guidewire, the physician made multiple attempts (with different support guidewires) to advance the impella 5.5 into the proximal subclavian artery. due to extremely narrowed and heavily calcified vessel, the impella 5.5 option was aborted. left axillary artery was noted to be diseased. the physician consulted with vascular surgery, and it was decided to place an intra-aorta balloon pump (iabp) instead of another impella device. the iabp was placed with some difficulty. however, the placement was completed and the patient was sent to the unit on support. the patient's status following the event was noted as unknown. however, it was noted there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the root cause of delivery issue was most likely patient condition since the pump implant was aborted due to extremely narrowed and heavily calcified vessel. d9 device returned to manufacturer date added.
8624 | 1220648-2025-29143 | 22174767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22174767&pc=OZD | 05/20/2025 | 09/09/2025 |  | 06/10/2025 |  |  | 05/20/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693473 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day 21 of impella cp support, motor current was high and the impella cp stopped. multiple attempts to restart the impella cp were unsuccessful, and a controller failure alarm occurred. the controller was exchanged but the pump did not restart. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) and aic ¿ pump stop has been completed. the console logs are consistent with complaint and show that 17 days since pump was transferred, a pump stop occurred (preceded by motor current spike) and ¿impella stopped. motor current high¿ alarm (13) presented. after 4 failed attempts to restart it, the pump was disconnected from the console and reconnected, but stopped again, and an ¿impella stopped. controller failure¿ alarm (138, due to pump speed deviation detected by pump motor driver (pmd)) presented, followed by pmd reset (as per design), and attempts to restart pump. the pump was then transferred to backup console, but issue remained unresolved and pump was unable to start. throughout the case, there were also multiple suction alarms (14, 73), and positioning alarms (115, impella position unknown). shortly before pump was disconnected a controller failure alarm (414, purge pressure sensor defective) presented. console was tested and was operating within specification, able to run an engineering pump without any stops. considering that the pump also malfunctioned on another console, it is most likely that the pump stops were due to the pump (product not returned, investigated under 1220648-2025-29258). purge system was tested, and found to be operating within specification. the controller failure alarm (414, purge pressure sensor defective) was most likely triggered due to purge cassette being disconnected prior to pump disconnection, while purge motor kept spinning. aic - controller failure (red alarm): the cause of the issue was most likely use related (timing of purge disconnection and pump disconnection). aic ¿ pump stop: pump stop due to aic failure - recovered without removing/replacing pump. d9 device returned to manufacturer date added.
8625 | 1220648-2025-29144 | 22174905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22174905&pc=OZD | 04/02/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/22/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1817137 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller experienced an optical sensor error. there was no patient involvement. during investigation and upon review of the data logs, it was believed that the device may have been swapped due to persistent optical sensor issues. | [5 empty] | The investigation has been completed. the console was returned for investigation. the logs end after the use with a test pump. there were no optical sensor errors on that day. the complaint, which was submitted on april 2, 2025 with no pump serial number, may be referencing february 28, 2025 when serial number pump (b)(6) was plugged in three times, and the first two times asked the user to reconnect the impella and the third time asked the user to switch consoles due to the persisting optical issue during case start. reviewing the real time logs, the first two had raw optical sensor signal of -3276.8 mmhg with near zero signal-to-noise ratio (snr) with symptoms similar to an air gap between the pump connection. the third connection had expected ranges of 5s values and did not have any noticeable issues in the real time logs. pump (b)(6) was transferred to a back up console with serial number (b)(6). the failure mode was not reproduced on march 3, 2025 with a test pump in the field during multiple tries and similarly was not reproduced during investigation. the optical system had a relatively high snr values across multiple pumps, excluding poor snr as a possible cause of the optical signal issues. the cause of the optical sensor system errors was not determined since the failure mode was not reproduced during testing. it is possible the first two instances were due to an air gap, resulting in low snr, but the third instance was not determined.
8626 | 1220648-2025-29145 | 22175051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175051&pc=OZD | 02/02/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/22/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468305 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal unreliable alarm appeared during impella cp support. the staff disabled the alarm and support with the implanted pump continued uninterrupted. there was no patient harm. upon analysis of the console logs during investigation, it was revealed that there were multiple instances of controller error (opm comm crc invalid) alarms throughout the case. the controller error was found to be due to an optical bench firmware communication protocol anomaly. | [5 empty] | The investigation has been completed. no product was returned. console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the following communication-related error "algs suspended opm signal invalid, and "invalid_bad_snr_sample_mask" due to snr falling below the minimum threshold. multiple instances of placement signal not reliable and controller error (opm comm crc invalid) [optical pump connected] alarms were observed through the case. the long term log data indicated that the unreliable placement signal persisted for half of the case, with varying intensity, which correlated with low signal-to-noise ratio (snr) and oscillating contrast. this led to placement signal not reliable alarms. real time log data also showed oscillating contrast with low snr values. the root of the controller error was due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller software operated as designed.
8627 | 1220648-2025-29146 | 22175069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175069&pc=OZD | 03/20/2025 | 09/23/2025 |  | 06/10/2025 |  |  | 03/20/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617098 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella in aorta alarm appeared during impella 5.5 support. positioning was verified to be correct and support continued without interruption. there was no patient harm. upon receipt of the device, a detached cannula and inflow cage were identified. the cannula was noted to be detached from the motor housing with a tear in the delo glue joint being evident. | [5 empty] | The impella 5.5 was returned for evaluation. no clinical details were provided on when the cannula detachment occurred. the root cause of the product damage could not be definitively determined as no clinical details were provided at time of damage occurring. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: upon visual inspection, no damage to the catheter was observed. pump passed laser continuity testing and all optical parameters were within normal limits. data logs were reviewed, and "impella in aorta" alarms were occurring intermittently during support. no hard failures of the optical sensor were observed in the returned data logs. clinical details confirm that pump was in proper position at time of alarms at beginning of case. the root cause of the optical signal issue was most likely patient condition related as no issues were found with the optical sensor on the returned pump and clinical details confirmed that it was a false "position in aorta" alarm. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated for optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-29146. d10 concomitant medical products and therapy dates.
8628 | 1220648-2025-29147 | 22175230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175230&pc=OZD | 05/13/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/13/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658383 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was placed on impella 5.5 as a bridge to transplant. implantation was difficult due to patient anatomy and the pump was placed in a suboptimal, shallow position. suction alarms occurred, so p levels were reduced and dobutamine was given. the pump was identified as sitting towards the mitral valve. attempts to reposition the pump were unsuccessful. the patient was successfully weaned from the pump and the patient had a successful bridge to transplant. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8629 | 1220648-2025-29148 | 22175440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175440&pc=OZD | 05/12/2025 | 07/02/2025 |  | 06/10/2025 |  |  | 05/12/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon starting up an implanted impella 5.5, the device was noted to be too deep in the ventricle. the doctor was made aware, but the patient went into ventricular fibrillation before the pump could be repositioned. the patient was shocked six times until stabilized. the pump was repositioned. once it was in the appropriate position, no arrhythmias were noted. additional information was added that the pump stopped weeks later after the patient was shocked by external implantable cardioverter-defibrillator 12 times. the pump was unable to be restarted. the patient was brought to the operating room for pump exchange. the procedure was successfully completed with the new device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Positioning issue: clinical states that the doctor placed the impella deep and patient went into vfib. repositioning solved the arrythmia. therefore, the root cause of the positioning issue was most likely use related to improper initial pump placement technique. vt/ vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. pump stop: clinical states that patient was shocked 6 times by external icd which led to pump stop. pump was not returned. data logs were not returned for investigation. without the pump or data logs the failure cannot be investigated. therefore, the root cause of the pump stop issue was not determined since product and data logs were not returned for investigation.
8630 | 1220648-2025-29149 | 22175485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175485&pc=OZD | 05/15/2025 | 07/02/2025 |  | 06/10/2025 |  |  | 05/15/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581603 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was positioned in the mitral. attempts were made to reposition to no avail. the pump was explanted and replaced with an intra-aortic balloon pump (iabp). | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation of positioning issues and product damage (guidewire damage - great vessel) has been completed. the impella device was not returned for evaluation. positioning issues: data logs were reviewed and there were no positioning alarms, specifically position in ventricle. product wasn't returned for investigation. since insufficient clinical details were provided regarding what made it difficult to position the pump and product wasn't returned for investigation, the root cause of the positioning issues could not be determined. product damage (guidewire): while trimming sutures, the 0.018" guidewire was mistakenly cut. it was replaced with a new one and the impella was then able to be advanced into the lv. product was not returned for investigation. though product was not returned for investigation, it can be concluded that the most likely cause of the product damage was a use issue based on the clinical details provided that the guidewire was mistakenly cut while trimming sutures.
8631 | 1220648-2025-29150 | 22175606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175606&pc=OZD | 05/18/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/18/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657780 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella support experienced venticular tachycardia and bleeding after coronary artery bypass graft. heparin was added to the impella purge the next day. systemic heparin was held as the patient experienced atrial fibrillation with rapid ventricular response overnight. amiodarone bolus times 3 was given and the patient continued on the amiodarone drip. the plan was to cardiovert that day. two days later, the patient went back into atrial fibrillation and was taken to the operating room for vascular removal due to the patient's low platelet levels and lack of systemic heparin. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8632 | 1220648-2025-29151 | 22175668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175668&pc=OZD | 05/15/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/15/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660299 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was placed, and flows were low. the 2nd cp was placed as replacement, and additionally the team made the decision to discuss rf flex placement to assist with the ventricular failure. the family made the choice to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8633 | 1220648-2025-29152 | 22175858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22175858&pc=OZD | 05/15/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/15/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661865 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Migration or Expulsion of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was not flowing properly. the c-arm images showed the impella was pulled back across aortic valve. the physician was unable to re-advance and the pump was explanted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8634 | 1220648-2025-29153 | 22176094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176094&pc=PYX | 05/15/2025 | 06/27/2025 |  | 06/10/2025 |  |  | 05/15/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575868 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, oozing was noted around the access site. additional hemostatic suture was placed. second suture held; angle matching performed, and the site was dry. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, additional hemostatic suture placed, 2nd suture held, angle matching performed to achieve hemostasis. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
8635 | 1220648-2025-29154 | 22176213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176213&pc=OZD | 05/15/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657975 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella cp was placed on a patient with cardiogenic shock. during support it was noted that the patient had bleeding at the left groin and was taken for vascular surgery to remove a hematoma. the patient remained on support until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8636 | 1220648-2025-29155 | 22176228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176228&pc=OZD | 05/17/2025 | 07/14/2025 |  | 06/10/2025 |  |  | 05/17/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed for support of a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump had a rapid changed in flows with suction. the pressure dropped to p-2. volume was given via albumin and lacatated ringer's boluses without resolution of suction. the device was repositioned and still there were poor flows. upon evaluation, the pump had extremely low overall flows throughout and there was a concern for a clot. the decision was to leave the pump in overnight at p-2 to trial wean. the patient was taken to the operating room the next day and the impella cp was explanted. at the time of explant, the surgeon did have to perform a two-part cut down to remove the device and close the arteriotomy due to an exceedingly high stick upon initial insertion. upon confirmation with the field representative, the low flows was due to thrombus seen at pigtail. the patient remained stable and survived the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of vascular damage - peripheral vessel, low pump flow, and thrombus has been completed. the impella device was not returned for evaluation. vascular damage - peripheral: based on the clinical details that a high stick angle was used during access, likely without ultrasound guidance, the root cause of the vascular damage was determined to most likely be a use issue. low pump flow: data logs were reviewed and several abnormalities were noted. on (b)(6) 2025, there were several spikes up in the ps that were followed by an initial set of motor current (mc) spikes to 1400 ma and speed deviations at p-9 that was followed by a slight step down in mc. however, shortly after, there was another spike to 1500 ma with speed deviation that was followed by a distinct drop in flows and loss of mc modulation. pump flows became low through the following day, when a third mc spike with speed deviation resulted in pump flows increasing back into specification for p-2. these are all indications of an object obstructing the impeller, which supports the report of potential biomaterial near the pump seen on echo. purge pressure was not affected at all during the case. product was not returned for investigation. based on the clinical details provided that neither volume or repositioning resolved the issue and data log review showing several instances of biomaterial ingestion signatures, the root cause of the low pump flows was most likely biomaterial. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined.
8637 | 1220648-2025-29156 | 22176288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176288&pc=OZD | 05/16/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579097 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient on impella cp escalated to the 5.5 pump for care escalation. it was reported that while on impella 5.5 a diffuse head bleed was noted on a computed tomography scan. anticoagulation was stopped. care was withdrawn two days later and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8638 | 1220648-2025-29157 | 22176368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176368&pc=PYX | 05/18/2025 | 07/15/2025 |  | 06/10/2025 |  |  | 05/18/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026636097 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and during support, while transferring the patient from the operating room table to patient's bed, the impella catheter and repo sheath were pulled out of position. multiple attempts were made to reposition the catheter but they were unsuccessful. the pump was explanted and the patient was placed on extracorporeal membrane oxygenation. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Reported pulling the pump out of position when transferring the patient from the operating room table to the bed. pump was removed and replaced with ecmo. the root cause of the positioning issue was most likely patient movement since the pump was pulled out when transferring the patient.
8639 | 1220648-2025-29158 | 22176430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176430&pc=OZD | 05/18/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/18/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633163 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was on impella cp support. during support, it was noted that patient¿s labs showed signs of hemolysis at p-level 8 and the medical doctor reduced support to p-7. a day later, the lactate was down to 1.8. the patient was making positive urine output (uop) on bumex drip and uop has cleared. the patient was later escalated to an impella 5.5 and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8640 | 1220648-2025-29159 | 22176573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176573&pc=OZD | 05/21/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/22/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628440 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient admitted in with st-segment elevation myocardial infarction. it was reported that there was distal limb ischemia on the right lower extremity (rle) (impella access site). faint pulses on the left lower extremity were reported but none on the rle per the medical doctor. a thrombectomy was done in the operating room and pulses returned. the patient was escalated to an impella 5.5 pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8641 | 1220648-2025-29160 | 22176663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176663&pc=OZD | 05/16/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656001 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support. it was reported that the patient had ischiemia on left hand on the side the impella 5.5 was implanted. no intervention was reported. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8642 | 1220648-2025-29161 | 22176816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176816&pc=OZD | 05/22/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/22/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653700 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 moved intoa patients ascending aorta. an attempt to advance was made at bedside with echocardiogram but was unsuccessful. the patient remained stable and the impella 5.5 was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8643 | 1220648-2025-29162 | 22176940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22176940&pc=OZD | 05/13/2025 | 06/25/2025 |  | 06/10/2025 |  |  | 05/13/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675325 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was placed on a patient to support until planned surgical mitral valve repair/replacement. during support, it was noted that the impella was on auto and set at p6 due to suction events. additionally, they were unable to flow greater than p4 without suction, receiving one of two units of packed red blood cells. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: the data logs from the case do not show any motor current spikes or positioning related issues. flows were within the appropriate range for matching p-level. there were suction alarms through support. clinical mention they were unable to flow greater than p4 without suction. there were discussions on patient volume after which volume was provided. the logs show p-level was increased beyond p4 after volume was given. the root cause of the low pump flow was most likely patient condition related as evidenced by no motor current spikes or positioning abnormalities in the data logs and clinical information on patient volume status and flow increase after volume administration.
8644 | 1220648-2025-29163 | 22177058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177058&pc=OZD | 12/20/2014 | 06/10/2025 |  | 06/10/2025 |  |  | 12/20/2014 | 06/10/2025 |  |  | 10/15/2016 |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1174105 |  | 0048-0003 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Fracture | Failure to Pump | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during the impella cp procedure the 0.018 guidewire severed while removing the guidewire during insertion. the impella cp would not start. the cp with the guidewire were removed intact. a new impella was inserted in the opposite groin due to the fact that the peel away sheath had already been removed. it was also reported that the patient expired. no other information is available. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8645 | 1220648-2025-29164 | 22177175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177175&pc=OZD | 05/18/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/18/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612165 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp to support an acute myocardial infarction/cardiogenic shock patient. the pump was placed via a 14 french sheath. the site was bleeding around the original impella peel away, which became progressively worse throughout the procedure. the decision was made to cannulate the patient for venoarterial extracorporeal membrane oxygenation. a total of six units of packed red blood cells and two units of fresh frozen plasma were given. the peel away was exchanged for a repo sheath, but bleeding continued. bleeding was still a concern at the impella site. the medical doctor made the decision to explanted and placed a non-abiomed 18 french sheath for the impella cp. no distal pulses were appreciable today and the foot was cool to touch. the implella cp was explanted and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8646 | 1220648-2025-29165 | 22177233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177233&pc=OZD | 05/16/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651213 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, the femoral sheath was fully occlusive. a fem-fem bypass was performed to assist with perfusion to the leg. shortly after, the patient's access site was oozing. the nurse took down the dressing for good angle matching with the forward tension sutures. the decision was ultimately made to explant the pump as a result of the bleed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8647 | 1220648-2025-29166 | 22177317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177317&pc=OZD | 05/18/2025 | 07/08/2025 |  | 06/10/2025 |  |  | 05/18/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658660 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during attempted delivery of an impella 5.5 pump there was difficult anatomy during insertion with steep takeoff from the graft anastomosis. aggressive force was used to attempt insertion but was unsuccessful. the surgeon pulled out to adjust the graft and artery for better angle. kink in the white impella catheter was noted. it is unclear if the pump would have issues if reinserted. the physician made the decision to get a new impella 5.5. and it was placed without issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the pump was not returned for investigation. a data log review was not required for this case run. according to the clinical report, the patient was noted to have difficult anatomy and tortuous vessel condition while attaching the graft, and the insertion angle was stiff during placement. excessive force was applied during insertion. the doctor noted a kink in the cannula and decided to replace the pump. a new pump was placed without issue. the cause of the delivery issue was most likely related to the patient's diseased or small vessels. the product damage (cannula kink) was most likely caused by the patient¿s diseased vessels.
8648 | 1220648-2025-29167 | 22177465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177465&pc=OZD | 05/16/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1130827 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that automated impella controller (aic) had a self-check failure during prep. the aic was removed from service. there was no harm to the patient. | [5 empty] | The investigation has been completed. the automated impella controller (aic) was returned for investigation. console logs showed a persistent communication failure with the pbm (power battery manager) during the initial bootup. due to this communication issue, the aic console rebooted automatically (as designed) to attempt another power-on. after rebooting, the console displayed "system self check failed" screen, and the console was then forcefully shut down. the issue was confirmed after console arrived for troubleshooting, as sau_powermod_1 would not load due to loss of communication with pbm1. visual inspection confirmed pbm contamination which has impacted a neighboring rs232 communication channel. additional analysis of the aic log revealed multiple failed handshakes with the rlm (unrelated to this complaint), caused by rl03136 unexpectedly rebooting due to its microsd card corruption (as evidenced by repeating ext4-fs error messages). the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
8649 | 1220648-2025-29168 | 22177727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177727&pc=OZD | 05/16/2025 | 07/14/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573408 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon insertion of impella 5.5, the impella cannula would not turn into the innominate artery despite several attempts. the physician believed it to be calcification; smaller artery size and "bovine" anatomy was to blame. decision was made to remove 5.5 and place impella cp. there was no harm to the patient reported. | [5 empty] | The investigation of delivery issues has been completed. product damage (cannula kink) added based on product evaluation. reported difficulty advancing pump into the innominate artery and patient had tortuous vasculature with calcification. impella 5.5 was returned and a cannula kink was found. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient had tortuous and calcified vasculature. cannula kink - the root cause of the product damage was most likely a kinked cannula due to patient condition since the patient had tortuous and calcified vasculature. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8650 | 1220648-2025-29169 | 22177735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177735&pc=OZD | 05/13/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/13/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658386 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient had a hematoma at the access site. the clinical representative confirmed that two units of blood products were given to the patient. additionally, the patient had ventricular tachycardia (vt) during support. the patient continued on support until the patient was bridged to transplant. the patient received a heart transplant and was reported to be doing well. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8651 | 1220648-2025-29170 | 22177753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177753&pc=OZD | 05/16/2025 | 07/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651647 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced ventricular fibrillation requiring medical intervention. the patient was found to have triple vessel disease involving the left main and left anterior descending arteries with thrombolysis in myocardial infarction ii flow and in-stent restenosis of the right coronary artery. the impella cp device was implanted, and the patient was immediately transported via helicopter to a different hospital. it was reported that while on the helicopter pad, the patient experienced brief ventricular fibrillation that converted after one round of cardiopulmonary resuscitation and defibrillation. once arrived at the receiving hospital unit, the impella cp pump was repositioned. formal echocardiogram showed ejection fraction now closer to 20%, and ongoing profound cardiogenic shock with respiratory acidosis. a central line and a-line were placed; pulmonary artery catheter was discussed with the team, but the team was unwilling to place at this time. the patient was taken to the operating room for escalation to an impella 5.5 which was successfully completed. the patient was reported to be stable following the event and currently remains on impella mcs as a bridge to recovery (lung, then potentially percutaneous coronary intervention versus coronary artery bypass graft surgery). | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Positioning issue - reported repositioning the pump after the patient was transported to another hospital via helicopter. cpr was performed on the helicopter pad. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical details were provided. vt/vf - reported patient had brief v-fib. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
8652 | 1220648-2025-29171 | 22177868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22177868&pc=OZD | 05/14/2025 | 07/02/2025 |  | 06/10/2025 |  |  | 05/14/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630042 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, a high purge pressure alarm was observed. the automated impella controller stated the purge pressure was less than two. tissue plasminogen activator was administered and the alarm was resolved. support was continued and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | High purge pressure (hpp): data log review shows purge pressure slowly increase from 450 mmhg during the first few days of support to 600 mmhg on the day of the complaint. a quicker but still gradual increase in purge pressure from 600 mmhg to greater than 1000 mmhg was seen over almost two hours on (b)(6) 2025. there were no motor current spikes seen prior to the rise. there was also purge flow variability and a gradual trend downward from 15 ml/hr starting in the first few days of support. purge flow began to decrease at a faster rate about 13 hours before the 2-hour purge pressure increase on (b)(6) 2025. the hpp sustained for 4 hours before resolving along with purge flow. motor current was stable throughout the case. there were also five instances of long bag/cassette changes prior to the hpp events. none of these instances occurred directly before the initial purge pressure increase or the 2-hour pp rise on (b)(6) 2025. no product was returned. the root cause of the high purge pressure was most likely biomaterial buildup as a gradual rise in purge pressure and no motor current spikes were seen in the data logs.
8653 | 1220648-2025-29172 | 22178153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22178153&pc=OZD | 05/16/2025 | 07/02/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. during implant, the peel away sheath was inserted and the 0.018 wire crossed the valve without issue. the pump was inserted into the 14x25 peel away sheath, however it was unable to advance into cfa due to extreme calcification and tortuosity. the physician eventually decided to abort the impella placement and proceed without it. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella was unable to advance into common femoral artery (cfa) due to extreme calcification and tortuosity.
8654 | 1220648-2025-29173 | 22178182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22178182&pc=OZD | 05/16/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629885 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had limb ischemia. bilateral upper and lower extremities were noted to be extremely mottled. order for bilateral arterial duplexes of upper and lower extremities placed. the patient continued to decline, and the family decided to withdraw care and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8655 | 1220648-2025-29174 | 22178239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22178239&pc=OZD | 05/17/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/17/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665788 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and a 6 french sheath had bleeding at the access site. the patient was transfused with six units of packed red blood cells and two units of fresh frozen plasma. the decision was made to explant the pump and implant a new impella cp with a 18 french sheath. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8656 | 1220648-2025-29175 | 22178306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22178306&pc=OZD | 05/16/2025 | 07/03/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. during implant, the impella was inserted and advanced to the distal aorta and unable to advance beyond. multiple wires and sheaths were attempted and unable to advance the impella cp beyond common iliac artery/distal aorta and the implant was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because patient had stenosis, tortuosity and an endograft that resisted to advance the device through and insertion was aborted.
8657 | 1220648-2025-29176 | 22178619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22178619&pc=OZD | 05/16/2025 | 07/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella cp and rp flex devices for mechanical circulatory bipella support and developed hemolysis requiring intervention. the patient would subsequently expire. the patient came out of the operating room post coronary artery bypass graft surgery on two to three vasopressors/inotropes. over the next two hours, the patient continued to decompensate requiring increasing inotropic support with worsening blood gases. the decision made to bring patient to catheterization lab for impella placement in the setting of worsening post-cardiotomy cardiogenic shock. the impella cp device was implanted and followed by the impella rp flex approximately 45 minutes later placed in the setting of cardiopulmonary resuscitation. the patient was stabilized and transported to unit in critical condition following 30min of advanced cardiac life support. however, post impellas implant the patient developed hemolysis. a bedside echocardiogram was performed to try to determine source of hemolysis. the impella cp was noted to be free floating mid left ventricle and free of all structures. the patient received multiple blood products: fresh frozen plasma, and red blood cells. however approximately six hours later, the patient expired. | [5 empty] | Additional information was received and noted the patient was severely acidotic and was chasing bleeding. the local staff and team understood the patient would probably not survive. minimum cardiac output patient needed was at least 4.1 l/min and was unable to be achieved; unable to get the volume to the left side to achieve adequate flows and patient¿s lactate continued to rise. medical safety officer reviewed the event and determined the patient on bipella support developed hemolysis. attempt made to determine the source for the hemolysis; no further information noted. the patient critically ill, expires after approximately eight hours after start of impellas mechanical circulatory support. with the inability to dissociate, the demise is to be filed against the impella cp (which was the first implant) and impella impella rp flex is a serious injury which is reported under manufacturing report number 1220648-2025-29109. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Complaint device not returned for analyzing. data log reviewed: no mc spike observed outside of p-level changes. ps was observed to be trending down. placement signal alarms observed followed by impella position in ventricle/ aorta at p-9 in the first hour of support after insertion. purge pressure spike observed followed by many suctions at p-3. drop in flow observed at -p4 to the end of cp support clinical stated: bedside echo performed to try to determine source of hemolysis. impella free floating mid lv and free of all structures. patient receiving multiple blood products, ffp, plasma, rbcs. multiple amps bicarbonate. red urine noted. concern for dic. good position confirmed. patient was severely acidotic and was chasing bleeding. hemolysis was not from the pump. hemolysis: the cause of the hemolysis issue most likely was due patient condition since patient was severely acidotic and was chasing bleeding. low pump flow: failure mode low pump flow added based on the observation from complaint pump log review. the cause of the low pump flow issue was most likely patient condition related since patient was severely acidotic, cardiac output was less than the minimum and unable to achieve adequate blood volume to the left side to achieve adequate flow.
8658 | 1220648-2025-29177 | 22178932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22178932&pc=OZD | 05/11/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/11/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655406 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella 5.5 support experienced suction events. two units of packed red blood cells were given. the next day heparin was on hold due to high partial thromboplastin time and low xa. the patient was brought to the operating room the following day for a washout and closure. red blood cells and platelets were given and labs were redrawn. an endoscopy was performed the next day for a gastrointestinal bleed. bleeding continued for days requiring platelets and red blood cells at every shift. the patient remained critically ill throughout support. it was noted that it was evident that the patient had multisystem organ failure and continued to have rectal bleeding. heparin was stopped. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8659 | 1220648-2025-29178 | 22179092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22179092&pc=OZD | 05/17/2025 | 08/12/2025 |  | 06/10/2025 |  |  | 05/17/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614143 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support, along with extracorporeal membrane oxygenation. support was ongoing for 11 days when there were signs of limb ischemia. the limb was accessed by the impella cp and a femorofemoral bypass was placed and ischemia improved. the team discussed escalating to an impella 5.5. four days later, a pump stop malfunction occurred. the team noted the pump stop/restart occurred 10 times. this was happening post indicated use of device time and the team made the decision to remove the impella cp the following day, on almost the 14th day of support. the patients outcome was stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of limb ischemia and pump stop has been completed. the impella device was returned for evaluation. limb ischemia - reversible: clinical details state that the patient had signs of limb ischemia on (b)(6) 2025 while on extracorporeal membrane oxygenation (ecmo) and high doses of vasopressors and catecholamines. a bypass was completed with arterial cannula, then ischemia resolved. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. pump stop: clinical details stated that the pump stopped for more than 10 times by high motor current at almost 14 days of cp support. pump was explanted. no data logs are available. the pump was returned and evaluated. the pump had a large purge gap, indicating bearing wear. no other abnormalities were found. the root cause of the pump stop was most likely bearing wear after 13.9 days of support, which is over the 8-day design life per design trace matrix.
8660 | 1220648-2025-29179 | 22179185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22179185&pc=OZD | 05/16/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/16/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639680 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that the patient is hemodynamically stable and lactic was normalized. it was decided to start a wean process and see if the patient would need to be transferred for escalation of care. after weaning, the patient required the need for increased vasopressors. in addition, there were continuous alarms and dropped to p-2. the patient was having frequent runs of v-fib. ultimately the patient coded and expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8661 | 1220648-2025-29180 | 22179284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22179284&pc=OZD | 05/13/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/13/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618448 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and developed hemolysis resulting in device explant with replacement to an impella 5.5. during the impella 5.5 device support, the patient experienced low pump flow, tachycardia, return of hemolysis and would subsequently expire. the patient presented with shortness of breath and diagnostics revealed an ejection fraction of 15% with multi-vessel disease. percutaneous coronary intervention (pci) was performed with two stents to the left anterior descending artery and one stent to the ramus artery which occurred approximately five days prior to impella cp device implant for hemodynamic support. the impella cp was successfully placed, and at this time coronaries were assessed, and it was revealed the recently placed ramus artery stent thrombosed. coronary angioplasty was performed with good return vascularization. after this pci, companion and peel-away sheaths were removed, repositioning sheath placed, angle matched with no oozing at the site. the patient was sent to the unit on impella mcs. approximately 16 hours after start of the impella cp mcs, the patient developed hemolysis managed with decrease in performance level and escalation to an impella 5.5 which was successfully placed. the patient survived the explant with replacement and was noted to be in stable condition. the patient's urine output increased and was clearing; next day, urine was clear/yellow. approximately, two days thereafter, during shock rounds, the physician attempted to reposition the impella 5.5 as an attempt to increase the performance level from p-4. the level was increased to p-5 and there was suction. therefore, the physician stated that the impella 5.5 was optimized as much as possible with p-4. two days later, the patient was reportedly having an increasingly tachycardiac event, and the following day hemolysis was noted: pink tinged urine with suction and increased performance level overnight. performance level was decreased to manage the hemolysis, and an echocardiogram was performed (results of the echocardiogram were unnoted). the patient continued support; four days thereafter, it was noted the patient began to decompensate throughout the day at a lowered performance level of p-2. the patient was placed on bilevel positive airway pressure overnight. venous oxygenation saturation was down to 26% and tachycardia worsened. three days later, the patient went into cardiac arrest and was subsequently intubated and started on high dose pressors post return of spontaneous circulation. overnight, the patient experienced additional cardiac events requiring defibrillation. the patient was noted not an extracorporeal membrane oxygenation candidate; urine output was down 10-15 cc/hour on lasix drip. continuous renal replacement therapy was noted as a possible need, but the patient would subsequently expire. | [5 empty] | Medical safety officer reviewed the event and determined the hemolysis occurring with the impella cp cleared after escalation and is a serious injury. however, impella 5.5-related events may have added to an already complex situation and there is not enough evidence to exclude the impella 5.5 device used as an associated factor to the patient's outcome. the impella 5.5 device is one of two devices associated with the event. refer to initial medical device report for impella cp serial number (b)(6) with date of event 11-may-2025 and patient identifier ending in (b)(6). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle - ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8662 | 1220648-2025-29181 | 22179559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22179559&pc=OZD | 05/11/2025 | 06/10/2025 |  | 06/10/2025 |  |  | 05/11/2025 | 06/10/2025 |  |  |  |  | 06/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577946 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed hemolysis resulting in device explant with replacement. the patient presented with shortness of breath and diagnostics revealed an ejection fraction of 15% with multi-vessel disease. percutaneous coronary intervention (pci) was performed with two stents to the left anterior descending artery and one stent to the ramus artery which occurred approximately five days prior to impella cp device implant for hemodynamic support. the impella cp was successfully placed, and at this time coronaries were assessed, and it was revealed the recently placed ramus artery stent thrombosed. coronary angioplasty was performed with good return vascularization. after this pci, companion and peel-away sheaths were removed, repositioning sheath placed, angle matched with no oozing at the site. the patient was sent to the unit on impella mcs. approximately 16 hours after start of the impella cp mcs, the patient developed hemolysis managed with decrease in performance level and escalation to an impella 5.5 which was successfully placed. the patient survived the explant with replacement and was noted to be in stable condition. the patient's urine output increased and was clearing; next day urine output was clear/yellow. however, the patient would experience other impella 5.5-related events and subsequently expired approximately 13 days later. | [5 empty] | Medical safety officer reviewed the event and determined the hemolysis occurring with the impella cp cleared after escalation and is a serious injury. however, impella 5.5-related events may have added to an already complex situation and there is not enough evidence to exclude the impella 5.5 device used as an associated factor to the patient's outcome. the impella cp device is one of two devices associated with the event. refer manufacturing report number 1220648-2025-29180 for the report on the impella 5.5 device. the impella cp device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8663 | 1220648-2025-29100 | 22165377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165377&pc=OZD | 05/12/2025 | 07/02/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557901 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella 5.5 was attempted to be pulled back for mitral valve replacment but in doing so the impella 5.5 quickly slid out of the ventricle to completely be in the aorta. the pump was explanted without issues. the patient survived. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. as per the clinical information provided, the root cause of the positioning issue was use related as the pump was pulled too far during repositioning. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8664 | 1220648-2025-29101 | 22165397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165397&pc=OZD | 05/13/2025 | 06/27/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644265 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the pump was in good position. the patient coughed and the pump moved into the aorta and kinked. the doctor attempted to unkink cannula but was unsuccessful. the decision was made to remove and replace the pump. the patient went into ventricular fibrillation and required about five minutes of cardiopulmonary resuscitation and defibrillation. a new pump was placed successfully. the patient survived the procedure. | [5 empty] | The investigation of positioning issues, product damage (cannula kink), and vf/vt/af/at has been completed. the impella device was not returned for evaluation. positioning issues/ product damage (cannula kink): based on the details provided, the root cause of the positioning issues is most likely due to patient movement as the patient had a coughing episode and moved, causing the pump to go into the aorta. the root cause of the product damage (cannula kink) is due to a kink in the cannula most likely due to patient movement based on the clinical details. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8665 | 1220648-2025-29103 | 22165485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165485&pc=OZD | 05/13/2025 | 06/25/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616883 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an attempted impella cp implant and during the implant, the impella cp was wired and unable to advance the impella motor. the impella cp was removed and proceeded with a 7.0 balloon. however, the staff was still unable to advance the impella. the implant was aborted. the patient's outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant was aborted due to groin access being calcified and small which made unable to implant the device. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8666 | 1220648-2025-29104 | 22165521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165521&pc=OZD | 05/13/2025 | 07/02/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656867 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 pump was placed on a male patient after the coronary artery bypass graft. the patient pulled the pump out of the ventricle while pushing himself up in the bed. the three point fixation was in place and the yellow repo pin was not in position. the pump was removed and replaced with a new impella 5.5. the patient survived the procedure. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical states that the pump was pulled out of the ventricle by the patient accidently while getting himself up on the bed. pump was not returned. the root cause of the positioning issue was patient movement as the patient pulled the impella out of the ventricle while getting himself up. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8667 | 1220648-2025-29106 | 22165634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165634&pc=OZD | 05/13/2025 | 07/07/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662823 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella cp support had no pulses to the lower extremity. duplex study was negative for flow. a distal perfusion catheter was placed. however, there was still no flow. additionally, upon intubation, the patient began to drop flows, having suction at every p level, requiring reduction to p3. the next day, extracorporeal membrane oxygenation circuit with chatter and impella with suction was noted. albumin was administered and hypotension, suction, and chatter improved. it was further reported that care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | Limb ischemia - irreversible: clinical details state that there were no pulses to lower extremity as of (b)(6) 2025. duplex study negative for flow. dpc placed however still no flow. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. low pump flow: clinical details state that on (b)(6) 2025 upon arriving to icu and being intubated, patient began to drop flows, having suction at every p level, requiring reduction to p3. echo at bedside, rv akinetic, lv with an ef of approx 7%. on (b)(6) 2025, ecmo circuit had chatter and impella had suction. albumin administered and hypotension, suction, and chatter improved. tte at bedside showing underfilled lv, adequate impella positioning. data log review showed a set of suction alarms on (b)(6) 2025 with varying flows with the several changes in p-level. there was also a general ps trend downward on this day. the next day, suction alarms aligned with drops in placement signal. product was not returned. the root cause of the low flow and suction was most likely patient condition since clinical details state the patient had an underfilled lv and data log review shows patient deterioration trends in ps. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8668 | 1220648-2025-29107 | 22165913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165913&pc=PYX | 05/15/2025 | 06/27/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561550 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. when attempting to gain access for the impella rp flex implant, the staff accidentally hit the femoral artery twice with a 18 gauge needle. activated clotting time was 280 at that time. once access was obtained, the impella rp flex was placed without issue. however, when ready to leave the catheterization lab, the patient's blood pressure began to decrease, and a hematoma was discovered via computed tomography. the patient received four units of packed red blood cells/plasma/platelets and manual pressure was held for two hours to manage the condition. the patients blood pressure stabilized. the patient passed away after support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding is determined to be use issue because the bleeding is caused form the 2 arterial sticks that happened during the procedure and manual pressure is held for 2 hours to control the bleeding. vascular damage- peripheral vessel: the root cause of vascular damage is determined to be use issue because the doctor hit the femoral artery 2 times with the 18g needle during the insertion. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
8669 | 1220648-2025-29108 | 22165969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22165969&pc=PYX | 05/17/2025 | 07/07/2025 |  | 06/09/2025 |  |  | 05/17/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560574 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on bipella support was bleeding from all access sites. blood products were ordered. the next day, bleeding from access sites improved and the patient started on systemic heparin. the patient survived the procedure. | [5 empty] | The root cause of the access site bleeding issue was most likely patient condition related since bleeding from all access sites. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
8670 | 1220648-2025-29109 | 22166045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166045&pc=PYX | 05/16/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/16/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575301 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex developed hemolysis. the impella was confirmed to be free floating mid left ventricle and free of all structures. the patient was transfused fresh frozen plasma and red blood cells. additionally, the patient's urine was red. the clinical rep confirmed that no bleeding occurred from the impella access site. the patient expired on support. the medical safety officer reviewed and determined "the patient on bipella support developed hemolysis. attempt made to determine the source for the hemolysis; no further information noted. patient critically ill expires after approximately 8 hours after start of support. unable to dissociate, the demise will be filed against the impella cp under complaint (b)(4) and impella rp flex is a serious injury." | [5 empty] | The device was not received for evaluation. if any new information or if the device is received, a supplemental will be filed.
8671 | 1220648-2025-29110 | 22166107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166107&pc=OZD | 05/14/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595063 |  | 1000080 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and post implant, same day, hematuria was present, and it was unknown if the hematuria was a result of foley insertion or impella support. labs hemolyzed and the staff was advised that the performance level could be reduced at the discretion of the physician to troubleshoot the issue. the decision was ultimately made to escalate to impella 5.5 support the following day. the latest update notes the patient remains on impella 5.5. mcs in stable condition. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8672 | 1220648-2025-29111 | 22166317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166317&pc=OZD | 05/14/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657642 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support experienced device malpositioning and ventricular tachycardia requiring device repositioning. approximately two days after start of impella mcs, the patient experienced arrhythmia. echocardiogram showed the device was deep. consequently, due to the ventricular tachycardia and the pump sitting against the wall, the impella 5.5 pump was repositioned, pulled back 3 to 4 centimeters. the arrhythmia noted to have resolved. three days later, the patient experience another run of ventricular tachycardia. however, echocardiogram showed the impella 5.5 device in good position (mid ventricle). a viability study had been completed, and the patient was not a candidate for surgery (surgical turndown); there were no other alternative options for the patient. the patient remained on impella mcs over the weekend. the next weekday, the impella device was therefore weaned and explanted with a survived patient outcome. the patient was stable and planned for discharge home. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8673 | 1220648-2025-29112 | 22166354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166354&pc=OZD | 05/10/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/11/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621510 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump support ongoing for a patient admitted in with need for coronary artery bypass graft x 4 and had the impella 5.5 pump placed for the support of the patient. the patient had ectopy and runs of ventricular tachycardia (vt) requiring defibrillation with 200 j. no chest compressions were required. the pump position was assessed. the left ventricle was noted to be small. the pump remained on for support despite the vt for five more days. | [5 empty] | The impella device was not returned, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the ectopy and ventricular tachycardia (vt) was not determined.
8674 | 1220648-2025-29113 | 22166450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166450&pc=OZD | 09/03/2024 | 06/09/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511324 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A literature title "spinal cord infarction as a complication of v-ecmo with microaxial flow pump in cardiogenic shock" was received that the patient on impella 5.5 support was diagnosed with spinal cord infarction (sci) based on clinical examination and maintained on systemic anticoagulation without thrombolytics. additional information was received that it is unclear to the team if it¿s caused by ecmo, impella, or another unknown source. at the time of the reported event, the medical doctor stated that it was not device related. the patient survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8675 | 1220648-2025-29114 | 22166576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166576&pc=OZD | 05/13/2025 | 07/02/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645610 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp, a patient developed hematuria. the patient had pink-tinged urine. it was unclear if the hematuria was hemolysis or due to a traumatic foley insertion. the pump's position was confirmed to be in the left ventricle. the patient was successfully weaned from the pump and survived. | [5 empty] | Data review: data logs showed pump was started and ran without issue. about 25.5 hours into support, pump stopped motor current (mc) spiked that triggered alarm 13 impella stopped-mc high. pump could be restarted and continued running on p-1 to the rest of the case. since purge and flow were also affected at the time of the pump stop, this is an indication of ingestion. product was not returned for review. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely biomaterial ingestion since the log meets the data pattern for biomaterial ingestion. the root cause of temporary pump stop was most likely biomaterial ingestion since the log meets the data pattern for biomaterial ingestion. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8676 | 1220648-2025-29115 | 22166651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166651&pc=OZD | 05/14/2025 | 06/18/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1457787 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a automated impella controller had a controller error that was discover in the cath lab. the issue was not related to a procedure therefore no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. console logs from the reported day of event (2025-05-14) are consistent with complaint and show controller error alarms (1039, defective optical sensor lamp), preceded by ¿5s¿ parameter measurements, which indicate very low, or no light, entering the optical bench ccd detector. the alarm was persistent, upon each console boot-up, and occurred for the first time on (b)(6) 2025 (pump 577012 was used), and was observed upon each boot-up between (b)(6) 2025. logs prior to (b)(6) 2025 show optical bench fully operational during preventive maintenance testing on (b)(6) 2025 (pm documentation shows ¿5s¿ within specification), and no alarms during first brief power-cycle on (b)(6) 2025, when console was already at the customer¿s site. inspection of the console performed and revealed that lamp assembly was producing light, and despite no visible contamination of optical connector fiber there was no light coming out of optical pump connector. after replacing the connector with a know-good one, the issue was resolved, and ¿5s¿ parameters were within specification. the controller error alarms was caused by damaged optical pump connector, however the root cause of the failure was not determined due to lack of evidence. | The impella device was returned, and an evaluation found is underway. upon completion of our analysis, a supplemental report will be filed. console logs from the reported day of event (b)(6) 2025 are consistent with complaint and show controller error alarms (#1039, defective optical sensor lamp), preceded by ¿5s¿ parameter measurements, which indicate very low, or no light, entering the optical bench ccd detector. the alarm was persistent, upon each console boot-up, and occurred for the first time on (b)(6) 2025 (pump 577012 was used), and was observed upon each boot-up between (b)(6) 2025 and (b)(6) 2025. logs prior to (b)(6) 2025 show optical bench fully operational during preventive maintenance testing on (b)(6) 2025 (pm documentation shows ¿5s¿ within specification), and no alarms during first brief power-cycle on (b)(6) 2025, when console was already at the customer¿s site. inspection of the console performed in danvers revealed that lamp assembly was producing light, and despite no visible contamination of optical connector fiber there was no light coming out of optical pump connector. after replacing the connector with a know-good one, the issue was resolved, and ¿5s¿ parameters were within specification. the controller error alarms was caused by damaged optical pump connector; however, the root cause of the failure was not determined due to lack of evidence.
8677 | 1220648-2025-29116 | 22166866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22166866&pc=OZD | 05/14/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479694 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, air was noted in lv/ascending aorta post impella 5.5 insertion. it was noted that the patient¿s arterial cannula from va ecmo had become dislodged while intra-aortic balloon pump (iabp) was still present. possibly allowing air to be sucked into vasculature. there was no reported patient harm or injury. | [5 empty] | D6 explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8678 | 1220648-2025-29117 | 22167223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22167223&pc=OZD | 05/14/2025 | 06/25/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025602696 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during the implant, resistance was met at the aortic arch and the pump jumped forward and the cannula kinked. the doctor advanced the device to position and started however, low flow and persistent suction were noted, kink clearly visible just below outflow cage, the impella turned to p0 and was swapped out for a new pump and the case was completed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. low pump flow: ¿clinical mentioned that resistance was met at aortic arch and impella cannula kinked. low flows were reported after pump start with persistent suction. it is likely the low pump flow was as a result of cannula kink. the root cause of the low pump flow was most likely patient condition related as evidenced by clinical detail of cannula kink when resistance was met passing pump through aortic arch and low flow after activation product damage (cannula kink): clinical mention cannula kink below outflow cage when resistance was met passing pump through aortic arch, confirmed via echocardiogram. the root cause of the product damage (cannula kink) was most likely due to a catheter kink as evidenced by clinical details of resistance was met passing pump through aortic arch.
8679 | 1220648-2025-29118 | 22167312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22167312&pc=OZD | 02/18/2025 | 08/13/2025 |  | 06/09/2025 |  |  | 05/26/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587220 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient went into hemorrhagic shock and passed away during impella cp support. it was noted that prior to impella cp implant, a retroperitoneal bleed had developed at the time of extracorporeal membrane oxygenation insertion. despite the ongoing condition, the patient was given anticoagulants following impella cp implant leading to the hemorrhagic shock. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8680 | 1220648-2025-29119 | 22167371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22167371&pc=OZD | 05/13/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660243 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Implant Pain | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, on the day after implant the patient's limb became numb/paresthesia despite intact pulses. limb was not ischemic. the team opened the axillary incision bedside and removed material that may have been causing the compression. the pump remained on for support for approximately two weeks when the patient expired after sepsis from hernia surgery and care/comfort measures. | [5 empty] | The investigation for the reported pain at insertion site and vascular damage peripheral vessel issue has been completed. the impella device was not received from the customer and the insufficient clinical details provided were not sufficient to determine a root cause. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8681 | 1220648-2025-29120 | 22167887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22167887&pc=OZD | 05/13/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/13/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with hypoxic and hypotensive with subsequent pulseless electrical activity arrest requiring one round of cardiopulmonary resuscitation and one epinephrine with return of spontaneous circulation achieved and was intubation completed. the decision for an impella cp was made. it was noted that placement of the impella cp was successful in the setting of acute myocardial infarction and cardiogenic shock stage e. during support, there was concern for more of a septic picture, likely a urinary tract infection. white blood cell was down to 30 (from 47) on broad spectrum antibiotics. cardiac output and index remained stable. additionally, it was noted that hemolysis had slightly improved after the pump decreased from p9 to p7. later, it was reported that the team was working to manage a gastrointestinal bleed. however, no information was provided on how the team was treating the bleed. it was noted that the hemolysis resolved after the impella cp was removed and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist: section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8682 | 1220648-2025-29121 | 22168150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168150&pc=OZD | 05/15/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/29/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587923 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and had a cerebral hemorrhage and gradually lost the ability to maintain their vital signs. the doctor did believe the impella could have been a factor. the patient expired due to the cerebral hemorrhage. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8683 | 1220648-2025-29122 | 22168320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168320&pc=OZD | 05/14/2025 | 07/14/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1712497 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. after implant of the impella 5.5 was started, the lv waveform became erratic on the aic. the team troubleshot and replaced the aic. the pump remained on for support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: the console logs from the reported day of the event align with the complaint. the case started on (b)(6) 2025 and pump started at (b)(6) 2025 - 18:02:00 and within minutes position unknown alarms were triggered. additionally, pump was turned off after 25 mins of the pump start. (b)(4). lt and rt logs shows negative spikes for placement signal at start of the pump but eventually resolved. note: please refer to the (b)(4) for pump related investigation. device analysis: the console (b)(6) was returned to field service for further investigation. a visual inspection of the internal circuitry and cable connections of the console did not reveal any issues. the console was tested with a known good pump and purge cassette, and no anomalies were observed during testing. additionally, visible light was detected from the optical pump connector, the analog output from the ccd array of the optical bench (fringe pattern) was seen without any distortion, and the "5s" parameter using optical cavity found that the values were within the specified values as per the pm procedure (0042-7316). root cause: the root cause of the " lv waveform became erratic" was not related to any issue with the aic. no issues were found with the aic. corrective actions: before returning the console to the customer, the following actions are instructed to perform: perform sections 23 of the preventative maintenance procedure (0042-7316) as precaution. perform a full functional check on the console and optical system (0042-7316).
8684 | 1220648-2025-29123 | 22168483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168483&pc=OZD | 05/14/2025 | 06/27/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640540 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had blood loss at surgical site for hip replacement and minimal bleeding at the impella site. the suture at the impella site came loose and had to be resutured. heparin was withheld, two units of packed red blood cells and one unit of platelets were given and support continued. the physician believes that the major cause of the blood loss is from a total hip replacement. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. it was noted that there was active bleeding from total hip replacement site along with the impella site. the root cause of access site bleeding is determined to be patient condition because of the bleeding at multiple sites.
8685 | 1220648-2025-29124 | 22168495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168495&pc=OZD | 05/17/2025 | 06/26/2025 |  | 06/09/2025 |  |  | 05/17/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663915 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was implanted to stabilize a patient during a high risk cardiac procedure. the procedure was performed but the patient became bradycardic with rate in the 20s and a temporary pacemaker was placed. the patient was also started on dopamine. during support, it was noted that the patient developed a hematoma in the right groin after an episode of nausea and vomiting. pressure was held and anticoagulation was stopped, and the hematoma was stabilized. the impella site was stable with reduced hematoma. distal pulses were present. the decision was made to wean off the impella cp and explant due to hematoma. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely anticoagulation management since the activated clotting time (act) level was 332 which is above the recommended level and out of range.
8686 | 1220648-2025-29125 | 22168585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168585&pc=OZD | 05/12/2025 | 07/07/2025 |  | 06/09/2025 |  |  | 05/12/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026675594 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported 5.5 pump support ongoing for a patient admitted in with cardiogenic shock. the pump had purge flow/pressure issues that were remedied by the addition of tissue plasminogen activator (tpa) lysis to the purge fluid. pump remained on to date despite malfunction. patient later expired when care was withdrawn after 11 days of support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of high purge pressure has been completed. the pump was returned for investigation. no physical damage was observed on the returned pump. a significant amount of biomaterial was noted on the outflow cage and impeller. the biomaterial was flushed out during the start of the test run, and the pump was run according to specifications in a bucket of water. there were no issues noted with the purge parameters during the test run. the data log shows a slow rise in purge pressure from 500 mmhg to 700 mmhg over a 20-hour period during the case on the 12th. no purge-related alarms were triggered during this time. purge flow was seen resolved with gradual decrease trend. as per the clinical, lysis protocol applied during purge pressure rise event. the cause of the purge pressure issue was most likely biomaterial buildup, based on data log review and returned product investigation. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8687 | 1220648-2025-29126 | 22168745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168745&pc=OZD | 05/19/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/30/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584255 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller blacked out during impella cp support. support stopped for about 20 to 30 seconds and then restarted. support continued and there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8688 | 1220648-2025-29127 | 22168873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22168873&pc=OZD | 05/14/2025 | 07/07/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616450 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, when repositioning the pump, the team accidentally pulled the pump across the valve and could not re-cross the pump. the hemodynamics were stable and ci was 3.1 with svo2 70. thus, the impella 5.5 was explanted. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the positioning issue was most likely use related since the pump was pulled back across the valve by the surgeon. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8689 | 1220648-2025-29128 | 22169060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22169060&pc=OZD | 05/14/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/14/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced runs ventricular tachycardia requiring intervention and would subsequently expire. the patient was admitted with plan for schedule percutaneous coronary intervention (pci) and decompensated throughout the day. the decision was made to place an impella cp after the pulmonary capillary wedge pressure of 30 mmhg. cardiac index of 1.31 l/min/m2 worsening lactate and ph. during insertion the patient pulseless electrical activity (pea) arrested. compressions initiated and the patient was intubated. persistent ventricular tachycardia/ventricular fibrillation was experienced requiring defibrillation along with advance cardiovascular life support medication. the decision was made to further escalated to extracorporeal membrane oxygenation (ecmo). the patient was eventually stabilized and returned to the unit on supports. later that evening, the patient experienced intermittent ventricular tachycardia. treating with meds and defibrillation. the patient continued on support with unnoted complications. on day seven of the support, the patient was noted to have expired. care was withdrawn. | [5 empty] | Although the decision was made to discontinue impella mechanical circulatory support, there is not enough evidence to exclude the impella device used as an associated factor to the overall outcome (given the arrhythmic events). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. however, an investigation was completed and noted the following: cardiac arrythmia can be reported during impella support and to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿ patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿ timing of the arrythmia event in relation to impella support (during or post-implant) ¿ electrocardiogram (ekg) recordings of the arrhythmia episodes ¿ evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿ assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿ presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿ response to any antiarrhythmic treatments or interventions during the event ¿ any documented device-related issues or complications during impella support ¿ evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿ review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. therefore, to determine the true root cause of the ventricular tachycardia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. in conclusion, as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia was not determined.
8690 | 1220648-2025-29129 | 22169520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22169520&pc=OZD | 05/14/2025 | 06/26/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655581 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Cardiac Tamponade | Thrombosis/Thrombus | Aortic Valve Insufficiency/ Regurgitation | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the purge pressure was high, which prompted the team to add tpa to the purge. the pump remained on for support. | Additional information that a recent transesophageal echocardiogram (tee) showed small clot in the left ventricle and the right ventricle completely collapsed. the patient was brought to the operating room another chest washout and leaky aortic valve replacement repair. |  |  |  |  | D6b explant date unknown the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. based on the data log analysis and clinical details, the root cause of the high purge pressure is most likely due to ingested biomaterial. the investigation of thrombus and heart valve damage has not been completed. should any new information or investigation results become available a supplemental report will be submitted. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter. ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ b1 adverse event was inadvertently omitted on the initial manufacturer device report number. b2 selection was inadvertently omitted on the initial manufacturer device report number. b5 additional information inadvertently omitted on the initial manufacturer device report number. h1 type of report was erroneously selected on the initial manufacturer device report number. h6 health effect - clinical codes and health effect - impact code were inadvertently omitted on the initial manufacturer device report number. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number.
8691 | 1220648-2025-29130 | 22169664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22169664&pc=OZD | 05/15/2025 | 07/08/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1574757 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facilities biomed reported that the display screen of the automated impella controller (aic) froze during prep and couldn't be shutdown. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. based on the data and device analysis there was no evidence in the logs to confirm the frozen screen, the difficulty in shutting down the console was most likely misreported as a frozen screen. upon power cycling the console 50 times, it was able to shut down properly every time. d8 serviced by a 3rd party was erroneously selected yes on the initial manufacturer device report number 1220648-2025-29130. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8692 | 1220648-2025-29131 | 22169906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22169906&pc=OZD | 05/15/2025 | 07/08/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656173 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump appeared to have a clot in the cannula and unable to achieve flows even when ecmo is turned down. the pump flows maxed out at .2 l/m. the pump was removed as a result of the noted issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of low pump flow and thrombus has been completed. low pump flows: data logs were reviewed, and pump flows were immediately low and remained low throughout the rest of the case. there were no motor current (mc) spikes, changes in purge pressure, or saturated flows to support that biomaterial was in the cannula and ultimately ingested. within the first 30 minutes of the case, there was a clear decrease of mc modulation as the mc max values dropped significantly. this triggered the first of many suction (i-min) alarms. there were a total of 62 suction alarms through support. it was also noted later in the case that the pump was repeatedly attempting to ramp up the ms from p-1, triggering a suction alarm, and dropping back down to p-1. the retrograde flow algorithm was likely what was causing the attempts to ramp up, though imc logs were not available to confirm. returned product was investigated and there was no biomaterial, but there were signs of cannula kinks. the pump was run in a bucket and low pump flows did not reproduce at p-9. this was expected because the cannula was in its normal shape during the run test. since there were no signs of biomaterial in the data logs or on returned product, that was likely not the cause of low pump flows. additionally, positioning and low volume likely weren't the cause either because repositioning, turning down ecmo flows, and administering volume did not help resolve low pump flows. therefore, the most likely cause of the low pump flows was the cannula kink found on returned product. insufficient clinical details were provided to determine when or how the cannula became kinked, though it was likely during delivery since pump flows were low during the entire case. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. product damage (cannula kink): upon inspection of returned product, a kink was found in the cannula. the kink was determined to be the most likely root cause of low pump flows. no clinical details were provided to explain how or when the cannula became kinked. based on returned product, the root cause of the product damage was determined to be a cannula kink. insufficient clinical details were provided to determine how the kink occurred. h10 narrative erroneously stated the device was not received on the initial manufacturer device report number 1220648-2025-29131.
8693 | 1220648-2025-29132 | 22169929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22169929&pc=OZD | 05/17/2025 | 09/25/2025 |  | 06/09/2025 |  |  | 05/17/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662779 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Device Sensing Problem | [5 empty] | Hemolysis | [12 empty] | Other | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complaint also a reported a placement signal not reliable alarm. the pump was not explanted due to the alarm. | The user facility reported an impella 5.5 was to be implanted in a patient admitted in with multiple cardiac and systemic comorbidities. the pump placement/delivery was complicated by a wire interaction. the wire snagged and was difficult to remove but support was placed. later, during pump support, the team observed signs of hemolysis in elevated free hemoglobin. the pump supported the patient for 10 days. |  |  |  | The investigation of delivery issues and hemolysis has been completed. thrombus failure mode added for biomaterial observed on impeller. the returned product was inspected and biomaterial was observed beneath impeller blades and in the cannula. the pump passed the light continuity test and the 5s parameters were within normal range. as the necessary information was not provided, the root cause of the thrombus was not determined delivery issues: clinical mentioned guidewire met resistance during removal and was snagged. the returned guidewire was inspected and damage was confirmed. the root cause of the delivery issues was most likely use related as evidenced by the damage on the returned product and the clinical mention of resistance met. hemolysis: biomaterial was observed wrapped around the impeller blades and in the cannula. clinical reported hemolysis on (b)(6). the data logs do not show any motor current spikes before reported hemolysis. there were no frequent or continuous suction alarms or position related alarms. clinical details also mention pump position was confirmed via imaging. the root cause of the hemolysis was most likely biomaterial as evidenced by the biomaterial wrapped around the impeller on the returned product b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29132. b5 patient outcomes was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29132. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29132. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29132. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue: the pump passed the light continuity test and the 5s parameters were within normal range. the data logs show a sudden loss of placement signal after which there was a 'placement signal not reliable' alarm. at the same instant there was a drop in signal-to-noise ratio (snr) and gain. the contrast increased in amplitude and was oscillating between positive and negative values and did not recover throughout support. upon review of the data logs and returned product, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-29132. d10 concomitant medical products and therapy dates updated. | The impella pump and guidewire were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8694 | 1220648-2025-29133 | 22169993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22169993&pc=PYX | 05/15/2025 | 06/25/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  | 06/30/2025 |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024360949 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella rp device for mechanical circulatory support. during implant, the physician had no issues gaining right femoral access for impella rp insertion but outlet unable to cross pv. the impella was pulled back and re-advanced, but cannula within right ventricle buckled over wire when advanced. after multiple attempts the rp insertion was aborted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The clinical details stated that the physician reported no issues gaining right femoral access. based on the clinical details the root cause of the product damage (cannula kink) is due to a cannula kink but the cause is unknown. the root cause of the delivery issues could not be determined as limited information was provided regarding the patient or why it was difficult to advance.
8695 | 1220648-2025-29134 | 22170183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22170183&pc=OZD | 05/17/2025 | 07/14/2025 |  | 06/09/2025 |  |  | 05/17/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp insertion, the left ventricle cavity was small and the patient had ectopy when the impella cp was advanced. additionally, during the cardiac procedure, the patient had one episode of ventricular fibrillation during stent placement and was successfully defibrillated once. it was further reported that the patient arrived from the catheterization lab with a hematoma to the right groin. the dressing was changed and manual compression was held until hematoma was soft to touch. heparin was held. the patients outcome was stable. | [5 empty] | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding is determined to be patient movement because the hematoma was noted after coming to icu from catheterization lab. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8696 | 1220648-2025-29135 | 22170213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22170213&pc=OZD | 05/14/2025 | 07/18/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662744 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. a fluid bolus was given to treat the condition. the following day, a purge system blocked alarm appeared coinciding with a high purge pressure. additionally, the provided flow readings showed signs of a saturated flow. the staff was advised that they could change the purge cassette or add tpa to the purge to troubleshoot the issue. the decision was ultimately made to remove the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | No product or data logs were returned for analysis. high purge pressure: clinical mentioned the following day after implant, a purge system blocked alarm appeared coinciding with a high purge pressure. the root cause of the high purge pressure was not determined as no product or data logs were returned for analysis. saturated flow: clinical mention the provided flow readings showed signs of a saturated flow following purge system blocked alarms. the root cause of the saturated flow was not determined as no product or data logs were returned for analysis. hemolysis: clinical mentioned patient experienced hemolysis. the root cause of the hemolysis was not determined as no product or data logs were returned for analysis.
8697 | 1220648-2025-29136 | 22170247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22170247&pc=OZD | 05/17/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/17/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657476 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | High Blood Pressure/ Hypertension | Tachycardia | Convulsion/Seizure | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and was escalated to impella 5.5 support. during support, the patient had a stroke. however, no hemorrhage was seen on the computed tomography (ct) brain scan. the patient was not moving their right arm or leg. the nurse weaned off the patient off of levophed and vasopressin infusions and added amiodarone infusion for short ventricular tachycardia. after the ct scan, the patient had a seizure. ativan was given and extubation was on hold. additionally, the patient had permissive hypertension due to ischemic stroke. heparin was being held due to worry of hemorrhagic conversion. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8698 | 1220648-2025-29137 | 22170494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22170494&pc=OZD | 05/15/2025 | 07/14/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657723 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, an alarm ¿impella stopped, motor current high¿ was noted. attempted to restart multiple times but unsuccessful. the surgeon opted to replace the pump with a new impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The pump was returned for investigation. data log was not provided for this case run. no physical damage was observed on the returned product. the pump passed the electrical testing. during the test in a bucket of water, the pump stop was reproduced in first attempt. the pump was able to restart after a few attempts, with reported flow saturation; however, the purge parameters were within specification range. the cause of the issue was not determined since data log was not conclusive and sufficient clinical information was not provided.
8699 | 1220648-2025-29138 | 22170815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22170815&pc=OZD | 05/15/2025 | 06/09/2025 |  | 06/09/2025 |  |  | 05/15/2025 | 06/09/2025 |  |  |  |  | 06/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764867 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic was not displaying ps/lv signal and had a yellow alarm that showed "controller error, switch to back up controller" during use. the aic was exchanged and the pump started without further issues. the impella was successfully weaned and explanted in the procedural area. there was no patient harm reported. | [5 empty] | The investigation for the reported issues has been completed. the product was returned. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the log entry processsamplingdatafromopticalbench: sentstopacquisitioncmd ack indicates the sentstopacquisitioncmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false communication stopped condition. per the results of the clinical review and investigation, the root cause was determined to be due to an aic software issue.
8700 | 1220648-2025-29055 | 22157058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157058&pc=OZD | 05/18/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/18/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026653999 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was placed on a patient with cardiogenetic shock. during support, it was noted that the pump ran well without any problems. extracorporeal membrane oxygenation was in weaning process and venting worked well. no escalation to an impella 5.5. initial indication was due to ischemia of the left leg. however, the problem was resolved with the help of an extracorporeal femorofemoral bypass. the next day it was noted that the patient was in critical condition overall. escalation to an impella 5.5 could be an option. however, they first wanted to check the patients neurological status in cause of prolonged resuscitation. the cranial computed tomography was positive for brain damage. the physician withdrew care in cause of diagnosed braindead and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8701 | 1220648-2025-29056 | 22157103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157103&pc=OZD | 05/16/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/16/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664104 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and soon thereafter the patient ventricular fibrillation arrested and expired. the patient was brought to catheterization lab, and initial angiogram showed the left anterior descending (lad) artery was totally occluded. epinephrine and levophed drips were initiated, and the decision was made to place an impella cp for the remainder of the percutaneous coronary intervention (pci) procedure. the lad was successfully opened percutaneous transluminal coronary angioplasty and drug-eluting stent. final angiograms showed improved flows to the treated area. drips were increased slightly over the course of the case due to significant drops in pressures. the decision was made to keep patient on impella mcs post pci while in the intensive care unit. during transfer to the unit, the patient became restless and began to significantly decline. code blue was called once on the floor, and it was noted that the patient was in ventricular fibrillation arrest. cardiopulmonary resuscitation and defibrillation were initiated, and intravenous pushes were administered over the course of 20 minutes. the decision was made to call the code after final pulse check was unsuccessful. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8702 | 1220648-2025-29057 | 22157145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157145&pc=OZD | 05/15/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651794 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the hemashield platinum graft started oozing a few minutes after impella implant. before the percutaneous coronary intervention, the surgeon medicated the graft with glue and tabotamp but the complication could not be managed. it was further reported that the coronary revascularization was successfully completed, and the graft oozing appeared controlled. during the transfer from the operating table to the transport bed, the patient developed ventricular arrhythmia with hemodynamic collapse, which was resolved after sinus rhythm restoration, adrenaline and fluid administration. a few minutes later, a similar episode occurred but the patient was unresponsive to treatment. echocardiography showed complete ventricular dysfunction and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8703 | 1220648-2025-29058 | 22157188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157188&pc=OZD | 05/15/2025 | 07/08/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump failed to restart following being cross clamped and may have possibly been clamped over the motor housing or outlet cage. the pump was restarted later but was replaced due to possible damage. there was no reported patient harm. | [5 empty] | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. temporary pump stop: data log review shows the pump going into surgical mode (p-0) and then when it was attempted to restart the pump, there was a motor current spike and "impella stopped - motor current high" alarm. the pump failed to restart this time. they attempted to start again 8 minutes later and the pump successfully restarted. no product was returned. the root cause of the temporary pump stop was most likely use-related as clinical details state that the pump failed to restart following being cross clamped at the motor housing/outlet cage. saturated flow: the failure mode saturated flow was added since data log review revealed saturated flow after the pump was successfully restarted. there are no clinical details describing events after pump was restarted or regarding the saturated flow. data logs show that after the pump was successfully restarted after being cross-clamped, the pump was running at p-1 for about 3 more hours with saturated flows at ~2.0 l/min. these high flows continued until the end of the data logs. there were no motor current spikes seen that indicated biomaterial ingestion. the root cause of the saturated flow was not determined since product was not returned and insufficient clinical details were provided. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8704 | 1220648-2025-29059 | 22157243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157243&pc=OZD | 05/14/2025 | 07/16/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266534 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) did not detect the purge cassette and disk. the aic was exchanged. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. data analysis: the logs from the complaint revealed that the console displayed ¿purge system failure (piston blocked) - alarm¿ event set: 417, likely due to dextrose deposit. the console also displayed ¿controller error (pbm voltage out-of-spec) - alarm,¿ event 501, due to vpurge voltage out of specification. device analysis: the console was examined and confirmed the dextrose deposit. after cleaning, upon running the optical test pump for 2 hours there was no event 417 or event 501. conclusion: the cause of the piston blocked was due to dextrose deposit. the reported purge disc not detected alarm issue was not determined due to inability to reproduce. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8705 | 1220648-2025-29060 | 22157300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157300&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642334 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 65 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported via protect iv study that the patient suffered from suspected transient global amnesia following a catheter intervention. the condition was believed to have a probable relationship to the impella procedure, and the patient was discharged without sequelae. | [5 empty] | Study name: protect iv, subject (b)(6). the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8706 | 1220648-2025-29061 | 22157332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157332&pc=OZD | 05/15/2025 | 07/02/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658511 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the team had suspicions that the aortic leaflet was either ¿impinged or perforated.¿ the physician stated that the patient¿s tissue was ¿friable.¿ the pump was explanted due to withdrawing care and the patient expired. transesophageal echocardiography (tee) showed probable perforation of the aortic leaflet. | [5 empty] | The investigation of heart valve damage has been completed. impella 5.5 was returned for evaluation. upon visual inspection of the returned product no damage was observed. biomaterial was present within the cannula at the impeller blade. data logs showed a few instances of "impella in aorta" alarms during support. no persistent positioning alarms noted during support. clinical details noted that patient had friable tissue and transesophageal echocardiogram (tee) showed perforation of the aortic leaflet after impella was removed. the root cause of the heart valve damage was most likely patient condition related as clinical details noted that patient had friable tissue and no abnormalities were found on returned product. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
8707 | 1220648-2025-29062 | 22157385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157385&pc=OZD | 05/15/2025 | 07/24/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635803 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the team observed that the luer-lock connection in front of the microfilter was broken, and they wish to do the purge bypass technique. the pump ws successfully weaned and the patient survived the procedure. | [5 empty] | Clinical details report that on the first day of impella cp support, a purge leak was noted at the luer-lock connection in front of the microfilter on the purge sidearm. the purge filter bypass protocol was done and reported to work well through explant on (b)(6) 2025. data logs were not available for review. returned product was investigated and the purge filter bypass was confirmed. when the tape was removed, it was confirmed that the entire purge sidearm had been cut off from the pump and not returned. therefore, the location and characterization of the damage to the sidearm could not be confirmed and investigated. based on the clinical details provided that the leak was noted to be from the purge sidearm and doing a purge filter bypass resolved the complication, the root cause of the purge leak was determined to most likely be a damaged sidearm. the exact location and characterization of the damage could not be determined since the sidearm was not returned for investigation. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
8708 | 1220648-2025-29063 | 22157439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157439&pc=OZD | 05/15/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/15/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, after insertion the impella flows were 0.5 liters per minute. the physician and the team felt that the pump was not adequately primed and requested another pump. the impella cp was removed and replaced with another impella cp. the patient's outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8709 | 1220648-2025-29064 | 22157475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157475&pc=OZD | 05/16/2025 | 09/17/2025 |  | 06/06/2025 |  |  | 05/16/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1768506 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The customer reported that the device exhibited an optical/placement signal issue. | During the investigation of the impella 5.5, automated impella controller - loss of opm data was discovered. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. throughout the case, there were instance of invalid placement signals though none that triggered alarms. the failure described in the complaint did not occur during surrounding usages of this console. the data logs show during the invalid placement signals. the console was returned but the failure mode was not reproduced. the 5s parameters were well above spec. the root cause of the placement signal low alarm was not determined since the failure mode was not reproduced. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. additional information was added to b5 and h10. | The investigation of automated impella controller (aic): loss of opm data has been completed. data analysis: throughout the case, there were instance of invalid placement signals though none that triggered alarms. on the complaint date of 3/29, there was placement signal low alarms confirming the complaint. the failure described in the complaint did not occur during surrounding usages of this console. the realt time (rt) logs show during the invalid placement signals, that the raw optical sensor was spiking to -3276.8 mmhg which intermittently triggered the calculated placement signal to 3002 mmhg. during the placement signal low alarm, the motor current is decreased just beforehand and at that time, signal not reliable (snr) and contrast both improve vastly as placement signal decreases and becomes steadier. device analysis: the console was returned but the failure mode was not reproduced. the 5s parameters were within specified limits. additionally, the optical pump connector was found to be clean and the ccd array analog output was good. a pump was run but did not create the placement signal low condition. ¿root cause: there was no loss of opm data found during the case. | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8710 | 1220648-2025-29065 | 22157577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157577&pc=OZD | 04/21/2024 | 06/06/2025 |  | 06/06/2025 |  |  | 05/14/2025 | 06/06/2025 |  |  | 09/30/2025 |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395945 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 30 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient had an impella 5.5 pump placed to support peri-partum cardiomyopathy. the delivery was difficult through the graft due to small stature and small vasculature. the pump supported for 17 days until explanted for a heart transplant. a publication on the event was located in may 2025 and in reviewing the literature there was a thrombus in the graft that required removal. the patient's outcome was stable. | [5 empty] | Literature citation: tsiouris, a., coleman, c. m., & jeyakumar, a. k. (2024). successful implantation and removal of impella 5.5 device via a 4-mm subclavian artery in an adult patient. indian journal of thoracic and cardiovascular surgery, 41(5), 587¿590. https://doi.org/10.1007/s12055-024-01856-w. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8711 | 1220648-2025-29066 | 22157687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157687&pc=OZD | 05/10/2025 | 07/07/2025 |  | 06/06/2025 |  |  | 05/12/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621540 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Cardiac Tamponade | Ectopic Heartbeat | Delirium | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with ventricular septal defect and was placed on venoarterial extracorporeal membrane oxygenation in addition to an impella 5.5 for unloading. during support, it was reported that the patient experienced cardiac tamponade. bedside washout and clot evacuation was performed. in addition, a purge blocked alarm resolved with administration of tissue plasminogen activator. it was also noted that the patient was extubated. however, the patient experienced delirium and reintubated. they also reported that transesophageal echocardiogram confirmed proper impella orientation and position, though slightly advanced at 5.4 centimeters. the patient optimized to p-8 shortly after and began to experience ectopy and suction events at p-8 which resolved at p-6. the impella was repositioned under echocardiogram for suction alarms. the patients outcome was stable. | [5 empty] | The investigation of low pump flow, ventricle perforation, vf/vt/af/at, and high purge pressure has been completed. low pump flow: data log review shows several suction alarms and low flow from (b)(6) 2025 to (b)(6) 2025 at various p-levels. suction alarms on (b)(6) 2025 when the pump was turned from p-8 to p-6 were triggered by diastolic suction. there were no positioning alarms associated with this time. suction alarms continued to be triggered intermittently throughout the rest of the case even after turning the p-level down. motor current was fairly stable throughout this time with no motor current spikes observed. there were some periods of low flow during the day of the complaint, particularly when suction alarms are triggered, but the low flow was not sustained for long periods of time. the investigated high purge pressure was resolved by the time of the suction reports. product was returned and evaluated. the pump was returned cut so it could be not be run in the lab. there was a kink found in the mid-section of the cannula on the returned pump. the correlation between the cannula kink and low pump flow is unknown as the low pump flow was inconsistent throughout the entire case and the cause of the kink is unknown. no biomaterial was observed in the cannula or on/near the impeller. no other abnormalities were found. the root cause of the suction was not determined since the suction alarms did not resolve completely after the repositioning event on (b)(6) 2025 based on data log review, preload was noted to be adequate at the time of the complaint, and the product was returned cut. ventricle perforation: the root cause of the ventricle perforation was not determined since insufficient clinical details were provided about the timing/cause of the perforation. vf/vt/af/at: the cause of the ectopy was unable to be determined due to insufficient clinical details. the failure mode vt/vf/at/af is used as a category for injuries including arrhythmia, asystole, bradycardia, cardiac arrest, heart block, paroxysmal atrial fibrillation, tachycardia, ventricular fibrillation, and ventricular premature complexes. high purge pressure (hpp): data log review shows purge pressure began to slightly/gradually increase and purge flow slightly/gradually decrease on (b)(6) 2025. then, on (b)(6) 2025, there's a 2 minute increase in purge pressure from 600 mmhg to greater than 1000 mmhg. at the beginning of this 2-min. rise, there is a shift up in motor current with a speed deviation (but no motor current spike). the high purge pressure was then sustained for about 10 hours before returning to 600 mmhg. there were also three instances of long bag/cassette changes prior to the hpp events. however, none of these instances occurred directly before the initial purge pressure increase or the 2-minute rise in pp on (b)(6) 2025. product was returned and evaluated. the pump was returned cut so it could not be tested for hpp reproduction. no biomaterial was observed in the cannula or on/near the impeller. no other abnormalities were found. the root cause of the high purge pressure was most likely biomaterial buildup as data log review showed a slight gradual rise in pp near the beginning of the case in combination with a 2-minute rise to hpp without motor current spikes. product damage (cannula kink): the failure mode (fm) product damage (cannula kink) was added since the returned product showed evidence of a cannula kink. there is no specific mention of a cannula kink in the clinical details, but there are reports of suction. data logs are not applicable. the returned pump had a kink in the mid-section of the cannula. the root cause of the product damage (cannula kink) was most likely a cannula kink since the returned product had a cannula kink but it is unknown how the kink occurred. instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ g1 reporting contact email revised on manufacturer device report 1220648-2025-29066 in accordance with updated procedures. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29066. h6 medical device problem code 1491 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29066. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-29066. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
8712 | 1220648-2025-29067 | 22157693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157693&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  | 08/31/2025 |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385358 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella cp support had bradycardic arrhythmia with flat heart rhythm and a line, but patient remained awake and alert. the physician took the patient for temporary pacemaker placement. additionally, the patient received two units of red blood cells (prbcs) for an unknown source of bleeding. patient required 2 more units of prbcs. no obvious signs of active bleeding but potentially left groin venous sheath. impella site did not have any oozing or hematoma. it was further reported that the patient had ventricular fibrillation arrhythmia on day two of impella support and heart rate sustaining in the 190s. the physician spoke with family and decided to withdraw care. the patient expired shortly after. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported ventricular fibrillation (vf) and unknown source of bleeding has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the bleeding nor the vf could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8713 | 1220648-2025-29068 | 22157824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157824&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648129 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, there was concern for hemolysis due to urine color. the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8714 | 1220648-2025-29069 | 22157915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22157915&pc=OZD | 06/03/2025 | 07/11/2025 |  | 06/06/2025 |  |  | 06/03/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1340487 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) had a yellow battery alarm and the aic was replaced. this issue occurred during patient support.the patient was eventually weaned and explanted. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) and aic - system self-check error has been completed. based on the clinical details, data and device analysis the following root causes were determined. no issues were found with the aic related to the reported "aic-battery comm failure" alarm. the root cause of the controller error (loss of comm(pbm1)) - alarm¿, event 1023 was determined to be power battery manger (pbm) corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8715 | 1220648-2025-29070 | 22158020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158020&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629964 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient's right arm became mottled during impella 5.5 support. vascular surgery was set to monitor the condition, however the patient passed away the following day. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8716 | 1220648-2025-29071 | 22158087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158087&pc=OZD | 05/11/2025 | 07/15/2025 |  | 06/06/2025 |  |  | 05/11/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562297 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp experienced cardiac arrhythmia. amiodrone was given. two days later, the patient was being weaned off the impella cp with a plan to explant. the patient was on heparin. after moving/repositioning the patient, bleeding inside the patients scrotum was noted. computed tomography scan was prompted. the scan showed a small retroperitoneal bleed in the groin. since the plan was to explant the impella cp, vascular surgery explanted the device successfully. the retroperitoneal bleed was believed to be a result of a high stick. the patient survived the procedure. | [5 empty] | Device recently returned on 07/01/2025 and a product evaluation and investigation are underway, may update record if any new/relevant information becomes available. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8717 | 1220648-2025-29072 | 22158208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158208&pc=OZD | 05/26/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/28/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617578 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced ventricular perforation leading to multiple organ failure and death. there were no problems with the impella insertion or operation and the attending physician stated the death was not due to the impella. abiomed's medical safety officer reviewed the event and determined the death cannot be dissociated and there is not enough evidence to exclude theimpella as an associated factor in the patient's outcome. | [5 empty] | A2, a3, and a4 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8718 | 1220648-2025-29073 | 22158229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158229&pc=OZD | 05/11/2025 | 07/02/2025 |  | 06/06/2025 |  |  | 05/11/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540304 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was implanted in a patient postoperative coronary artery bypass graft on four vessls. the pump was supporting for over five hours when the pump was seen to be in poor position. during repositioning, the impella cp came back across the aortic valve. the medical doctor was unable to re-advance the pump into the left ventricle. the decision was made to remove the pump. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issues most likely was use related due to improper technique since during repositioning. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8719 | 1220648-2025-29074 | 22158519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158519&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620206 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an implanted impella cp pump was found to be in the papillary muscle during support. the pump was explanted the following day and the patient was escalated to impella 5.5 support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8720 | 1220648-2025-29075 | 22158754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158754&pc=OZD | 05/11/2025 | 10/01/2025 |  | 06/06/2025 |  |  | 05/11/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026649911 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump began to experience high purge pressure during patient support. a lysis protocol was sent to the to the staff in response to the decrease in purge flow. support continued for another five days, at which point care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. high purge pressure (hpp): the root cause of the high purge pressure was most likely a kinked purge line due to the quick rises and falls of purge pressure with no evidence of biomaterial ingestion seen in the data logs along with clinical details noting patient transfer at the time of one hpp report. product damage (purge tubing kink): the failure mode product damage (purge tubing kink) was added since the root cause of the high purge pressure was found to most likely be a kinked purge line, and clinical details report multiple phases of hpp and note that the patient was moved from the icu to the or during one instance. data log review showed quick rises and falls of purge pressure throughout the case and minimal changes to motor current and pump flow. the root cause of the product damage (purge tubing kink) was most likely patient movement since clinical details note transferring the patient at one instance of high purge pressure. g1 reporting contact address was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29075. g1 manufacturing site facility name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-29075. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8721 | 1220648-2025-29076 | 22158822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158822&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026648140 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient had impella cp support and up explanting, 4 proglides, a balloon and two angioseals were needed to achieve hemostasis at the access site. no blood products were used for the oozing. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a2, a5 are unknown.
8722 | 1220648-2025-29077 | 22158936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22158936&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675420 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient lost all distal pulse during impella cp support. due to the patient's underlying condition, the patient's family withdrew care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8723 | 1220648-2025-29078 | 22159136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22159136&pc=OZD | 05/07/2025 | 07/07/2025 |  | 06/06/2025 |  |  | 05/09/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1707157 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, the pump stopped four times by itself. the pump would drop to p0 and the nurse manually restarted the pump to p9. the automated impella controller (aic) was then replaced. the patient survived the procedure. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - pump stop has been completed. data analysis: the console logs from the reported day of the event align with the complaint. the case started on (b)(6) 2025 and first impella stopped (current mismatch event 119)" - alarm was issued after 21 days of pump start. additionally, multiple "impella stopped" alarms were seen. device analysis: the console was for further investigation. a visual inspection of the internal circuitry of the console did not reveal any issues. the console was tested with a known good pump and purge cassette, and no anomalies were observed during testing. root cause: the root cause of the pump stop was not related to any issue with the aic. no issues were found with the aic.
8724 | 1220648-2025-29079 | 22159398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22159398&pc=OZD | 06/03/2025 | 07/11/2025 |  | 06/06/2025 |  |  | 06/05/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316946 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) there had an alarm for the control unit. it was reported as a battery issue. the aic was forcibly shut down. it was stated the patient's hemodynamics were stable. the patient remained on support. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues and aic - shutdown or restart issue has been completed. data analysis: the console logs showed the qnx process io-graphics crashed which caused the aic software to reboot. this reboot was accompanied by an unexpected shutdown alarm, and the pump flow decreased to zero before it was resumed at the previous p-level pump flow. therapy continued for next 15 mins before the pump was unplugged, and the console was then powered down. device analysis: the failure mode was not reproduced during investigation on engineering lab bench. the io-graphics was force shutdown using a telnet command and the console responded the same as in the field. additionally, the optical bench was examined for 5s parameters, which were found to be within specification. aic was run with know good pump and purge cassette for a few hours, and no issues were found. root cause the cause of the aic issue (unexpected shutdown alarm) was the 3rd party hardware/software that runs io-graphics. root cause aic - battery charging/other issues: there was no issue found during the case. the device functioned/operated to specification. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8725 | 1220648-2025-29080 | 22159551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22159551&pc=OZD | 05/09/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/09/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654634 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient with an occluded left main artery received a coronary stent and was cannulated for extracorporeal membrane oxygenation. the patient was implanted with an impella 5.5 and during support, the patient experienced a gastrointestinal bleed. heparin was withheld. eleven days later, the pump into the aorta was unable to recross the aortic valve and the pump was explanted. the patients outcome is unknown. | [5 empty] | The impella device was not returned, and the bleeding investigation has been completed. the root cause of vascular damage peripheral vessel could not be determined since the product was not returned and insufficient clinical details were provided. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8726 | 1220648-2025-29081 | 22159629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22159629&pc=OZD | 05/12/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/12/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1439249 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported impella 5.5 support was ongoing when the automated impella controller (aic) screen went blank. the aic screen did restore itself and to date the console has not been replaced. no harm or injury was reported. the pump remained on for support. the patients outcome is stable. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8727 | 1220648-2025-29082 | 22159753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22159753&pc=OZD | 05/14/2025 | 07/07/2025 |  | 06/06/2025 |  |  | 05/14/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643373 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed some hemolysis overnight during impella cp support. the p-level was lowered and the pump was explanted the following day in response to the condition. the procedural outcome was the patient survived surgery. | [5 empty] | Clinical stated: patient had severe aortic valve insufficiency and valvular disease. patient quickly decompensated from severe aortic regurge and patient was with placement signal alarms all day and audio off. hemolysis observed and urine darken and impella partially removed due to hemolysis. data log reviewed: no mc spike observed outside of p-level changes. suctions occurred, impella position in aorta alarmed followed by placement signal low alarms around 2hrs after support initiated. hemolysis: the cause of hemolysis issue most likely was patient condition related since patient had severe aortic valve insufficiency and valvular disease. aortic regurgitation (leaking valve) can cause turbulent blood flow resulting in hemolysis, which damages red blood cells as they pass through the abnormal valve. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8728 | 1220648-2025-29083 | 22159775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22159775&pc=OZD | 05/12/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/12/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Renal Failure | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, there was difficulty finding distal pulses bilaterally. the patient's leg was warm, pink and dry but their toes were dusky bilaterally. additionally, the patient developed hemolysis. the patient was anuric with hemolyzed labs. the position was verified. this led to the patient having organ damage and dialysis was initiated. two units of blood products were given to the patient. the patient's outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8729 | 1220648-2025-29084 | 22160015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160015&pc=OZD | 11/04/2024 | 11/07/2025 |  | 06/06/2025 |  |  | 11/04/2024 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568002 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Hemoptysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5. the pump generated suction alarms which were resolved by transfusion of red blood cells and fluid bolus. the patient had atrial fibrillation and oozing at the access site. the patient had been anemic upon admission and was being monitored for potential bleeding from the cholecystitis drain and positive guaiac stool tests, and therefore continued to receive transfusions of red blood cells when the hemoglobin dropped below 7mg/dl. the patient exhibited hemoptysis and was treated with heparin and 3% nebulizer. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of the reported failure modes has been completed. no product was returned. data logs were returned for analysis. intermittent suction alarms were seen throughout the case and only resolved after the morning of (b)(6) 2024 around 8am. no motor current spikes were seen. there were intermittent placement signal spikes which corresponded with impella position unknown alarms. no purge pressure spikes were seen outside of purge bag changes. the pump flow was variable for a given p-level throughout the case. low pump flow: the root cause of the low pump flow was most likely patient condition based on clinical details and data log analysis. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: hemoptysis reported. blood in stool reported. the root cause of vascular damage - peripheral vessel was not determined due to lack of sufficient clinical details. g2 report source; study was missing from manufacturer device report 1220648-2025-29084. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: isection: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8730 | 1220648-2025-29085 | 22160187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160187&pc=OZD | 06/12/2024 | 06/06/2025 |  | 06/06/2025 |  |  | 11/24/2024 | 06/06/2025 |  |  | 02/28/2026 |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471134 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhagic Stroke | [11 empty] | Required Intervention | Death | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Abiomed legal originally notified of an allegation regarding a patient who was implanted with an impella 5.5 for mechanical circulatory support and during support, the patient developed a hematoma at the impella insertion site, a massive subdural hematoma, and subarachnoid hemorrhage. no further details surrounding the event were provided. the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8731 | 1220648-2025-29086 | 22160268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160268&pc=OZD | 03/02/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/21/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593536 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 63 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported through abiomed¿s long-term outcome and quality indicator impella registry that a patient implanted with an impella 5.5 support experienced atrial fibrillation (af) and rapid ventricular response (rvr) when the pump was weaned to p 2. the p level was increased and support continued. the patient also displayed hemolysis and was transfused packed red blood cells. the patient was successfully weaned from the pump and survived. | [5 empty] | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
8732 | 1220648-2025-29087 | 22160641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160641&pc=OZD | 05/09/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/30/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616991 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp to support the patient with myocarditis. after 7 days the patient expired from multiple organ failure and the relationship to impella is unknown. additional information provided from the doctor that the extracorporeal membrane oxygenation circuit came loose, causing the patient's overall condition to deteriorate, leading to disseminated intravascular coagulation and subsequent death from multiple organ failure. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. abiomed's medical safety officer review the patients outcome and determined there is not enough evidence to exclude impella as an associated factor in the patient's outcome. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8733 | 1220648-2025-29088 | 22160727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160727&pc=OZD | 06/03/2025 | 08/29/2025 |  | 06/06/2025 |  |  | 06/03/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638131 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was not able to be implanted. the patient¿s vessel was tortuous, and the arch was severely calcified. ecmo was placed. patient outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller states the following : section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a1-a5 are unknown d.4 serial number, lot number, expiration date and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown. | The investigation has been completed. no product was returned. the root cause of delivery issue is determined to be patient condition because the pump was unable to be implanted due to severely calcified and tortuous vessels.
8734 | 1220648-2025-29089 | 22160847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160847&pc=OZD | 05/14/2025 | 07/08/2025 |  | 06/06/2025 |  |  | 05/14/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1255435 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during prep of an automated impella controller (aic) an error message reading "system self check failed. change console immediately" displayed on start screen. the console was turned on and off several time but the error continued to display. the aic was replaced and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the console (ic4602) logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, this confirms the reported failure mode. the console was then forcefully shut down. device analysis: this console was tested and upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm corrosion. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. h6 type of investigation code 4114 was erroneously entered on the initial manufacturer device report number 1220648-2025-29089. h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-29089. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8735 | 1220648-2025-29090 | 22160875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160875&pc=OZD | 05/07/2025 | 07/03/2025 |  | 06/06/2025 |  |  | 05/07/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519927 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the a patient endured ventricular tachycardia (vt) and hemolysis during impella 5.5 support. the condition of hemolysis was said to have been present for the entirety of pump support. the patient labs confirmed the condition and volume status was low. red blood cells and volume were given. the patient was also scheduled to start dialysis around this time. it is unknown if the hemolysis contributed to the renal deficiency. roughly five days into support, the patient had eleven runs of vt. the patient was shocked eleven times in response to the episodes. plans were subsequently made to wean and explant the pump as a result of the conditions. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of vf/vt/af/at, hemolysis, and renal failure has been completed. hemolysis: the data logs do not show any motor current spikes outside p-level changes. the logs show placement signal trending upwards and noisy and clinical mention patient was on dialysis. visual inspection found biomaterial inside the pump housing but not around the impeller. no other issues were found on the rest of the pump. pump passed hemolysis test mih 6.99. the root cause of hemolysis was unable to be determined as no motor current spikes were seen in the logs, no definitive patient related abnormalities were noted and returned pump passed hemolysis testing with no visual abnormalities. vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. renal failure: the cause of the renal failure was not determined due to insufficient clinical details were provided.
8736 | 1220648-2025-29091 | 22160984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22160984&pc=PYX | 03/14/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/02/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1115979 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the sensor of an implanted impella rp flex pump was not zeroed because the prompt did not appear on the automated impella controller (aic). there was no resulting patient harm and support continued uninterrupted. upon review of the logs during the investigation, it was believed that this event may be a aic display related error. | [5 empty] | The investigation has been completed. no product was returned. analysis: case start wizard was started up, and steps 1-3 were performed as the dextrose was spiked, followed by the purge cassette inserting and pump inserted as expected. at 9:30:59, case start is switched from step 3 to 4 then 4 to 7. the user is then prompted to insert purge fluid information (step 6). at 9:31:30, the case start screen is completed, and the pump is started but the logged gui screen appeared: "imcgui: casestart: msg c:90 e:3071 0x82937f0 screen curr:8 prev:8 (3066)". this is not the expected gui screen which would have asked for the zero sensor prompt as indicated by this log entry:"imcgui: casestart: msg c:90 e:3011 0x0 screen curr:8 prev:8 (3009)". the root cause of the incorrect case start display screen issue was not determined since the console was not returned for investigation.
8737 | 1220648-2025-29092 | 22161022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161022&pc=OZD | 05/12/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/12/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610285 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with a left anterior descending artery (lad) lesion. the decision for an impella cp was made before intervention to the coronary. the impella cp was placed and the lad artery was stented. the patient was sent to the unit to recover. during support, it was noted that pulses to the right distal extremity were not present. the patient was taken to cardiac catheterization lab where a femorofemoral bypass utilizing two sheaths was performed. the patient was weaned successfully and the impella cp was explanted. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8738 | 1220648-2025-29093 | 22161040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161040&pc=OZD | 05/14/2025 | 06/27/2025 |  | 06/06/2025 |  |  | 05/14/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320433 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller failure. the aic was removed from service and returned for investigation. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: the complaint is most likely referring to the controller error (not controller failure) alarm (opm comm crc error ¿ event set 1045) the user received on 5/7 that resolved a minute and a half later. the real time (rt) logs confirm during this time all of the optical signals go to -16384 due to the crc error. device analysis: the failure mode was not reproduced during testing. root cause: there was no evidence of a controller failure. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8739 | 1220648-2025-29094 | 22161236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161236&pc=OZD | 05/13/2025 | 07/01/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient experienced runs of ventricular tachycardia when positioned to the right side. | The complainant had an impella cp placed to support a patient admitted in with a st-segment elevation myocardial infarction. throughout the support, the team observed the pump to be out of optimal position. the pump was shallow at the papillary muscle and mitral valve. the malposition persists despite all efforts. the patient's outcome is stable. |  |  |  |  | The investigation of positioning issues and vt/vf has been completed. the impella device was not returned for evaluation. positioning issues: the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. vt/ vf: clinical states that patient had runs of vt when positioned on right side. as the necessary information was not provided, the root cause of the vt/vf was not determined. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29094. b2 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29094. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29094. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29094. h6 health effect - clinical code 2095 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29094. h6 health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2025-29094. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8740 | 1220648-2025-29095 | 22161255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161255&pc=OZD | 05/09/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/09/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648025 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Atrial Fibrillation | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. prior to impella placement, the patient was being treated for atrial fibrillation rapid ventricular response. the patient became hypotensive and levo was started. it was further reported that the patient coded and had around 30 minutes of cardiopulmonary resuscitation before return of spontaneous circulation was achieved. the patient was emergently taken to cath lab for impella placement. during placement, when pigtail was advanced through the peel away, the sheath had a kink in it. a decision was made to exchange long peel away for the short one and impella was placed. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8741 | 1220648-2025-29096 | 22161269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161269&pc=OZD | 05/13/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/13/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646989 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and exhibited access site bleeding. the peel away sheath was explanted and surgical closure was performed. two units of red blood cells were transfused to the patient. the patient was successfully weaned from the device and survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8742 | 1220648-2025-29098 | 22161578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161578&pc=OZD | 05/07/2025 | 07/14/2025 |  | 06/06/2025 |  |  | 05/14/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619785 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support for just over 8 days when the pump stopped due to thrombus at the inlet of the pump. the 5.5 was removed and replaced. the patient survived. | [5 empty] | The pump was returned for investigation. the provided data log did not record any pump stops. the returned pump had free-floating biomaterial in the cannula. the biomaterial was removed from the pump, and it was able to run according to specifications in a bucket of water. the purge parameters were within specifications during the test run. the cause of the pump stop issue was most likely biomaterial ingestion, based on returned product investigation. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8743 | 1220648-2025-29097 | 22161595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161595&pc=OZD | 01/20/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 01/20/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600555 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Thrombocytopenia | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had ectopy. the pump was found fairly deep and the impella was repositioned. the ectopy resolved. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured premature ventricular contractions with a "definite" relationship to the impella device used: ¿on (b)(6) 2025, left cp impella placed for hemodynamic support during emergent pci. left impella has been leading to new pvc¿s post placement. shortly after arriving to the icu started having high burden of pvcs and then non-perfusing nsvt. impella looked deep on ultrasound, was retracted which did improve pvc burden, but ectopy was still frequent enough that lidocaine infusion was started. still deep on repeat limited echo but looks like the pigtail catheter somewhat caught up in some of the muscle trabeculae. he is having hemolysis but this is actually improving after retracting the impella. plan: left impella readjust as needed by interventional cardiology due to pvc's. amiodarone 1.8 mg/ml and cangrelor 50 mg. monitor telemetry. on k+, mg+, ca2+ replacement protocol. repeat limited echo (b)(6) 2025 showed impella in appropriate position in the lv cavity. impella removed and amio and lidocaine dc¿d (b)(6) 2025. no ectopy or sustained arrythmias on tele.¿. the patient recovered with no sequelae. | [5 empty] | The investigation has been completed. no product was returned. as the necessary information was not provided, the root cause of the ectopy was not determined. clinical reported pump repositioned after found to be deep in ventricle. information not provided as to how mispositioning occurred. the root cause of the positioning issue was not determined as no product was returned and insufficient information related to failure was provided. no product was returned for analysis. the data logs do not show any positioning related issues or motor current spikes. clinical confirmed pump was mispositioned and pigtail caught in muscle trabeculae. clinical noted that repositioning resolved hemolysis. the root cause of the hemolysis was most likely due to improper positioning as evidenced by clinical mention of hemolysis resolution after retracting and repositioning pump from contact with muscle trabeculae. the cause of the thrombocytopenia was not determined due to insufficient clinical details. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8744 | 1220648-2025-29099 | 22161802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22161802&pc=OZD | 05/09/2025 | 06/06/2025 |  | 06/06/2025 |  |  | 05/09/2025 | 06/06/2025 |  |  |  |  | 06/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653591 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support had a thoracic washout due to infectious pockets in the chest cavity. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use:
8745 | 1220648-2025-29010 | 22148392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148392&pc=OZD | 05/31/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 06/02/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647352 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Shock | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp pump was placed to support a patient with acute myocardial infarction. after cardiopulmonary resuscitation, an impella cp was inserted to stabilize the hemodynamics. bleeding at the site of chest compressions was difficult to control. a blood transfusion was performed. the bleeding lead to hemorrhagic shock and the patient died. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8746 | 1220648-2025-29011 | 22148494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148494&pc=OZD | 05/19/2025 | 09/29/2025 |  | 06/05/2025 |  |  | 05/19/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1695927 |  | 0042-0040-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient underwent elective percutaneous coronary intervention (pci) supported by impella cp. the nurse observed aorta and left ventricle (lv) signals were abnormal with artifact lines after starting the pump and running at "auto". they tried to adjust the p level down to p4 then p2, but the lv was still showing a white line at p2. alarm "placement signal not reliable" appeared. this happened right after the pump was started before the pci procedure. no guiding catheter was near to the optical sensor. no other equipment was near to the automated impella controller (aic) or the impella pump. the aic was swapped to another aic using the same pump. the signal became normal without any artifacts. the patients blood pressure kept stable during the procedure. the impella cp was weaned after the pci and the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. device analysis: although the 5s parameters (contrast, snr, and gain) of the original optical bench were within specifications, the psnr alarm was triggered when tested with the demo pump. after the optical bench was replaced, no psnr alarms were reproduced, with an snr of 16,243, and both the contrast and raw optical sensor signal remained stable. note: the claim of 'showing a white line at p2' shall be investigated in the disposable complaint (b)(4). root cause: the root cause of the 'ao and lv signal with artifact lines' and the 'placement signal not reliable' issue was most likely an optical bench malfunction. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
8747 | 1220648-2025-29012 | 22148526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148526&pc=OZD | 05/16/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655560 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and developed a gastrointestinal bleed. the team held the systemic heparin as a result. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8748 | 1220648-2025-29013 | 22148729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148729&pc=OZD | 05/12/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025571593 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient with cardiogenic shock was placed on impella cp. during support the physician reported that the leg was ischemic. there was no blood flow detected, and the physician declined a fem-fem bypass and was transferring the patient to a university hospital. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8749 | 1220648-2025-29014 | 22148730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148730&pc=OZD | 05/19/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/19/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662688 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 05/19/2025 |  |  |  |  | The complainant reported a patient endured a ventricular tachycardia (vt) storm. the patient was shocked several times during the procedure for vt. the supported percutaneous coronary intervention was successful and the impella site remained unremarkable. three days later, the pump was successfully weaned and explanted. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8750 | 1220648-2025-29015 | 22148831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148831&pc=OZD | 05/19/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/19/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via publication that a patient was implanted with an impella cp device after a patient was found unresponsive at home, suffering from ventricular tachycardia/ventricular fibrillation. following insertion, hemolysis occurred. a drop in hemoglobin was observed and blood transfusion was given. the pump was weaned and explanted. the patient's outcome was stable. | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8751 | 1220648-2025-29016 | 22148912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22148912&pc=OZD | 09/06/2024 | 06/05/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501322 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed's impact study team that a patient endured thrombocytopenia during impella 5.5 support. the condition was marked with a drop in platelets and was believed to have a probable relationship with impella support. support continued uninterrupted and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8752 | 1220648-2025-29017 | 22149010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149010&pc=OZD | 05/12/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026652493 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient suffered a cerebral embolism during impella 5.5 support. care was withdrawn due to the noted condition. the patient later passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8753 | 1220648-2025-29018 | 22149057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149057&pc=OZD | 05/16/2025 | 07/08/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665517 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after an impella cp was implanted there was an attempt to reposition the impella cp as it was sitting shallow. a thrombus was seen near the post wall on echocardiogram after cannulation. there was also bleeding from all access sites of cannulation. the patient was given multiple blood products. the patient also developed compartment syndrome and a left atrial clot was seen via transesophageal echocardiography. after all conditions were addressed, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of limb ischemia, positioning issues, vascular damage - peripheral vessel, and thrombus has been completed. the impella device was not returned for evaluation. positioning issues: clinical details noted that after patient transferred hospitals, pump was measuring at 2.4cm and was to be repositioned to 4cm. the root cause of the positioning issues was most likely due to patient movement as the pump was observed to be shallow after patient hospital transfer occurred. vascular damage - peripheral vessel: clinical details noted that patient was bleeding from all access sites of cannulation and was given multiple blood products. the root cause of the vascular damage - peripheral vessel could not be definitively determined as limited clinical details were provided. limb ischemia - reversible: as all the necessary information was not provided, the root cause of the limb ischemia was not determined. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8754 | 1220648-2025-29019 | 22149190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149190&pc=OZD | 06/24/2024 | 06/05/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  | 01/31/2026 |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461250 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) study that a subject experienced cardiogenic shock. loqi reports: patient admitted on venoarterial extracorporeal membrane due to cardiogenic shock complicated by a st-segment elevation myocardial infarction and underwent right gastric artery revascularization which was complicated by complete heart block. the patient had lmpella placement for cardiac support. hospital stay was complicated by subarachnoid hemorrhage, spinal ischemia, pneumonia, and chronic respiratory failure. ultimately, the patient deteriorated, and the family opted to proceed with hospice care. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8755 | 1220648-2025-29020 | 22149322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149322&pc=OZD | 05/16/2025 | 07/14/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced hemolysis and device positioning issues which led to intervention. a patient was admitted as non-st-segment elevation myocardial infarction with multivessel disease and ischemic cardiomyopathy heart failure exacerbation. percutaneous coronary intervention (pci) was attempted. however, the pci was aborted due to hypotension. an intra-aortic balloon pump (iabp) was placed, and the patient was sent for surgical work up. the patient was accepted for a coronary artery bypass graft but had since decompensated on the unit with worsening acute kidney injury and heart failure signs and symptoms. the patient was taken to the catheterization lab in acute decompensated cardiogenic shock for iabp removal and impella cp insertion. post impella insertion, there was concern for hemolysis with lactate dehydrogenase of 2200. the pump was shallow at one centimeter and directed posteriorly with possible papillary contact. the pump was advanced two centimeters. lactate dehydrogenase was still around 3000. consequently, the patient was escalated to an impella 5.5 due to need for more support and hemolysis concern. the patient was noted to be in stable condition following the event and latest update notes the patient continues on impella 5.5 mcs. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. hemolysis: the cause of the hemolysis issue most likely was improper positioning related due to improper technique since the issue resolved after repositioning. positioning issues: the cause of shallow (1cm) positioning issues most likely was patient condition related due to small lv since after repositioning, the pump was still at 2cm which is less than the recommended 3.5cm. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8756 | 1220648-2025-29021 | 22149427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149427&pc=OZD | 05/12/2025 | 06/23/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656759 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision for impella 5.5 was made on a high-risk surgery patient with postcardiotomy cardiogenic shock/low cardiac output syndrome. it was reported that the left axillary cutdown, for left ventricle (lv) access, intra-aortic balloon pump was in place, the chest was open, bypass flowing 2.2 lpm, al-1 / 035 to access lv, however the wire was swapped for v-18. the impella 5.5 wouldn¿t advance into ascending aorta. after multiple attempts the 5.5 was swapped for the impella cp however the patient expired in the operating room. abiomed's medical safety officer review the patient's outcome and it was determined there is not enough information to exclude the impella 5.5 device as an associated factor to the patient's outcome. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely use issue since non-abiomed product was used which is not recommended per instructions for use. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. abiomed's medical safety officer review the patients outcome and it was determined there is not enough information to exclude the impella 5.5 device as an associated factor to the patients outcome. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8757 | 1220648-2025-29022 | 22149551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149551&pc=PYX | 05/16/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  | 01/31/2026 |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517206 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support (mcs) and during insertion an issue with the dilator was encountered. subsequently, the patient experienced cardiac arrhythmic events and expired. the patient underwent surgical unroofing of the right coronary artery. post open heart surgery while on the unit, the patient went into acute right ventricular failure. the patient was returned to the operating room, and the decision was made to place an impella rp flex. upon insertion into the right internal jugular, the dilator of a 23 french peel away sheath split at the tip. a new impella rp flex sheath kit was opened, and a new dilator and sheath were placed. following successful pump implant, the patient endured runs of ventricular tachycardia for which multiple shocks were required. support continued with the current setup following the intervention. the patient remained critically ill with very poor prognosis. the patient's chest was open, maxed on multiple inotropes/pressors on continuous renal replacement therapy. approximately three days later, the patient developed ventricular fibrillation arrest in the evening, was shocked, and converted to asystole. paced at a backup rate of 40. advanced cardiac life support protocol was followed minus chest compressions due to open chest status. the family was notified of critical condition. one hour later, the patient ventricular fibrillation arrested again and was shocked one time at 200 joules. blood pressure was 67/41. the family was at bedside and requested to stop all measures, and the patient subsequently passed away. | [5 empty] | There is not enough evidence to exclude the impella device used as an associated factor to the overall outcome (given the arrhythmic events requiring advance cardiac life support). however, it should be noted the patient¿s prognosis was very poor. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
8758 | 1220648-2025-29023 | 22149684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149684&pc=OZD | 05/12/2025 | 06/23/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659445 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump being placed to support a patient admitted with acute myocardial infarction and cardiogenic shock. when having pci the patient deteriorated and cp placed. the delivery failed as the cp kinked and was unable to be delivered via the short 14fr sheath. the product was replaced by long 14fr and the new cp pump. no harm or injury noted. | [5 empty] | The investigation of product damage (cannula kink) and delivery issues has been completed. the impella device was not returned for evaluation. product damage (cannula kink): the root cause of the product damage (cannula kink) issue was most likely a catheter kink, pump inserted and kinked on tortuous anatomy of the left common iliac. delivery issue: the root cause of the delivery issue was most likely use related, kinked pump removed, 13cm peel away swapped out for 25cm peel away and this resolved the issue. new pump inserted successfully. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8759 | 1220648-2025-29024 | 22149728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149728&pc=OZD | 05/14/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/14/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815751 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facilities biomedical engineer reported that the automated impella controller (aic) has a damaged power cord and requested a replacement. a new power cord was sent to the biomedical engineer. there was no report of patient involvement. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8760 | 1220648-2025-29025 | 22149775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22149775&pc=OZD | 05/12/2025 | 06/25/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655068 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was primed and ready for placement. the short peel away inserted without issue. the left ventricle was accessed and pump inserted. pump inlet would not pass right iliac lesion despite multiple attempts and multiple shockwave attempts. the impella cp insertion was aborted. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because of severe right iliac lesion and was unable to pass through despite multiple attempts and the insertion was aborted.
8761 | 1220648-2025-29026 | 22150085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22150085&pc=OZD | 05/11/2025 | 07/02/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666578 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp pump to support a st-segment elevation myocardial infarction patient who had been cardioverted and coded. the site was in the process of getting dressed and it was noticed that pump was deep in the left ventricle. attempt was made to reposition impella and catheter was inadvertently pulled into aorta. multiple attempts were made to recross atrioventricular valve that were unsuccessful. the medical doctor made the decision to open new pump to place patient on support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was inadvertently pulled out of the left ventricle (lv) during repositioning of the pump.
8762 | 1220648-2025-29027 | 22150160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22150160&pc=OZD | 05/11/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/11/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632953 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an implanted impella cp generated suction alarms and came out of position into the aorta. the medical doctor attempted to reposition but was unable to do so. the pump pulled into the descending aorta. the pump was explanted and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8763 | 1220648-2025-29028 | 22150625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22150625&pc=OZD | 05/12/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559749 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp to support a patient in shock and getting cardiopulmonary resuscitation. it was reported that during support the patient has bilaterally cold legs, which the rep states may have been caused by high catecholamines. later it was reported that the patient expired while on support. the abiomed's medical safety officer review the patient's outcome and there is not enough evidence to exclude the impella used as an associated factor to the patient outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. the abiomed's medical safety officer review the patient's outcome and there is not enough evidence to exclude the impella used as an associated factor to the patient outcome. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8764 | 1220648-2025-29029 | 22150949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22150949&pc=OZD | 05/12/2025 | 06/27/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642172 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a significant hematoma following impella cp implant. the site was re-sutured and the perclose were tightened to manage the condition. pressure was noted to help dissipate the hematoma and the site was redressed. eptifibatide was also turned off in response to the hematoma. the hematoma / oozing remained for roughly three days following the initial intervention. support continued uninterrupted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: access site bleed / hematoma. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related since blue suture hub was into the skin up to the first rib. g1 reporting contact email revised on manufacturer device report 1220648-2025-29029 in accordance with updated procedures.
8765 | 1220648-2025-29030 | 22151180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22151180&pc=OZD | 05/12/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623844 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | Asian | Death | Yes | No | Malposition of Device | [6 empty] | Cardiac Perforation | Intraoperative Cardiac Injury | [11 empty] | Life Threatening | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complainant in india reported the patient was on impella cp support and during support, there was a left ventricle perforation. the clinical rep confirmed that the perforation occurred at the end of the procedure and it was caused by the impella. the patient expired within 3 hours of unit admission. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8766 | 1220648-2025-29031 | 22151641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22151641&pc=OZD | 05/12/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632581 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Failure to Sense | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn |  |  | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support and during support, there were suction alarms. albumin was given and support continued. additionally, the patient small ventricular tachycardia (vt) runs during support. furthermore, the patient began having left-sided muscle weakness to upper and lower extremity. the patient had a right pca per ct. symptoms were improving. echo was done and no clots were visible and position was verified. support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8767 | 1220648-2025-29032 | 22151652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22151652&pc=OZD | 08/24/2023 | 10/01/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  | 06/30/2025 |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357378 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | [11 empty] | Life Threatening | Required Intervention | Other | Death | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient endured a gastrointestinal bleed, but no other details were provided. | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening cardiogenic shock with a "unknown (undetermined)" relationship to the impella device used: ¿patient was admitted with severe cardiomyopathy, ejection fraction of <20%, congested hepatopathy. patient in multi-system organ failure. patient had a bronchoscopy, ecmo placement, hemodialysis, patient made dnr by family, time of death 1258 on (b)(6) 2023.¿. the patient outcome fatal. |  |  |  |  | Vascular damage - peripheral vessel: gastrointestinal bleed (gi) bleed was noted in clinical details but no other details were provided to investigate further to determine the root cause. b5 gi bleed was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29032. g1 reporting contact email revised on manufacturer device report 1220648-2025-29032 in accordance with updated procedures. h6 health effect - clinical code 1888 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29032. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: impella 5.5with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8768 | 1220648-2025-29033 | 22151832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22151832&pc=OZD | 03/06/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/11/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620203 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Asystole | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that shortly following impella cp insertion, a patient developed a complete heart block with intermittent very low and wide escape, and was eventually asystolic. the patient required emergent right common femoral vein transvenous pacing (tvp). this was removed after iterative decremental testing demonstrated persistence of this rhythm following percutaneous coronary intervention (despite impella repositioning and removal). the right internal jugular was accessed and used to deliver a screw in tvp in the 7 french merit heath. a few days following explant, the patient again went into asystole without capture of their pacemaker and was agonally breathing. code blue was called and anesthesia intubated for airway protection. the code lasted for 45 minutes with multiple rounds of cardiopulmonary resuscitation, epinephrine, calcium, bicarbonate and attempts to transcutaneously pace which was unfortunately unsuccessful. the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8769 | 1220648-2025-29034 | 22151876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22151876&pc=OZD | 05/08/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/12/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627324 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Other | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured renal failure with a "possible" relationship to the impella device used: ¿the patient was admitted with a history of chronic kidney disease type iii. the patient developed acute-on-chronic kidney insufficiency during index admission due to poor renal perfusion. the patient was started on continuous renal replacement therapy on (b)(6) 2025. admission creatinine 1.90 mg/dl; peak creatinine 4.67 mg/dl ((b)(6) 2025).¿ the patient outcome not recovered/not resolved. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: impella 5.5with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8770 | 1220648-2025-29035 | 22151948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22151948&pc=OZD | 05/11/2025 | 07/02/2025 |  | 06/05/2025 |  |  | 05/11/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628504 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp. the peel away sheath was removed and the repo sheath was advanced but bleeding started where the blue suture pad and white connection meet. the impella cp was removed and not replaced. the patient survived. | [5 empty] | Reported blood leaking from the repositioning unit where the blue butterfly and white repo hub meet and impella cp was removed. impella cp was returned and no damages or abnormalities were found. the blue butterfly is not a hemostatic seal and not intended to be inserted into the skin. the root cause of the access site bleeding - major was most likely use related since there was reported bleeding between the blue butterfly and white repo hub. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8771 | 1220648-2025-29036 | 22152275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152275&pc=OZD | 05/13/2025 | 07/14/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636260 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support and during support, permanent suction alarms were visible on the aic and the flow never get higher than 0.5 l/min independent from which p-level was chosen. fluid treatment was given and reposition the device was attempted, but even at p2 level it did not change the issue. the impella was explanted. | [5 empty] | The impella device received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | Returned compliant cp pump visually inspected. biomaterial around impeller and cannula kink in bending point observed. no broken blade found. biomaterial was soaked in water and removed to run the pump and reproduce reported low pump flow issue. impella flow was within specified limit and no low flow reproduced. little biomaterial still observed on the impeller blades even after soaking and impella ran. data log reviewed: no mc spike outside of p-level changes observed. no placement signal issue seen. no positioning issue observed. many suction alarms occurred started at p_9 following by drop in flow and motor current. low impella flow seen in all following p-levels during the 9hours support and not higher than 0.5 l/min by end of support. since there was no mc spike at the time of the occurring low pump flow, cannula kink was likely the cause of the low pump flow. failure mode product damage (cannula kink) added as a result of cannula kink observation during device evaluation. low pump flow: the cause of low pump flow issue most likely was a cannula kink. product damage (cannula kink): the cause of the cannula kink was not determined since insufficient clinical details were provided.
8772 | 1220648-2025-29037 | 22152337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152337&pc=OZD | 05/13/2025 | 06/10/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129490 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during preventive maintenance the automated impella controller (aic) generated a system check failure alarm. there was no patient involved in the event. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. h10 investigation results was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29037. | No patient was involved in the reported event. therefore, section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8773 | 1220648-2025-29039 | 22152439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152439&pc=OZD | 05/13/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643464 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 93 KG | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that impella cp was implanted to support the percutaneous coronary intervention procedure. upon using single access with companion sheath, the orange ring was cracked on the posterior side. the physician was able to reposition the sheath and complete the procedure. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8774 | 1220648-2025-29038 | 22152448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152448&pc=PYX | 05/18/2025 | 07/18/2025 |  | 06/05/2025 |  |  | 05/18/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025602705 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp flex for support. it was reported that the device failed. placement signal unreliable along with unable to calculate flows were noted. flashing yellow triage on the automated impella controller and no waveforms occurred when the pump turned on. the pump was replaced and there was no patient harm. | [5 empty] | The investigation of optical signal issue has been completed. placement signal issue: the failure mode was updated from optical signal issue to placement signal issue since the investigation revealed that the dp sensor caused the issue. clinical details stated that the console alarmed placement signal not reliable (psnr), unable to calculate flows, and had no waveforms when the pump was turned on. the pump was replaced due to this issue. data log review showed pap out of range at 300 mmhg upon pump connection then stable around 70 mmhg for 2 minutes before drifting to zero and then out of range. there were no abnormalities with the optical parameters during the case. motor current was stable but flows never appear in the logs. product was returned and evaluated. the pump had an open electrical line between sensor and an out-of-spec reading of 0.946 mohms for v+/v-. the sensor face and delo around the sensor did not appear to have any abnormalities. the blue pump plug was opened and the dp sensor cable was looped inside, appearing to be kinked and twisted in spots. the bonds on the blue pump plug were intact. the pump was run and had intermittent placement signal and "flow calculation disabled" message. the sensor was electrical tested multiple times while manipulating the catheter, patient cable, and gently moving the dp sensor cable inside the blue handle and the s+/- line remained open. the cause of the placement signal issue was most likely an electrical open since the returned product had an open line in the dp sensor cable and the cable appeared to be kinked and twisted inside the blue pump plug; however, the exact cause of the electrical open was unable to be determined. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8775 | 1220648-2025-29040 | 22152504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152504&pc=OZD | 05/13/2025 | 06/30/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634547 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 113 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia requiring defibrillation. defibrillation caused the pump to come out of position. the pump could not be repositioned so it was explanted. the patient survived. | [5 empty] | The investigation of positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. positioning issues: pump was placed successfully. as the physician was about to start working on the left anterior descending artery (lad), patient went into sustained ventricular tachycardia (vtach). patient was not sedated and patient was shocked. they moved around rigorously, violently jolted repeatedly and due to his movement the pump became dislodged into the aorta. unable to recross and decision was made to remove pump. the cause of the positioning issue was most likely patient movement as the pump migrated into the aorta when the patient moved around rigorously and wasn't fully sedated. vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. b1 product problem was was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29040. health effect - clinical code 1721 and health effect - impact code 4628 were erroneously entered on the initial manufacturer device report number 1220648-2025-29040. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
8776 | 1220648-2025-29041 | 22152640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152640&pc=OZD | 05/18/2025 | 07/15/2025 |  | 06/05/2025 |  |  | 05/18/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637874 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient that was a bridge to left ventricular assist device has a pump exchanged due to persistent positioning alarms. it was reported that during support on the second impella 5.5 that there was an impella stopped alarm. the automated impella controller (aic) was exchanged without success. the cannula remained across the aortic valve. however, repeated attempts to restart were unsuccessful. the patient remained stable hemodynamically and mentating at baseline throughout the pump stop period. the patient was taken to the operating room for impella 5.5 explant. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of pump stop has been completed. the impella device was not returned for evaluation. the data logs show that the motor current and pump flow were very stable throughout the case. on the 18th, there is suddenly a large spike in motor current to 30000 and an impella 119 stopped alarm. the pump is plugged into another console and attempted to be restarted but the motor current would spike to 30000 and issue a 115 alarm. based on the clinical details and data log analysis, the root cause of the pump stop is most likely due to open electrical lines from excessive force. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8777 | 1220648-2025-29042 | 22152651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152651&pc=OZD | 05/19/2025 | 07/11/2025 |  | 06/05/2025 |  |  | 05/30/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1404666 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella aic (console) blacked out during impella support. support stopped for about 20 to 30 seconds. there was no impact to the patient. | [5 empty] | Additional information provided in h3, h6 and h10. the investigation for the console pump stop and shut down issue has been complete. console logs are consistent with complaint and show that with pump was running at p-6. the console unexpectedly shut down and rebooted. the pump was stopped during start-up communication initialization, but then was detected by the console, recognized as previously running, and automatically restarted at p-6. the pause in hemodynamic support was approx. 29 seconds. during the boot-up sequence console¿s software detected a software process that had previously failed, but this does not necessarily indicate that the process caused the shutdown. an alarm ¿unexpected controller shutdown¿ presented, and data streaming had been disabled, as per design. after pump restart, the support continued without issues for the next 2 days. console was returned for evaluation but the unexpected shutdown and pump stop was not reproduced, therefore root cause could not be determined. the most likely components that could have contributed to the issue were the 4-in-1 assembly and power battery manager assembly, as these are responsible for power delivery and distribution as well as software execution. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
8778 | 1220648-2025-29043 | 22152728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152728&pc=OZD | 05/16/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/26/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP -JP |  |  | ABIOMED, INC. |  |  | US |  | 2025603850 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed to support the patient found down outside with ventricular tachycardia (vt) and placed on cp-ecmo support. the pump was placed but the patient still expired after 3 days. patient had cerebral edema and expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8779 | 1220648-2025-29044 | 22152927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22152927&pc=OZD | 05/13/2025 | 09/17/2025 |  | 06/05/2025 |  |  | 05/16/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567737 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. | The complainant had an impella 5.5 pump placed for a cardiogenic shock patient. the pump had decreasing purge flow and increasing purge pressure. two days later, lyse therapy was performed. purge flow and purge pressure were in range now. additionally, the team stopped heparin systemically due to colon bleeding. the procedural outcome was the patient survived the surgery. |  |  |  |  | The investigation of high purge pressure and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. high purge pressure: the data logs show a gradual rise in purge pressure about 138 hours into support. purge pressure rises for over an hour before high purge pressure alarms while purge flows decrease. clinical mention lysis protocol was followed. purge pressure remained elevated for about 4 or 5 hours before returning to normal values. the root cause of the high purge pressure was most likely biomaterial as evidenced by the pattern observed in the data logs and clinical mention of lysis resolving the issue. vascular damage - peripheral vessel: clinical mention a gi bleed with colon bleeding that led to pause of heparin. no additional information was provided. the root cause of the vascular damage - peripheral vessel was not determined as insufficient clinical information was provided. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data logs do not show any log abnormalities. the 5s parameters were stable and there were no 'impella position unknown' alarms before placement signal monitoring was disabled. the logs do not show any false impella position in aorta or in ventricle alarms. the root cause of the optical signal issue was not determined as there were no issues or alarms related to pump positioning abnormalities, limited clinical information was provided and no product was returned for analysis a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-29044. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-29044 | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8780 | 1220648-2025-29045 | 22153001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153001&pc=OZD | 05/16/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 06/03/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025611201 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient experienced a cardiac tamponade and developed a bleeding tendency during impella cp support. the relationship between impella support and the noted conditions was unknown. after seven days of support, the patient passed away on support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8781 | 1220648-2025-29046 | 22153256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153256&pc=OZD | 05/18/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/18/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613906 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support. after impella cp implantation, the leg became ischemic. distal pulses were not palpable. the leg was cold and high doses of catecholamines were noted. the patient was resusicitated almost the whole night. the decision was made for ecmella. the patient had ventricular fibrillation several times and defibrillation was performed. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8782 | 1220648-2025-29047 | 22153432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153432&pc=OZD | 05/18/2025 | 07/14/2025 |  | 06/05/2025 |  |  | 05/18/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574767 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, a large hematoma formed at the access site. no bleeding appreciated at the skin, but the hematoma extended into the lower abdominal area after the repo sheath was placed and dressed with forward tension sutures and the angle of sheath was maintained. the staff noted that the site looked and felt good, then proceeded to set up for arterial line insertion. the hematoma was noticed and manual pressure was held. the patient had bent their legs multiple times. blood pressure dropped to systolic in the 40s with reduced heart rate. the doctor returned and the patient re-prepped, and performed up and over balloon tamponade from left femoral. an angiogram revealed no dissection or perforation, but confirmed rapid bleeding at arteriotomy. the decision was made to quickly wean the impella and place an intra-aortic balloon pump in the contralateral groin. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. impella cp was returned and no damages or abnormalities were found. the root cause of the access site bleeding/hematoma was not determined due to insufficient clinical details provided. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29047. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8783 | 1220648-2025-29048 | 22153453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153453&pc=OZD | 05/13/2025 | 06/25/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610511 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 90 KG | Hispanic |  | Malfunction | No | Yes | Material Puncture/Hole | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while prepping the patient for impella 5.5 placement in a three point fixation, it was noted that there was blood leaking from the red plug. a new impella 5.5 was implanted to successfully support the patient. | [5 empty] | The investigation of product damage (hole in catheter) has been completed. the returned product was inspected and blood was found in the red pump plug. a hole in the catheter was observed around the 35cm marking. the root cause of the product damage (hole in catheter) was most likely use-related as evidenced by the hole observed on the returned product. h6 medical device problem code 4008 was erroneously entered on the initial manufacturer device report number 1220648-2025-29048. | A5, race, is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8784 | 1220648-2025-29049 | 22153693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153693&pc=OZD | 05/13/2025 | 06/27/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658317 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 61 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that multiple attempts to implant an impella 5.5 failed due to the patient's anatomy. instead, an impella cp was implanted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the impella 5.5 was unable to pass through the vasculature due to small vessel size and late cp was implanted successfully and the support was initiated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
8785 | 1220648-2025-29050 | 22153715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153715&pc=OZD | 05/09/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610727 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported through abiomed's protect iv study that a patient was implanted with an impella cp and exhibited a hematoma at the groin site. a computed tomography scan was performed showing a new non-occlusive thrombus. no patient harm was reported. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8786 | 1220648-2025-29052 | 22153994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22153994&pc=OZD | 05/16/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 05/22/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573959 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the support was ongoing when the limb became ischemic and the team placed an antegrade sheath for perfusion. the limb did not improve and and the patient expired. it was noted that there were no problems with impella operation throughout management. however, the lower limb compartment syndrome was likely caused by the impella. it was further reported that the cause of death was multiple organ failure due to lower limb compartment syndrome. therefore, there is not enough evidence to exclude the impella used as an associated factor to the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8787 | 1220648-2025-29051 | 22154256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22154256&pc=OZD | 05/12/2025 | 06/27/2025 |  | 06/05/2025 |  |  | 05/13/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618613 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported tdelivery issues with an impella cp. the doctor tried to sent the left anterior descending artery and the vessel shut down. the right femoral artery was too small for the impella. the doctor ended up stenting to the left main artery and aborted the impella procedure due to the patient's anatomy. tno patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the impella cp insertion was aborted due to patient anatomy. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿.
8788 | 1220648-2025-29053 | 22154553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22154553&pc=OZD | 03/15/2025 | 08/11/2025 |  | 06/05/2025 |  |  | 05/30/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1578302 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | Unexpected Shutdown |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the patient was on impella 5.5 support when an impella defective alarm occurred, and the pump stopped. the impella was exchanged. several days later, the patient expired from multiple organ failure. | [5 empty] | The investigation has been completed. the impella device was returned for evaluation. the console logs from the reported day of event are mostly consistent with complaint and show that after over four days of uninterrupted support with the pump, a software process (imcgui) became unresponsive, which was detected by software watchdog and a soft reset was initiated as a mitigation response. during system restart, the pump was stopped when pump motor driver (pmd) communication was reinitialized, but prior to pump self-restart, reading of pump eeprom had failed, which resulted in ¿impella defective¿ alarm and inability to restart the pump. console was tested, but issue was not reproduced. reported issue of pump stop on ic8893 was concluded to have been pump-related. the issue only manifested itself after console performed soft-reset, in response to unrelated software process crash, which was followed by failure to read pump eeprom (due to pump failure). the cause of unresponsive software process was not determined, due to inability to reproduce. there were no issues with console's ability to run a good pump. the cause of the software process failure that led to unexpected automated impella controller restart was not determined since issue could not be reproduced. there was no issue regarding console's ability to read eeprom of a good pump. the device functioned/operated to specification. d.9 device returned to manufacturer date was added. b.3 revised date of event as it was entered incorrectly on initial manufacturer device report number 1220648-2025-29053. d.3 manufacturer name was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29053. d.4 lot number was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29053. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-29053 as it is unknown if the initial reporter also sent the report to the food and drug administration. h. 5 revised as it was entered incorrectly on the initial manufacturer device report number 1220648-2025-29053. h.8 added usage of device as it was inadvertently omitted on initial manufacturer device report number 1220648-2025-29053. h.10 instructions for use were reported incorrectly on the initial manufacturer device report number 1220648-2025-29053. | This complaint was open to investigate the impella aic. the death was captured under manufacturer mfr report #1220648-2025-27395. the impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. ifu states: impella cp with smartassist for use during cardiogenic shock. section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
8789 | 1220648-2025-29054 | 22154600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22154600&pc=OZD | 06/02/2025 | 06/05/2025 |  | 06/05/2025 |  |  | 06/02/2025 | 06/05/2025 |  |  |  |  | 06/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1359283 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported cp pump placed to support a cardiomyopathy patient being bridged to vad. the pump was supporting for just under 6 hours when it was explanted and replaced due to sensor failure. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. d4-udi is unknown h4 mfr date is unknown.
8790 | 1220648-2025-28962 | 22139867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22139867&pc=OZD | 05/27/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/28/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574995 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support with intra-aortic balloon pump (iabp). during support, the patient had frequent fatal arrhythmias. the continuation of treatment was abandoned and the patient expired. the patient's overall condition was poor to begin with, and the patient had a strong tendency toward collapse and infection prior to impella being inserted. the patient had been supported with iabp for a long period of time. it was stated that the arrhythmia was not caused by the impella, however it cannot be confirmed that the impella did not contribute to the patient's arrhythmias. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8791 | 1220648-2025-28963 | 22139982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22139982&pc=OZD | 05/09/2025 | 09/25/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660465 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported intermittent "position in aorta" alarms. the pump¿s position was verified via echocardiography (echo). the clinical team considered the alarms likely to be false due to the patient's high end-diastolic pressure and low pulsatility. the alarms were silenced. daily echos were planned to be performed to confirm pump placement. | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient reportedly called emergency medical services due to nausea and vomiting. the patient as taken to the nearest emergency room. code st- segment elevation myocardial infarction was called on the patient. the patient deteriorated and the decision was made to place an impella cp. during impella cp support, the patient experienced atrial fibrilation that was in the 120's. additional information noted care was withdrawn and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and cardiac arrythmia. no product was returned for evaluation. data logs were reviewed and "impella in aorta" alarms were triggered during support. "suction" alarms were triggered concurrently with the "impella in aorta" alarms. all optical parameters were stable, and no hard failures of the optical sensor were observed during support. no motor current spikes were present prior to alarms. the root cause of the optical signal issue was most likely patient condition related. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28963. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | Medical safety review of the event noted there is not enough information to exclude the impella device used as an associated factor. it appears some time after the wire was introduced, the impella was implanted, and the patient developed atrial fibrillation. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8792 | 1220648-2025-28964 | 22140069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140069&pc=OZD | 05/16/2025 | 07/14/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  | 07/31/2025 |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435493 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | Device-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp was inserted via the left femoral artery without issue. very little pulsatility was noted. single access with avanti 7 french sheath was made. percutaneous coronary intervention was completed. bleeding from the diaphragm of the 14 french impella sheath was observed, and the decision was made to peel away. forward tension sutures were placed. the device was locked and dressed. the patient was reported to be in stable condition following the event and as of the latest update continued on impella mcs. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding - major: during implant procedure, impella was inserted without issue with very little pulsitility noted. 14frx13cm diaphragm of sheath bleeding after single access with avanti 7fr sheath. the introducer and pump were not returned. the cause of the access site bleeding was most likely use related due to best practices not being followed with the abiomed 7fr companion sheath not being used. d6b explant date added. | The device remains implanted, supporting the patient, and therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8793 | 1220648-2025-28965 | 22140155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140155&pc=OZD | 05/09/2025 | 06/24/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626857 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the device was unable to advance through the innominate artery. after multiple failed attempts, the implant was aborted and an impella cp pump was placed for patient support. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since in spite of multiple attempts made, surgeon was unable to pass through innominate artery. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller. section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8794 | 1220648-2025-28966 | 22140169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140169&pc=OZD | 05/16/2025 | 07/14/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532470 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient began having dark red urine with hemolysis suspected. the physician ordered the impella cp device to be weaned to performance level p-6. the dark red urine persisted as well as labs hemolyzing. the impella cp device position was verified, and the pump was weaned to p-4. urine began clearing and was pale yellow 30 minutes after change. the patient continued on support for an additional three days, and as planned the impella cp was successfully weaned and explanted without incident with a survived outcome. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the data logs showed pump ran without issue during support. based on clinical description & data analysis, the root cause of hemolysis was most likely patient condition related since the patient had high afterload and hemolysis was resolved after p level was turned down. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8795 | 1220648-2025-28967 | 22140211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140211&pc=OZD | 05/09/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660118 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. after support, there was an uneventful impella cp wean and explant. however, there was persistent ooze from the groin site. hemostasis was achieved by perclose x1, angioseal x1, manual pressure x10 minutes followed by femstop. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8796 | 1220648-2025-28968 | 22140299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140299&pc=OZD | 05/16/2025 | 07/14/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681683 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Defective Device | Patient-Device Incompatibility |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful. during the first attempt, a 14 x 25 introducer was placed. however, resistance was met and was removed. a 14 x 13 peel away sheath was then placed, and once again resistance was encountered and was removed. upon removal, damage was noted at the tip of both sheaths. a third introducer was then placed successfully. however, the staff was unable to advance the pump through the patient's vasculature. calcification was noted in the vessel, and the decision was ultimately made to abort the impella cp device implant. the physician discussed options with the family. however next steps to treat the patient was unnoted, and the patient's status following the event was "unknown." | [5 empty] | The impella pump, purge cassette and introducer were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of introducer damage ¿ mechanical, delivery issues, and difficulty inserting/removing introducer has been completed. the pump was returned with a 14fr x 25cm introducer. no physical damage was noted on the returned pump. the 14fr x 25cm had damage at the tip. the cause of the 14fr x 25cm introducer damage issue was a sheath tip split. the cause of the delivery issue was most likely patient condition related to diseased vessels. the cause of the difficulty inserting/removing 14fr x 25cm introducer issue was most likely patient condition related.
8797 | 1220648-2025-28969 | 22140415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140415&pc=PYX | 05/10/2025 | 06/24/2025 |  | 06/04/2025 |  |  | 05/10/2025 | 06/04/2025 |  |  | 10/31/2025 |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024418612 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp - impella rp flex bipella support was placed to support a patient who was suffering from acute myocardial infarction. the impella rp flex pump was removed after 24 hours. the pump had low flows and the patient experienced hemolysis. the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of hemolysis and low pump flow has been completed. the impella device was not returned for evaluation. hemolysis: the data logs show instances of suction alarms which resolved. after 16 hours, there was a motor current spike with speed deviation and a shift in purge flow and purge pressure. after about 3 hours, there was another motor current spike with a shift in placement signal and onset of reduced flows but flows remained within range for p-level. the motor current spikes are indicative of potential biomaterial interaction. it is unknown when hemolysis onset was. the root cause of the hemolysis was unable to be determined as no product was returned for analysis and specific timing of onset of hemolysis was unknown in addition to limited clinical information provided related to failure. low pump flow: the root cause of the low pump flow was most likely due to biomaterial as evidenced by the motor current spike causing reduced flows in the data logs. thrombus: based on data analysis thrombus was added as a failure mode. as the necessary information was not provided, the root cause of the thrombus was not determined.
8798 | 1220648-2025-28970 | 22140585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140585&pc=OZD | 05/10/2025 | 07/14/2025 |  | 06/04/2025 |  |  | 05/10/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650695 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock was on impella cp support. it was reported that while on support, there was high purge pressure (pp) and purge flow (pf) was low. the purge cassette was replaced and a short time after, the pf and pp were back in range. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of high pressure purge has been completed. the impella device was not returned for evaluation. high purge pressure: data logs were reviewed and a motor current spike and a immediate purge pressure rise of 5 minutes was observed. purge pressure remained high for approximately 1 hours before slowly resolving. clinical details noted a sudden increase in purge pressure that was able to be resolved after tissue plasminogen activator (tpa) was added to the purge solution. the root cause of the high purge pressure was most likely due to biomaterial ingestion based on returned data log analysis.
8799 | 1220648-2025-28971 | 22140824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140824&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487302 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with heart failure/cardiogenic shock was placed on impella cp for support through a planned valve replacement. during support, it was noted that a patient had an arrest. in addition, heparin was stopped due to a gastrointestinal bleed. the family decided on comfort care and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8800 | 1220648-2025-28972 | 22140917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22140917&pc=OZD | 05/09/2025 | 07/11/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571395 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. after placing the 0.018 wire in the left ventricle, placement of the cp was not possible. fluoroscopy revealed severe aortic stenosis (aorta) and it was not possible to advance the impella further. therefore, the implantation was aborted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to severe aortic stenosis.
8801 | 1220648-2025-28973 | 22141114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22141114&pc=OZD | 05/09/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616873 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's right leg became molted and cold. the patient became mottled in all extremities and also chest and back. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8802 | 1220648-2025-28974 | 22141243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22141243&pc=OZD | 05/11/2025 | 07/07/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636175 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Cardiac Arrest | Fever | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella 5.5 placed for support during high-risk cardiac procedure. prior to placement, the patient had an intra-aortic balloon pump placement due to low volume in the left ventricle (lv). however, the patient was eventually implanted with an impella 5.5. during repositioning, the impella 5.5 perforated the lv wall, requiring the chest to be opened for subsequent lv repair. in addition, during support, oozing was noted from multiple sites. the patient received four units of blood. during this time, the patient was in atrial fibrillation in addition to being febrile with last temp 100.9 fahrenheit. labs were drawn. it was also noted that the patient coded but was stabilized and is currently back on triple vasopressor support. the patient was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Ventricle perforation: during repositioning, impella perforated lv wall, required chest to be opened for subsequent lv repair. data logs were reviewed and low ps pulsatility and mc pulsatility observed on 5/12. ps and mc pulsatility recovers on 5/13. the cause of the ventricle perforation was most likely improper positioning of the device as the lv was punctured during repositioning of the pump during the implant procedure. access site bleeding - major: after 1 day on support, pt oozed from multiple sites but not on anticoagulants. no treatment noted and bleeding not noted for remainder of case. the cause of the access site bleeding - major was most patient condition related due to oozing from multiple sites and the bleeding resolving without known treatment. vf/vt/af/at: after 2 days on support, pt experienced afib. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. infection/sepsis: after 3 days on support, pt was febrile with last temp 100.9f. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined
8803 | 1220648-2025-28975 | 22141409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22141409&pc=OZD | 05/09/2025 | 06/23/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661194 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. during the implant, it was noted that the patient had distal aortic calcification so the impella cp would not pass the calcific lesion in distal aorta. the physician decided that the patient did not need the pump, and the cp insertion was aborted. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, patient had distal aortic calcification so pump would not pass the calcific lesion in distal aorta. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8804 | 1220648-2025-28976 | 22141516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22141516&pc=OZD | 05/14/2025 | 07/29/2025 |  | 06/04/2025 |  |  | 05/14/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615543 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the impella 5.5 came out of position after the peel away sheath was used. the same impella 5.5 was re-implanted with another axillary kit. there was no reported harm to the patient. | [5 empty] | The introducer was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the root cause of the positioning issue was most likely use related since the pump was dislodged during sheath removal.
8805 | 1220648-2025-28977 | 22141702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22141702&pc=OZD | 05/14/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/14/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625113 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, leaking was noted at the hub of the impella cp sheath and the companion sheath was repositioned. the procedure was completed and the impella cp was weaned and removed. two perclose and an angioseal could not obtain hemostasis. an up and over balloon was inflated at site to tamponade. covered self-expanding stent was placed and post dilated for hemostasis. the patient outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8806 | 1220648-2025-28978 | 22141824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22141824&pc=OZD | 05/09/2025 | 07/02/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella flows were less than optimal. a kink was subsequently discovered in the catheter that could not be removed through manipulation due to the patient's very short ascending aorta. the staff was advised on the option to replace the pump; however, the decision was made to proceed with the current setup. there was no reported patient harm. | [5 empty] | The investigation of low pump flow and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. low flow: data logs confirmed low pump flow when pump started & remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely patient condition related as cannula kink lead to reduced flows due to anatomy complications. product damage: the root cause of kinked cannula was most likely patient condition related since the patient has a very short ascending aorta resulted in difficult delivery the pump to left ventricle and cause kink. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8807 | 1220648-2025-28979 | 22142058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142058&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  | 07/31/2025 |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379756 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient flipped into atrial fibrillation and cardiac index dropped. overall stable lasic drops were increased in addition to impella flows. the patient had an episode of ventricular tachycardia (vt) a few days later and was cardioverted. the impella 5.5 was repositioned. they pulled back approximately two centimeters (cm) to see if it was contributing to the vt. measuring 4.1 cm and mid left ventricle (lv) space. after another run of vt, the team was concerned it was from poor coronary flow. the impella 5.5 was explanted at bedside. the patient was surgically closed, hemostasis was achieved, and the patient remained stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8808 | 1220648-2025-28980 | 22142065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142065&pc=OZD | 05/09/2025 | 07/01/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636384 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. after support, the impella was successfully weaned. however, during removal the angioseal failed, and manual pressure was applied. the patient continued bleeding after manual pressure. a dissection was found and the vascular surgeon took the patient to the or and a covered stent was placed. the patient was noted to be in stable condition. | [5 empty] | The impella device was discarded, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleed: based on the clinical details provided that the closure device (angioseal) failed, the root cause of the access site bleed was determined to most likely be a use issue. vascular damage - peripheral: since product wasn't returned for investigation, and clinical details are limited and did not report when the damage to the vessel occurred, the root cause of the vascular damage could not be determined.
8809 | 1220648-2025-28981 | 22142155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142155&pc=OZD | 05/13/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/13/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631165 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient endured suspected sepsis during impella 5.5 support. support continued with the implanted pump despite the noted condition. care was withdrawn roughly six days later and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist, section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8810 | 1220648-2025-28982 | 22142365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142365&pc=OZD | 05/14/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/14/2025 | 06/04/2025 |  |  | 11/30/2025 |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024429713 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced some oozing of the axillary site. per the doctor, the bleeding is from the pectoral muscle bleeding. the oozing slowed and multiple blood products were given to treat the condition. the patient outcome was noted to be stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8811 | 1220648-2025-28983 | 22142487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142487&pc=OZD | 05/10/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/10/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643681 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported 70 year old male patient on the impella cp and during support, the patient developed hemolysis with tinged urine. when the hemolysis started to improve and the urine was clearing up, the following day 1 unit of blood was given for hemoglobin of 7.6. it was further reported that the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude the impella device used as an associated factor to the patient outcome. the patient developed hemolysis; severe hemolysis can cause multi-organ failure. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8812 | 1220648-2025-28984 | 22142570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142570&pc=OZD | 05/14/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/14/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1290419 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facilities biomedical engineer reported that during prep, the automated impella controller (aic) system self-check failed. troubleshooting was unsuccessful and the aic was changed immediately. the console was turned on and off. the switch/button was pressed underneath on and off. there was no reported patient involvement. | [5 empty] | The impella device was not returned, and the investigation has been completed. console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the automated impella controller rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen. the issue was confirmed after the console arrived for troubleshooting, as sau_powermod_1 would not load due to loss of communication with pbm1. visual inspection confirmed pbm contamination which has impacted a neighboring rs232 communication channel. it¿s been also observed that both batteries were fully depleted, as the result of pbm failure. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
8813 | 1220648-2025-28985 | 22142793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142793&pc=OZD | 05/10/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/10/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656936 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 support was having ectopy. the impella was repositioned with results of less ectopy. the patient bridged to transplant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8814 | 1220648-2025-28986 | 22142803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142803&pc=OZD | 05/09/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656937 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp and experienced limb ischemia and renal failure. additionally, there was unrelated bleeding from the nasogastric tube and the endotracheal tube and heparin was withheld. ultimately care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8815 | 1220648-2025-28987 | 22142958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22142958&pc=OZD | 05/09/2025 | 07/01/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607789 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a patient and during patient transport, the patient sat up and pulled at the impella catheter. an ¿impella in aorta¿ alarm immediately occurred. the pump was repositioned but the inlet appeared very close to the mitral valve. it was later noted that the patient was stable and the impella cp was successfully weaned and explanted. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of positioning issues most likely was due to patient movement since impella dislodged when patient sat up and pulled at impella catheter. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8816 | 1220648-2025-28988 | 22143007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143007&pc=OZD | 05/11/2025 | 06/25/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643948 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp the patient developed bleeding around the sheath. heparin was withheld and femstop was applied to achieve hemostasis. the impella cp was explanted and replaced with an impella 5.5. the patient survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient movement because patient attempted to sit up 3 times and the site began to bleed. femstop placed to control the bleeding. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8817 | 1220648-2025-28989 | 22143079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143079&pc=PYX | 05/15/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/15/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025602730 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella rp flex support and during support, another suture was applied to assist with oozing. the patient continued to decline so the family opted to withdraw support and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not returned, and the investigation has been completed. the root cause of access site bleeding was most likely use issue since an additional suture was needed to assist the bleed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
8818 | 1220648-2025-28990 | 22143102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143102&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657458 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced cardiac arrythmias. the patient was defibrillated, and was transfused amiodrone. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8819 | 1220648-2025-28991 | 22143260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143260&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the lower right extremity of the patient became cool to the touch with no pedal pulses during impella cp support. the patient was escalated to impella 5.5 support in response to the ischemic limb. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8820 | 1220648-2025-28992 | 22143556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143556&pc=OZD | 05/09/2025 | 06/25/2025 |  | 06/04/2025 |  |  | 05/09/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657653 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 had falsely high ao signals overnight, but a-line was accurate. aortic signal self-resolved. | The complainant reported that after an impella 5.5 was implanted and prior to leaving the operating room, the patient had a sustained run of ventricular tachycardia. one shock was required and the patient's outcome is unknown. |  |  |  |  | B1 revised to reflect additional information received from the clinical site. b5 revised to reflect additional information received from the clinical site. d6b explant date added. h6 medical device problem code revised to reflect additional information received from the clinical site. | The impella device was not received from the customer as it is still supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8821 | 1220648-2025-28993 | 22143719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143719&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/10/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella pc lost pulses in the leg and exhibited hemolysis. the pump was repositioned but the hemolysis did not resolve. reperfusion of the cannula resolved the limb ischemia. support was withdrawn and the patient expired. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8822 | 1220648-2025-28994 | 22143757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143757&pc=OZD | 05/15/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/15/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653654 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had atrial fibrillation. the patient was given aminodarone and support was continued. additionally, there were diastolic suction alarms. blood products were given to the patient. support was continued until weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8823 | 1220648-2025-28995 | 22143809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143809&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658500 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5 the patient experienced ventricular tachycardia. the patient was defibrillated and received cardiopulmonary resuscitation until return of spontaneous circulation was obtained. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8824 | 1220648-2025-28996 | 22143994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22143994&pc=OZD | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656732 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 80 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Gastrointestinal Hemorrhage | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp the patient experienced atrial fibrillation with rapid ventricular response. three failed cardioversion attempts were made. the patient was transfused a bolus of amiodarone and the impella was repositioned. the patient also sustained a gastrointestinal bleed and critically low hemoglobin levels. the patient's family chose to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8825 | 1220648-2025-28997 | 22144301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144301&pc=OZD | 05/15/2025 | 08/12/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786119 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller pulled from service due to a controller failure. there was no patient use at the time they observed the controller failure. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: console logs from the reported day of the event are consistent with the complaint, showing an controller error [pump not connected] - alarm, just after the pump was disconnected. notably, the pump was connected and reconnected twice prior to the alarm. device analysis: the console was returned to field service for further investigation but the issue could not be reproduced. during the inspection, visible light was observed coming out from the external optical pump connector. the voltage supplied to both the optical bench and the bulb was confirmed to be functioning correctly. the 5s parameters were checked and all were within the specified value as per the preventative maintenance (pm) procedure. root cause: no issue were found within the console related to the controller failure. the root cause of the controller error (opm comm stopped - event # 1035) was not determined due to inability to reproduce. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8826 | 1220648-2025-28998 | 22144327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144327&pc=OZD | 05/12/2025 | 07/03/2025 |  | 06/04/2025 |  |  | 05/12/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578034 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was inserted and after starting the cp, high motor current, high lv pressures, and approximately equal flows in systolic, diastolic and medium occurred. the impella position was checked, and it was free in the left ventricle. the issue was not resolved, and cp was exchanged for a new one. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of saturated flow has been completed. there were no data logs retuned. the pump was returned and there were no kinks seen in the catheter or cannula and there was no biomaterial found in the impeller or in the purge gap. the pump was run in a bucket of water and saturated flow reproduced. the motor current spiked to 1500 a few times as well, but there was no material seen in the bucket or around the impeller. the pump was torn down and in the process the bearing became slightly damaged. however, there were no abnormalities seen with the ball bearing, magnet or shaft. due to lack of clinical details and no data logs returned, the true root cause of the saturated flow cannot be determined.
8827 | 1220648-2025-28999 | 22144519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144519&pc=OZD | 05/12/2025 | 07/31/2025 |  | 06/04/2025 |  |  | 05/12/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638092 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was inserted and after starting the cp, high motor current, high lv pressures, and approximately equal flows in systolic, diastolic and middle occurred. the values normalized over time. the intervention was successfully completed. | [5 empty] | The investigation of saturated flow has been completed. the returned product was inspected and no biomaterial was observed in the purge gap. the pump was run in the lab and saturated flows reproduced. the pump was torn down and there was no biomaterial seen under the impeller, in the bearing or around the motor no data logs were provided for the case. the root cause of the saturated flow was unable to be definitively determined as no data logs were provided for the case and no conclusive results from returned product evaluation were obtained. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
8828 | 1220648-2025-29000 | 22144587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144587&pc=OZD | 05/15/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/15/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656874 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and experienced a hematoma which resolved. the patient also had numbness in the hand that remained. the patient was off heparin for the last two days. bival was scheduled to be restarted that day. four days later, a hematoma was noted which required some pain medications. it was noted that the patients site started oozing that night serosangeous fluid. the site was tender but that has been the same since the implant. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8829 | 1220648-2025-29001 | 22144653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144653&pc=OZD | 05/10/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/12/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607810 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the impella cp was placed on a patient with cardiogenic shock for hemodynamic support. however, the pump was removed due to flows not being able to achieve higher than 1.9 l/min on p9. it was noted that the interventional cardiologist believed that the impella catheter got kinked during insertion. the introducer was left in place to insert a new cp. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a1-a5 are unknown.
8830 | 1220648-2025-29002 | 22144786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144786&pc=OZD | 05/11/2025 | 06/27/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648141 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted in a patient, the impella cp stopped multiple times overnight. the pump was able to be restarted, the heart team was made aware. overnight the decision was made to reduce the impella to p-4 and bival was stopped. the patient was hemodynamically stable on dobutamine at 5. the following morning the bedside nurse called stating the pump had stopped again. this time they were unable to restart. multiple attempts were made. the heart team was made aware again. the impella first was pulled back across the aortic valve into the iliac. once the patient was properly sedated, the pump was explanted. post explant, pulses were heard on doppler. the patient was placed on low dose levophed and remained on dobutamine. the patient was hemodynamically stable and the patient left intubated and sedated. the patient survived the explant. | [5 empty] | The impella cp was returned for evaluation. no physical abnormalities were found on the returned product, and the pump passed electrical testing. upon further inspection, a large purge gap was found, which is attributed to bearing wear. the data log shows the pump stopped on day 7 with a motor current high alarm. the pump was able to restart with saturated pump flow, and a purge pressure hike was noted within the specification range. the pump stopped several times and was unable to restart 10 hours after the first temporary pump stop. saturated pump flow was identified as the investigated failure mode based on the data log review. the cause of the pump stop was bearing wear on day 7 of pump use, beyond the pump design life. the cause of the temporary pump stop was bearing wear on day 7 of pump use, beyond the pump design life. the cause of the saturated pump flow was bearing wear, beyond the pump design life. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
8831 | 1220648-2025-29003 | 22144950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144950&pc=OZD | 05/12/2025 | 08/26/2025 |  | 06/04/2025 |  |  | 05/12/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1199628 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon inserting an impella 5.5 pump, a controller error appeared on the aic coinciding with a prompt to switch to a backup controller. the aic was swapped to resolve the issue. there was no patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation has been completed. the impella pump was returned for investigation. on the complaint date, one hour after the support was initiated with pump sn: (b)(6), the epm communication issue led to a reset of the epm. after the reset, due to the persistent epm communication issue, the console displayed ¿controller error (loss of comm (epm)) ¿ alarm,¿ event #1021 (imc_logs_ic3818.pdf). this confirms the reported failure mode. the console then displayed ¿controller error (pbm voltage out-of-spec) ¿ alarm,¿ event #501, due to the abnormal vpch1 voltage, indicating that the epm is drawing more current beyond the threshold of 209ma. because of the epm communication issue, the pump disconnection was not recognized, and the user could not shut down the console in the normal way. hence, the console was shut down by disengaging the battery switch when there was no ac power. the console was tested after arriving at fs. upon the initial bootup, the console rebooted automatically due to an epm communication issue. after the reboot, because of the persistent epm communication issue, the console displayed the "system self check failed" screen. the ribbon cable between the impellatronic board and the 4-in-1 was connected properly. the firmware check via telnet confirmed that the epm was not communicating with the 4-in-1. upon replacing the original impellatronic board with a known good one, the console booted up as intended. while running the optical test pump, no controller error was reproduced. the root cause of the controller error was an impellatronic board malfunction. .
8832 | 1220648-2025-29004 | 22144977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22144977&pc=PYX | 05/11/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/11/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026640586 |  | 1000323 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella rp flex and exhibited low pump flow. the surgeon stated that the flows were down to 1.7 on p-8 and the patient was hemolyzing. the surgeon called again and the patient had been subtherapeutic for anticoagulation and flows had not gotten any better during most of that time. the pump was exchanged for a new impella rp flex and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8833 | 1220648-2025-29005 | 22145117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22145117&pc=OZD | 05/16/2025 | 07/09/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645970 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during an impella cp implant, the pump position was being optimized and the impella cp came out of the ventricle which necessitated removal and eventual recrossing of the aortic valve. the impella cp was successfully removed. however, the impella wire looped back on itself and kinked preventing removal through the sheath. the sheath and wire were pulled as a whole after another standard wire had been placed through the hemostatic valve. pressure was reassessed and had normalized without pharmacological support. the implant was aborted and there was no patient harm. | [5 empty] | The investigation of positioning issues and delivery issues has been completed. positioning issues: clinical details report that after initial access of the lv, pump position was being optimized by the physician and it was pulled back into the aorta mistakenly. returned product was investigated and there were no abnormalities noted with the repositioning unit. data logs were not available for investigation. based on the clinical details provided that the pump was inadvertently pulled back into the aorta, the root cause of the positioning issues was determined to most likely be a use issue. delivery issues: once the pump was mistakenly pulled back into the aorta, the decision was made to restart the delivery procedure. the pump was successfully removed, but when attempting to pull back the 0.018" guidewire, it became kinked. for this reason, it was unable to be removed through the introducer. the both parts were pulled out together, and at this point, the decision was made to not retry delivery because the patient's condition had improved. there were no reports of abnormal patient anatomy or vessel conditions. returned product was investigated and the kinks on the distal part of the catheter were confirmed. there was no split, unraveling, or detachment noted. returned product was kinked and clinical details reported that it became kinked when trying to remove it without any reports of abnormal patient anatomy. for these reasons, the root cause of the delivery issues was determined to most likely be a use issue. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number. | The impella pump and guidewire were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8834 | 1220648-2025-29006 | 22145147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22145147&pc=OZD | 05/29/2025 | 08/26/2025 |  | 06/04/2025 |  |  | 05/29/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP -JP |  |  | ABIOMED, INC. |  |  | US |  | 2025587224 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and on the 9th day of support, there was a pump stop. the pump was explanted. | [5 empty] | The investigation of pump stop has been completed. the impella device was not returned for evaluation. the data logs show that everything was stable up until about 38 hours before the pump stopped. during this time there is a significant drop in purge pressure and purge flow became erratic. there were no motor current spikes or increases in purge pressure before the pump stopped. at the time of the stoppage there was a motor current spike to 1500 triggering alarm #13 and #115. due to lack of product returned, the root cause of the pump stop cannot definitely be determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
8835 | 1220648-2025-29007 | 22145174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22145174&pc=OZD | 05/30/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/30/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026647315 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed to support a but there was an aortic dissection observed and patient expired. the patient received e-cpr on admission and his prognosis was poor so the decision made to stop the pump support. patient expired from the multi-organ failure and aortic dissection. no pump malfunction. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8836 | 1220648-2025-29008 | 22145178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22145178&pc=OZD | 05/16/2025 | 06/04/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620690 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 support developed ectopy overnight. an echocardiogam revealed the pump was at 6.2 centimeters. the pump was repositioned to 5 centimeters. the procedure was successfully completed and the patient outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8837 | 1220648-2025-29009 | 22145303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22145303&pc=OZD | 05/16/2025 | 07/07/2025 |  | 06/04/2025 |  |  | 05/16/2025 | 06/04/2025 |  |  |  |  | 06/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637913 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5. heparin was held until the incision sit was more stable. four days later. heparin infusion had been stopped temporarily due to a small hematoma suspected at the insertion site. the patients outcome is unknown. | [5 empty] | The root cause of the access site bleeding issue was anticoagulation management due to high patient act values; heparin was stopped to achieve hemostasis as per the clinical description. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8838 | 1220648-2025-28917 | 22130922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22130922&pc=OZD | 05/13/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/14/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624100 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was in acute myocardial infarction/cardiogenic shock amd was implanted with an impella 5.5 device for mechanical circulatory support (mcs). the next day after start of impella mcs, the patient was extubated and experienced frequent episodes of ventricular fibrillation throughout the course of the day, requiring repeated cardioverting. the decision was made to reintubate the patient, and further action was unclear. two days later, it was noted the patient was planned for transfer to an outside hospital for a left ventricle assist device transplant and a statement of "electrical storms are much less than in the last few days." no further updates were noted. however, the latest information indicates the patient remains on impella mcs. | [5 empty] | The device remains implanted supporting the patient so therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8839 | 1220648-2025-28918 | 22131151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131151&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026639149 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, there was high purge pressure. lysis was added to the purge. the purge pressures and flows started to normalize. support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8840 | 1220648-2025-28919 | 22131191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131191&pc=OZD | 05/14/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/14/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced access site bleeding, hemolysis and positioning issue. after the impella cp device was implanted, the patient was sent to the unit on support and approximately 12 hours later, it was noted that labs kept hemolyzing. there was a concern for hemolysis. a foley catheter was placed to monitor the urine. the patient's urine started yellow and eventually turned to dark red. discussions with the physician were made and the patient was rapidly weaned off epinephrine. performance level was decreased to p-7 from p-8. an echocardiogram was completed. the pump measured at 2.5 centimeters, and it appeared to be interfering with mitral leaflets. the pump was repositioned, advanced now measuring 3.8 centimeters at 93 centimeters. the hemolysis signs persisted and the impella cp site and all others were oozing. the start of heparin remained postponed. heparin-induced thrombocytopenia and disseminated intravascular coagulation panels were sent. however, approximately 16 hours after start of mcs, the impella cp device was explanted and replaced with an intra-aortic balloon pump for mcs. the urine started to clear. the patients' outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8841 | 1220648-2025-28920 | 22131295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131295&pc=OZD | 05/14/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/14/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651845 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and encountered a low pump flow after approximately a couple of hours on support. the impella cp device had been placed and was at performance level p-7 of approximate flows of 3.2 liters. the patient required bilevel positive airway pressure to maintain adequate oxygenation. coronary artery bypass graft procedure was started once the patient was placed on the pump. the impella cp device was placed in surgical mode. two-vessel bypass was performed without any noted or apparent complications. at completion of the procedure and just prior to weaning from by pump, the physician noted the patient appeared to be deteriorating. impella therapy was resumed and initially was able to achieve performance level p-4 with flows of approximately 2.5 liters. however, shortly after initiation of the impella, the automated impella controller (aic) waveforms quickly separated, and suction was noted. it was advised that the patient could possibly benefit from additional volume. however, placement observation was necessary. echocardiogram was requested and was delayed due to the need for a cardiologist presence. while waiting for echocardiogram, the physician attempted to pull the impella cp back very slowly while monitoring aic waveforms. unfortunately, waveforms remained unchanged, and suction persisted. the discussion was then made regarding impella 5.5 placement. however, this would still require echocardiogram. the physician opted to remove the impella cp device and place an intra-aortic balloon pump due to time constraints and lack of echocardiogram support. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8842 | 1220648-2025-28921 | 22131441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131441&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656872 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the closure of the access site failed following explant. after three hours of manual compression, the bleeding persisted. vascular surgery was subsequently required to manage the condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8843 | 1220648-2025-28922 | 22131450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131450&pc=OZD | 05/10/2025 | 07/17/2025 |  | 06/03/2025 |  |  | 05/12/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588938 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, a thrombus was observed on echocardiogram. few moments later, the aspiration alarm appeared even in p2 and the flow from impella was at 0.1 liters per minute. the pump was explanted and replaced. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Clinical summary: low flow occurred during case. thrombus observed on echo. pump then was replaced by a new cp. data review: data logs showed pump was started & run without issue for about 5 hours & 45 minutes, then flow suddenly dropped to 0.5 l/min that triggered auto suction response & suction alarms. mc spiked and was then dampened with flow dropped to 0.1 l & remained to the rest of the case. product was not returned for review. based on clinical description & data review, the root cause of low flow was most likely due to biomaterial ingestion since the data log shows mc spike followed by dampened mc and low flows & biomaterial observed under echo per clinical team.¿ thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
8844 | 1220648-2025-28923 | 22131616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131616&pc=OZD | 05/14/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/14/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649044 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support developed limb ischemia requiring device explanation. the patient was at a hospital for a colonoscopy and had been having chest pain for the previous two hours. the patient was sent to the emergency department and was found to be having a st-segment elevation myocardial infarction. the patient was taken to the catheterization lab for percutaneous coronary intervention (pci). an impella cp was placed for single access pci. the patient had chest pain relief upon placement of the impella cp and the pci was completed. thereafter, the patient was sent to the unit in stable condition. the following day, the patient was being weaned to possibly explant the impella cp the next morning. however, the nurse noted distal pulse present upon earlier assessment was no longer present. the decision was made to explant the impella cp this day related to loss of distal pulse. during manual pressure hold to the access site post explant, blood pressure decreased and levophed was ordered. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8845 | 1220648-2025-28924 | 22131667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131667&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662677 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, the patient had no distal pulses in their lower extremities. heparin was ordered and distal perfusion catheters were placed to treat the condition. two days into support, heparin was put on temporary hold due to thrombocytopenia and oozing from the access site. the pump was eventually weaned and explanted after five days of support. a dual fasciotomy was performed following pump explant due to compartment syndrome and the non-impella leg was found to have deep vein thrombosis. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8846 | 1220648-2025-28925 | 22131865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131865&pc=OZD | 05/14/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/14/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646670 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and percutaneous coronary intervention (pci). a stent was paced to the left anterior descending artery and percutaneous transluminal coronary angioplasty to the circumflex. after the peel-away sheath was removed, a repositioning sheath was sutured in place. however, at the end of the procedure a hematoma was noted. consequently, the site had to have the sutures redone and angle matching. following, the pump was successfully weaned and explanted from the patient with a survived outcome and the site was closed with perclose. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8847 | 1220648-2025-28926 | 22131910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22131910&pc=OZD | 05/09/2025 | 10/17/2025 |  | 06/03/2025 |  |  | 05/11/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655290 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp the patient had ectopy with any movement. additionally, the patient's right groin developed a hematoma. one unit of blood was transfused to the patient. the patient was successfully weaned from the pump and survived. | Additional information received from the long-term outcome and quality indicator (loqi) impella registry: "patient was seen on (b)(6) 2025 for post hospital follow up. he notes ongoing lump and drainage at the site of his impella placement and angiogram. right femoral artery access/impella site with erythema and drainage on exam, and small nodule palpated with some tenderness on exam. patient and wife note this looks similar from hospital discharge, but given concern for infection will treat with keflex 500 mg qid x 7 days. will also obtain us arterial le pseudoaneurysm to rule out av fistula or pseudoaneurysm. arterial us of le on (b)(6) 2025 showed no evidence of pseudoaneurysm or arteriovenous fistula. on (b)(6) 2025, cardiac therapy note stated that patients groin site better now, and patient stated just 1 small scab on it and no pain." |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B5 updated with additional information receive from the clinical site. h6 health effect - clinical code and health effect - impact code updated according to the additional information received from the clinical site. instructions for use: cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. g2 report source; study was missing from manufacturer device report 1220648-2025-28926.
8848 | 1220648-2025-28927 | 22132008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132008&pc=OZD | 05/14/2025 | 07/02/2025 |  | 06/03/2025 |  |  | 05/14/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633346 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella cp device to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful. the patient was found to have significant distal right coronary artery disease distal to the right internal mammary artery bypass. the decision was made to attempt percutaneous coronary intervention (pci). multiple guides and guidewires were used to attempt crossing of the very calcified lesion. during the pci attempt, lack of distal flow with corresponding st changes on electrocardiogram were noted. the decision was therefore, made to place an impella cp at such time. an arteriogram of right femoral artery revealed significant calcification and tortuosity was noted. the long peel-away sheath was inserted but would not advance past the lesion. the short peel-away sheath was then inserted. left ventricular access was then obtained with pigtail catheter and abiomed .018" wire was inserted. the impella was inserted, advanced, but was unable to cross calcified peripheral vasculature. the physician decided to not attempt percutaneous transluminal angioplasty and the impella implant was aborted. the peel-away sheath was exchanged. an intra-aortic balloon pump (iabp) was inserted. the patient was transferred to the unit in stable condition with plans to transfer to a high level of care facility. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. product damage (cannula kink) added as a failure mode upon product evaluation. reported left ventricle (lv) accessed with guidewire but pump unable to cross calcified peripheral vasculature and noted significant calcification and tortuosity. impella cp was returned and a cannula kink was identified. no other damages or abnormalities found. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient had significant calcification and tortuous vasculature. cannula kink - the root cause of the issue was most likely a kinked cannula due to patient condition since the patient had significant calcification and tortuous vasculature.
8849 | 1220648-2025-28928 | 22132090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132090&pc=OZD | 05/09/2025 | 06/27/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 61 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a patient experiencing cardiac arrest. during support it was noted that the groin site was oozing. packed red blood cells were transfused and the pump position was verified. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. patient had groin oozing. bleeding was resolved and hemostasis was achieved by angle matching. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved through angle matching.
8850 | 1220648-2025-28929 | 22132180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132180&pc=OZD | 05/12/2025 | 06/27/2025 |  | 06/03/2025 |  |  | 05/13/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634281 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support when the impella cp was pulled back too far during mitral valve replacement. there was no patient harm and the pump was subsequently explanted. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issue was user related, as it was accidentally pulled out. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8851 | 1220648-2025-28930 | 22132307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132307&pc=OZD | 05/12/2025 | 06/27/2025 |  | 06/03/2025 |  |  | 05/12/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632704 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant and during the implant, the patients anatomy was tortuous and unable to pass the impella cp pump into the position. the implant was aborted and there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through previous stenting of illiac artery. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system : section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8852 | 1220648-2025-28931 | 22132470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132470&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1760215 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The automated impella controller (aic) had a ¿connect grey to grey connectors¿ error after the purge cassette was inserted, during priming. the screen displayed pictograms related to previous retired impella models. task aborted, and different aic utilized for priming and subsequent impella support without issue. | [5 empty] | The investigation into the reported issue has been completed. console logs were consistent with what was reported in the complaint. during case start step 3, the user was instructed to connect the impella catheter. 20 seconds later, the epm reset because the console did not detect a pump. after the epm reset, the user was prompted with a different iteration of step 3 instructing the user to connect grey connectors. console was tested after arriving in field service. the reported issue was unable to be reproduced while testing on an engineering lab bench. console was able to detect the pump without any issues. the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a mitigation in place.
8853 | 1220648-2025-28932 | 22132519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132519&pc=OZD | 05/12/2025 | 07/02/2025 |  | 06/03/2025 |  |  | 05/12/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp insertion, the left iliac had significant tortuosity and the catheter kinked. the impella cp was exchanged for a new one. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings; ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | Delivery issue: the root cause of delivery issue is determined to be most likely use issue because the new pump was inserted successfully without any issues later. product damage(cannula kink): the root cause of product damage catheter kink is determined to be use issue because the new pump was inserted successfully without any issues later.
8854 | 1220648-2025-28933 | 22132663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132663&pc=OZD | 05/08/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/08/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646048 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced a brief episode of pulse less ventricular tachycardia (vt). one round of cardiopulmonary resuscitation was performed. the impella cp was replaced the next day. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8855 | 1220648-2025-28934 | 22132737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22132737&pc=OZD | 05/08/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/08/2025 | 06/03/2025 |  |  | 05/31/2025 |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353100 |  | 0550-0002 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced heparin-induced thrombocytopenia. the purge fluid was switched to bicarbonate and support was continued. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8856 | 1220648-2025-28935 | 22133168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133168&pc=OZD | 05/13/2025 | 11/17/2025 |  | 06/03/2025 |  |  | 05/13/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654671 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Failure to Sense | Adverse Event Without Identified Device or Use Problem |  |  |  | Fever | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while implanted with an impella 5.5, the patient had a fever. cultures were collected from the access site and the results are pending. additionally, the patient had oozing at the access site. intervention for the bleeding is unknown. impella 5.5 support continues. | Additionally, the patient had intermittent runs of ventricular tachycardia notably with positional changes such as leaning forward. support continued. the pump also had placement signal issues. staff was aware of ¿placement signal not reliable¿ and ¿placement monitoring disabled¿ alarms. staff monitored. the patient eventually stepped on the impella line and the pump stopped. the pump was explanted. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for this issue has been completed. no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed placement signal not reliable (psnr) alarms and placement signal low alarms, both starting around noon of (b)(6) 2025. psnr alarms #130, #1048, and #1036 were triggered. flows were variable for a given p-level, with it dropping to 0 around 8:50am on 7 june 2025. this drop in flow level corresponded with a motor current spike and a purge pressure drop, triggering a motor current high alarm. after this event, the flow rate became more variable with a larger amplitude. no motor current spikes outside of this event and towards the end of the case when the pump stop occurred were seen. no purge pressure spikes were seen outside of purge bag changes. snr dropped to 0 after the motor current high alarm, light remained stable, lamp increased to 100, contrast dropped slightly, and gain increased slightly. access site bleeding: (b)(6) 2025, oozing reported from impella site. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. infection/sepsis: patient with recent positive wound cultures from impella exit site on (b)(6) 2025. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. vf/vt/af/at: patient experienced vtach on (b)(6) 2025. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. optical signal issue: rn and staff aware of ¿placement signal not reliable¿ and ¿placement monitoring disabled¿ alarms. the root cause of the optical signal issue was most likely sensor wear based on the data log analysis and clinical details provided. optical signal issue. optical signal issue. pump stop: pump stop after patient reportedly stepped on impella line and pulled on it when he did. the root cause of the pump stop was most likely electrical open due to patient movement based on the clinical information provided. low pump flow: biomed agard called to informed that this console has a suction alarm, console also due for pm. (b)(4) console complaint. the root cause of the low pump flow was not determined due to insufficient clinical details and inconclusive evidence from data log analysis. b5: added information regarding low flow. h3: added information regarding the status of the device investigation. h6: added codes for low flow and the results of the investigation. h10: added investigation information. | B1 updated to reflect addition of product problem. b5 updated to reflect additional information received from the clinical site. d6b explant date added. h6 updated to reflect the additional information. iimpella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8857 | 1220648-2025-28936 | 22133192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133192&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631949 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 101 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient implanted with an impella 5.5 had two episodes of ventricular fibrillation and was defibrillated. it was further reported that the patient had a gastrointestinal bleed and heparin was withheld. the patient survived. | [5 empty] | A4, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8858 | 1220648-2025-28937 | 22133269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133269&pc=PYX | 05/08/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/08/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590637 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella rp flex and an impella cp. the patient experienced low flow and ventricular tachycardia (vt). the patient was unable to tolerate flows above p5. the patient experienced runs of vt. the plan was to slowly give volume bolus and an attempt again was made. it was emphasized for the need of therapeutic anticoagulation especially below p5. the patient was successfully weaned from the pump and the pump was explanted four days later. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
8859 | 1220648-2025-28938 | 22133404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133404&pc=OZD | 05/08/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/08/2025 | 06/03/2025 |  |  | 05/31/2025 |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359801 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp, the patient had bleeding issues. the physician had nicked the patient causing skin track bleeding. the medical doctor came in early and removed the pump. the patient was transfused with two units of blood after the pump was explanted. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8860 | 1220648-2025-28940 | 22133510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133510&pc=OZD | 05/09/2025 | 06/27/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075752 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella support. during support, there was a controller error and shutdown that occurred on the automated impella controller (aic) that happened during transport. there was no report of patient harm or injury. | [5 empty] | The investigation of automated impella controller (aic) - shutdown or restart issue and aic - controller error (yellow alarm) has been completed. the console logs from the reported day of event are consistent with complaint and show that the unit unexpectedly shut down, then promptly restarted, recognized connected pump, and automatically resumed hemodynamic support. console was but issue was not reproduced, and there was no conclusive evidence in data logs to determine root cause. most likely contribution to the event might have been from the 4-in-1 assembly (including cf cards with software) and pbm assembly (power manager). automated impella controller (aic) - shutdown or restart issue: the cause of the issue was not determined since issue could not be reproduced. aic - controller error (yellow alarm): there were no controller error alarms associated with investigated event d2 common device name revised on manufacturer device report 1220648-2025-28940 in accordance with updated procedures. | The impella automated controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8861 | 1220648-2025-28939 | 22133511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133511&pc=OZD | 05/08/2025 | 07/01/2025 |  | 06/03/2025 |  |  | 05/08/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661901 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp, the patient had oozing at the patient's groin and hemolysis. heparin was withheld and the patient was transfused volume. the pump position was verified and the p level was lowered. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. access site bleeding: clinical states that the patient was transported via flight and oozing was noticed right after that. pump was also found shallow on echocardiogram. as per the clinical information provided, the root cause of the access site bleeding issue was most likely patient movement since oozing was observed after the patient was transported via flight. hemolysis: clinical states that, patient had hemolysis and position was verified on (b)(6) 2025 and volume was given to the patient. data logs were investigated. placement signal was spread and episodes of overload were observed. on (b)(6) 2025, there were no positioning alarms, suction or motor current spike events. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs were inconclusive.
8862 | 1220648-2025-28941 | 22133651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133651&pc=OZD | 05/09/2025 | 09/17/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417342 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | Unintended Application Program Shut Down |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the auomated impella controller (aic) screen would briefly shut down after powering on. the screen went black. it would then return to the home screen. the impella cp stopped. the impella cp and aic were exchanged. there was no patient injury. the patients outcome was noted to be good and the procedure was successful following device exchange. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the shutdown or restart issue, pump stop, and optical signal issue has been completed. the data logs were returned and the placement signal disappears and the motor current flatlines. the console was returned. the failure mode was not reproduced in two runs with a test pump in the field afterwards. the cause of the unexpected restart was the calculator subprocess crashing but the cause of the calculator crashing was not determined since the failure mode was unable to be reproduced. the cause of the pump stopping was not determined since the failure mode was unable to be reproduced. there was no optical signal issue found during the case. the placement signal disappeared because it was no longer able to be calculated due to the calculator crash. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28941.
8863 | 1220648-2025-28942 | 22133864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22133864&pc=OZD | 05/09/2025 | 07/02/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, IN.C |  |  | US |  | 2026612906 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon implantation of an impella cp a purge leak occurred which required replacement of the purge cassette. support continued for five days. upon explant of the pump the patient lost blood flow to the leg. during intervention to restore flow a fibrin clot was found in the pump sheath. blood products were transfused to the patient. the patient was successfully weaned from the pump and survived. | Patient running temps of 101.6. patient was on intravenous ancef and vancomycin and that has been switched to merrem and vancomycin. patient getting intravenous tylenol for fevers. urine, sputum, and blood cultures have been sent. the patient also went into atrial fibrillation (afib) with rapid ventricular response (rvr) . amio, lido, and digoxin given and now also on a amio drip. hear rate in the 60s now with lots of ectopy. patient still febrile; although down from yesterday. patient remains on vancomycin and merrem for antibiotics, so far blood cultures are negative. the physician wanted explant the cp due to fevers. after 5 days of support the pump was explanted. when explanted the team observed the limb to have lost flow. the leg needed intervention to restore blood flow and gave blood products due to blood loss during thrombectomy. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the pump was returned for investigation and there were no visual abnormalities were noted with the pump. access site bleeding: during explant procedure, after clot removal from limb via mechanical thrombectomy to remove fibrin pieces, rcfa starting oozing. this occurred after the cp was removed. the pump was returned where no visual abnormalities were noted. the cause of the access site bleeding - major was most likely use related due to bleeding occurring after pump removal. limb ischemia: as all the necessary information was not provided, the root cause of the limb ischemia was not determined. investigation of vt/af and sepsis/infection have not been completed. should new information or investigation results become available a supplemental report will be submitted. instructions for use: cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28942. h6 health effect - clinical codes 2067, 1729, 1721, 1888, health effect - impact codes 4627, 4644, 4624, and component code 3038 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28942. h6 health effect - impact code and 4629 medical device problem code 1250 were erroneously entered on the initial manufacturer device report number 1220648-2025-28942. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28942.
8864 | 1220648-2025-28943 | 22134055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134055&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The automated impella controller (aic) had a purge disc detection issue. this issue occurred during preparation for impella support. there was no patient injury. | [5 empty] | The impella console was received from the customer and an investigation regarding the returned device has been completed. logs from the complaint date revealed that the console issued purge disc not detected alarm multiple times. the console was examined. a broken purge disc flag was identified. the purge unit driver connector lock was broken. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
8865 | 1220648-2025-28944 | 22134147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134147&pc=OZD | 05/09/2025 | 06/27/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653588 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during site preparation for impella 5.5 implant, bleeding was noted at the anastomosis of the graft. an additional suture was placed to manage the bleed. the subsequent implant was successful. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. there was some bleeding at the anastomosis of the graft. the root cause of the access site bleeding was most likely issue was use related, there was some bleeding at the anastomosis of the graft that required additional sutures to resolve the bleeding.
8866 | 1220648-2025-28945 | 22134151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134151&pc=PYX | 05/09/2025 | 07/07/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561517 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. prior to the impella implant, the patient had a thrombectomy. the doctor was aware of the chance of clot ingestion from thrombus particles being dislodged. the impella rp flex was inserted and flows of 1.1 liters was achieved and then abruptly dropped to zero. the impella rp flex was removed. the clot was visualized in the cannula, and a new device was inserted without complication. the patients outcome was no harm. | [5 empty] | The product was not returned. low pump flow and pump stop. the data logs show that the pump was attempted to be started and the motor current ramps up and it was ramping up for about 23 seconds but no flows were seen and there was spikes in motor current. the flows reach around 1.1 l per minute for about 10 seconds and then the motor current spikes to 30,000 and the pump stops triggering a high motor current pump stop alarm but no electrical alarms. ps remained steady and no purge abnormalities. the clinical details state that there was a clot visualized in the cannula. however, due to lack of product returned the root cause of the low pump flow cannot be confirmed. the root cause of the pump stop cannot be confirmed. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8867 | 1220648-2025-28946 | 22134309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134309&pc=OZD | 05/09/2025 | 06/24/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679978 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp for mechanical circulatory support. during the implant, the left femoral artery was difficult to obtain due to the patient's size and vasculature. after a stiff wire (0.035 wire) was placed in the vessel, fluoroscopy showed that the vessel access was too deep, supposedly in the superficial femoral artery. the wire was moved laterally and down into the deep femoral artery. the wire was eventually removed, and the implant was aborted. no patient harm was reported. | [5 empty] | The investigation of delivery issues and difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. delivery issues: clinical mentions impella cp opened and primed but unable to be placed due to patient anatomy - tortuosity and obesity. the root cause of the delivery issue was patient condition as evidenced by clinical details of abortion due to patient size and anatomy. delivery issues: clinical mentioned fluoroscopy showed that the vessel access was too deep for 0.035 wire, supposedly in the sfa. the wire moved laterally and down into the deep femoral artery. the wire was eventually removed. the root cause of the delivery issue was most likely patient condition related as evidenced by clinical detail of resistance during guidewire placement due to patient anatomy difficulty inserting/removing introducer: the root cause of the difficulty inserting introducer was most likely patient condition related as evidenced by clinical details of implant complications due to patient size and anatomy. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
8868 | 1220648-2025-28947 | 22134537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134537&pc=OZD | 05/09/2025 | 07/02/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656627 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. during delivery of the impella 5.5 pump, resistance was met at the innominate resulting in the catheter bending and buckling. delivery was ultimately successful, and support was initiated for the patient. no patient harm was reported. | [5 empty] | The investigation of product damage (cannula kink) has been completed. reported difficulty passing anatomy since patient had sever kyphosis limiting manipulation and the catheter bending and buckling. impella 5.5 was successfully implanted. impella 5.5 was returned and two cannula kinks were noted. there was a catheter kink noted below the motor housing. catheter kink: the root cause of the product damage was a catheter kink most likely due to patient condition since the patient had severe kyphosis. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8869 | 1220648-2025-28948 | 22134557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134557&pc=OZD | 05/10/2025 | 07/02/2025 |  | 06/03/2025 |  |  | 05/10/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582186 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74 year old patient on an impella cp and during support, blood was continuously dripping from the blue suture collar transition to the white rotatable portion of the repositioning sheath following implant. the pump was removed and a new impella cp device was placed for ongoing support. | [5 empty] | The investigation of access site bleeding has been completed. reported blood dripping from repositioning sheath between blue butterfly and white repo hub. impella cp was returned and no damages or abnormalities were found. guidewire repositioning unit leak testing was performed and no leak was found. blue butterfly is not a hemostatic seal and not intended to be inserted into the skin. the root cause of the access site bleeding was most likely use related since bleeding was reported between the blue butterfly and white repositioning hub. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8870 | 1220648-2025-28949 | 22134599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134599&pc=OZD | 05/09/2025 | 06/27/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that there was oozing and bleeding from the femoral impella site during impella access. manual pressure was held in addition to femoral stop placed. additionally, there was bleeding from other line sites. the decision was made to swap the impella out for intra-aortic balloon pump (iabp) due to bleeding at groin site. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. there was oozing and bleeding from right femoral impella site. the root cause of the access site bleeding issue was most likely patient condition related, there was oozing and bleeding from right femoral impella site during impella access and additionally there was bleeding from other line sites as well. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions. for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8871 | 1220648-2025-28950 | 22134811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134811&pc=OZD | 05/09/2025 | 06/25/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613501 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella 5.5 device for mechanical circulatory support. during the implant, there were no issues with access to the right axillary artery and graft placement, but due to tortuosity in the aorta, they could not pass the impella 5.5 into the left ventricle. the team instead successfully placed an impella cp for support. the patient's outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to ass through vasculature due to tortuosity in aorta and the impella 5.5 insertion was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8872 | 1220648-2025-28951 | 22134882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134882&pc=OZD | 05/10/2025 | 06/18/2025 |  | 06/03/2025 |  |  | 05/10/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622589 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump being placed to support a patient admitted in with multiple cardiac and systemic comorbidities. the pump placement was going in via the axillary artery graft but delivery failed and in the failure the team observed a tear which prompted repair. the graft was partially torn away from the artery. the team placed new pump via a direct approach. | [5 empty] | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue was most likely use issue since non-abiomed product was used which is not recommended per instructions for use (ifu) along with excessive force by the doctor. vascular damage peripheral vessel: the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h10 incomplete instructions for use were provided on the initial manufacturer device report number 1220648-2025-28951. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
8873 | 1220648-2025-28952 | 22134911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22134911&pc=OZD | 05/13/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/13/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649496 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the pump position was verified and it was found that the position was causing mitral regurgitation. the pump was replaced and patients' outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8874 | 1220648-2025-28954 | 22135119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135119&pc=OZD | 05/06/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/06/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532966 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, the sheath was found to be occlusive and bleeding was noted at the access site. a surgical cutdown was performed and the site and blood products were given to treat the condition. the patient remains on support. | [5 empty] | A2 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8875 | 1220648-2025-28953 | 22135121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135121&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626254 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp, the patient exhibited signs of hemolysis. the patient was transfused two units of packed red blood cells. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8876 | 1220648-2025-28955 | 22135360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135360&pc=OZD | 05/06/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/06/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp implant, low flows were encountered. the pump was replaced in response to the failure. there was no indication of a clot on removal. the patient's outcome was updated to expired following the events. additional information was requested but not provided. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8877 | 1220648-2025-28956 | 22135372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135372&pc=OZD | 05/09/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/09/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655580 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and had some sustained ventricular tachycardia. the patient was treated with amiodarone and support was continued for 23 more days. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8878 | 1220648-2025-28957 | 22135521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135521&pc=OZD | 05/13/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/13/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655057 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support when there was a gastrointestinal bleed. support was continued. care was eventually withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8879 | 1220648-2025-28958 | 22135546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135546&pc=OZD | 05/24/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/26/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513969 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump being placed/primed for the support of the patient admitted in with acute myocardial infarction and cardiogenic shock. during priming, no purge fluid was observed flowing from the impella motor area, so insertion of the impella was abandoned. an intra-aortic balloon pump was instead placed. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8880 | 1220648-2025-28959 | 22135632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135632&pc=OZD | 05/13/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/13/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657838 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and went into sustained symptomatic ventricular tachycardia at a rate of 220 with mean arterial pressures in the 40¿s. one shock was given at 200 joules and impella p-level was increased to p5 from p4. the patient converted to sinus tachycardia in the 120¿s. support was continued until weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8881 | 1220648-2025-28960 | 22135762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22135762&pc=OZD | 05/10/2025 | 07/09/2025 |  | 06/03/2025 |  |  | 05/10/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579151 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after placing an impella 5.5 via sternotomy with direct approach on a 86 year old female patient, around 30 to 40 minutes after the insertion of impella, flows began to drop. blood products were being administered during this interim period to address the existing volume deficit. however, flows did not improve and progressively got worse. anesthesia was asked to reassess the position of pump. anesthesia conveyed concerned that the impella was only measuring 4.86cm below the aortic annulus and believed based on the assessment of the transesophageal echo (tee) that the impella catheter was too shallow. it was reported that the catheter appeared deep despite measurement due to the patient¿s small lv. anesthesia disagreed and made the recommendation to md to advance the impella catheter 1-2cm to see if flows would improve. md advanced impella catheter 2cm and pt's hemodynamics began to deteriorate rapidly. impella flows became significantly worse, and motor current became less pulsatile with complete separation in placement signal waveform and lv signal waveform. recommendation was made to anesthesia to reassess the position and overall contractility of the heart following advancement of catheter. it was noted on tee that the patient had an acute pericardial effusion. bleeding in the chest noted. an incomplete perforation of the left ventricle (lv) was noted. the impella was pulled back 3.5cm, and the perforation was repaired. surgeon noted that the patient's tissue was extremely friable in area of perforation due to an lad infarct. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of low pump flow, positioning issues, and ventricle perforation has been completed. the impella device was not returned for evaluation. low pump flow: the pump was not returned. data log review showed suction events and low flow occurring prior to noted device repositioning that did not occur for the remainder of the case. the cause of the low pump flow was most likely improper positioning of the device due to the low flow and suction events occurring prior to repositioning of the device when additional volume give was noted to not resolve the issue and no low flows occurred for the remained of the case. positioning issues: the cause of the positioning issue was most likely use related due to the pump being inserted too deep into the left ventricle (lv) after initial implant. ventricle perforation: after pump was pushed deeper then 4.86cm by an additional 1-2cm, it was noted on tee that the patient had an acute pericardial effusion. bleeding in the chest noted. an incomplete perforation of the lv was noted. patient had existing friable tissue condition. the root cause of the ventricle perforation issue was most likely patient condition related as patient had existing friable tissue and impella positioning deep led to vessel perforation later per clinical review.
8882 | 1220648-2025-28961 | 22136025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22136025&pc=OZD | 05/10/2025 | 06/03/2025 |  | 06/03/2025 |  |  | 05/10/2025 | 06/03/2025 |  |  |  |  | 06/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582188 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, the patient had oozing with a hematoma. dressing was changed and heparin was withheld. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8883 | 1220648-2025-28877 | 22122068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122068&pc=OZD | 05/10/2025 | 07/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642188 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | Positioning Problem |  |  |  |  | Atrial Fibrillation | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient was sedated in the catheterization lab for the duration of the procedure. as the sedation wore off, the access site started to bleed. it was noted that the patient would tighten his abdomen and move his impella leg along with hypertension which caused the site to bleed. when asked, the patient stated that he has severe restless leg syndrome. manual pressure was held, femostop was placed, and heparin was held overnight to manage the condition. the following day, the patient's urine output was red tinged and their labs came back hemolyzed. plans were made to explant the pump in response to the condition. however, the patient went into atrial fibrillation the following morning. amiodarone bolus was given. the pump was repositioned and the patient's condition was monitored. explant was subsequently postponed until the patient could be cardioverted to a normal sinus rhythm. as support continued, the impella alarmed position in aorta and almost simultaneously a high purge pressure alarm triggered. the registered nurse denied any kinks in the tubing. the purge cassette was replaced but the issue persisted. the impella was advanced two centimeters. therefore, motor current improved immediately along with flows. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding, hemolysis, vf/vt/af/at, positioning issues, and high purge pressure has been completed. the impella device was not returned for evaluation. positioning issues: failure mode optical signal issue was replaced with positioning issues because there was no reported issue with the optical sensor waveform and the data log optical signal metrics were stable throughout support with no psnr alarms. the cause of the pump position in aorta was patient movement since the positioning issue occurred when the patient rolled in bed. access site bleeding: the cause of the bleeding was most likely patient movement since bleeding occurred when patient would tighten his abdomen and move this impella leg in conjunction with hypertension. hemolysis: the cause of the hemolysis was unable to be determined since product was not returned and due to insufficient clinical details. high purge pressure (hpp): data logs were reviewed and immediately following a impella position in aorta alarm, a sharp rise in purge pressure (3 min) was observed with resulting purge pressure high alarm. pump flow because saturated 3.5 l/min at 39,000 rpm (exceeds the cp threshold of 2.9 l/min at 39,000 rpm on p6) despite no change in p-level and motor current increased with a loss of waveform pulsatility. no motor current spike indicative of an ingestion event. 2 hours later, after p-level was momentarily reduced, normal flows of 3.0 l/min at p7 were achieved and high purge pressure resolved. motor current pulsatility resumed. the cause of the high purge pressure was most likely biomaterial since the saturated flow persisted after repositioning resolved the hpp. saturated flow: fm saturated flow was added per data log review findings. per data low review, pump flows was normal at p6 (2.5 l/min) but upon the spike in pp, p2 flows become saturated (2.7 l/min) but then saturated flow briefly resolves at p7 (3.0 l/min) but then comes back again while still on p7 (3.5 l/min). that indicates loose material since the resistance on the motor doesn't permanently resolve upon repositioning. the cause of the saturated flow was most likely biomaterial since purge, motor current (mc), and flow were affected simultaneously. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8884 | 1220648-2025-28880 | 22122146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122146&pc=OZD | 05/10/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503614 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on day two of impella cp support and developed hematoma in the access site. one unit of packed red blood cells was administered. it was further reported that the patient expired while on support. there was no other information provided. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8885 | 1220648-2025-28882 | 22122198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122198&pc=OZD | 05/04/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/04/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654239 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an emergency chest exploration was required during impella 5.5 support to control a bleed. the pump was repositioned following the intervention and support continued with the implanted device. during support the patient well supported on impella 5.5 on p9 with 5.0 lpm noted. patient had 6 beats of ventricular tachycardia overnight. the family withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8886 | 1220648-2025-28883 | 22122262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122262&pc=OZD | 05/10/2025 | 07/17/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609672 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed some bleeding from their groin following impella cp implant. it was noted that a vein was accidently punctured during insertion. two units of blood products were given and a vessel repair was performed. it was then documented that the pump had no flow at various levels. the position was verified and large amounts of volume were adminstered. however, the impella still showed no flow. the pump was subsequently replaced in response to the noted issues. | [5 empty] | E.1 name prefix/title, first name and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Complaint device not returned for analyzing. data log reviewed: shows low flows from start of case. suction alarms near end of case. no mc spike observed. no positioning alarms. clinical: a vein was accidentally punctured during insertion, resulting in bleeding. over 10 liters of volume were administered and yet the impella still has suction alarms and no flow at various levels. the position was checked which did not resolve issue. no cannula kinks observed. low pump flow: the cause of low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. ¿vascular damage - peripheral vessel: the cause of vascular damage - peripheral vessel issue most likely was use related as a vein was accidentally punctured during insertion resulting in bleeding.
8887 | 1220648-2025-28885 | 22122310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122310&pc=OZD | 05/10/2025 | 06/30/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Difficult to Remove | Deformation Due to Compressive Stress | Positioning Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed for a patient due to myocardial infarction. the impella was successfully placed. it was noted that an echocardiogram showed a clotted effusion over the right ventricle. in addition, there was a notification of ongoing suction alarms. a decision was made to take the patient to the operating room (or) to remove the clot in the pericardium. while in the or, the impella cp got pulled back and was attempted to recross. the impella did recross, however, flows would not come up. a decision was made to place an extracorporeal membrane oxygenation and remove the impella cp. it was noted that the impella cp got caught in the femoral and came apart and vascular was called to remove the impella. it was noted that there was a cannula kink. however, the pump was successfully removed. no patient harm was reported. | [5 empty] | The investigation of removal issues, low pump flow, product damage (cannula kink), positioning issues, and product damage (detached inflow/outflow ¿ peripheral vessel) has been completed. the impella device was not returned for evaluation. positioning issues: the clinical stated that while in the operating room to remove a clotted effusion over the right ventricle (rv), the cp got pulled back and then was attempted to recross. recrossing was successful but there was low pump flows. due to lack of clinical details, the root cause of the positioning issue of the pump getting pulled back, cannot be determined. product damage (cannula kink): a kink was seen on the cannula post explant. the root cause of the product damage (cannula kink) is most likely due to a cannula kink that occurred while repositioning the pump. low pump flow: the data logs show the pump flows drop rapidly while the p-level remains on 9. there are suction and positioning alarms. this shows that the cannula is most likely completely occluded. based on the clinical details and data log analysis the root cause of the low pump flow is most likely due to a cannula kink. removal issue: during removal, the cp got caught on the femoral artery, and the cannula separated. the root cause of the removal issue is most likely due to use as it was caught on the femoral artery and then separated with further attempts to pull. product damage (detached inflow/outflow): during removal, the cp got caught on the femoral artery, and the cannula separated. the root cause of the product damage (detached inflow/outflow) is most likely due to use as it was caught on the femoral artery and then separated with further attempts to pull. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28885. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8888 | 1220648-2025-28887 | 22122367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122367&pc=OZD | 05/09/2025 | 06/10/2025 |  | 06/02/2025 |  |  | 05/09/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585398 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Stroke/CVA | Ischemia | Tachycardia | Thrombocytopenia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient on impella 5.5 support experienced limb ischemia. the nurse reported was unable to doppler a right radial pulse, but right brachial pulse was present. vascular ultrasound (us) performed at bedside showed arterial disease in the right leg and slow flow in both legs. it was also noted lower pressure on right arm compared to left arm, right fingers discolored. both feet look discolored. additionally, the patient was taken to the operating room as impella was not able to be repositioned at bedside. transesophageal echocardiogram (tee) and x-ray in or used to flip 5.5 back away from septal wall. | The user facility reported that an impella 5.5 was placed to allow the patient¿s heart to rest and re-evaluate for coronary artery bypass graft. the patient remained only oriented to self and neurology evaluated and reported possible stroke or shower emboli. also, heparin infusion was turned off due to decreased platelets and the patient was not on systemic anticoagulation. thrombocytopenia panel came back as weak positive so hematologist made the decision for argatroban. on trans-esophageal echocardiogram, large right atrium thrombus was noted and several days later, a possible new clot was seen in the left ventricle that was not present at the beginning of the case. the physician is aware and wanted to continue with the current plan. the patient also experienced four runs of ventricular tachycardia while on impella 5.5 support. |  |  |  |  | D6b explant date added b1 product problem was inadvertently omitted on the initial manufacturer device report number. b5 additional information from the clinical site was inadvertently omitted on the initial manufacturer device report number. h6 health effect - clinical code 1942, health effect - impact code 4628, medical device problem code 3009, and component code 3038 were inadvertently omitted on the initial manufacturer device report number. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8889 | 1220648-2025-28888 | 22122421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122421&pc=OZD | 05/08/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626292 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited bleeding from the radiofrequency ablation site and the patient's hemoglobin dropped from 12 to 9. one unit of packed red blood cells were transfused. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8890 | 1220648-2025-28890 | 22122508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122508&pc=OZD | 05/08/2025 | 09/24/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631901 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | Stroke/CVA | Thrombosis/Thrombus | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implantation a thrombus was discovered in the left ventricle. the pump was monitored for evidence of clot ingestion and showed lower than expected flows and de-coupled lv and ao waveforms. additionally, the patient developed a gastrointestinal bleed. one unit of packed red blood cells was transfused and an esophagogastroduodenoscopy was performed. it was further reported that a head ct showed a possible infarct but the angiography was negative for thrombus. the patient was successfully weaned from the pump and survived. after explant, a thrombus was noted in the pump outflow, and an mri showed scattered embolic throughout brainstem, cerebellum, and cerebrum. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. vascular damage - peripheral vessel: after 4 days on support, gi bleed occurred. 1 unit prbcs given, pt remained supratherauetic on heparin, dose adjusted, no further bleeding episodes. the pump was not returned. the cause of the vascular damage - peripheral vessel was not determined since product was not returned and insufficient clinical details. low pump flow: after implant, concern for lower than expected flows and de-coupled lv and ao waveforms. team concerned with clotting. the pump was not returned. data log review showed first half of the case had low snr/contrast along with ps that was very noisy before improving around the fifth day of support. flows at p-9 (expected to be 5.0-5.5l/min) trend down to 4.9l/min on (b)(6) 2025 just before 0600. the rt logs at this time were not returned and clinical details do not account for why these low flows occurred. the cause of the low pump flow was not determined since product was not returned, inconclusive data log review, and insufficient clinical details provided. optical signal issue: after implant, concern for lower-than-expected flows and de-coupled lv and ao waveforms. the pump was not explanted at time of noted ao waveform de-coupling. the pump was not returned. data log review showed first half of the case had low snr/contrast along with ps that was very noisy before improving around the fifth day of support. the cause of the optical signal issue was not determined since product was not returned, inconclusive data log review, and insufficient clinical details provided. thrombus: during support at an unknown time, cta came back negative for thrombus. however, upon pump explant, bedside rn noted thrombus on pump outflow cage. mri of the head was done which showed scattered embolic throughout brainstem, cerebellum, and cerebrum. the pump was not returned. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. stroke/cva: during support on 5.5 while pt sedated and weaning off ventilation, concern for cva arose. ct on day 8 was negative for cva. on explant on day 9, mri of the head was done which showed scattered embolic throughout brainstem, cerebellum, and cerebrum. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates.
8891 | 1220648-2025-28891 | 22122576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122576&pc=OZD | 05/08/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639554 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Shock | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that upon removal of the peel-away sheath, the patient began to bleed from the access site of the impella cp. manual compression and femostop were utilized to manage the condition. four units of blood were also given. however, the bleed persisted and the patient went into hemorrhagic shock. the patient passed away the following day. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8892 | 1220648-2025-28893 | 22122653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122653&pc=OZD | 05/08/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617730 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that prior to the impella 5.5 being inserted, the patient went into continuous ventricular tachycardia (vt). after the impella 5.5 was placed this episode continued. multiple shocks, amiodarone boluses/drips, lidocaine drip and electrolytes replacements were given. this lasted for approximately 30-45 minutes. paddles were used which were ineffective. new pads were placed in which a nice shock was delivered converting the patient to sinus rhythm. the patient continued with vt during impella 5.5 support. it was further reported that the patient had extensive hemorrhagic stroke with midline shift and expired. there is not enough information to exclude the impella used as an associated factor to the event outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
8893 | 1220648-2025-28894 | 22122737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122737&pc=OZD | 05/08/2025 | 06/23/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592846 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following setup, a low purge pressure alarm appeared on the console. there were no observable leaks or loose connections at the time of the failure. the purge cassette was replaced, but the staff was unable to reinitiate the priming process. the power on the console was cycled, but once again, the console went directly to the main placement screen with no option to go to the case start steps and priming. a new pump and cassette were subsequently opened and successfully utilized for implant/support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | Purge leak: due to lack of product returned, the root cause of the purge leak - pump cannot be determined. priming issues: the clinical stated that the impella set up completed without any issues priming. device fully primed and fluid seen exiting catheter. low purge pressure alarms occurred and due to lack of resolution from troubleshooting the issue, the pump was unplugged and the console was shut down. after the console was shut down and the impella was re-connected, the impella wouldnt go through the case start steps. the real time (rt) logs show that the pump was originally primed and the imc logs show the pump fully went through case start. when it was connected, the impella jumped right to the "start impella screen." due to the pump already being primed, it will not go through case start again, so the there was no problem found regarding the priming issue.
8894 | 1220648-2025-28895 | 22122968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22122968&pc=OZD | 05/07/2025 | 07/02/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644917 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the implant was subsequently aborted. additional information from a study was provided. positive bleeding on the right groin percutaneous coronary intervention and mechanical circulatory site was noted. an episode of vomiting occurred. 150 milliliters of frank red blood was noted. pressure was held on the groin for 40 minutes. the patient was started on levophed and dopamine was discontinued. packed red blood cells were transfused. the patient recovered and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and access site bleeding has been completed. the impella device was not returned for evaluation. delivery issue: cp pump was unable to advance through the patient's vasculature. the root cause of the delivery issue was most likely patient condition related, pump was unable to advance through the patient's vasculature during attempted delivery. access site bleeding: bleeding right groin. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided.
8895 | 1220648-2025-28876 | 22123274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22123274&pc=OZD | 05/03/2025 | 06/27/2025 |  | 06/02/2025 |  |  | 05/03/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645788 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Puncture/Hole | Device Sensing Problem | Device-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was placed for the support of a patient admitted in with acute myocardial infarction. the pump was placed via single access but, the team punctured a hole in the pump. the placement signal (ps) and left ventricle waveform were lost and the pump was removed and replaced. it was reported that while on the replacement pump, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the patient's outcome was reviewed by abiomed's medical safety officer. the impella cp pump 1 malfunctioned during use and was able to be exchanged to successfully complete an impella cp placement for mechanical circulatory support without complications. while on impella cp pump 2, a left ventricle perforation occurred and there is not enough information to exclude the impella cp pump 2 as an associated factor to the patient's demise outcome (which is captured under complaint file (b)(4)). | The investigation of optical signal issue and product damage (hole in catheter) has been completed. the pump was returned for investigation. no evidence of physical damage was found in the returned product. all the 5s optical parameters were within the specification range. the pump passed the laser continuity test for the optical line. the pump was further tested in a bucket of water at different p-levels, and no issues were noted with the placement signal. the catheter was thoroughly checked, and no holes or damage were observed. also, the red handle was opened and verified to had no blood in it. there was no product damage (hole in catheter) issue found on the returned product. total of 51 minutes of case run time was recorded on the data log. the data log shows one instance of a motor current spike without significantly affecting the 5s optical signal parameters. no "placement signal not reliable" alarm was reported. the cause of the optical signal issue could not be determined, as the issue could not be reproduced on the returned product, and the clinical information was insufficient to derive the root cause. d9 device returned to manufacturer date added
8896 | 1220648-2025-28879 | 22123406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22123406&pc=OZD | 05/10/2025 | 09/17/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654221 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with cardiogenic shock was placed on impella 5.5. during support, it was noted that heparin was being stopped due to hemolysis and hematuria. the swan was removed as well and support was continued. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned for investigation. data logs review confirmed low pulsatility, and the concentration of alarms aligned with the patient updates that reported low pulsatility. based on this information, the root cause of the optical signal issue was determined to most likely be due to patient condition. since limited clinical details were provided regarding what troubleshooting method resolved the issue, product wasn't returned, and data logs did not show any relevant patterns pertaining to hemolysis, the root cause of the hemolysis could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28879.
8897 | 1220648-2025-28878 | 22123430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22123430&pc=OZD | 05/10/2025 | 07/07/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675590 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced a short episode of ventricular fibrillation during impella 5.5 support. the pump was repositioned and support was continued uninterrupted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease, timing of the arrythmia event in relation to impella support (during or post-implant), electrocardiogram (ekg) recordings of the arrhythmia episodes, evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances, assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements, presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities, response to any antiarrhythmic treatments or interventions during the event, any documented device-related issues or complications during impella support, evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand, and review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. the device was returned for evaluation, but no issues were found with the device. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided, the root cause of the vt/vf was not determined.
8898 | 1220648-2025-28881 | 22123646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22123646&pc=OZD | 05/13/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/13/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625332 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp developed retroperitoneal bleed. anticoagulation was withheld. the procedure was successful. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8899 | 1220648-2025-28884 | 22123813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22123813&pc=PYX | 05/08/2025 | 07/03/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025540863 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a rp flex being placed for a patient with heart failure. the rp flex was unable to be delivered and was removed, pump and sheath and all as 1 system. the team made decision not to try again to deliver the rp flex. the patient outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | Additional information provided in d9, h3, h6 and h10. the investigation into the reported issue has been completed. impella rp flex, introducer, and guidewire (gw) were returned for inspection. the device was returned stuck with the flex, gw, and patient line all in the introducer. upon further inspection and removing the devices from the introducer, the patient line was found stuck under the guidewire and around the inflow cage. the root cause of the delivery issue was an interaction with another device during subclavian access as stated in the clinical the issue coincided with another venous line and was confirmed through returned product investigation. corrected information provided in d4.
8900 | 1220648-2025-28886 | 22123868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22123868&pc=OZD | 05/08/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/08/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632881 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp was moved by nursing staff and subsequently, the pump's lv waveform flipped. the patient's family decided to withdraw care and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8901 | 1220648-2025-28892 | 22124236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22124236&pc=OZD | 05/09/2025 | 06/13/2025 |  | 06/02/2025 |  |  | 05/09/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122977 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | 0the complainant reported that the impella automatic impella controller (aic) was unable to detect the purge disc after the cassette was changed. a new aic was used to provide support. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-28892. h6 type of investigation, investigation findings, and investigation conclusions were erroneously entered on the initial manufacturer device report number 1220648-2025-28892. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28892. | A4 and a5, ethnicity, is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8902 | 1220648-2025-28896 | 22124335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22124335&pc=OZD | 04/22/2025 | 06/18/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631838 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hematoma | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support a patient with known cardiac history. the patient was in cardiogenic shock/hypovolemic shock. the pump was placed and there was an observation made of swelling at the site. we received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured left axillary impella site hematoma with a "possible" relationship to the impella device used: ¿patient had worsening pain over impella site. l axillary hematoma was monitored but worsened. chest cta obtained. patient had debridement, irrigation, evacuation of l axillary hematoma. jp drain inserted. removed (b)(6) 2025. deep wound culture was positive for moderate wbcs, light staph epidermis. fungus culture." the patient did not recover/resolve. additionally a month later, the team observed a purge sidearm leak which caused no harm. no changes in the purge flow or purge pressure. the team continued to monitor. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of access site bleeding, and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the cause of the purge leak - pump was most likely a leak from the sidearm due to clinical details stating as such; however, the exact location and cause of the leak was not determined due to no product returned. d6b explant date added.
8903 | 1220648-2025-28889 | 22124482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22124482&pc=OZD | 05/09/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/09/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656897 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was taken to the operating room for escalation of therapy from the impella cp to an impella 5.5. the axillary was obtained with 10 millimeter graft followed by left ventricle access with wire and catheter. the impella 5.5 was advanced and the impella cp was withdrawn into the descending aorta. however, the patient went into sustained ventricular fibrillation. repeated defibrillation, and chemical support was provided. transesophageal echocardiogram showed worsening right ventricle dilation with worsening ph and lactic acid. the physician elected to support the patient with peripheral venoarterial extracorporeal membrane oxygenation (ecmo). antegrade access in the right superficial femoral artery was obtained for distal flow to ecmo. once in the cardiac catheterization lab and after time on full support, the patient was once again defibrillated and broke to normal sinus. additionally, it was noted of an impella in position aorta alarm. the device position was confirmed in the ascending aorta on echocardiogram. attempts to re-advance were unsuccessful. the patient remained hemodynamically stable on ecmo. there was no change in pressor/inotrope requirements. the decision to explant at bedside was made and plans to wean the ecmo flows and for decannulation. ejection fraction was 40% at time of explant and right ventricle function was restored. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8904 | 1220648-2025-28899 | 22124987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22124987&pc=OZD | 05/10/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597610 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp and during support, bleeding noted around the placement sheath, staff decided to not remove the peel away and patient continued to bleed at bedside. blood products were given to the patient. additionally, computed tomography scan showed large amounts of swelling in the patient's brain and brain bleeding. the patient was taken off of heparin at that time. support continued until care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8905 | 1220648-2025-28900 | 22125141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125141&pc=OZD | 05/07/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582402 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient suffered ventricular fibrillation while implanted with an impella 5.5. cardiopulmonary resuscitation was attempted but the patient ultimately passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8906 | 1220648-2025-28901 | 22125291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125291&pc=OZD | 05/10/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643378 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp prep, complications were encountered with a utilized prostyle resulting in manual pressure being held and two units of blood being given. following implant, a fluoro revealed no distal flow past the impella sheath. a fem-fem bypass was performed to allow reperfusion of the extremity. once the positioning sheath was placed, oozing continued at the site. manual pressure, femostop, and a hemostatic patch were utilized to achieve hemostasis. two days into support, heparin was stopped due to concerns for low platelets and thrombocytopenia. support continued for one more day until the decision was made by the patient's family to withdraw care and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8907 | 1220648-2025-28898 | 22125311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125311&pc=OZD | 05/07/2025 | 06/25/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639791 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while an impella cp was implanted, the pump flows dropped from 3.4 to 0.6. the pump position was checked. the pump was removed, inspected, and re-implanted with no change to the low flows. the impella cp was then explanted. the doctor was offered to place a second pump to continue with support for the case but declined as the pressures were good and one drip had already been weaned off. the patient survived. | [5 empty] | The investigation of low pump flow has been completed. returned complaint cp pump visually inspected. broken blades found in impeller. no cannula kink or biomaterial observed. low pump flow: the cause of the low pump flow issue was damaged impeller due to interaction with another device since the flow dropped after fixing the lad with stent placement. the characteristic of the impeller damage is consistent with an interaction with another device potentially a stent. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8908 | 1220648-2025-28903 | 22125546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125546&pc=OZD | 05/07/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629950 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Renal Failure | [10 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and developed a hematoma at the access site. heparin was withheld and support was continued. the patient had atrial fibrillation that resulted in significant deterioration/worsening acute kidney injury. attempts to cardiovert were unsuccessful. the impella 5.5 was removed and a washout of the insertion site was performed. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8909 | 1220648-2025-28902 | 22125583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125583&pc=OZD | 05/05/2025 | 07/01/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646613 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was implanted with an impella cp and experienced arrhythmia and bright red urine. the impella's position was optimized and flows were reduced to resolve these symptoms. while repositioning the pump resistance was felt, and the pump was pulled back across valve when the resistance released. the next day the urine was clearing. the impella cp was replaced a day later with an impella 5.5. the patient survived. | [5 empty] | The investigation of positioning issues, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: according to clinical notes, hemolysis was reported on the evening of the pump implant day and was seen to resolve later the following day. the data log shows that a few suction alarms were reported at the start of the case on the morning of (b)(6). no abnormalities were found related to motor current spikes or positioning issues during the reported hemolysis event and the days later when it was resolved. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. positioning issue: the root cause of the positioning issue was most likely use related since doctor pulled back the pump during reposition of the device. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28902. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | H6 2558 was erroneously entered on the initial manufacturer device report number 1220648-2025-28902.
8910 | 1220648-2025-28904 | 22125735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125735&pc=OZD | 05/10/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661910 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced ventricular tachycardia during impella cp implant. the patient was defibrillated to treat the condition. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
8911 | 1220648-2025-28905 | 22125798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125798&pc=OZD | 05/09/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636849 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and endured a short episode of ventricular tachycardia. one round of cardiopulmonary resusitation was performed coinciding with two shocks to achieve a return of spontaneous circulation. after six days of support, the patient was successfully weaned from the impella. upon explant, it was noted that the patient had lost of pulses in their leg. a doppler was performed and showed absent distal flow. vascular surgery was consulted and the patient was sent to the operating room for clot removal. the patient survived. | [5 empty] | The impella device was not returned, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia, limb ischemia, and thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8912 | 1220648-2025-28906 | 22125957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125957&pc=OZD | 05/07/2025 | 06/23/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654483 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced low pump flow and alveolar bleeding. the endotracheal tube was clotted off and a bronchoscopy was performed. heparin was withheld. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of vascular damage - peripheral vessel and low pump flow has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: the root cause of the vascular damage (peripheral vessel) issue was not determined since product was not returned and insufficient clinical information was provided. the relation to impella is unknown. low pump flow: data logs were investigated. there were suction events at the pump start and the pump was not able to generate sufficient flows on p8 to p4 levels for the initial hour. no motor current spikes or positioning issue observed. placement signal was pulsatile. later suction alarms resolved and impella was running at the lower flow range on p8. venoarterial extracorporeal membrane oxygenation (va ecmo) was also running parallelly with impella. therefore, based on the data log review and clinical information provided, the root cause of the low pump flow issue was most likely patient condition related since the patient was on va ecmo along with impella and had several valve issues. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8913 | 1220648-2025-28907 | 22125976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22125976&pc=OZD | 05/10/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/10/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658518 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support, and during support the patient had ventricular tachycardia. the patient was defibrillated. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
8914 | 1220648-2025-28908 | 22126153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22126153&pc=OZD | 05/07/2025 | 06/19/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636182 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision was made to replace an implanted impella 5.5 due to loss of placement signal and the patient was prepped in surgery. however, the pump was unable to pass through the patient's vasculatrure. the existing pump was left in place. the patient was brought back to the intensive care unit without complications. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to patient anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8915 | 1220648-2025-28909 | 22126273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22126273&pc=OZD | 05/11/2025 | 07/14/2025 |  | 06/02/2025 |  |  | 05/11/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1429529 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during prep the automated impella controller (aic) failed to purge in a emergency situation. a second aic was used. there was no harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. data analysis: on the complaint date, the logs show that the console had piston blocked as the console is constantly trying to complete purging. this behavior also reproduced on 5/9 and 5/14 but the console is able to purge by 6/12. the real time (rt) logs from the complaint date confirms that the purge ticks only went to the 20-30 range which shows that the console was not able to have the full range of piston. device analysis: the console was returned for investigation and the symptoms from the complaint were not reproduced as the purge motor was able to full rotate. however, the motor shaft had signs of glucose ingress. root cause: the cause of the piston block was most likely a glucose ingress that was worn away by the console's use on 6/12 which would then continue to perform as expected once the console was returned for investigation. d9 device returned to manufacturer date added.
8916 | 1220648-2025-28910 | 22126353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22126353&pc=OZD | 05/07/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/07/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656864 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and had ventricular tachycardia (vt) that required cardioverting. the patient expired days later. there is not enough evidence to exclude the impella cp device with its related event as an associated factor. the impella cp related event may have added to an already complex situation and is more of an impact to the demise outcome, which is captured under impella cp complaint (b)(4). | [5 empty] | Abiomed's medical safety officer review: there is not enough evidence to exclude the impella cp device with its related event as an associated factor. the impella cp related event may have added to an already complex situation and is more of an impact to the demise outcome, which is captured under impella cp complaint (b)(4). the impella device was not returned, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular fibrillation was not determined.
8917 | 1220648-2025-28911 | 22126473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22126473&pc=OZD | 05/12/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/12/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849905 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, there was a pump stop. attempts to manually restart the pump were unsuccessful. a new automated impella controller (aic) was attempted to no avail. the pump was removed. this complaint was created to document the change in consoles. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8918 | 1220648-2025-28912 | 22126574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22126574&pc=OZD | 05/13/2025 | 12/09/2025 |  | 06/02/2025 |  |  | 05/13/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534505 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female | 100 KG |  |  | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump stopped unexpectedly during patient support. the pump was restarted on p-4, but the mean, max, and minimum flows were all at 3.4 l/min. the pump was subsequently explanted as a result of the noted failures. later, care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | Data logs were reviewed and the temporary pump stop was confirmed on the reported date. pump flows were noted to become initially saturated on (b)(6) 2025 and continued to be for the rest of the case. there were no abnormal mc spikes until (b)(6) 2025. leading up to the spikes, there was distinct variability in the purge pressure. there were a total of 18 high motor current pump stops over the course of 19 minutes before the pump was restarted. it ran for 2.5 more hours before the user disconnected the pump. during this time, flows remained saturated. returned product was investigated and it had an enlarged purge gap. the pump was run in a bucket and pump flows were immediately saturated for the entire run. there were several motor current spikes throughout the run, but only after 2 hours did a mc spike trigger a high mc pump stop. the pump was not able to be restarted. after the pump stop, purge pressure quickly rose over 1000 mmhg. the pump was torn down and confirmed wear on the inner race of the ball bearing. temporary pump stop: per data log review, the initial pump stop occurred on the 6th day of support. returned product inspection found bearing wear, confirming that the cause of the pump stop was due to the wear. this is prior to the 8 day design life of the product, per design trace matrix 0046-9910. based on data log review and inspection of returned product, the cause of the temporary pump stop was determined to be bearing wear. saturated flows: per data log review, pump flows became saturated on the second day of support and continued to be through the rest of the case. there were no abnormal motor current spikes prior to the flow saturation and purge pressure remained generally stable until the end of the case. on the last day of support, there were several high motor current pump stops with variability in the purge pressure. returned product confirmed that the ball bearing was worn. based on the review of data logs and confirmation of bearing wear on returned product, the cause of the saturated flows was determined to most likely be bearing wear. this is prior to the 8 day design life of the product, per design trace matrix 0046-9910. d10 concomitant medical products and therapy dates updated. | A5 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8919 | 1220648-2025-28913 | 22126706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22126706&pc=OZD | 05/12/2025 | 06/18/2025 |  | 06/02/2025 |  |  | 05/12/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133366 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preventative maintenance, the automated impella controller displayed a "controller error" on display on reboot. troubleshooting steps were noted as reseated all cables and performed two restarts. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data review and device analysis: console logs from the reported day of event are consistent with complaint and show controller error alarm #1039 (defective optical sensor lamp) presenting upon boot-up, preceded by ¿5s¿ parameters indicating no light captured by ccd detector in optical bench: console troubleshooting indicated that despite operational lamp assembly and light visible at the optical pump connector, analog output of ccd detector was flat, indicating no light received. further inspection of the optical bench revealed broken fiber at the entry of the optical cavity of optical bench, which houses ccd detector. since the issue did not present prior to preventative maintenance (pm), it is most likely that it was due to technician error. there were no controller failure alarms of any type occurring at pm. | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8920 | 1220648-2025-28914 | 22127135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22127135&pc=OZD | 05/04/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/04/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535639 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced some oozing around the access site following removal of the peel-away sheath. the perclose was tightened; however, the issue persisted. the groin was injected with epi/lido and manual pressure was held to manage the bleeding. two days later, however, the patient continued to have oozing at the access site. the repositioning sheath was advanced to the skin and forward tension sutures were placed. mattress sutures were also placed with angle matching and sterile dressings applied. the complication was noted to have resolved and support continued with the implanted pump for approximately three days before being escalated to an impella 5.5 device. the patient would expire approximately 7 days later; care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted the impella cp related event is a serious injury. the patient experienced a bleed that required minimal intervention; sutures were tightened to address the bleeding. there was no known transfusion or adverse impact associated with the demise outcome. the patient continued mcs with an impella 5.5 (after escalation in care) and subsequently expired 7 days later due to respiratory complaint; care was withdrawn. of note, based on information at hand the escalation in care impella 5.5 pump appears to have provided support as intended. there is no information that directly or indirectly indicates otherwise. the impella cp device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
8921 | 1220648-2025-28915 | 22128326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22128326&pc=OZD | 05/26/2025 | 06/02/2025 |  | 06/02/2025 |  |  | 05/26/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a blackout and had to be restarted. the blackout kept happening so the aic was exchanged. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8922 | 1220648-2025-28916 | 22128688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22128688&pc=OZD | 05/18/2025 | 07/15/2025 |  | 06/02/2025 |  |  | 05/18/2025 | 06/02/2025 |  |  |  |  | 06/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025611195 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump was unable to advance through the patient's tortuous vasculature during attempted delivery. the implant was subsequently aborted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's blood vessels were too tortuous and pump could not be passed.
8923 | 1220648-2025-28829 | 22113048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113048&pc=OZD | 05/22/2025 | 07/23/2025 |  | 05/30/2025 |  |  | 05/25/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611264 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Shock | Cardiac Tamponade | Cardiac Perforation | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during insertion of an impella cp smartassist in a patient, a left ventricle (lv) perforation occurred. device placement was checked and found the device to be approximately seven centimeters from the inlet area to the aortic valve. it was confirmed that the pigtail was pressed against the left ventricular wall. the 0.018-inch wire was placed facing the posterior wall and threaded between the healthy chodare tendineae. the lv perforation resulted in cardiac tamponade and shock vital signs. free wall was repair was performed in the operating room. cerebral air embolism occurred afterwards. the pericardium was opened under impella drive, and the pigtail was protruding, and it repeatedly dislocated. there was no insertion of a transesophageal echocardiogram (tee). after the repair surgery, a tte was performed to check the inside of the heart and found that it was filled with air. after checking the light reflex, the pupils were dilated and it was determined that the patient had an air embolism. the patient's outcome is serious. | The patient outcome was later shared, as death. the patient expired from the hypoxemia due to shock caused by cardiac tamponade at the time of left ventricular perforation, or air embolism during free wall repair surgery. the true cause of death is not known. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ | The investigation of ventricle perforation and embolism has been completed. the pump was not returned for investigation. a data log review was not required for this case. clinical details confirm that the pump was placed deep into the ventricle and perforated the lv during delivery. there was no indication of a filter leak during the case run. the heart was found to be filled with air after the left ventricle (lv) repair. the cause of the ventricle issue was most likely use related since pump was inserted deep into lv. the cause of the embolism issue was most likely use related since air was found after lv was repaired after perforation. b2 updated to reflect the patient outcome death. b5 updated to include patient outcome. h1 type of report updated to death. h6 health effect - impact code updated to include death.
8924 | 1220648-2025-28830 | 22113125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113125&pc=OZD | 05/26/2025 | 10/02/2025 |  | 05/30/2025 |  |  | 05/26/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645585 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp implanted. immediately after pump inertion, the pump was started and a pump stop occurred. a thrombus or left ventricular structure adhered near the impeller. no abnormalities such as thrombus were found on the preliminary computed tomography or ultrasound. there was clearly interference from the thrombus or a structure within the left ventricle. tissue analysis will be performed at the hospital. the impeller was removed and a new impella was placed. the procedure was completed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Low pump flow was added based on the investigation. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. pump stop and low pump flow: the data logs were reviewed and right after starting the pump there was a small motor current spike followed by pump flows of almost 0. motor current remains lower and the p-levels are changed, but then there are continuous motor current spikes followed by changes in flow and then a high motor current pump stop. the pump is then removed. the pump was returned and there was a slight amount of thrombus that remained wrapped around the impeller and some in the purge gap. no cannula kinks were seen. when running the pump the purge pressure was slightly high and saturated flows were produced as there was thrombus on the impeller but a pump stop was unable to reproduce. according to the clinical, histology was going to analyze the thrombus so likely more was removed from the pump, which is indicative in the data logs that more was present based on the lack of flows and significant motor current spikes. based on the clinical details, data log and product analysis, the root cause of the pump stop is most likely due to biomaterial ingestion. the root cause of the low pump flow is most likely due to biomaterial. d9 device returned to manufacture date added.
8925 | 1220648-2025-28831 | 22113278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113278&pc=OZD | 05/27/2025 | 08/22/2025 |  | 05/30/2025 |  |  | 05/27/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622035 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was inserted in a patient for a high-risk cardiac procedure. the impella cp started in auto mode, but both the max and min values were 0.0, so the team was thinking that something was wrong. the position was checked using echocardiography and fluoroscopy to make sure it was not getting caught on the chordae tendineae, and it was confirmed that the left ventricular volume was adequate. although the pump was pushed and pulled to change its position, the pump flow rate did not increase. the suction alarm continued to sound, and the motor current consumption value did not increase as usual and was around 200. (the exact p value is unknown, but it seems to have been raised to around p6-8.) taking into consideration the patient's hemodynamics, the impella cp was replaced with a new one and inserted it in the same location. the new cp worked without any problems. no harm was reported. there were also some strange phenomena, such as no left ventricle waveform even at p4 or above, and the position waveform appearing as a left ventricular pressure waveform even though it was still within the aorta. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for low pump flow has been completed. during implant procedure, when pump was turned on, both the max and min values were 0.0 despite echo showing proper positioning. no lv waveform even at p4 or above, and the position waveform appearing as a left ventricular pressure waveform even though it was still within the aorta. troubleshooting efforts yielded no success and the pump was exchanged for a new one which was successful and had no issues. data log review showed no mc spike and ps lower than expected. flows hovering at 0l/min. when at p-9 to start the case, the pump was in suction with mc at 212ma. for the pump to calculate flow rate at p-9, it is expected the mc be no lower than 470ma, which explains why the flows are listed at or near zero as the pump is unable to compensate for a blockage. the pump was returned and inspected. biomaterial was found wrapped around the impellor blades. the cause of the low pump flow was most likely biomaterial due to biomaterial found wrapped around the pump and data log review showing lower than expected mc values at p-9 indicating flow blockage.
8926 | 1220648-2025-28832 | 22113442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113442&pc=PYX | 04/08/2025 | 06/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025611211 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Patient Device Interaction Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | Unspecified Kidney or Urinary Problem | Hemodynamic instability | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that after explant of an impella rp flex, a patient experienced hemolysis coinciding with an acute kidney injury. protek duo placement was performed due to the noted issue. the patient was noted to be stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | Complaint/patient coding was updated/corrected and therefore, section h6 (health effect ¿ clinical code, health effect ¿ impact code, and medical device problem code) has been fully repopulated and should be regarded as the final coding that reflects the reported event. this report is being submitted pursuant to the provisions of 21 cfr, part 803. this report may be based on information which has not been investigated or verified prior to the required reporting date. this report does not reflect a conclusion by abiomed inc, or its employees that the report constitutes an admission that the product, abiomed inc, or its employees caused or contributed to the potential event described in this report. if information is obtained that was not available for the initial report, a follow-up report will be filed as appropriate. | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: no product was returned for analysis. clinical details mention ldh greater than 1400 on explant coinciding with acute kidney infection. the data logs show a motor current spike with speed deviation and placement shift on 4/6, about 20 hours after implant. at the same time there was a slight rise in purge pressure and slight decrease in purge flow. the root cause of the hemolysis was most likely biomaterial as evidenced by the ingestion pattern observed in the data logs occurring before reported hemolysis. based on the data log review the failure mode thrombus was added. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined.
8927 | 1220648-2025-28833 | 22113543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113543&pc=OZD | 02/01/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 02/01/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598179 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient with heart failure was implanted with an impella 5.5. during support the pump stopped for 15 minutes. there was no adverse impact to the patient reported. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
8928 | 1220648-2025-28834 | 22113560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113560&pc=PYX | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  | 02/28/2026 |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478617 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was taken to the cardiac catheterization laboratory for thrombectomy followed by rp flex placement. however, during the thrombectomy, the patient coded. the impella rp flex had been prepped and a 23 french sheath was placed. however, the decision was made to abort the impella rp flex and go with venoarterial extracorporeal membrane oxygenation. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
8929 | 1220648-2025-28835 | 22113683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113683&pc=OZD | 03/25/2025 | 10/27/2025 |  | 05/30/2025 |  |  | 03/25/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599867 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Swelling/ Edema | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the patient had access site swelling and there were suction alarms. the impella was explanted in place for an axillary insertion. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow and pain at insertion site has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
8930 | 1220648-2025-28836 | 22113756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113756&pc=OZD | 05/08/2025 | 07/01/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656784 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and during support, there was a hematoma after the repositioning sheath was inserted. the doctor forcefully inserted the sheath and this is thought to be the cause of the hematoma. manual pressure was attempted to stabilize the hematoma. this attempt was unsuccessful, and the impella cp was removed. after waiting for the doctor to decide how best to approach for achieving hemostasis, they got contralateral access and deployed a balloon at the site accompanied by manual external pressure and securement of their prior placed perclose sutures. due to the delayed contralateral access and ballooning, the patient went into cardiac arrest and was unable to be resuscitated. the patient expired. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related. the doctor forcefully inserted the sheath and this is thought to be the cause of the hematoma. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8931 | 1220648-2025-28837 | 22113895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113895&pc=OZD | 04/12/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 04/12/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637095 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp the patient experienced multiple instances of ventricular tachycardia. correct placement of the impella was confirmed and support continue until the patient was successfully weaned from the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
8932 | 1220648-2025-28838 | 22113905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22113905&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628179 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed hemolysis during impella cp support. continuous renal replacement therapy was started due to an acute kidney injury and the patient was ultimately escalated to impella 5.5 support. the patient's urine output improved following impella 5.5 placement. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8933 | 1220648-2025-28840 | 22114061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114061&pc=OZD | 05/07/2025 | 07/01/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613432 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after impella cp was placed, on fluoroscopy the cannula showed kinked at the aortic arch and required removal and replacement. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The returned product was inspected and a kink in the cannula was observed close to the outlet cage. clinical mentioned tavr in use in the vessel and that when advancing repo-sheath the pump moved forward. the root cause of the product damage (cannula kink) was determined to be a cannula kink most likely occurring due to patient condition as evidenced by clinical detail of kink at the aortic arch and product inspection.
8934 | 1220648-2025-28839 | 22114101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114101&pc=OZD | 05/07/2025 | 09/17/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678568 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the pump was getting suction alarms so the staff decided to wean the pump early. there was no reported patient harm. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. product was not returned. data log review shows adequate placement signal. the placement signal improves once the pump p-level is decreased and there are no issues found during the rest of the case. the cause of the optical signal issue was not determined since the product was not returned and insufficient clinical details were provided. the root cause of the suction was not determined as data log review was inconclusive without sufficient clinical details. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-28839. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28839.
8935 | 1220648-2025-28841 | 22114235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114235&pc=OZD | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653690 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and had a few episodes of pulseless ventricular tachycardia overnight and amiodarone was started. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8936 | 1220648-2025-28842 | 22114312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114312&pc=OZD | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563752 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump support placed for the patient admitted in with known cardiac history and a prior coronary artery bypass graft. the patient was being escalated from an intra-aortic balloon pump and being bridged to the left ventricular assist device. the patient had multiple ventricular tachycardia runs that prompted 20 defibrillation shocks. the pump remains and the team has not replaced product. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8937 | 1220648-2025-28843 | 22114501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114501&pc=OZD | 05/07/2025 | 07/28/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273699 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump support placed for the patient admitted in with known cardiac history and a prior coronary artery bypass graft. it was reported that the aic (console) did not prompt step two during the de-airing to disconnect yellow to yellow luer lock. there was an "air in purge" alarm. the problem was eventually resolved and the procedure was completed with this console. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. data analysis: console logs were consistent with what was reported in the complaint. device analysis: console was tested and the reported issue was unable to be reproduced unless it was induced by disconnecting the luers prior to disconnect luer step of the deairing procedure. conclusion: the root cause of this issue was likely use related. it is likely that the luers were already disconnected before reaching this stage of the deairing procedure. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the impella pump has been reported on mfr report: 1220648-2025-28842.
8938 | 1220648-2025-28844 | 22114769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114769&pc=OZD | 05/08/2025 | 07/08/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1739661 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility had a 5.5 pump support ongoing for approximately 3 weeks for the cardiogenic shock patient when there was an aic (console) alarm. the battery high temperature resolved within a minute. no harm or injury was reported. | [5 empty] | The investigation of automated impella controller (aic) - battery overheating has been completed. data analysis: console logs from the complaint date revealed the ¿battery temperature high¿ alarm (49, due to battery temperature over 60 degrees celsius) that self-resolved in 10 seconds. there was no preceding battery temperature alarm 48, which occurs when battery temperature is between 50 degrees celsius and 60 degrees celsius. battery and power data embedded in logs do not show any temperature rise or elevated battery temperatures, as log data is recorded every 60 seconds. device analysis: the console was returned to field service for further investigation. the reported complaint could not be reproduced. the console¿s batteries and power management system were tested, and no issues were observed. root cause: the root cause of the high battery temperature alarm was due to a momentary communication glitch between the batteries and power battery manager (pbm), where a single sample from battery data (in this case value of battery temperature) was corrupt. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
8939 | 1220648-2025-28845 | 22114822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114822&pc=OZD | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient was taken to have a sepsis workup. the location of the sepsis is unknown and therefore it is unknown if the impella was related to the sepsis. additionally, the patient had arrhythmias. the p level was switched to p 8. support was continued. care was withdrawn two days later and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not returned, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the arrythmia was not determined. evidence has not been provided, the root cause of the sepsis cannot be determined. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8940 | 1220648-2025-28846 | 22114894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22114894&pc=OZD | 04/19/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636069 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiogenic Shock | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The impact study reported that the patient on impella 5.5 support had cardiogenic shock. patient was treated with inotropes/pressors, epicardial av pacing and two transfusions. the outcome was recovered/resolved. the pi relatedness, ¿possible¿ related to the impella procedure and remote(unlikely) to the impella device. additionally, the patient also experienced atrial fibrillation. treated with heparin and amiodarone which converted to normal sinus rhythm with treatment. outcome, recovered/resolved with sequelae. the pi relatedness, ¿possible¿ related to the impella procedure and remote(unlikely) to the impella device. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8941 | 1220648-2025-28847 | 22115004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115004&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621520 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an impella cp was implanted, and percutaneous coronary intervention was completed but a hematoma and bleeding/oozing was noted around the peel-away sheath. it was managed with manual pressure but due to the hematoma, a decision was made to remove the impella. the patient tolerated a slow wean and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8942 | 1220648-2025-28848 | 22115148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115148&pc=OZD | 05/08/2025 | 06/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for the support of a patient admitted with cardiac history and st-segment elevation myocardial infarction. while on support, there was intermittent suction alarms to which albumin was being given. support was continued and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 adverse event was erroneously selected on the initial manufacturer device report number 1220648-2025-28848. b2 should have been left blank on the initial manufacturer device report number 1220648-2025-28848. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28848. h6 codes 4445 and 4640 were was erroneously entered on the initial manufacturer device report number 1220648-2025-28848. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8943 | 1220648-2025-28849 | 22115247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115247&pc=OZD | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656115 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump supporting a patient with acute myocardial infarction and cardiogenic shock. the pump supported for 92 hours when the pump was explanted. there was ischemia at the time of the explant due to thrombus burden. the thrombus required thrombectomy. pulses returned and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8944 | 1220648-2025-28850 | 22115716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115716&pc=OZD | 05/10/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/10/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640397 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following transfer to the intensive care unit, a hematoma was noted at the right femoral impella cp site. pressure was held at the site and femostop was placed to manage the condition. the patient received one unit of red blood cells and multiple albumin to replace the lost volume. the following day, the patient's lower extremities were dusky and mottled. the decision was subsequently made to escalate to impella 5.5 support with veno-venous extracorporeal membrane oxygenation. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8945 | 1220648-2025-28851 | 22115778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115778&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1494885 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us based hospital team was having a training when there were 2 battery failure alarms. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8946 | 1220648-2025-28852 | 22115920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115920&pc=OZD | 05/08/2025 | 08/16/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633273 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a cp pump placed to support the acute myocardial infarction patient. the pump supported until wean/explant after 45 hours of support. a right common femoral artery cut down with vascular report was performed. additionally, when explanted there was biomaterial observed. the physician observed the material came with the pump. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Vascular damage code has been removed. the cutdown and vascular repair was just how the physician chose to close the impella access site. h6 health effect - clinical code 1888 was was erroneously entered on the initial manufacturer device report number 1220648-2025-28852-1. | The investigation of thrombus has been completed. the pump was returned and evaluated. there was blood seen around the impeller that came off after soaking in water. there was also dried blood inside the cannula. no thrombus was observed on the returned pump and no other abnormalities were found. data logs are not applicable. the investigation of vascular damage - peripheral vessel has not been completed. should additional information or investigation results become available a supplemental report will be submitted. instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28852. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-28852. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28852. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28852. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-28852. the device was received at time of report h6 health effect - clinical code 4582, health effect - impact code 2199, and medical device problem code 1528 were erroneously entered on the initial manufacturer device report number 1220648-2025-28852. h10 erroneously stated the device will not be received for investigation on the initial manufacturer device report number 1220648-2025-28852 h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28852.
8947 | 1220648-2025-28853 | 22115969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22115969&pc=PYX | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548245 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hypovolemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella rp flex support would not stop bleeding from all sites. leading to hypovolemia and finally resulting in the passing of the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
8948 | 1220648-2025-28854 | 22116165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22116165&pc=OZD | 05/08/2025 | 06/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563577 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed bright red urine and the impella was repositioned several times. additionally, when changing the disposable absorbent pads under the patient, ectopy occurred, and bleeding was noted at the access site. manual pressure was held for the bleeding and heparin was paused. | [5 empty] | The investigation of hemolysis, access site bleeding, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. hemolysis: data logs were reviewed, and there were no motor current spikes, positioning alarms, or suction alarms. additionally, pump flows were within range for the respective p-level throughout the case. however, it was noted that at the beginning of the case, pump flows were variable while at p-9. additionally, when the pump was running at p-7, there was a clear improvement in pump flows after dropping to p-2. this aligns with the patient update around that time that the pump was noted to be out of position, so they dropped to p-2 to reposition. no product was returned. clinical details reported the pump had to be repositioned several times during the first 24 hours of support while the patient was experiencing hemolysis symptoms, data logs showed pump flows to improve after a repositioning event, and hemolysis symptoms had resolved within 24 hours of said repositioning. for these reasons, the root cause of the hemolysis was determined to most likely be improper positioning. positioning issues: product was not returned for investigation. the cause of the positioning issue was not determined due to lack of clinical information and product not being returned. access site bleeding: based on the clinical details provided, the patient's activated clotting time (act) was super therapeutic, and for that reason, the root cause of the access site bleed was most likely anticoagulation management. the movement of the patient could have been a contributing factor as well, but the bleed is less likely to occur if acts are within therapeutic range. ectopy: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8949 | 1220648-2025-28855 | 22116200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22116200&pc=OZD | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660303 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and after the pump removal, there was a clot in the leg where the impella cp was and penumbra was used to remove significant clot. the patient's leg was cold. after the clot removal, flow was restored and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8950 | 1220648-2025-28856 | 22116419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22116419&pc=OZD | 04/30/2025 | 06/19/2025 |  | 05/30/2025 |  |  | 04/30/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600833 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 123 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient implanted with an impella cp developed hemolysis and ectopy which resolved without intervention on the following day. the patient also exhibited consumptive coagulopathy due to bleeding/drainage from a sternotomy. multiple blood products were given, and a washout was performed. two days later the patient was escalated to an impella 5.5 and the patient expired. abiomed's medical safety officer found there is not enough evidence to exclude the impella cp device as an associated factor in the patient's outcome. | [5 empty] | The investigation of vf/vt/af/at, vascular damage - peripheral vessel, and hemolysis has been completed. the impella device was not returned for evaluation. the cause of the vascular damage - peripheral vessel was not determined since product was not returned and insufficient clinical details. hemolysis: data logs were recovered. data logs show positioning issues that occur on the day hemolysis was noted, with impella position in ventricle and suction alarms. the cause of the hemolysis was most likely improper positioning of the device due to data log review showing pump out of position and clinical details stating pump required repositioning that both occur on day hemolysis was noted which resolved within 24 hours. as the necessary information was not provided, the root cause of the vt/vf was not determined b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28856. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28856. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
8951 | 1220648-2025-28857 | 22116493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22116493&pc=OZD | 05/01/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/01/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492729 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a hematoma at their impella cp access site following implant. pressure was initially held and the site was marked to monitor the condition. the following day, the patient was transferred and suction was noted at support level p5. the hematoma had grown and hardened at the site. four units of blood were given and the patient was subsequently taken to the operating room for impella removal and arterial repair. the patients outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8952 | 1220648-2025-28858 | 22116866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22116866&pc=OZD | 05/04/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/04/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628614 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of an impella cp in a patient in cardiogenic shock from an acute myocardial infarction was unsuccessful. there was a delay in starting support. another impella cp would be used. however, the patient would subsequently expire. at the start of the procedure, the patient received contrast to review vessels. ejection fraction at such time was approximately 45%. the patient went into flash pulmonary edema requiring emergent intubation and post intubation. ejection fraction was less than 20%. the initial impella cp placement was attempted. however, there was an issue delivering the impella catheter over a 0.018 guidewire. it could not be determined if the issue was with the guidewire or the impella catheter. consequently, another impella cp device was used and placed pre percutaneous coronary intervention (pci) due to the severity of grafts and reduced ejection fraction. the pci (drug-eluting stenting) was completed, and the patient's mean arterial pressure remained greater than 65 throughout case. the decision was made to leave the impella cp device in place overnight and the patient was transferred to the unit on mechanical circulatory support. however, approximately five hours after the start of the support, the patient expired. no further details were noted. | [5 empty] | Medical safety reviewed the event and noted limited details surrounding the demise outcome, and there is not enough evidence to exclude the impella cp pump 1 as an associated factor to the patient's outcome. an issue with delivering pump 1 over the guidewire; the device was replaced and patient expired approximately five hours after the start of impella mechanical circulatory support on the replacement device. based on the information at hand, the impella cp pump 2 appears to have provided support as intended. there is no information that directly or indirectly indicates otherwise. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8953 | 1220648-2025-28859 | 22116925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22116925&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636272 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp procedure was aborted after multiple attempts failed to pass through due to the patient having heavy calcification of the aortic arch. the decision was made to switch to an intrac-aortic balloon pump, which was successfully placed. the patients condition was stable with no complications. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8954 | 1220648-2025-28861 | 22117149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117149&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654470 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp placed for a patient admitted in with cardiogenic shock/acute myocardial infarction who had percutaneous coronary intervention (pci) done and during the pci, the patient had ventricular fibrillation (vf) which was shocked out. the patient had recurrent ventricular tachycardia (vt) degeneration into the vf and cardiopulmonary resuscitation and levophed were started. after return of spontaneous circulation, lidocaine was started. the medical team assessed and the patient continued to have vt/vf throughout the day. the patient was brought back to the cardiac catheterization lab for a relook and insertion of the impella for hemodynamic stabilization. the patient was escalated to an impella 5.5 axillary. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8955 | 1220648-2025-28862 | 22117167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117167&pc=OZD | 05/10/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/10/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650203 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an acute myocardial infarction and cardiogenic shock patient admitted in with known history and coding with shocks was being applied. the impella cp pump was placed in the right femoral artery and the patient had ventricular tachycardia treated by shock. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8956 | 1220648-2025-28863 | 22117290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117290&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636846 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 support developed bleeding from the right axillary pocket and was taken back to the operating room to intervene and clean out. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8957 | 1220648-2025-28864 | 22117519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117519&pc=OZD | 05/07/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630104 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp, a patient lost distal pulses in the right lower extremity. contrast showed no blood flow past the impella sheath. the impella was explanted the next day and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8958 | 1220648-2025-28865 | 22117546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117546&pc=OZD | 03/27/2025 | 07/01/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured retroperitoneal hematoma, left groin hematoma, and hemolysis with a "possible" relationship to the impella device used: ¿cta ap on (b)(6) 2025 showed asymmetric fullness in l psoas muscle which suggests developing retroperitoneal hematoma. h/h baseline 13-15. hemoglobin dropped from 14.6 ((b)(6) 2025 at 2200) to 13.6 ((b)(6) 2025 at 0227) following impella placement associated with hypotension. 2 units of prbc given. nadir h/h was 9.1/26.1 ((b)(6) 2025). vascular following. ctm but no surgical intervention. the patient recovered with no sequelae. cta done for the left groin hematoma, ¿cta ap on (b)(6) 2025 showed distal left iliopsoas, adductor and partially imaged anterior compartment musculature expansion and heterogeneity reflect intramuscular hemorrhage. the patient recovered with no sequelae. the patient developed also hemolysis. ¿baseline of haptoglobin, ldh and pfhgb unknown. haptoglobin nadir was <30 on (b)(6) 2025. peak for pfhgb was 22.5 on (b)(6) 2025. peak ldh was 533 on (b)(6) 2025 at 1055. team believed hemoglobin drop contribution from hemolysis d/t impella cp. patient improved safely enough for impella removed (b)(6) 2025.¿ the patient recovered with no sequelae. | [5 empty] | The investigation of vascular damage, access site bleeding, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. vascular damage: from a gi case the clinical states that retroperitoneal hematoma. 2 units of blood was given. pump was not returned. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. access site bleeding: from a gi case the clinical states that there was a left groin hematoma and 2 units of blood was given. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. hemolysis: from a gi case the clinical states that the patient endured hemolysis and was possibly related to impella. pump was not returned. data logs were investigated. there was no issues observed in the data logs. no positioning alarms, no suction alarms, no motor current spikes. therefore, the root cause of the hemolysis issue was not determined since product was not returned, data logs were inconclusive, and insufficient clinical information was provided. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8959 | 1220648-2025-28866 | 22117758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117758&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella cp was placed on a patient that presented with altered level of consciousness, acute respiratory failure requiring intubation, and stress-induced cardiomyopathy with reduced ejection fraction. the impella cp was successfully placed. a distal perfusion angiography performed for confirmation and the groin site had a small hematoma. fem stop was applied. during support, it was noted that fluids and red blood cells were given for suction and low central venous pressure. groin site dressing was taken down and manual pressure was held. the plan was to consult vascular for repair and to stabilize/optimize patient. the impella cp was removed in operating room. a hole, three times the size of impella was noted. the hole was transversely with an interrupted 5-0 prolene suture. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8960 | 1220648-2025-28867 | 22117785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117785&pc=OZD | 05/07/2025 | 06/30/2025 |  | 05/30/2025 |  |  | 05/07/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639732 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp and had been intermittently confused and agitated. the patient pulled the impella back fully into the aorta. the patient remained hemodynamically stable, but the impella cp was unable to advance back into the left ventricle. the impella cp was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on clinical description, the root cause of positioning issue was most likely patient movement as the patient was agitated and pulled pump back fully into the aorta.
8961 | 1220648-2025-28868 | 22117953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22117953&pc=OZD | 05/08/2025 | 06/27/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667205 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 179 KG | Non Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional details mentioned that the nurse had turned the patient and redressed the groin site when it started oozing around the peel-away that was left in from insertion shortly after. no other details known about the timing or treatment of the bleed. | It was reported that a patient implanted with an impella cp support. during redressing of the access site the patient was turned and the impella pulled back completely into the aorta. the impella was explanted, and the patient survived. |  |  |  |  | Positioning issues: data log review confirmed an impella in aorta alarm about 5 hours into the case. the root cause of the positioning issue was most likely patient movement since the pump was pulled into the aorta when the nurse had turned the patient per clinical details. access site bleeding - major: clinical details state that the pump was pulled back completely into the aorta after the nurse had turned the patient and redressed the groin site when it started oozing around the peel-away that was left in from insertion shortly after.¿ no other details known about the timing or treatment of the bleed. no product was returned and logs are not applicable. the root cause of the access site bleeding was not determined since insufficient clinical details were provided. instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28868. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28868. b5 additional details regarding bandage change were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28868. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28868. h6 health effect - clinical code 1888 and health effect - impact code 4614 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28868. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28868. | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8962 | 1220648-2025-28869 | 22118252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22118252&pc=OZD | 05/08/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/08/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609899 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 59 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed limb ischemia. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8963 | 1220648-2025-28870 | 22118910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22118910&pc=OZD | 05/21/2025 | 07/21/2025 |  | 05/30/2025 |  |  | 05/21/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599650 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient had a percutaneous coronary intervention performed under intra-aortic balloon pump support, the iabp was then replaced with impella cp. while on support a black waste fluid was observed from the nasogastric tube, and gastrointestinal bleeding was suspected and heparin was in the purge fluid. it was noted that the relationship to impella was unclear, however the effects of anticoagulant therapy are thought to be one of the causes. additionally, there seemed to be a decrease in platelet count, so platelet therapy (units unknown) was administered and the support level was adjusted (lowered). the patient's condition improved after that. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | Vascular damage peripheral vessel : gastrointestinal bleeding was suspected. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided. thrombocytopenia : there seemed to be a decrease in platelet count, so platelet therapy (units unknown) was administered. the cause of the thrombocytopenia was not determined due to insufficient clinical details. h6 4431 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28870. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8964 | 1220648-2025-28871 | 22118985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22118985&pc=OZD | 05/26/2025 | 08/21/2025 |  | 05/30/2025 |  |  | 05/26/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025603903 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the purge side arm of the implanted pump was found to be damaged during patient support. the pump was subsequently replaced in response to the noted damage. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of purge leak ¿ pump has been completed. impella cp was returned for evaluation. upon visual inspection, complete separation between top of filter and purge reservoir. white discoloration and cracking present. data logs were reviewed and a sudden drop in purge pressure with a spike then drop in purge flow to 0 at time of reported issue. purge metrics remain at 0 for reminder of case. clinical details noted that the purge sidearm disconnect from the chamber on day 3 of support. no additional details provided when purge filter became disconnected. the root cause of the purge leak - pump was due to damage at the joint reservoir to filter based on returned product analysis. d9 device returned to manufacturer date added e4 should have been left blank on the initial manufacturer device report number 1220648-2025-28871 as it is unknown if the initial reporter also sent the report to the food and drug administration.
8965 | 1220648-2025-28873 | 22119070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22119070&pc=OZD | 05/27/2025 | 09/26/2025 |  | 05/30/2025 |  |  | 05/27/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025604892 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, there were placement signal issues. there was no left ventricle signal was displayed even at level p-7. there was a possibility of a left ventricular thrombus. the pump was explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. optical signal issue. the data logs show that the pump was attempting to be started and increasing speed, it took about 15 seconds for the left ventricle (lv) signal to appear. about 13 seconds later there was a dip in motor current, flow and placement signal. the lv signal was present for approximately one minute before it was turned off due to the pump being in suction. the lv signal displays one more time about 35 minutes later as the suction alarms turn off but shortly turn back on and the lv is not displayed. the 5s parameters are stable throughout. the return product had no signs of thrombus, the 5s parameters were stable and no issues running. the root cause of the optical signal issue is most likely due to the pump in a suction condition. thrombus: no thrombus on returned product. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates. h2 type of reportable event. h6 type of investigation.
8966 | 1220648-2025-28874 | 22119152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22119152&pc=OZD | 05/29/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/29/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 -JP |  |  | ABIOMED, INC. |  |  | US |  | 2026616169 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient who underwent coronary artery bypass graft after impella insertion due to acute myocardial infarction. when weaning from ecmo, the device was escalated to impella 5.5. however, it was subsequently discovered that the patient had a cerebral hemorrhage, and impella was removed. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8967 | 1220648-2025-28872 | 22119259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22119259&pc=OZD | 05/26/2025 | 08/12/2025 |  | 05/30/2025 |  |  | 05/27/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026638860 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery due to the patient's anatomy. the catheter kinked and the decision was made to abort the implant. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the root cause of the delivery issue was most likely patient condition related, impella cp pump was unable to advance through the patient's vasculature during attempted delivery due to the patient's anatomy. product damage (catheter kink): the catheter kinked and the decision was made to abort the implant. the root cause of the product damage catheter kink issue was most likely patient condition related; cp pump was unable to advance through the patient's vasculature during attempted delivery due to the patient's anatomy and the catheter kinked during the procedure. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28872. d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2025-28872. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28872.
8968 | 1220648-2025-28875 | 22120282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22120282&pc=OZD | 05/04/2025 | 05/30/2025 |  | 05/30/2025 |  |  | 05/04/2025 | 05/30/2025 |  |  |  |  | 05/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623973 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella 5.5 support and during support, the patient had ventricular tachycardia (vt). the patient was shocked and support continued. the vt occurred again later in support. impella placement was verified. additionally, the patient had a hematoma around the axillary pocket site. one unit of packed red blood cells was given to the patient. furthermore, care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8969 | 1220648-2025-28800 | 22103620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22103620&pc=OZD | 05/08/2025 | 07/16/2025 |  | 05/29/2025 |  |  | 05/08/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617123 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp purge line was leaking. the purge cassette was exchange and the purgeline was still leaking. the leakage was noted to be between the purge filter and the pressure reservoir. purge filter bypass and pump exchange were recommended as an intervention. additionally, the patient was on inotropes due to septic shock. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist: section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump and infection/sepsis has been completed. the purge sidearm had been cut off the pump and was not returned, so the exact location of the leak could not be determined. purge leak - pump: data logs were reviewed, and signs of a leak were noted on the reported date. over the course of 1 minute, purge pressure dropped from 500 mmhg to 100 mmhg and triggered "purge pressure low." the purge cassette was replaced, but purge pressure remained low until the de-air procedure was completed about 1.5 hours later. purge pressure remained stable at ~500 mmhg for the remainder of the case. returned product was investigated and it was confirmed that a sidearm bypass had been done. the purge sidearm had been cut off the pump and was not returned, so the exact location of the leak could not be determined. based on the clinical details provided and data log review, the root cause of the purge leak was determined to most likely be a damaged sidearm. the exact location and cause of the leak could not be determined because insufficient product was returned. infection/sepsis: without further clinical details the root cause of the infection cannot be determined. d9 device returned to manufacturer date added.
8970 | 1220648-2025-28801 | 22103722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22103722&pc=OZD | 05/08/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/08/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636897 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella 5.5 pump was placed to support a patient admitted in with cardiogenic shock. approximately 10 days later, a hematoma developed at the impella site with computed tomography showing an active bleed at the axillary artery. the patient was taken back to the operating room for hematoma evacuation and axillary exploration. one unit of packed red blood cells was given. the pump was successfully weaned and the patients outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8971 | 1220648-2025-28802 | 22103801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22103801&pc=OZD | 05/07/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598929 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, there was no distal flow to the right lower extremity. the extremity was noted to be cool and mottled. the decision was made to have the impella cp removed and escalated to an axillary impella 5.5. surgical repair to right femoral artery was performed with fasciotomies performed to the affected extremity. despite repair, no doppler flow was noted to the right foot at this time. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8972 | 1220648-2025-28803 | 22103884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22103884&pc=OZD | 05/07/2025 | 07/01/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574893 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while an impella 5.5 was implanted for support there were four pump stops. the pump would drop to p0 and the nurse manually restarted the pump to p9. the pump was replaced with another impella 5.5 to resolve the issue. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smart assist). section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. based on the clinical details and the data log analysis, the root cause of the temporary pump stop is due to intermittent open electrical lines.
8973 | 1220648-2025-28804 | 22103948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22103948&pc=OZD | 05/04/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/04/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659502 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella cp was placed on a patient with multi vessel disease. during support, it was noted that upon assessment of the patient's right groin, the site looked intact, but the right leg was cooler than the left and only a femoral pulse was present. the tuohy burst (tb) valve was not locked on arrival, so the impella tb valve was tightened down and echocardiogram was completed to verify that the impella was in the correct place. the impella cp was explanted and the left femoral artery was opened by surgeons with an artery cut down and repair. the groin was sutured closed. an impella 5.5 was placed and it was noted that after upgrade, the left leg regained popliteal and posterior tibial pulses. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8974 | 1220648-2025-28805 | 22104003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104003&pc=OZD | 05/04/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/04/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656793 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted to support a patient with ventricular tachycardia. the pump generated low flow alarms and was explanted. a new impella cp was implanted and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8975 | 1220648-2025-28806 | 22104137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104137&pc=OZD | 05/04/2025 | 06/23/2025 |  | 05/29/2025 |  |  | 05/04/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494786 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was initially inserted with difficulty optimizing position. several attempts were made including taking the device out, reaccessing the ventricle and rewiring the impella insertion. the positioning was still not optimal and flows were reduced. inlet was suspected to still be close to the papillary. the provider pulled the device back and it went above the aortic valve. attempts were made to advance and the cannula bent/kinked on itself. the cannula unkinked without additional intervention but the decision was made to utilize a new device due to concern for the pumps integrity. while explanting the device, damage was done to the peel away sheath and the sheath was disposed of. a new sheath was placed for the second impella placement. upon doing so, a large hematoma was noted to have grown and the patients hemodynamics became very unstable. pressure was held to the groin. the patient was taken to the operating room for vascular cut down. two purse strings were placed to control the bleeding. the patient required eith units of red blood cells. the impella remained in place during this interventio and the patient was taken to the intensive care unit (icu) on support where the groin cdi. bilateral doppler pulses were present in the icu. the patients outcome was stable. | [5 empty] | Returned complaint cp pump visually inspected. slight cannula kink near outlet observed. no other abnormality observed. positioning issues: the cause of reported positioning issues most likely was use related due to improper technique since impella inlet stacked in pap and went above the aortic valve during repositioning. product damage (cannula kink): the cause of product damage (cannula kink) was cannula kink due to improper technique since wireless re-access across the av was attempted after the pump pulled out. access site bleeding - major: the cause of access site bleeding - major cannot be determined since the sheath not returned for analyzing and insufficient clinical data provided. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8976 | 1220648-2025-28807 | 22104245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104245&pc=OZD | 05/03/2025 | 06/26/2025 |  | 05/29/2025 |  |  | 05/04/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636383 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient had an impella cp placed due to myocardial infarction. during support, the first cp was punctured, it was successfully weaned and removed through the peel away sheath. pigtail re-inserted into the left ventricle, and the second cp was primed and backloaded over the 0.018 wire and placed across hemostatic valve into proper position under fluoroscopy without issue. the patient was moved over to the icu bed and shortly after became nauseous and diaphoretic. patient voiced increasing nausea and shortness of breath. respiratory took bipap off and placed on non rebreather. o2 saturations decreased and patient tried to sit up in bed having increased shortness of breath. patient then became severely hypotensive and code was called. suction alarms present on aic (console) and continued without resolution to p1. patient was intubated and taken for a repeat angiogram and right heart cath. echo was performed in cath lab and revealed large pericardial effusion. the physician then performed a pericardiocentesis while cpr was being performed. suction alarms were still persistent and p level was maintained on p1. patient expired in cath lab while on impella support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude this impella as an associated factor in the patient's outcome. the left ventricular perforation is a severe injury, considered to be a life-threatening condition, and therefore, the patient demise outcome is more likely associated with this event involving the impella cp pump. *this reportable event is for the same patient in manufacturer mdr report #1220648-2025-28876. however, this manufacturer report captures reportable events as it relates to the replacement impella cp, serial number (b)(6). | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. ventricle perforation: data log review shows the placement signal begin trending downwards after the pump was turned down from p-8. there were also impella in ventricle and aorta alarms near the end of the case while the pump was at p-1. it is unclear when the alarms occurred in relation to the clinical timeline. no product was returned. the root cause of the ventricle perforation was most likely patient movement since the clinical details state that the patient's condition began deteriorating after transfer to icu and further worsened after the patient sat up. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28807. h6 health effect - clinical code 2135 was erroneously entered on the initial manufacturer device report number 1220648-2025-28807.
8977 | 1220648-2025-28808 | 22104286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104286&pc=OZD | 05/04/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/04/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622516 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp, a patient developed an abdominal hematoma. the pump was explanted and a femoral cutdown was performed to surgically close the femoral artery. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8978 | 1220648-2025-28809 | 22104388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104388&pc=OZD | 05/06/2025 | 06/02/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1717178 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  | 0700250000-2025-8006 | Yes | Initial submission | Followup | [8 empty] | The complainant reported that a purge disc not detected alarm appeared on the automated impella controller (aic) during patient support. the purge cassette was replaced, but the issue remained. the registered nurse was walked through an aic-to-aic transfer and the issue resolved. there was no patient harm. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. the logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) multiple times. device analysis: a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. es2023-0090 documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the real time (rt) log data associated with pump 557054 on ic3502 indicates that at this point purge line was already primed, which explains why steps 4 and 5 of case start wizard were skipped. most likely because operator expected to see prompts to prime the pump, a purge change wizard was manually launched and pump was unplugged. console was then power-cycled, case start wizard was launched and pump was connected. the cause for skipping purging prompts of case start wizard was the fact that purge line was already primed, and pump was initialized. there was no ac power issue found during the case. the device functioned/operated to specification. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27821.
8979 | 1220648-2025-28810 | 22104425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104425&pc=OZD | 05/06/2025 | 09/25/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654204 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that there were placement signal issues. motor current and flows remained stable. once the patient was in the intensive care unit bed, the placement signal returned to normal. | The complainant reported the patient was on impella 5.5 support and during support, the patient had episodes of ventricular tachycardia. the patient was shocked out of it. support continued until care was withdrawn and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for cardiac arrythmia and optical signal issue. as the necessary information was not provided, the root cause of the vt/vf was not determined. no product was returned. data log review showed an abrupt placement signal (ps) jump the root cause of the optical signal issue was not determined since product was not returned and inconclusive evidence based on log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28810. d10 concomitant medical products and therapy dates updated. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
8980 | 22104523 | 22104523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104523&pc=OZD | 05/05/2025 | 05/06/2025 | 05/12/2025 | 05/29/2025 | 05/12/2025 | 05/29/2025 |  | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0040-US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report |  | Non-Healthcare Professional | [6 empty] | 41 YR |  | Male |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Defective Device | [5 empty] | Anxiety | Low Pulmonary Arterial Wedge Pressure | [36 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Overnight during impella purge bag and cassette change impella would not let rn proceed forward with the cassette change after completing the outlined steps. impella console stated that disk was not locked in place after countless attempts to locking disc in place. two different impella cassettes attempted, both times the cassette stopped at the same step. second rn called in to help evaluate and contact impella rep and aprn [advance practice nurse] brought to bedside. per impella rep after troubleshooting disc more to swap consoles due to us unable to exit out of screen to see impella wave forms, flows, pressures and more. second console was obtained and ok per aprn to swap consoles. during console swap pt had drop in bp [blood pressure] with map's [mean arterial pressures] in the 30s. pt became acutely anxious and felt unwell. after impella driveline was inserted into new impella console pts bp recovered.
8981 | 1220648-2025-28811 | 22104661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104661&pc=OZD | 05/06/2025 | 09/17/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643229 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient developed a hematoma at their impella cp access site following implant. manual pressure was initially utilized to manage the condition. however, the bleed persisted, prompting forward tension sutures and an angle matching dressing to be applied. the procedural outcome was the patient survived surgery. | Additional information indicated that, during patient support, the pump experienced placement signal low alarms began to appear. repositioning was completed, the alarm remained intermittent. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed, and the placement signal low alarms were confirmed. optical signal metrics were all within specification throughout the case, indicating a reliable signal. optical signal was reliable at this time, and both alarms were triggering correctly based on data log review. since there are several different clinical factors that could have been contributing to the alarms/patterns seen in data logs that went unresolved, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-28811. h6 medical device problem code revised to include 4628 device repositioning 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28811. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. since the clinical reported that the patient's activated clotting time (act( was 297 at the time the repositioning sheath was being advanced, the root cause of the access site bleed was determined to most likely be anticoagulation management with high acts.
8982 | 1220648-2025-28812 | 22104903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104903&pc=OZD | 05/06/2025 | 06/30/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610333 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during an emergent percutaneous coronary intervention an impella cp was attempted to be placed, but unable to cross highly calcified atrioventricular probably due to aortic stenosis. an intra-aortic balloon pump was placed however the patient expired in procedure area. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of delivery issues has been completed. no damages or abnormalities were found on the returned pump. the root cause of the delivery issue was most likely patient condition related since the patient had a highly calcified atrioventricular (av). | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8983 | 1220648-2025-28813 | 22104925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22104925&pc=OZD | 05/07/2025 | 07/01/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653688 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a high purge pressure alarm appeared during impella 5.5 support. the purge cassette was replaced but the issue remained. the catheter was noted to be frayed/kinked. the device was explanted in response to the issue. there was no patient harm. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure and product damage (cannula kink) has been completed. high purge pressure: the returned product was inspected and a substantial kink and twist of the catheter was observed at around the 40cm marking. a window in the catheter was cut and purge line kink was confirmed. there was no biomaterial observed in the purge gap. the pump was run in the lab while the kink was unraveled and straightened and high purge pressure did not reproduce but there were changes in flow rate of the purge further confirming issue with purge line. the data logs from the field show a steady and gradual rise in purge pressure simultaneous with a gradual decrease in purge flow. there was elevated purge pressure for about 11 minutes during which there were several troubleshoot attempts of turning the pump down to p0 and back up to p8. there were also purge system blocked alarms. the root cause of the high purge pressure was determined to be due to a kinked purge line as evidenced by the returned product inspection and testing. product damage (cannula kink): the returned product was inspected and a substantial kink and twist of the catheter was observed at around the 40cm marking. clinical details noted kink formed when physician tried rotating the device. no information on patient anatomy complications were reported. the root cause of the product damage (cannula kink) was determined to be due to a catheter kink most likely due to a use-related issue as evidenced by returned product inspection and clinical details.
8984 | 1220648-2025-28814 | 22105070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22105070&pc=OZD | 05/06/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  | 02/28/2026 |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467026 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during the impella 5.5 implantation the 5.5 was back loaded onto wire and advanced into graft but unable to fully advance into the artery due to vessel size. multiple attempts were made but unsuccessful. the 5.5 placement was aborted and the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8985 | 1220648-2025-28815 | 22105090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22105090&pc=OZD | 05/07/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583118 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a computed tomography scan that resulted in the presence of transtentorial herniation, indications of hypoxia/ischemic encephalopathy and edema. care was withdrawn and the patient expired. it is unknown if the patient's brain injuries were related to the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8986 | 1220648-2025-28816 | 22105330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22105330&pc=OZD | 05/06/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595054 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had ventricular tachycardia. the patient was shocked multiple times. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
8987 | 1220648-2025-28817 | 22105579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22105579&pc=OZD | 05/06/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623483 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Fever | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was placed on impella cp as a bridge to coronary artery bypass graft. during impella support, it was reported that heparin was paused due to the patient coughing up blood. the patient was also febrile and lactate had increased. antibiotics were been started along with cultures. ultimately the patient was taken to cath lab for right heart catheterization and impella explant due to bacteremia. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8988 | 1220648-2025-28818 | 22105970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22105970&pc=OZD | 05/06/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645879 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 80 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced some oozing from their access site during impella cp support. two units of blood were transfused and femostop was applied to treat the condition. the suture hub was repositioned and the angle was matched. the ooze resolved and support continued with the implanted pump. the patient was successfully weaned from the pump and survived. | [5 empty] | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8989 | 1220648-2025-28819 | 22106486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22106486&pc=OZD | 02/28/2025 | 06/23/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607255 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 59 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Disability | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient endured limb ischemia and hemolysis with a "possible" relationship to the impella cp. the patient experienced mottled toes/foot of the right lower extremity and no signals distal to the popliteal artery. the right leg was amputated. blood products were transfused as the patient's hemoglobin was 8.8 but there were no other signs of bleeding. the impella was successfully explanted and the patient continued on ecmo support. | [5 empty] | The investigation of limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 required intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28819. b2 other serious injury was erroneously selected on the initial manufacturer device report number 1220648-2025-28819. g2; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28819. h6 health effect - clinical code 1888 was erroneously entered on the initial manufacturer device report number 1220648-2025-28819. h6 health effect - impact code 4627 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28819. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8990 | 1220648-2025-28820 | 22106681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22106681&pc=OZD | 02/11/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607858 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp exhibited was concerns of a thrombus in the left ventricle and ventricular tachycardia. the tachycardia was treated with lidocaine and amiodarone. the patient recovered with no sequelae. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
8991 | 1220648-2025-28821 | 22106993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22106993&pc=OZD | 05/06/2025 | 06/26/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 123 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Deformation Due to Compressive Stress | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | It was reported that a patient was undergoing escalation from an impella cp to an impella 5.5. removal of the guidewire was difficult as it was kinked. oozing was noted from the access site so a jackson-pratt drain was added, in addition to a graft cutdown with suture. five units of blood products were transfused. |  |  |  |  | The investigation of product damage (guidewire damage - great vessel), positioning issues, and access site bleeding has been completed. the impella device was not returned for evaluation. product damage (guidewire damage - great vessel): upon wire removal, wire became kinked distal to impella inlet and was unable to be removed. clinical detail state the gw unraveled but do not mention the state of the vessels. the wire was not returned. the cause of the product damage (guidewire damage) was not determined since product was not returned and insufficient clinical detail. positioning issues: after the guidewire (gw) became kinked distal to impella inlet, impella 5.5 and wire and come back into the aorta and the impella 5.5 catheter was wrapped below the impella outlet. catheter was manipulated up and down and torqued to relieve the malpositioned catheter. the physician stated pump was "fine," no kinks, bends, or abnormalities noted on inspection. the pump and gw were not returned. the cause of the positioning issue was most likely an interaction with another device being the kinked gw that lodged onto the impella inlet cage and pulled the pump into the aorta during gw removal attempts. access site bleeding - major: after placing the pump some oozing noted from impella pocket and jp drain added prior to closing the site. patients act was noted to be at 308 after total heparin of 28000u given. the pump was not returned. the cause of the access site bleeding - major was most likely high patient activated clotting times (act) values of 308 after 28000u of heparin given. b2 death and date of death were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28821. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28821. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28821. h6 health effect - impact code 1802 and medical device problem code 3009 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28821. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8992 | 1220648-2025-28822 | 22107175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22107175&pc=OZD | 03/26/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618542 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced thrombocytopenia that could have possibly been caused by the impella. the patient's platelet count was at 100k prior to implantation and was at 96k after implantation. the patient was transfused two units of platelets. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8993 | 1220648-2025-28823 | 22107356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22107356&pc=OZD | 05/17/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/19/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553408 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella cp support that a ventricular septal perforation (vsp) was confirmed a few days after percutaneous coronary intervention. the doctor believed that the vsp was not caused by impella, but the perforation occurred during support. later the patient expired from multiple organ failure. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8994 | 1220648-2025-28824 | 22107737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22107737&pc=OZD | 05/06/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650324 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO, IMPELLA 5.5 | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and an impella cp. after ecmo cannulas were removed and cutdown repairs were done on both arteries, there was a significant amount of bleeding and the patient's platelet count dropped to 43. the patient was transfused with six units of packed red blood cells and cryo. additionally, heparin was reversed. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8995 | 1220648-2025-28825 | 22108345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22108345&pc=OZD | 05/06/2025 | 05/29/2025 |  | 05/29/2025 |  |  | 05/06/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654232 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 101 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 for transplant evaluation. during support it was reported that the purge solution was changed due to high pressure alarms. tissue plasminogen activator was started. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8996 | 1220648-2025-28826 | 22108577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22108577&pc=OZD | 05/07/2025 | 06/30/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600499 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented for a planned percutaneous axillary cp for high risk cardiac procedure. it was reported that after the impella cp was in, they noticed bleeding from the membrane in the 14 french by 30 cm sheath. placing the repositioning sheath through the 14 french sheath resolved the bleeding. upon removal the 14 french peel away sheath was inspected and noticed a slash in the membrane of the sheath. there was no patient harm as a result. | [5 empty] | The investigation of introducer damage ¿ mechanical and access site bleeding has been completed. the impella cp and 14fr x 13cm introducer were returned for evaluation. introducer damage - mechanical: 14fr x 13cm introducer was returned for evaluation. upon visual inspection, no abnormalities were observed on the introducer. the introducer was able to hold pressure up to approximately 380mmhg. the introducer was also leak tested with a catheter inserted into the hemostatic valve and the introducer was able to hold pressure up to approximately 215mmhg. clinical details noted that bleeding around catheter was observed from the introducer valve and when the repositioning sheath was inserted through the introducer valve over the catheter, the bleeding stopped. it was also noted that the catheter was observed to be inserted slightly above the center of the sheath membrane. no problem was found with the returned introducer as no sheath damage was found and the introducer operated to specification during leak testing. access site bleeding - major: clinical details noted that patient was bleeding from the hemostatic valve with pump inserted. once repositioning sheath was inserted through the hemostatic valve, the bleeding subsided. it was also noted that the catheter was observed to be inserted slightly above the center of the sheath membrane. the root cause of the access site bleeding was use related as the bleeding occurred from a biased catheter. d9 device returned to manufacturer date added h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28826. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8997 | 1220648-2025-28827 | 22108834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22108834&pc=OZD | 05/07/2025 | 08/11/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1845417 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | Unexpected Shutdown |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a console failure occurred. the aic was removed from cath lab service. no noted harm or injury. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation has been completed. the console successfully underwent preventative maintenance on april 8, 2025. on may 2, 2025, likely at the customer site, after the console was initially booted, it displayed ¿battery failure (transport connection blown/missing) ¿ alarm,¿ indicating that a battery switch was most likely disengaged. the console was then improperly shut down, likely by disconnecting the alternating current (ac) power. after a manual reboot, alongside the previously mentioned alarm, there was also an ¿unexpected controller shutdown ¿ alarm¿. a couple more times, the user performed improper shutdowns by disconnecting the ac power while the battery switch was disengaged, and after every manual reboot, the console displayed both alarms. there was no controller error alarm in the logs. most likely, either the 'battery failure' or the 'unexpected controller shutdown ¿ alarm' was misreported as a controller error alarm. this console was tested and upon running an optical test, there were no battery failure or unexpected shutdown alarms reproduced. the battery was able to charge up to 100%. the root cause of the ¿battery failure¿ alarm was most likely that the battery switch was disengaged, indicating a use issue. the root cause of the ¿unexpected controller shutdown ¿ alarm¿ was the improper shutdown, most likely due to disconnecting the ac power while the battery switch was disengaged, indicating a use issue. there was no controller error alarm in the logs. either the ¿battery failure¿ or the ¿unexpected controller shutdown ¿ alarm¿ was misreported as a controller error alarm. e.1 email address was previously entered incorrectly on initial manufacturer device report number 1220648-2025-28827. phone number was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-28827. h.6 codes 3221 and 4315 were reported incorrectly on initial manufacturer device report number 1220648-2025-28827. codes 3233 and 11 should of been reported on initial manufacturer device report number 1220648-2025-28827.
8998 | 1220648-2025-28828 | 22109059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22109059&pc=OZD | 05/07/2025 | 07/15/2025 |  | 05/29/2025 |  |  | 05/07/2025 | 05/29/2025 |  |  |  |  | 05/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026645471 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the patient developed hemolysis. the pump was found deep in the left ventricle. the pump was pulled back and repositioned. volume was added. hemolysis labs started to decrease and support continued. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues and hemolysis has been completed. the pump was returned. the data logs were unavailable. positioning issues: after 1 day on support, echocardiogram showed impella in too deep into left ventricle (lv), pulled impella to correct position and verified by echo and signals, volume added and p level back to 9. the cause of the positioning issue was most likely use related due to the pump requiring repositioning 1 day after implantation. hemolysis: the urine output became red over time despite starting as yellow at time of pump start. after 1 day on support, echocardiogram showed impella in too deep into lv, pulled impella to correct position and verified by echocardiogram and signals, volume added and p level back to 9, urine check regularly. update 8 hours later states hemolysis labs decreased. urine back to yellow 2 days later. data logs were not recovered. the pump was returned and inspected. no visual abnormalities were noted. the cause of the hemolysis was most likely improper positioning of the device due to hemolysis beginning to resolve within 8 hours of repositioning of the device. d9 device returned to manufacturer date added.
8999 | 1220648-2025-28751 | 22094578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22094578&pc=OZD | 04/01/2025 | 08/15/2025 |  | 05/28/2025 |  |  | 05/21/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612157 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after ascending a replacement, impella cp insertion was difficult. snare technique was performed but the impella cp could not be placed. the patient returned supported only on extracorporeal membrane oxygenation. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be use issue because the pump was unable to pass through vasculature due to resistance at aortic arch. h5 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28751.
9000 | 1220648-2025-28752 | 22094696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22094696&pc=OZD | 05/09/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/09/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647180 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. some time after the impella cp was implanted, during the high-risk percutaneous coronary intervention and balloon intervention, a vascular rupture with pericardial effusion occurred (not impella related), and it was successfully treated with a covered stent. the patient continued on support until the device was successfully weaned. the patient's outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9001 | 1220648-2025-28753 | 22094759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22094759&pc=OZD | 05/09/2025 | 09/17/2025 |  | 05/28/2025 |  |  | 05/09/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630909 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during insertion of the impella cp, a placement signal was received. after starting the impella cp, the placement signal disappeared, and the motor current curve was flat. the impella cp was immediately removed and a new impella cp was implanted. the patient continued on support until the device was successfully weaned. there was no patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. based on the data and device analysis there was no problem found regarding the pump for the pump stop. | H6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number. this report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. all pertinent information available to abiomed has been submitted at this time.
9002 | 1220648-2025-28754 | 22094836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22094836&pc=OZD | 05/09/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/09/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1236849 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that an automated impella controller (aic) brought up a self-check failed message. the aic was immediatley changed. there was no patient involved. | [5 empty] | The investigation into the reported issue has been completed. the console was returned. console logs were consistent with what was reported in the complaint. numerous communication issues with the power battery manager (pbm) were observed in the logs. the console was tested and the reported system self check failure was able to be reproduced while testing the console on an engineering lab bench. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring communication channel. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
9003 | 1220648-2025-28755 | 22094909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22094909&pc=OZD | 05/04/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/09/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626542 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Embolism/Embolus | Thrombosis/Thrombus | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. embolism and thrombosis of the left femoral artery were noted after impella supported percutaneous coronary intervention. post-interventional impella support was reduced over two days. acute ischemia of the left leg was reported after impella cp removal. embolectomy and thrombectomy of the left femoral artery was performed. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9004 | 1220648-2025-28756 | 22094982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22094982&pc=OZD | 05/07/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492213 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella 5.5 was placed and during the cardiac intervention case, after fixing the left anterior descending artery, the pump flows dropped from 3.4 to 0.6. the pump position was checked. the pump was taken out, inspected, and then put back in with no change to the flows. the impella 5.5 was then removed. a second impella pump was not used as the pressures were good and one drip had already been weaned off. the case continued and the circumflex artery was fixed without an issue. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9005 | 1220648-2025-28757 | 22095049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095049&pc=OZD | 04/29/2025 | 09/25/2025 |  | 05/28/2025 |  |  | 04/29/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632539 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem | Infusion or Flow Problem |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for approximately 26 hours before being implanted with an impella 5.5 device for escalation in mechanical circulatory support (mcs). the next day after start of the impella 5.5 mcs, continual issues with ¿in aorta¿ alarm throughout night was noted despite proper placement on echocardiogram. the impella 5.5 pump was manipulated without improvement in flows. chest x-ray showed redundant slack in aorta prior to repositioning. a repeat chest x-ray was planned with repositioning of the device if required. the following day, however, the impella 5.5 device was weaned to performance level p-0 and the device was explanted without an issue. after removal a clot was noted in the inlet. the pocket was closed; jackson-pratt drain remained in place. the site was dressed, and the patient reportedly remained hemodynamically stable and returned to the unit with plans to extubate. | Us complainant reported that the device exhibited an optical signal issue. |  |  |  |  | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of low pump flow, positioning issues, and thrombus has been completed. the impella device was not returned for evaluation. low pump flow: data logs were investigated and there 2 incidents of low flows at p4 and p5 with "impella in aorta" alarms. since clinical said that the position alarms are not true and repositioning the pump did not help get better flows, the root cause of the low pump flow issue was not determined. positioning issue: data logs were returned and numerous "impella in aorta" alarms were observed. no other supporting information was provided. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. saturated flows: failure mode has been added as it was found during investigation. clinical states that there was a clot noted at inlet. data logs were investigated. on (b)(6) 2025 around 11am, at p4, the low flows turned to saturated flows with active ¿impella in aorta¿ alarms. the flow gradually reduced and rose again to saturated flows from 4pm to 6am, (b)(6) 2025. a motor current spike corresponding to the ¿impella in aorta¿ alarms observed followed by saturated flows. placement signal was seen trending upwards during the saturated flows. therefore, the root cause of the saturated flow issue was most likely biomaterial since data logs confirmed a motor current spike followed by saturated flows. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: clinical states that there were false "impella in aorta" alarms. pump was not returned for investigation. data logs were investigated and " impella in aorta" alarms were confirmed along with low flows. the 5s parameter were all in specification. placement signal was seen to be slightly drifting up after the motor current spike. rt logs relative to failure event were not available. without the return product the failure cannot be further investigated. therefore, the root cause of the optical signal issue was not determined since product was not returned, clinical information was insufficient and data logs were inconclusive. low pump flow: clinical states that there was false impella in aorta alarms and pump was manipulated without improvement in flows. pump was not returned for investigation. data logs were investigated and there 2 incidents of low flows at p4 and p5 with "impella in aorta" alarms. since clinical said that the position alarms are not true and repositioning the pump did not help get better flows, the root cause of the low pump flow issue was not determined. positioning issue: clinical states that the pump had false "impella in aorta" alarms. cxr showed redundant slack in aorta prior to repositioning. pump was not returned for investigation. data logs were returned and numerous "impella in aorta" alarms were observed. no other supporting information was provided. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. saturated flows: failure mode has been added as it was found during investigation. clinical states that there was a clot noted at inlet. pump was not returned for investigation. data logs were investigated. on (b)(6) 2025, around 11am, at p4, the low flows turned to saturated flows with active ¿impella in aorta¿ alarms. the flow gradually reduced and rose again to saturated flows from 4pm to 6am, on (b)(6) 2025. a motor current spike corresponding to the ¿impella in aorta¿ alarms observed followed by saturated flows. placement signal was seen trending upwards during the saturated flows. therefore, the root cause of the saturated flow issue was most likely biomaterial since data logs confirmed a motor current spike followed by saturated flows. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. added code 4251 and 4310 which are the investigation results for the optical signal issue.
9006 | 1220648-2025-28758 | 22095126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095126&pc=OZD | 05/04/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608266 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female | 85 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5 pump the patient had positive blood culture results indicating sepsis. the pump was replaced with another impella 5.5 and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & the instructions for use for the cp with smartassist states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9007 | 1220648-2025-28759 | 22095229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095229&pc=OZD | 05/05/2025 | 06/20/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594037 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 57 KG |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was implanted in a patient admitted in with cardiogenic shock and cardiomyopathy. a ct scan of the head showed a new infarct although the patient showed no neurological defects. the patient survived. | [5 empty] | The investigation of stroke/cva and positioning issues has been completed. the impella device was not returned for evaluation. stroke/ cva: as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. positioning issue: clinical states that there were suction alarms and "impella position unknown" alarms while moving the patient. the pump was found to be angled towards mitral valve at 4.5cm which is not the initial position. pump was not returned for investigation. as per the clinical information provided, the root cause of the positioning issues was most likely patient movement since the pump was found to be out of initial place and facing mitral valve after moving the patient. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28759. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28759. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9008 | 1220648-2025-28760 | 22095360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095360&pc=OZD | 05/07/2025 | 06/05/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1281141 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the silver button on the automated impella controller was nonresponsive when clicking it. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - kbd/rotary knob issues has been completed. the console logs from the reported day of event are consistent with complaint and show that after boot-up, only soft-key presses were recorded, but none of the on-screen menu items were confirmed via rotary push button ¿click¿. console was unable to be manually shut down, and eventually turned itself off after 45 minutes of no used interaction. console was tested and the issue was reproduced - the ¿click¿ (push) of the rpb was not recorded by the device, although the rotation was. further testing of the rpb encoder, impellatronic board (passing rpb signals to 4-in-1 carrier), and pbm (providing power to rpb) revealed failure of impellatronic board, and temporarily replacing the impellatronic assembly resolved the issue.
9009 | 1220648-2025-28761 | 22095405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095405&pc=OZD | 05/01/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642648 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the day after explant of an impella cp, the patient had a ct in preparation for a transcatheter aortic valve replacement. the ct scan showed the patient had descending aortic dissection. it is unknown whether the injury was caused by the impella device. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9010 | 1220648-2025-28762 | 22095446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095446&pc=OZD | 05/07/2025 | 07/16/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620827 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 pump was placed for a patient admitted in with cardiomyopathy and cardiogenic shock. the impella 5.5 had a purge leak that prompted the team to explant after six hours. the team placed a new pump and support proceeded without harm or injury from interruption. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the pump was returned for investigation. no damage was observed visually. the pump was started for priming, and a leak was visible inside the sidearm cage. upon opening the sidearm clear case, it was found to be broken at the connection between the filter and reservoir joint. the data log shows that the purge pressure was trending downward for around 3.5 hours before the damage affected the open purge line. according to the clinical notes, a purge leak was noted from the connection of the purge filter and reservoir, and during tightening, the connection completely cracked into two pieces. the cause of the purge leak was damage to the reservoir-to-filter joint, as confirmed by the doctor, with the data log showing evidence of a slow purge leak trend.
9011 | 1220648-2025-28763 | 22095538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095538&pc=OZD | 05/07/2025 | 07/08/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568900 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and after the implant, there were low pump flows. no thrombus was observed. the pump was explanted and a new pump was implanted. the patients outcome is stable. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the pump was returned for investigation. a data log was not provided, and it was not retrieved from the remote server. no physical damage was observed on returned product, except the kinked cannula. the impeller was free of biomaterial. the pump was tested in a bucket of water without any abnormalities. the cause of the low pump flow issue was most likely due to a kinked cannula; however, the reason for the cannula kink remains unknown. product damage (cannula kink) failure mode was added as an investigated failure mode. the cause of the product damage issue was kinked cannula, however the cause of the cannula kink was unknown without sufficient clinical information.
9012 | 1220648-2025-28764 | 22095562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095562&pc=OZD | 05/04/2025 | 09/25/2025 |  | 05/28/2025 |  |  | 05/04/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572150 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Positioning Failure | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing along with extracorporeal membrane oxygenation for a patient admitted in with cardiogenetic shock and cardiomyopathy. it was reported that the pump is in poor position and had to be adjusted frequently. it was also noted when repositioning that the tail end of the repo sheath was torn. it is unknown if the tear in the sleeve occurred during use or not. there was no patient harm. | The customer reported that the device exhibited an optical signal issue. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product damage (sterile sleeve damage). it was noted that during repositioning the tail end of the repositioning sheath was torn. it was unclear if it got damaged during multiple attempts of repositioning, during insertion or if it was a manufacturing complication. due to lack of product returned, the root cause of the product damage (sterile sleeve damage) could not be determined. positioning issues: a few hours after support, there was concern for improper impella placement, due to no ao signal and the impella was at 87cm. repositioning was conducted and also volume was given to the patient. additionally, there was repositioning done after an impella position in aorta alarm on (b)(6) 2025. due to lack of clinical details, the root cause of the positioning issues cannot fully be determined as it is unclear if the pump was placed in an incorrect position. optical signal issue: it was stated that the pump was not in the correct place and repositioning was done. there was also low placement signals. the patient was also on ecmo and patient began auto-diuresing and education was done on volume status as there was poor lv and rv function. after volume was given, the pulsatility and placement signal returned to normal. the data logs show a significant increase in placement signal (ps) from 80mmhg to ~ 130mmhg over the course of 18 mins, and the p-level was decreased. the placement started trending down from 130mmhg to about 40mmhg with a trend in flows. ps low alarms resolved shortly after. based on the clinical description and data logs, the root cause of the optical signal issue is most likely due to the patients condition as the ps low alarms improved with adding volume. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. added codes for investigation findings and conclusion based off the investigation for the optical signal issue.
9013 | 1220648-2025-28765 | 22095808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095808&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653492 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 suffered a puncture of the impella through the left ventricle wall. the perforation was repaired with a stitch. care was withdrawn and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9014 | 1220648-2025-28766 | 22095948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095948&pc=OZD | 05/16/2025 | 08/01/2025 |  | 05/28/2025 |  |  | 05/16/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026663557 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted and an impella cp device was placed for planned support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. product damage (cannula kink) and product damage (catheter kink) added based on product evaluation. reported aborting impella 5.5 delivery due to anatomical reasons and impella cp was placed. the patient had a stenosis in his vessels. impella 5.5 was returned and multiple cannula kinks were found as well as catheter kinks. delivery issues - the root cause of the delivery issue was most likely patient condition related since the patient had a stenosis in his blood vessels. cannula kink - the root cause of the product damage (cannula kink) was most likely patient condition related since the patient had a stenosis in his blood vessels. d9: device returned to manufacturer date added. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with automated impella controller section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28766. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9015 | 1220648-2025-28767 | 22095979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22095979&pc=OZD | 05/22/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/22/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026617724 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that impella cp was inserted in a patient with acute myocardial infarction. fluoroscopy confirmed that the blood vessels at the puncture site were tortuous. there was no calcification or stenosis. a hematoma was observed a while after returning to the icu. the access site was sutured and the impella was removed. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9016 | 1220648-2025-28768 | 22096010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096010&pc=OZD | 05/04/2025 | 06/23/2025 |  | 05/28/2025 |  |  | 05/04/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621642 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a .018 guidewire kinked during insertion for impella cp implant. the wire was replaced with another guidewire from a second set and the subsequent implant was successful. oozing was noted at the access site leading to femostop placement to manage the condition. suction alarms appeared the following day and the pigtail was found to be under the papillary muscle. troubleshooting resulted in suboptimal positioning with on the lower end for the set p-level. due to the patient's stability, the decision was made to continue with the current positioning/flows. although the oozing had stopped, a resulting hematoma had doubled in size. five units of blood were given to treat the hematoma, however significant bleeding then developed at the site following additional repositioning of the device. suction also remained intermittent, for which volume and positioning were managed. plans were made to escalate to impella 5.5 support, but the patient continued to deteriorate and later passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding, positioning issues, delivery issues, product damage (guidewire damage - peripheral vessel, and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data logs were reviewed and lt logs show suction alarms occurring intermittently during support. multiple instances of placement signal (ps) increases with drop in motor current (mc) and pump flow. no positioning alarms or motor current spikes to indicate ingestion were observed. clinical details noted that suction alarms occurred multiple times throughout support. clinical states good flows initially, but then later on day 1 of support the pump was repositioned into less than ideal position with flows on low end for the p-level. on day 2 of support patient had suction alarms, was given volume, and the pump was repositioned with improvement in flows. suction alarms stated on day 3 with pump stated to be shallow. the root cause of the low pump flow was most likely improper position based on clinical details indicating low flow issues occurring concurrently with positioning issues. positioning issues: the root cause of the positioning issues was most likely due to patient movement as the pump was noted to be out of position after patient was transferred to unit and after patient was transferred hospitals. access site bleeding: the root cause of the access site bleeding could not be definitively determined as multiple factors, including patient movement and transfer and guidewire kinking, occurred prior to the oozing and hematoma. product damage (guidewire damage): clinical details noted that first 0.018 guidewire attempted to be inserted was kinked and guidewire was replaced with another 0.018 guidewire. the root cause of the guidewire damage could not be definitively determined as no product was returned for evaluation and lack of clinical details. delivery issues: clinical details noted that guidewire kinked during insertion. guidewire was removed and a new guidewire was used for pump insertion. the root cause of the delivery issues was most likely due to a guidewire kink per clinical details. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9017 | 1220648-2025-28769 | 22096053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096053&pc=OZD | 05/07/2025 | 07/08/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570082 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the purgeline was leaking at the red impella plug. the pump was explanted in response to the leak. the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The investigation of purge leak ¿ pump has been completed. data log review confirmed leaking occurred on day 5 with decreased purge pressure and increased purge flow with no motor current (mc) spikes or flow changes outside of p-level changes. the pump was returned and inspected. the red pump plug leak was able to be reproduced with purge fluid leaking in the catheter line on the pump side of the plug. there was no further damage noted to the pump. the cause of the purge leak - pump was damaged purge tubing causing the purge fluid to leak into the red pump plug. however, the cause of this leak is not known due to insufficient clinical details and no further damage noted to the pump. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9018 | 1220648-2025-28770 | 22096262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096262&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633307 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female | 82 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient's urine was red tinged and their labs were hemolyzed. as the patient would not tolerate a lower p-level, the decision was made to continue with support despite the condition. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
9019 | 1220648-2025-28771 | 22096523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096523&pc=OZD | 05/05/2025 | 07/11/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629717 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a 63 year old male patient experienced some oozing at their access site during impella cp support. heparin levels were reduced and the dressing was changed to manage the condition. support continued following the intervention. three days later, care was withdrawn and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | Additionally ventricular tachycardia (vt) was added due to cardioversions were required for an unknown indication. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28771. h6 health effect - clinical code 2095 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28771. | The root cause of access site bleeding is determined to be anti-coagulation management because the hemostasis was achieved by reducing the heparin and dressing change.
9020 | 1220648-2025-28772 | 22096533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096533&pc=OZD | 05/06/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532039 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient experienced a run of ventricular tachycardia (vt) post impella cp implant. the patient did not lose consciousness or pulse and came out of it. amiodarone bolus and drops were started. impella support was continued until the patient was successfully weaned off the impella cp. the patient survived. additional information was received that the impella was not the cause of the vt. the position was good and not interacting with any structures. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9021 | 1220648-2025-28773 | 22096658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096658&pc=OZD | 05/06/2025 | 07/01/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658400 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced oozing at the access site. new forward tensions sutures and angle matching dressings were applied to obtain homeostasis. support was continued until the patient was successfully weaned from the pump. the patients outcome was stable. | [5 empty] | The root cause of access site bleeding was determined to be use issue because the bleeding was resolved by new forward tension sutures and angle matching dressing at the bleeding site. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9022 | 1220648-2025-28774 | 22096913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096913&pc=OZD | 05/06/2025 | 06/30/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523533 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an "impella in aorta" alarm appeared during impella cp support. an echocardiogram revealed that the pump's pigtail was still across the valve and allowed for readvancement of the pump. at the same time, a left ventricle thrombus was confirmed. the pump was elected to be removed for an intra-aortic balloon pump. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: the cause of positioning issues most likely was due to patient movement as the pump became out of position after transfer. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined.
9023 | 1220648-2025-28775 | 22096952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22096952&pc=OZD | 05/05/2025 | 06/23/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615294 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Migration or Expulsion of Device | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed on a patient with lateral wall mi. the impella was successfully placed. during support it was reported that the patient went into cardiac arrested after arrival to the unit. during cpr the impella migrated into the aorta and while using echo guidance they could not cross back into the valve. the impella was placed in the descending aorta and placed on p1 as the team tried to connect with family to discuss goals of care. ultimately, the patient expired on support. | [5 empty] | The investigation of positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issues was most likely use related as the pump came out of position during cardiopulmonary resuscitation (cpr). vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9024 | 1220648-2025-28776 | 22097119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097119&pc=OZD | 05/02/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/02/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597636 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and the following day, the patient developed decreased pulses in the right lower extremity. the impella cp was explanted and an impella 5.5 was implanted for escalation of care. during this time a fasciotomy was performed to resolve the limb ischemia, and two units of blood were given to the patient following the intervention. strong pulses were noted status post fasciotomy. the patient was noted to be in stable condition following the event and continued on impella 5.5 mcs. however, the patient's neurological status was not fully intact, patient confused and was given haldol. four days later, the patient experienced heparin induced thrombocytopenia (hit), and the next day a run of ventricular tachycardia, which was resolved without intervention. however, the patient was still hit positive. the patient's neurological status was still not fully intact; the patient remained confused. approximately nine days later the patient would expire. care was withdrawn. | [5 empty] | Medical safety reviewed the event and given the details of event, impella 5.5 related events there is not enough evidence to exclude the impella 5.5 as an associated factor to the patient's outcome. the impella 5.5 is one of two devices associated with the reported event. refer to medical device report for impella cp serial number (b)(6) with date of event 38-apr-2025 and patient identifier ending in 3700. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9025 | 1220648-2025-28777 | 22097279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097279&pc=OZD | 05/06/2025 | 06/19/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619239 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant and the pump was unable to pass the left ventricle. the abdominal aorta was calcified and the pump was unable to pass. the medical doctor aborted the insertion and proceeded with percutaneous coronary intervention using the right femoral access. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, pump inserted via sheath but unable to access lv and aorta calcified and unable to pass.
9026 | 1220648-2025-28779 | 22097423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097423&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593527 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported the 64 year old male patient was on impella cp support and during support, the access site had bleeding. the site was resutured and resecured at bedside, and a femstop placed. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9027 | 1220648-2025-28780 | 22097454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097454&pc=OZD | 04/28/2025 | 05/29/2025 |  | 05/28/2025 |  |  | 04/29/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614226 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support, escalated to an impella 5.5 device and experienced a hematoma at both access sites with persistent bleeding. the patient would subsequently expire. the patient presented with chest pain and left heart catheterization revealed multi-vessel coronary artery disease and elevated end diastolic pressure. the decision was made to place impella cp via the right femoral artery, which was successfully completed. the patient was placed on p-6 support level, noted to be hemodynamically stable awaiting life flight for transfer to a higher level of care facility and cardiovascular consultation. approximately, 16 days later, the impella cp device was explanted; the patient was escalated to an impella 5.5 via the left axillary artery for escalation in therapy. following, a large hematoma to left chest wall encompassing entire pectoral and shoulder areas and extending to left upper arm had developed. a right femoral hematoma as well had developed at site of impella cp explant. hemostatic agents were used in the pockets while in the operating room and long periods of manual pressure were also applied. four "jumbo" platelets were given while in the operating room and as well protamine was administered. the patient was continued on therapy with the impella 5.5 with possible sepsis indicated approximately six days later. the next day, there was no cough, no gag, pupils unreactive, and hemodynamics were declining. the patient would expire the following day; care was withdrawn. | [5 empty] | Medical safety review of the event noted large hematoma at femoral access site requiring intervention is serious injury for the impella cp; there was escalation of care to impella 5.5 and adverse events occurring (that includes possible sepsis) contributed more to the patient¿s demise. there is not enough evidence to exclude the impella 5.5 device used as an associated factor to the patient's outcome. the impella 5.5 is one of two devices associated with the event. refer to medical device report for impella cp serial number (b)(6) with date of event 28-apr-2025 and patient identifier ending in (b)(6). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of the reported issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9028 | 1220648-2025-28782 | 22097485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097485&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567264 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump placed to support a patient admitted in with cardiogenic shock, cardiomyopathy, and muscular dystrophy. the pump had suction alarms, positioning issues, and a hematoma which was surgically treated. the pump remains on for support despite this as the team awaits heart transplant. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9029 | 1220648-2025-28778 | 22097492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097492&pc=OZD | 04/28/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 04/28/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626221 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and the following day, the patient developed decreased pulses in the right lower extremity. the impella cp was explanted and an impella 5.5 was implanted for escalation of care. during this time a fasciotomy was performed to resolve the limb ischemia, and two units of blood were given to the patient following the intervention. strong pulses were noted status post fasciotomy. the patient was noted to be in stable condition following the event and continued on impella 5.5 mcs. however, the patient's neurological status was not fully intact, patient confused and was given haldol. four days later, the patient experienced heparin induced thrombocytopenia (hit), and the next day a run of ventricular tachycardia, which was resolved without intervention. however, the patient was still hit positive. the patient's neurological status was still not fully intact; the patient remained confused. approximately nine days later the patient would expire. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted although the patient experienced limb ischemia and underwent fasciotomy the issue resolved. strong pulses were noted the next day. serious injury for impella cp, and the demise outcome is associated with the related complaint for the impella 5.5 device reported under manufacturing report number 1220648-2025-28776. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9030 | 1220648-2025-28783 | 22097559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097559&pc=OZD | 04/28/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 04/29/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614149 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support, escalated to an impella 5.5 device and experienced a hematoma at both access sites with persistent bleeding. the patient would subsequently expire. the patient presented with chest pain and left heart catheterization revealed multi-vessel coronary artery disease and elevated end diastolic pressure. the decision was made to place impella cp via the right femoral artery, which was successfully completed. the patient was placed on p-6 support level, noted to be hemodynamically stable awaiting life flight for transfer to a higher level of care facility and cardiovascular consultation. approximately, 16 days later, the impella cp device was explanted; the patient was escalated to an impella 5.5 via the left axillary artery for escalation in therapy. following, a large hematoma to left chest wall encompassing entire pectoral and shoulder areas and extending to left upper arm had developed. a right femoral hematoma as well had developed at site of impella cp explant. hemostatic agents were used in the pockets while in the operating room and long periods of manual pressure were also applied. four "jumbo" platelets were given while in the operating room and as well protamine was administered. the patient was continued on therapy with the impella 5.5 with possible sepsis indicated approximately six days later. the next day, there was no cough, no gag, pupils unreactive, and hemodynamics were declining. the patient would expire the following day; care was withdrawn. | [5 empty] | Medical safety review of the event noted large hematoma at femoral access site requiring intervention is serious injury for the impella cp; there was escalation of care to impella 5.5 and adverse events occurring (that includes possible sepsis) contributed more to the patient¿s demise. there is not enough evidence to exclude the impella 5.5 device used as an associated factor to the patient's outcome. refer to manufacturing report number 1220648-2025-28780 for the impella 5.5 medical device report. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9031 | 1220648-2025-28781 | 22097673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097673&pc=OZD | 05/06/2025 | 06/30/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656869 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after multiple attempts at attempting to advance the impella guidewire/catheter, implantation of the impella 5.5 was aborted due to the patient¿s vascular anatomy/stenosis. an impella cp was implanted via the axillary. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since even after multiple attempts the pump was unable to advance impella wire/catheter due to patient¿s vascular anatomy/stenosis.
9032 | 1220648-2025-28784 | 22097726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097726&pc=OZD | 05/06/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  | 01/31/2026 |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461167 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Female |  |  |  | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient had an attempted impella 5.5 implant and during the implant, a 10 millimeter graft was sewn to the innominate artery. when the impella 5.5 was inserted, it partially tore the graft from the artery resulting in bleeding. the decision was made to abort the impella 5.5 implant at this time. surgical repair of the innominate artery was done. the patient was converted to central venoarterial extracorporeal membrane oxygenation and sent back to unit. the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9033 | 1220648-2025-28785 | 22097908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22097908&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657549 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump supporting a patient admitted in for high risk cardiac procedure. the pump was supporting for a week when due to rapid atrial fibrillation the team made choice to restart heparin systemically and remove the pump. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9034 | 1220648-2025-28786 | 22098198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22098198&pc=OZD | 05/06/2025 | 06/25/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1762205 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported from a biomedical engineer that when an automated impella controller (aic) was being prepped for use, a controller error was received. an alarm occurred before the aic was put into use. the aic was removed from the room and another aic was used. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: on 4/14, before the complaint date, a 5s test was run on the device with passing metrics. no pump is used until the complaint date on 4/23 where on boot up there is a 1039 controller error for a defective optical sensor lamp due to low light voltage (0.73 and 0.74v). this was reproduced over multiple boot ups. around 0.5v would have implied no light so 0.7-0.8v hints towards a disruption of the light. device analysis: the failure mode was reproduced. upon further inspection of the console, the optical bench was found to be heavily contaminated. the contamination was able to be cleaned and the failure mode was resolved. root cause: the controller error was due to a contaminated optical connector that was resolved after cleaning the optical connector. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9035 | 1220648-2025-28787 | 22098536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22098536&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656676 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had ectopy and dobutamine discontinued with improvement in rhythm. additionally, the pump's inflow was rotated towards septum but free of mitral valve apparatus at 5.1 cm inflow to aortic valve. the pump was not repositioned. staff monitored. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9036 | 1220648-2025-28788 | 22098583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22098583&pc=PYX | 05/06/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026636099 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp flex being placed to a support a post transplant patient. once placed, the impella rp flex failed to start and was replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9037 | 1220648-2025-28789 | 22098712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22098712&pc=OZD | 04/30/2025 | 06/23/2025 |  | 05/28/2025 |  |  | 04/30/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656632 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Hemolysis | Sepsis | Gastrointestinal Hemorrhage | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 52-year-old male patient with continued monophasic flow to right lower extremity (impella cp leg) and continued pressor/inotropes/low ejection fraction, low pulsatility. decision was made to escalate to impella 5.5 given persistent cardiogenic shock. the 5.5 was successfully implanted and cp explanted. during 5.5 support, it was noted of some suction events and pink-tinged urine. in addition, the patient experienced atrial fibrillation with rapid ventricular response (rvr). however, it was documented that after repositioning, that hemolysis and suction resolved. the patient received 2 units of blood for a gastrointestinal (gi) bleed and heparin was placed on hold. during this time, the patient experienced intermittent fever. due to the septic component, the patient was given addition of vasopressin and lactic acid. family elected to withdraw care and the patient expired on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the medical safety officer (mso) reviewed the patient's outcome and determined impella 5.5 related events (including but not limited to low flow, sepsis etc.) are of a more direct issue and there is not enough information to exclude this device used as an associated factor to the patient's outcome. | The investigation of hemolysis, low pump flow, vf/vt/af/at, vascular damage - peripheral vessel, and infection/sepsis has been completed. the impella device was not returned for evaluation. hemolysis: the cause of the hemolysis issue was most likely patient condition related, based on data log review and given clinical details. low pump flow: the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. vascular damage - peripheral vessel: the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was provided. the relation between gi bleed and impella was unknown. infection/sepsis: because sufficient evidence has not been provided, the root cause of the infection cannot be determined. instructions for use: hemolysis: section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ infection/sepsis: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28789. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-28789 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-28789.
9038 | 1220648-2025-28790 | 22098997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22098997&pc=OZD | 05/07/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632657 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support and had bloody urine and lactate dehydrogenase of 2100. after escalating to an impella 5.5, the urine start to clear. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9039 | 1220648-2025-28791 | 22099223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099223&pc=OZD | 04/29/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 04/29/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634620 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experience limb ischemia prompting escalation in care to an impella 5.5 device. the patient presented with non-st segment elevation myocardial infarction eight days prior; non-compliant to medication. catheterization revealed left main, left anterior descending and circumflex arteries disease. protected percutaneous coronary intervention (pci) with an impella cp was planned but the patient left the hospital against medical advice. the patient would return 5 days later and presented with chest pain and increase breathing difficulty. the impella cp device was implanted for cardiogenic shock and pci implanting drug-eluting stents. at the end of the intervention, distal runoff showed no flow below impella access site. the impella cp was at performance level p-9 with minimal pulse pressure difference. right heart catheterization was performed, and swan-ganz catheter was left in place. the physician consulted for a higher level of care, and the patient was transferred out. the patient was with continued monophasic flow to right lower extremity (impella leg) and continued pressor/inotropes/low ejection fraction, and low pulsatility. the decision was made to escalate to impella 5.5 given the persistent cardiogenic shock. while on the impella 5.5 device the patient would experience further complications and expire after 10 days on impella 5.5 mcs. | [5 empty] | Medical safety reviewed the event and noted limb ischemia can lead to more complications down the line. however, it is the impella 5.5 related events that are of a more immediate/direct issue and cannot be excluded as an associated factor to the patient's outcome. refer to manufacturing report number 1220648-2025-28789 for the medical device report on the impella 5.5. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9040 | 1220648-2025-28792 | 22099241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099241&pc=OZD | 05/07/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/07/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627541 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during removal of the .018 guidewire, the wire got stuck in the pigtail of the impella cp. both the pump and guidewire were removed and a new pump was placed for planned support. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9041 | 1220648-2025-28793 | 22099366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099366&pc=OZD | 05/05/2025 | 07/01/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597969 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had episodes of non-sustained ventricular tachycardia. the patient was shocked three times. support continued. additionally, during support, there was a pump stop. attempts to manually restart the pump were unsuccessful. a new automated impella controller was attempted to no avail. the pump was removed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use. impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of vf/vt/af/at and pump stop has been completed. the impella 5.5 was returned for evaluation. pump stop: upon visual inspection of the pump, no abnormalities were observed. pump was electrical tested and all three motor cable lines were open. red handle of the pump was opened and all three motor cable lines were found to be necking and broken just past the glue joint inside catheter. data logs were reviewed and lt log shows alarm 115 pump stop alarm, with a motor current spike to 3000. pump was attempted to be restarted multiple times without success. clinical details noted that a pump stop occurred on day 8 of support and was unable to be restarted. the root cause of the pump stop was due to an electrical open of the motor cable wires within the red handle based on returned product and data log analysis. vf/vt: as the necessary information was not provided, the root cause of the vt/vf was not determined.
9042 | 1220648-2025-28794 | 22099615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099615&pc=PYX | 05/06/2025 | 06/19/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502659 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that at the beginning of impella rp flex placement, the physician had venous and pulmonary artery access with wire and was just about to start with the impella introducer kit when the patient coded and expired. per medical safety officer review: there is not enough information to exclude the impella device [guidewire] used as an associated factor in the patient's outcome. given the information at hand, access was made with the guidewire, used at the time of the patient's demise and no alternative cause for the event is identified as a likely cause of such event. information is limited and cause of death details are unnoted. | [5 empty] | A.1, a.4 and a.5 are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. per medical safety officer review: there is not enough information to exclude the impella device [guidewire] used as an associated factor in the patient's outcome. given the information at hand, access was made with the guidewire, used at the time of the patient's demise and no alternative cause for the event is identified as a likely cause of such event. information is limited and cause of death details are unnoted. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because patient unfortunately expired before device and the insertion was aborted.
9043 | 1220648-2025-28795 | 22099653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099653&pc=OZD | 05/05/2025 | 06/17/2025 |  | 05/28/2025 |  |  | 05/12/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569329 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ventricular Fibrillation | Embolism/Embolus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with decompensated heart failure in need of impella 5.5 support as a bridge to transplant. during support, it was noted of some suction episode which 1 albumin was administered. the patient also was on bed rest, and experienced ventricular fibrillation/ventricular tachycardia (vt) storm and went on extracorporeal membrane oxygenation support emergently. the patient is listed status 1 for heart transplant. it was noted that upon removal impella 5.5 pump, physician inspected the inlet and outlet and visualized what he thought was a thrombus on the inlet. the physician was suspicious that there may be thrombus because of events that occurred during the course of treatment. patient went into vt again. suction events before during and after this event occurred while the pump was in proper position according to echo. flows were not optimal according to physician at this time and had difficulty unloading the left ventricle and optimizing fluid status. patient was successfully transplanted, however, physician wanted the pump to be saved and sent back to company. | It was additionally reported that during support, it was noted of some oozing from jp insertion site. also, the patient continued to spike low grade fevers. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of vf/vt/af/at, infection/sepsis, access site bleeding, and low pump flow has been completed. there was only the cannula of the pump returned for investigation. thrombus was added based on the investigation. as the necessary information was not provided, the root cause of the thrombus was not determined access site bleeding: the clinical stated that there was oozing at the jp site. due to lack of clinical details provided in regards to why the bleeding occurred and how it was resolved, the root cause of the access site bleeding cannot be determined. low pump flow: there was a small piece of biomaterial found underneath the impeller. there was no biomaterial found in the purge gap. the data logs show that suction alarms were correctly triggered throughout the case. there was an increase of suction alarms on the 28th as well as the 5th during the reported events. the flows throughout support were slightly on the lower end at a few suction points. however, flows are suboptimal on 5/5 but slightly improve throughout the rest of the case. during this time period, the placement signal becomes very flat and there is a trend with flows. however, on 5/6 there is a spike in motor current causing the flows to increase from their low state but not become saturated. this is most likely due to a biomaterial ingestion causing additional resistance on the impeller in terms causing the increase in flows as your motor current remains slightly elevated. based on the clinical information and data log analysis, the root cause of the low pump flow throughout the majority of the case is most likely due to the patients condition. vt/vf and sepsis: as the necessary information was not provided, the root cause of the vt/vf and sepsis was not determined. b5 revise with additional information inadvertently omitted on the initial manufacturer device report number 1220648-2025-28795. h6 health effect - clinical code 1930 and 1888 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28795.
9044 | 1220648-2025-28796 | 22099771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099771&pc=OZD | 05/05/2025 | 05/28/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637962 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support developed gastrointestinal bleeding and heparin was held. additionally, the patient required intubation due to recurrent ventricular tachycardia. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9045 | 1220648-2025-28797 | 22099869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22099869&pc=OZD | 04/28/2025 | 06/17/2025 |  | 05/28/2025 |  |  | 04/28/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610431 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  | MW5169972 | Yes | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and bridge to transplant. approximately 13 days after start of the impella mcs, frequent pump stops "every 5 minutes" requiring restart were encountered. this continued through the next morning. the device, therefore, was explanted and replaced. following, a kink at the red plug was noted. the latest update received noted the patient, in stable condition, remains on the replacement impella 5.5 device for mcs while awaiting transplant. a voluntary medwatch was subsequently received, and additional information was noted as the patient was emergently taken to the operating room for the device exchange and the issue with the pump was assumed to be associated with the power wire damage within the driveline. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, entering cardiac output (impella cp with smart assist), section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella 5.5 with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of temporary pump stop and product damage (cannula kink) has been completed. the data logs show frequent electrical 119 pump stops on (b)(6) with no other changes in pump metrics. the returned product had a significant kink in the catheter near the red pump plug. electrical testing was conducted and no wires were open. the pump was able to run but when it was bent at the location of the kink, the motor current would increase and create a pump stop. the catheter was cut open and the jacket around the motor cable lines was kinked and the lines were stretched. the root cause of the temporary pump stop is most likely due to open electrical lines based on the clinical details provided, data log analysis and returned product. the root cause of the product damage (cannula kink) is most likely due to a kink in the catheter due to use as the pump was not fixated using 3-point fixation and caused the pump to kink upon patient movement. e4 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28797. h9 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28797.
9046 | 1220648-2025-28798 | 22100060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22100060&pc=OZD | 05/05/2025 | 06/27/2025 |  | 05/28/2025 |  |  | 05/05/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631227 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had a gastrointestinal bleed and blood products were given. anticoagulation was also withheld. additionally, there were suction events during ambulation. albumin was given to the patient and support continued. furthermore, there was moisture inside the purge filter with a brown substance. there were no issues with the pump/purge flows. the exterior had a tacky substance (presumably purge fluid). the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of purge leak ¿ pump, hemolysis, and low pump flow has been completed. the impella device was not returned for evaluation. purge leak - pump: data logs were reviewed, and there were no abnormalities to suggest a leak. throughout the entire case, there was a slight trend down in purge flows, rather than up. there were 3 purge system open alarms, however, they each coincided with the de-air procedure being done, and resolved within seconds. based on the clinical details reporting moisture build up in the sidearm retainer, the root cause of the purge leak was most likely a damaged sidearm. since product wasn't returned for investigation, the cause of the damage to the sidearm could not be determined. vascular damage - peripheral: since product wasn't returned and clinical details did not report details pertaining to the source/cause of the bleed, the root cause of the vascular damage could not be determined. low pump flow: data logs were investigated, and there was a large concentration of suction alarms around 14:00 on the reported date. the ps did also a trend down in the 24 hours leading up to the suction but recovered after. the suction alarms were triggering because motor current (mc) min values were too low compared to the placement signal (ps) value at the same time in each cardiac cycle. the optical signal was reliable at this time. that being said, the threshold for the alarm was barely being triggered, and motor current values were generally within specification for the p-level at this time. additionally, pump flows were well within specification on this date. this suggests that volume was not an issue, but rather the suction alarms were most likely occurring because the pump was being shifted into a unideal position because the patient was ambulating. this is further supported by the fact that the pump was repositioned and the alarms did not trigger when the patient ambulated later that day. this is not to say that there was a general positioning issue though, because suction alarms were not an issue when the patient was not ambulating and there were no positioning alarms. based on the clinical detail provided that the patient was ambulating at the time of the alarms, and mc/pump flow values were generally within specification, the root cause of the low pump flows was determined to most likely be improper position due to patient movement.
9047 | 1220648-2025-28799 | 22100315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22100315&pc=OZD | 05/06/2025 | 06/20/2025 |  | 05/28/2025 |  |  | 05/06/2025 | 05/28/2025 |  |  |  |  | 05/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1731299 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) would not recognize the purge disc. the aic was replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue h6 type of investigation, investigation findings, and investigation conclusions were erroneously entered on the initial manufacturer device report number 1220648-2025-28799. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28799. h10 narrative was erroneously entered on the initial manufacturer device report number 1220648-2025-28799, as the investigation was completed at time of initial report.
9048 | 1220648-2025-28708 | 22086651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22086651&pc=OZD | 05/01/2025 | 09/25/2025 |  | 05/27/2025 |  |  | 05/01/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637984 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant also reported that the impella alarmed multiple times that the pump is in the aorta. under echocardiography guidance, it was verified multiple times that the pump was in the correct position. | The complainant reported that a patient was implanted with an impella 5.5 and experienced bleeding at the impella access site. heparin was held. the next day heparin was started and it was noted that the patient coded twice over the night but return of spontaneous circulation was able to be returned quickly. three days later, the patient went into ventricular tachycardia after turning the transvenous pacemaker down to 50. the patient briefly coded and return of spontaneous circulation was achieved. the patient unfortunately had no bridge to left ventricle assist device or transplant. the family made the decision for comfort care three days later, and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: as per the clinical information, the root cause of the access site bleeding issue was improper anticoagulation management of the patient since the patient was on the blood thinners with high activated clotting time (act) values at implant which was 262. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump was not returned for investigation. the data logs show that the signal-to-noise ratio (snr), contrast, and lamp were decreased and light and gain were stable. the 5s parameters were stable otherwise throughout case. flows were stable at the same p-level with no changes to it. however, the clinical confirmed the pump was in good position. therefore, the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28708. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B2 selection of required intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28708.
9049 | 1220648-2025-28709 | 22086772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22086772&pc=OZD | 05/02/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658462 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was connected to an automated impella controller (aic) and skipped the flushing process immediately going to the start screen. the aic was changed and the same thing occurred. a new impella cp was opened and connected to the first aic and went through the appropriate steps without any issue. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9050 | 1220648-2025-28710 | 22086831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22086831&pc=OZD | 03/24/2025 | 06/03/2025 |  | 05/27/2025 |  |  | 03/26/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026624110 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day two of impella 5.5 support the pump stopped. attempts to restart were unsuccessful. the patient is stable. the impella was explanted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. after 2 days on support, pump stop occurred. attempts to restart it were unsuccessful. data logs were not recovered. the pump was returned and inspected which found extensive biomaterial wrapped around the impellor blades and blocking the outflow cage. additionally, a hole was found in the catheter at the 33cm mark with blood in the catheter and blood in the red handle. electrical testing on the returned pump showed electrical shorts. the cause of the pump stop was most likely a hole in the catheter leading to blood in the catheter and red handle causing electrical short based on returned product investigation. g1 reporting contact first and last name was erroneously entered on the initial manufacturer device report number 1220648-2025-28710.
9051 | 1220648-2025-28712 | 22086992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22086992&pc=OZD | 05/01/2025 | 06/23/2025 |  | 05/27/2025 |  |  | 05/01/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631329 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had an additional stitch added to the chest incision while in icu. it appeared that the stitches applied in the or weren¿t holding the incision closed completely. additionally, the patient with a history of ventricular tachycardia (vt), experienced vt while on support, requiring cpr, epinephrine and shock x1. the impella was pulled back about 0.5 cm under echo guidance. electrophysiologist was consulted and the plan was to interrogate implantable cardioverter-defibrillator and plan for vt ablation. it was further reported that the family decided to withdraw care, and the patient expired. there is not enough information to exclude the impella used as an associated factor to the event outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel, vf/vt/af/at, and positioning issues has been completed. product was returned for review. visual inspection found no issues on the pump or issues with repositioning unit. positioning issues: the data logs show positioning alarms throughout the case. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no abnormalities found on returned pump. vascular damage: the root cause of vascular damage was unable to be determined as limited clinical details were provided. vt: the root cause of vt was unable to be determined as limited clinical details were provided. d9 device returned to manufacturer date added.
9052 | 1220648-2025-28713 | 22087056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087056&pc=OZD | 05/05/2025 | 06/25/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629217 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was placed on a patient and percutaneous coronary intervention was completed. the patient was taken to the intensive care unit post case. an echocardiogram noticed effusion. emergent intubation and pericardial drain was placed with 180 cubic centimeters of blood removed. once the blood was removed, pulsatility returned. the patient was taken back to the operating room for pericardial window. at this time, upon transesophageal echocardiogram (tee), heart function was improved but apex was akinetic. pericardial window was completed. once the drain was placed, the physician removed the impella cp as the heart had recovered. once the removal was completed, a large blood formation around the heart was identified on tee. the physician returned to open the chest and explore where the blood was coming from. a perforation was noted at the apex of the heart with some darkened necrotic tissue identified. the hole was patched and the patient¿s chest closed. post case, the physician reviewed films and stated there was a silhouette sign behind the heart showing the patient had an effusion prior to impella placement. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ | The investigation of ventricle perforation has been completed. ventricle perforation: clinical details state that echocardiogram noticed effusion after transfer to intensive care unit post impella implant. patient was taken back to operating room for pericardial window. perforation was found at the apex of the heart with some darkened necrotic tissue identified. later, the physician reviewed films and stated there was a silhouette sign behind the heart showing the patient had an effusion prior to impella placement. methamphetamines was listed as a contributing factor. product was returned and evaluated. there was evidence of a cannula kink found. no other abnormalities were found. data log review did not show any abnormalities with positioning. there was no issue found with the impella during the case as the clinical details state that the perforation was present prior to impella implant according to review of imaging. the device functioned/operated to specification. d9 device returned to manufacturer date added.
9053 | 1220648-2025-28714 | 22087088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087088&pc=OZD | 05/02/2025 | 09/24/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580249 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 109 KG |  | White | Injury | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | Device Sensing Problem |  |  |  |  | Arrhythmia | Sepsis | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an impella 5.5 was implanted in a patient with non st elevated myocardial infarction. the pump had positional issues and was repositioned. the patient suffered from ventricular tachycardia and a gi bleed and was transfused multiple units of red blood cells. the patient was being assessed for possible infection as he was weaned from the impella. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs were reviewed and oscillating contrast failure observed relative to optical 5s failure pattern. pump was returned and optical 5s parameters were verified to be in normal range. light continuity test was done and light exited the sensor face showing no abnormalities. there was no problem found during the case relative to pump failure causing optical signal issue. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information h.6. medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28714. | The investigation of positioning issues and vf/vt/af/at has been completed. the returned compliant 5.5 pump visually inspected. cannula kink and catheter kink near outlet observed and was added as failure mode. clinical stated: on (b)(6) echo showed measuring at 3.7 cm. they had heavily diuresed him, and with his thickened left ventricle (lv) wall thought with low volume it was causing impella to interact with lv wall. positioning issues: the root cause of positioning issues most likely was patient condition related due thickened lv wall leading to low volume causing impella to interact with lv wall. failure mode product damage (cannula kink) since cannula kink was observed during product evaluation. product damage (cannula kink): the root cause of product damage (cannula kink) was cannula kink due to interaction with thickened lv wall. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. the investigation results for infection/sepsis and vascular damage - peripheral vessel has not been completed. should additional information become available a supplemental report will be filed. d9 device returned to manufacturer date added h6 health effect - clinical code 2067 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28714.
9054 | 1220648-2025-28715 | 22087115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087115&pc=OZD | 05/05/2025 | 06/26/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657495 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on the impella 5.5 for hemodynamic support. during support, it was reported that blood tinge purge fluid was leaking back into the purge filter and out around the red plug. purge pressure was running around 350 and the flow was between 18-25 milliliters. the patient was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. purge leak - pump: clinical details state that blood-tinged purge fluid was leaking back into the purge filter and out around the red plug. purge pressure was running around 350 mmhg and flow was between 18-25 ml. product was returned and evaluated. there was blood-tinged fluid residue in the red handle. there was also a puncture hole in the catheter found within 1 cm of the motor housing. the catheter jacket was removed and there was a puncture in the purge tubing just under the puncture mark found on the catheter. when purging the returned pump, there was purge fluid confirmed to be coming out of the puncture in the purge tubing. data log review showed purge pressure dropping from 450-500 mmhg to around 350 mmhg about an hour into the case. purge flow also increased from 10-15 ml/hr to 20-30 ml at this time. on day 4 of the case, purge pressure slightly increased and purge flow decreased below 20 ml/hr for the remainder of the case. the cause of the pump purge leak was damaged purge tubing as a puncture was found in the purge tubing inside the catheter just above the motor housing of the returned product. product damage (hole in catheter): the failure mode product damage (hole in catheter) was added because a hole in the catheter was found on the returned product. clinical details state that blood-tinged purge fluid was leaking back into the purge filter and out around the red plug. the pump was implanted with a direct aortic approach. product was returned and evaluated. there was blood-tinged fluid residue in the red handle. there was also a puncture hole in the catheter found within 1 cm of the motor housing. clamp marks were also found between the 16 -30 cm catheter markings. the cause of the hole in the catheter was most likely use-related as a puncture hole was found in the catheter just above the motor housing, which led to blood to ingress through the catheter to the red handle. d9 device returned to manufacturer date added.
9055 | 1220648-2025-28716 | 22087277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087277&pc=PYX | 05/05/2025 | 06/19/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540898 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Material Fragmentation | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and during support, the pump position changed from 38.5 to 41.5 centimeters. the patient developed hemolysis. the pump was repositioned and the patient's urine cleared. additionally, the impella rp flex sheered off the distal peripherally inserted central catheter upon explant. the doctor suspected it may have gone into the impella rp flex cannula. the patient was stable with no harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of hemolysis, positioning issues, and removal issues has been completed. no product was returned for analysis. thrombus failure mode added due to ingestion pattern observed. hemolysis: clinical detail mentions the pump position changed from 38.5 to 41.5cm. the pump was repositioned and the patient's urine cleared. the data logs do not show any positioning related alarms. the data logs show a motor current spike with speed deviation and placement signal shift on 05/04 before reported hemolysis. the root cause of the hemolysis was most likely due to biomaterial as evidenced by the motor current spike in the logs before reported hemolysis. the root cause of the positioning issue was not determined as limited clinical information was provided on why pump was repositioned. removal issues: rp flex sheared off distal picc catheter upon explant. ¿the root cause of the removal issues was most likely due to interaction with other devices as the picc line was sheared by pump inlet during explant. as the necessary information was not provided, the root cause of the thrombus was not determined.
9056 | 1220648-2025-28717 | 22087347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087347&pc=OZD | 11/25/2024 | 05/27/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  | 10/31/2025 |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024417555 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp experienced a gi bleed and ventricular tachycardia. the patient was treated with epinephrine, amiodarone, and defibrillation. the impella cp was explanted and replaced with an impella 5.5. the patient recovered with no sequelae. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
9057 | 1220648-2025-28718 | 22087420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087420&pc=OZD | 05/05/2025 | 06/30/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella cp pump placed via the axillary artery for cardiogenic shock. there was an access site bleed at the axillary that prompted two units of blood to be given to the patient. the next day the family/medical team made the choice to withdraw care and the patient expired. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. it was noted that during the time of bleeding, the activated clotting time (act) value was 233 which is above the recommended range during the use of impella. the root cause of access site bleeding was determined to be anticoagulation management because the act were not maintained in the recommended therapeutic range during the time of bleeding. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9058 | 1220648-2025-28719 | 22087504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087504&pc=OZD | 05/02/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634719 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support had a left ventricle thrombus. the staff was concerned about the thrombus and started anticoagulation until the pump was explanted and the patient was placed on a balloon pump. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9059 | 1220648-2025-28720 | 22087518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087518&pc=OZD | 05/04/2025 | 06/26/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636207 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was also reported that the patient had a run of ventricular tachycardia (vt). the team believes that arrhythmia is related to prolonged qt correlating with tikosyn dosing overnight. medication dosage was reduced. | The user facility reported that an impella 5.5 anti-contamination sleeve torn at the distal end. it had come completely detached. it was cleaned and tegaderm was placed. the patient continued on support and no harm was reported. |  |  |  |  | The investigation of product damage (sterile sleeve damage) and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28720. b2 outcome selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28720. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28720. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28720. h6 health effect - clinical code 2095 and health effect - impact code 4641 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28720. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9060 | 1220648-2025-28721 | 22087591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087591&pc=OZD | 05/03/2025 | 07/08/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645308 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed some oozing at the pump access site after implantation of an impella cp. manual pressure was applied and blood products were given to manage the condition. the following day, an alarm appeared indicating a leak in the purge system. the purge fluid was increased and the issue resolved. the patient survived. | [5 empty] | A2 and a4 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and purge leak ¿ pump has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding major was most likely a use-related issue since the repo sheath was not properly placed with angle matching per clinical details. purge leak - pump: the root cause of the purge leak - pump was not determined since no product/logs were returned.
9061 | 1220648-2025-28722 | 22087705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087705&pc=OZD | 05/03/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565606 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted in a patient after they experienced several episodes of ventricular tachycardia. after implantation a femoral perforation and bleeding was observed. the patient developed and hematoma and the vessel was treated by a covered stent. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 and a5, race, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high-risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9062 | 1220648-2025-28723 | 22087714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087714&pc=OZD | 05/05/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570552 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on day three of impella 5.5 support and was tachycardic. on day five of impella 5.5 support, the patient had cardioversion. additionally, on day nine of support axillary site was noted to be bleeding. the plan was to take the patient for an axillary washout. one unit of packed red blood cells was given. additionally, the impella pocket was explored and clots were found and removed. later, there was exploration of the axillary artery and evacuation of a hematoma. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9063 | 1220648-2025-28724 | 22087808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22087808&pc=OZD | 05/02/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594143 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 86 KG |  | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5 the patient suffered reoccurring ventricular standstill. when this occurred suctioning took place, and one unit of packed red blood cells was given to the patient. the patient survived and was successfully weaned from the impella. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9064 | 1220648-2025-28725 | 22088026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088026&pc=OZD | 05/03/2025 | 06/26/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502210 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was attempted but was unsuccessful so implantation of an impella 5.5 was attempted via axillary. the impella 5.5 passed an anastomosis, but stopped just a little after. implantation of the impella 5.5 was aborted and a new impella cp was implanted via axillary instead. | [5 empty] | A4 and a5, ethnicity, are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. reported having issues delivering impella 5.5. it passed the anastomosis but stopped just a little bit after. impella 5.5 was aborted and cp was implanted successfully. the patient had small vessels. impella 5.5 was returned and no damages or abnormalities were found. the root cause of the delivery issue was most likely patient condition related since the patient had small vessels.
9065 | 1220648-2025-28726 | 22088164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088164&pc=OZD | 05/03/2025 | 06/25/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560093 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER SN (B)(6) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal unreliable alarm appeared during impella support. both the waveforms and positioning appeared to be correct at the time of the failure. the pump was replaced, but the issue persisted. the automated impella controller was also then replaced, and the issue resolved. | [5 empty] | The investigation of optical signal issue has been completed. clinical details state that a placement signal not reliable (psnr) alarm was triggered after pump insertion. the pump was removed as a result after the procedure was performed. the replacement pump had the same issue, which was eventually resolved by replacing the console. the first pump used (sn (b)(6)) was returned and evaluated. no abnormalities were found with the 5s parameters or laser continuity test. there were no kinks or biomaterial found. the pump was run in the lab with no psnr alarms or optical parameter abnormalities. no problem was found with the returned pump. upon data log review and returned pump evaluation, it is apparent that the console is the potential cause of the issue. please refer to manufacturer device report number 1220648-2025-28728 for an investigation into the console. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9066 | 1220648-2025-28727 | 22088224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088224&pc=OZD | 05/06/2025 | 06/25/2025 |  | 05/27/2025 |  |  | 05/06/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644327 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed an impella cp for support of a patient admitted in with gastrointestinal issues and was being worked up for surgery. the impella cp was placed but during the support there was suction treated by pump position assessement and giving blood/volume. the patient had the surgical procedure and the pump was weaned and explanted. | [5 empty] | The investigation of low pump flow has been completed. the returned complaint cp pump visually inspected. no biomaterial observed. no cannula kink or broken blade found. impella cp connected to automated impella controller (aic) to to reproduce reported low pump flow issue at it run for more than 10 minutes. flow was with specified limit at p-9. no low flow can be reproduced and no other issue observed. data log reviewed: no motor current (mc) spike observed outside of p-level changes. no positioning issue or placement signal trend observed. suctions occurred on 06-may as reported. low pump flow: the cause of reported low pump flow issue cannot be determined since low flow cannot be reproduced in investigation lab and no evidence found. access site bleeding - minor: impella placed with routine fashion and easy way access. bleeding from site after peel-away sheath removed and repo sheath placed. manual pressure and femstop placed with good control. the cause of access site bleeding - minor most likely was lack of vessel recoil related. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Device returned 13-may-2025, and investigation is underway. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28727. h10 narrative erroneously stated the "device was not returned" on the initial manufacturer device report number 1220648-2025-28727.
9067 | 1220648-2025-28728 | 22088396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088396&pc=OZD | 05/03/2025 | 06/30/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1079576 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal unreliable alarm appeared during impella support. both the waveforms and positioning appeared to be correct at the time of the failure. the pump was replaced, but the issue persisted. the automated impella controller was also then replaced, and the issue resolved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue has been completed. based on the data and device analysis the root cause for the placement signal not reliable (snr) alarms was due to the defective optical bench with low snr.
9068 | 1220648-2025-28729 | 22088398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088398&pc=OZD | 05/06/2025 | 07/03/2025 |  | 05/27/2025 |  |  | 05/06/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614416 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a console undergoing a preventative maintenance and the team discovered the automated impella controller to have a flickering display. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Additional information provided in h3, h6 and h10. the investigation into the reported issue has been completed. console logs from the reported day of event do not contain any information relevant to this failure mode. the console was sent back to field service for further investigation and was tested. the reported issue could not be reproduced, and no flickering was observed on the lcd screen. further testing was conducted with both the 4-in-1 assembly and the lcd and no problems were observed. furthermore, a visual inspection of the internal circuitry revealed no abnormalities. the reported issue of the "aic display is flickering" was not determined due to inability to reproduced.
9069 | 1220648-2025-28730 | 22088672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088672&pc=OZD | 05/03/2025 | 09/24/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635359 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 44 year old male patient with an impella 5.5. during support, a nurse called in for continuous suction alarms however placement had been confirmed. cardiothoracic surgeon requested the alarm audio be disabled. suction and placement alarms continued and the position was verified via echo. all alarms were disabled. flow was within normal limits. patient bridged to transplant. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the clinical details state that the patient was having suction alarms on 5.may and position was verified. there were no other troubleshooting details provided. the data logs show that there instances of suction triggered throughout the case as well as impella position unknown alarms. the instances of the suction were triggered during the time frame of the ps declining. the suction alarms met triggering criteria but are not true events of suction as your ps was reading abnormal values, falsely triggering those alarms. flows were stable. there was no problem found in regards to the low pump flow based on data log analysis. d6a implant date was erroneously entered on the initial manufacturer device report number 1220648-2025-28730. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the data logs show that the placement signal (ps) was stable for about, 562 hours until it began to drift down for about 15 hours until it went out of range and triggered ps unreliable. the 5s parameters remained relatively stable. the pump was returned cut, so it was unable to run. laser continuity did not reveal any abnormalities on the catheter until the cut portion. based on the data logs, the root cause of the optical signal issue is most likely due to sensor durability wear. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d6b updated with additional information: d9 device returned to manufacturer date added. h3 device evaluated by manufacturer updated as the initial manufacture report stated no product return which is incorrect. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28730. d10 concomitant medical products and therapy dates updated.
9070 | 1220648-2025-28731 | 22088689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088689&pc=OZD | 05/05/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  | 11/30/2025 |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434162 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO, IABP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient had an impella 5.5 pump implanted to support a bridge to left ventricle assist device after escalation from intra-aortic balloon pump and extracorporeal membrane oxygenation. the impella 5.5 generated high purge pressure alarms which were resolved with a new cassette and adding tissue plasminogen activator to the purge. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9071 | 1220648-2025-28732 | 22088900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088900&pc=OZD | 05/06/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/06/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607579 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient's leg became numb and cold. the impella cp was explanted and flow was restored post explant. no patient harm was noted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9072 | 1220648-2025-28733 | 22088947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22088947&pc=OZD | 05/03/2025 | 06/26/2025 |  | 05/27/2025 |  |  | 05/04/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed hemolysis during impella cp support. it was noted that the patient had a small left ventricle (lv) and lactate levels were rising. the patient was escalated to impella 5.5 support in response to the condition. patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. ifu states: section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: clinical states that patients labs were hemolyzing and doctor decided to reposition the pump after found out the pump was shallow in the left ventricle (lv). the pump was repositioned which cleared the urine. pump was not returned for investigation. data logs were investigated, and "impella in aorta" alarm was observed. placement signal was mildly trending down and then stabled out. no motor current (mc) spikes outside p level changes, no suction alarms. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device as mentioned in clinical that repositioning the pump cleared the urine.
9073 | 1220648-2025-28734 | 22089089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089089&pc=OZD | 05/06/2025 | 06/23/2025 |  | 05/27/2025 |  |  | 05/06/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626223 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support for a high-risk percutaneous coronary intervention patient. when the impella cp was started on auto it was followed by suction alarms and unable to get flows above 1.3 liters per minute. the medical doctor stated there was a kink in the cannula. there were attempts to pull the impella c p back however, the kink was still noted. there was an attempt to reposition to get the kink out of the cannula but unsuccessful. ultimately, it was decided to replace the impella cp. the patient remained stable. | [5 empty] | The investigation of product damage (cannula kink) and low pump flow has been completed. product damage (cannula kink): clinical mentioned that low flows were observed at p9 and a kink was observed under fluoroscopy. pump was returned for investigation and a kink was confirmed near the center of the cannula, towards the motor housing side. therefore, the root cause of the product damage issue was a kinked cannula confirmed during the product investigation but the reason for the kink is unknown. low pump flow: clinical mentioned that low flows were observed at p9 and a kink was observed under fluoroscopy. pump was returned for investigation and a kink was confirmed near the center of the cannula, towards the motor housing side. data logs were investigated and there was suction alarms and low flows were confirmed. the pump ran for less than 5 minutes and pump was explanted. therefore, the root cause of the low pump flow issue was a cannula kink confirmed during the product investigation but the reason for the kink is unknown. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9074 | 1220648-2025-28735 | 22089344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089344&pc=OZD | 05/05/2025 | 07/16/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630897 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Mitral Perforation | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and exhibited access site bleeding. hemostasis was achieved by repositioning the sheath. support was continued. also, there was low pump flow and the patient's plasma free hemoglobin spiked. the low pump flow error resolved without intervention. the patient then experienced mitral valve rupture. it is unclear if the rupture was a result of impella use or an acute myocardial infarction. also, a bleed occurred which appeared to come from the impella sideport. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿. section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation of heart valve damage, hemolysis, access site bleeding, and low pump flow has been completed. heart valve damage: the root cause of the heart valve damage was not determined as limited clinical information related to failure was provided. hemolysis: the returned product was inspected and there were no physical abnormalities related to hemolysis such as biomaterial wrapped around the impeller or damaged impeller blades. the data logs from the case do not show any motor current spikes or positioning issues. clinical details note suction alarms pfhgb spiked to 300 but decreased to 13 within 24h. clinical details do not confirm pump out of position. the pump was hemolysis tested and passed. the root cause of hemolysis was not determined as there were no data log abnormalities, abnormalities on returned product and pump passed hemolysis testing; limited clinical information was provided. access site bleeding: ¿clinical mention a bleed with no apparent focus (large volume of bright red blood in the bed. the root cause of the access site bleeding was not determined as limited clinical information was provided. low pump flow: the returned pump did not have any biomaterial or cannula kinks. the data logs confirm suction alarms. there were no motor current spikes or position related issues. clinical mention suction alarms (diastolic) and this was associated with intravascular hypovolemia and sirs. the pump was tested in the lab and low pump flow did not reproduce. the root cause of the low pump flow was most likely patient condition as no log abnormalities were observed, no issues or reproduction on the returned product and clinical details mention suction was associated with intravascular hypovolemia and sirs. d9 device returned to manufacturer date added.
9075 | 1220648-2025-28736 | 22089477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089477&pc=OZD | 05/03/2025 | 07/02/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that difficulty was encountered during attempted delivery of an impella cp pump. the pump was initially inserted via right femoral artery; however severe calcification of the artery was noted at the bifurcation. the impella would not advance further and attempts were then made via left femoral artery. the guidewire was able to advance; however resistance was still noted and the decision was made to abort the implant. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of delivery issue was most likely patient condition since implant was aborted due to calcified artery at bifurcation and torturous anatomy.
9076 | 1220648-2025-28737 | 22089548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089548&pc=OZD | 05/04/2025 | 06/25/2025 |  | 05/27/2025 |  |  | 05/04/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626302 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after a second impella cp was implanted in a patient, a large hematoma grew and the patient hemodynamics became unstable. the patient underwent a vascular cut down. two purse strings were placed to help control the bleeding, and eight units of packed red blood cells were transfused. when attempts were made to advance impella flows dropped significantly with suction alarms, so the device was left in a shallow position. the patient was successfully weaned off the pump and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel, positioning issues, and low pump flow has been completed. the pump was not returned. data logs were not recovered. the cause of the reported issues was not determined since product was not returned and insufficient clinical details.
9077 | 1220648-2025-28738 | 22089647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089647&pc=OZD | 05/07/2025 | 06/25/2025 |  | 05/27/2025 |  |  | 05/07/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614177 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that several episodes of suction alarms occurred while an impella 5.5 was implanted in a patient. the alarms were resolved with fluid bolus. the patient also did not tolerate weaning of extracorporeal membrane oxygenation and was shocked twice in response to ventricular tachycardia. the patient was also treated with norepinephrine and neo-synephrine. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and low pump flow has been completed. the pump was not returned for investigation. the data log shows several suction alarms throughout the case run. pump flow trended low with an associated drop in the placement signal while running on a steady p-level. there was no motor current spike, nor was any positioning issue noted in the data log. clinical information suggests that the patient was on extracorporeal membrane oxygenation (ecmo) support, and the suction condition resolved with a fluid bolus. low pump flow: the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined.
9078 | 1220648-2025-28739 | 22089885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089885&pc=OZD | 05/03/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645573 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported cp pump was placed to support a patient admitted in for surgery and preoperatively the cp was placed. after 34 hours of support the patient went into pulseless vt and the patient expired when care was withdrawn. the patient expired on the cp support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9079 | 1220648-2025-28740 | 22089889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22089889&pc=OZD | 05/04/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/04/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625124 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp and developed a gastrointestinal bleed. the patient was transfused one unit of packed red blood cells and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9080 | 1220648-2025-28741 | 22090018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090018&pc=OZD | 05/04/2025 | 06/27/2025 |  | 05/27/2025 |  |  | 05/04/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658064 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there were continuous suction alarm. around the same time, the patient had a hematoma with oozing. manual pressure was held and two units of packed red blood cells were given to the patient. support continued until weaning and explant was successful. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding was most likely patient condition related since the access site was initially stuck several times prior to implanting the pump and since multiple access sites were bleeding. low flow: data logs show suction and low flows in the early morning of 4-may-2025. no motor current spikes, placement signal (ps) trend or other log pattern identified. the root cause of the low pump flow was most likely patient condition related since the patient was on venoarterial extracorporeal membrane oxygenation (va ecmo) and patient was also coughing experiencing low flow on all devices. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9081 | 1220648-2025-28742 | 22090019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090019&pc=OZD | 05/03/2025 | 06/24/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620134 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient was found to not have obstructive disease. the decision was made to place impella cp for hemodynamic support. during support, the patient was still having atrial fibrillation with rapid ventricular response. levophed was held and they were using neo-synephrine for blood pressure support the rate was improving with adjustment. also the patient was given one and a half liter of fluids for hemolysis and declining renal function. they also reported that heparin was on hold due to bleeding at the site. also the patient was given albumin and a 500 milliliter bolus to resolve suction. additionally, performance level was decreased due to ventricular ectopy and concerns for suction/hemolysis. the patient remained on support for 7 days in total. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis". "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure." | The investigation of access site bleeding, hemolysis, vf/vt/af/at, and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding - major: the root cause of the access site bleeding was most likely due to anticoagulation management as heparin was paused and bleeding was able to resolve without any other troubleshooting. hemolysis: data logs were reviewed and suction alarms occurred only on first day of support. no motor current spikes or persistent positioning alarms were noted in logs. clinical details noted that the patient was experiencing red urine after pump insertion. patient received multiple rounds of fluid and boluses and red urine cleared up. the root cause of the hemolysis was most likely patient condition related as data logs did not show any signs of ingestion or improper positioning of the pump and clinical details noted that additional fluids help reduce red urine. low pump flow: data logs were reviewed and the log showed suction alarms occurred on first day of support on (b)(6) 2025 with slight increase in ps. pump flows prior to suction alarms showed flows lower than set p-level of p-9. when pump was turned down to p-8, pump flows were within normal p-level range for a few days before pump flows again dropped below set p-level range on (b)(6) 2025. when pump was turned down to p-7 on (b)(6) 2025, pump flows were below set p-level range for reminder of data logs. no motor current spikes noted at time of low flows. clinical details noted that patient had suction alarms occurring and fluids were given and suction alarms resolved. the root cause of the low pump flow was most likely patient condition related as clinical details noted that additional fluids resolved suction alarms. vf/vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. b2 outcome death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28742. b5 patient information regarding outcome of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28742. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28742. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28742.
9082 | 1220648-2025-28743 | 22090124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090124&pc=OZD | 05/03/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574817 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 using best practices. during support, it was noted that the ventricular tachycardia requiring one shock and required reintubation. additionally, the patient had a gastrointestinal bleed. one unit of blood was given overnight, in addition to another unit given during the day along with two units of plasma. it was noted that the patient was experiencing intestinal issues. however, a full body computed tomography scan did not reveal any significant findings. they were evaluating for liver versus bowel/gastric ischemia due to persistent lactate spill. ultimately, care was withdrawn due to ischemic bowel and stomach issues and the patient expired shortly after support was ramped down. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9083 | 1220648-2025-28744 | 22090281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090281&pc=OZD | 05/03/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593635 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on a second impella cp pump and cardiopulmonary resuscitation (cpr) was on going. the impella cp at p8 had a constant suction alarm. p level was weaned until suction resolved. only able to flow 2.0 liters per minute with the impella cp due to suction at higher p level. team blousing fluids and verifying the position. no return of spontaneous circulation with holding cpr. systemic pressures were not adequate despite impella cp support and advanced cardiac life support medications. the doctor made the decision to stop resuscitation efforts. the patient expired in the procedural area. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9084 | 1220648-2025-28746 | 22090499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090499&pc=OZD | 05/03/2025 | 06/19/2025 |  | 05/27/2025 |  |  | 05/03/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had an impella cp placed to support a patient admitted in with acute myocardial infarction. the patient had a thoracentesis done and this meant an adjustment to the anticoagulation, but subsequent to that there were purge alarms that prompted tissue plasminogen activator to be added to the purge. the next day, the pump was out of position and was adjusted. the pump remained on for support for 6 days in total. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of high purge pressure, positioning issues, and hemolysis has been completed. no product was returned for analysis. thrombus was added based on the investigation findings. high purge pressure: the data logs show a motor current spike followed by a rise in purge pressure and low purge flows. the pump flows became saturated indicating biomaterial ingestion and position on the impeller. the flows remain elevated and logs show another purge pressure rise after attempted troubleshoot (change p-level from p9 to p0 to p9). clinical mention tissue plasminogen activator (tpa) was added to the purge which resolved the issue. the root cause of the high purge pressure was most likely biomaterial as evidenced by the ingestion pattern observed in the data logs. as the necessary information was not provided, the root cause of the thrombus was not determined. the root cause of the positioning issue was most likely due to a use-related issue of positioning the pump too deep inadvertently by physician as evidenced by clinical details. the data logs show an ingestion with motor current spike about 24 hours before reported hemolysis. there were no suction alarms or improper positioning alarms in the logs around the time of reported hemolysis. the root cause of the hemolysis was most likely due to biomaterial as evidenced by motor current spike in data logs before reported hemolysis as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9085 | 1220648-2025-28745 | 22090505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090505&pc=OZD | 05/05/2025 | 06/25/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604923 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 30 year old patient placed on impella 5.5 due to cardiogenic shock. during support, they reported suction events, however both events resolved. additionally, it was reported that the pump was explanted due to hemolysis. patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and low pump flow has been completed. returned complaint 5.5 (cut) visually inspected. biomaterial observed in the outflow. no cannula kink, sharp edge or broken blade found. clinical stated: during support they reported suction events. suctions resolved and support continued and pump was explanted for de novo hemolysis. data log not available to review. hemolysis: the cause of the hemolysis was not determined since data logs were not returned and insufficient clinical details were provided. it was unknown if the biomaterial in the returned pump was due to an ingestion event. low pump flow: the cause of the low pump flow was not determined since data logs were not returned and insufficient clinical details were provided. it was unknown if the biomaterial in the returned pump was due to an ingestion event.
9086 | 1220648-2025-28747 | 22090684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090684&pc=OZD | 02/18/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  | 11/30/2025 |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432612 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed's long-term outcome and quality indicator (loqi) impella registry that a patient was implanted with an impella cp and experienced atrial arrythmia "possibly" related to the pump. the loqi states: "after pci with impella patient was found to be in new onset atrial fibrillation with rvr via ekg. patient was started on an amiodarone drip transitioning to an oral dose. plan was for the patient to undergo a cardioversion with anesthesia on (b)(6) 2025 but patient expired the day before due to cardiac arrest." there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9087 | 1220648-2025-28748 | 22090743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090743&pc=OZD | 05/05/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/05/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608917 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp. after explant of the pump pulses were absent in the patient's foot. an angiogram revealed slow flow down the superficial femoral artery due to a thrombus in the profunda. intervention was performed to restore blood flow to the foot. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9088 | 1220648-2025-28749 | 22090791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22090791&pc=OZD | 05/02/2025 | 05/27/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  | 11/30/2025 |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425839 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | Vascular Dissection | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed for an acute myocardial infarction/cardiogenic shock patient. after the impella cp was placed, the patient consistently went into ventricular tachycardia and ventricular fibrillation (vt/vf) requiring a total of 17 shocks. the support was for 19 hours. the pump access site had a bleed that required three perclose and they then discovered a dissection. the dissection was treated by stenting. afterwards the bleed had resolved and pulses were present. the patient was in stable condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9089 | 1220648-2025-28750 | 22091084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22091084&pc=OZD | 05/01/2025 | 06/23/2025 |  | 05/27/2025 |  |  | 05/02/2025 | 05/27/2025 |  |  |  |  | 05/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657984 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic |  | Malfunction | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support when there was leaking of blood at the interface of the orange ring and the white ¿wing¿ of the introducer sheath when single access with 7fr companion sheath introduced at the ¿2 o¿clock¿ position, unresolved with manipulation. companion sheath removed, and reintroduced at 10 o¿clock position with resolution of leaking from prior side of sheath wing. hemostasis achieved in this position. the impella was successfully weaned and explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of introducer damage ¿ mechanical and access site bleeding has been completed. the clinical details state that there was blood leaking at the interface of the orange ring and white "wing" of the introducer sheath. the product was returned and the introducer was leak tested. a leak was noted to be coming out from under the orange valve. it leaked around 22 mmhg which lower than the 180 mmhg specifications the root cause of the introducer damage is most likely due to a damaged valve based on clinical details and returned product testing. the root cause of the access site bleeding is most likely due to a damaged valve based on the clinical details and returned product testing. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9090 | 1220648-2025-28671 | 22077206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077206&pc=OZD | 04/28/2025 | 06/27/2025 |  | 05/23/2025 |  |  | 04/28/2025 | 05/23/2025 |  |  | 11/30/2025 |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433138 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | Ischemia | Hematuria | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs. after the implant, same day, the patient had suction alarms. the performance level was turn down to rectify the alarm. the patient was given fluid and continued to monitor. thereafter, a small amount of blood was in the patient's urine, along with the patient now having a cold leg and inability to feel pulses (that were present earlier). a discussion was made to remove the impella cp device at bedside and use a femstop to create hemostasis. subsequently, the impella cp was removed, and at some point, thereafter, cardiopulmonary resuscitation had been initiated with return of spontaneous circulation achieved. the next steps for the patient were noted to be unknown. however, the patients outcome noted the patient had eventually expired after a total of approximately five hours on impella mcs. | [5 empty] | The investigation of low pump flow, limb ischemia, and hemolysis has been completed. the impella device was not returned for evaluation. low pump flow: data logs confirm suction alarms and low flows even after reducing the p-level. flows 2.8l/min at p7. once suctions alarms were off, p-level increased to p8 and still low flows of 3.0l/min were observed. placement signal was also seen to be trending upwards after suction events. therefore, the root cause of the low pump flow issue was not determined since insufficient clinical information was provided and data logs were inconclusive. limb ischemia: due to insufficient clinical details and inconclusive data log investigation, the root cause of limb ischemia was not determined. hemolysis: data logs were investigated. ¿impella in aorta¿ alarms were observed at 10:10pm right before the pump was explanted. no suction or motor current spikes around that time. therefore, due to insufficient clinical details and inconclusive data log investigation, the root cause of the hemolysis issue was not determined. | There is not enough information to exclude the impella related events as an associated factor. very limited details surrounding the patient¿s demise. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9091 | 1220648-2025-28672 | 22077306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077306&pc=OZD | 05/16/2025 | 07/14/2025 |  | 05/23/2025 |  |  | 05/16/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390016 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported there was an automated impella controller (aic) alarm when the pump was attempted to be started. the alarm stated "self-test failed, replace the control unit immediately". percutaneous coronary intervention (pci) was performed while supporting with intra-aortic balloon pump (iabp) until a replacement aic arrived. the control device arrived immediately after pci, and the iabp was replaced with an impella cp. the patient is hemodynamically stable. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. device analysis: the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9092 | 1220648-2025-28673 | 22077393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077393&pc=OZD | 05/02/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649449 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted in a patient with st-elevated myocardial infarction and left heart catheterization. the patient experienced ventricular fibrillation that progressed to pulseless electrical activity. the patient was treated with inotropes and vasopressors, cardiopulmonary resuscitation, and defibrillation but ultimately expired on support. additional information was received. the patient expired on support so the pump is not available. the patient had ventricular issues to pulseless electrical activity prior to the impella as well as after. the patient was in e shock. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9093 | 1220648-2025-28674 | 22077429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077429&pc=OZD | 01/15/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561478 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry, that a patient endured ventricular tachycardia during impella 5.5 support. the condition was believed to have a possible relationship to the device used. lidocaine, amiodarone, and cardioversion were necessary to treat the vt. the patient recovered from the event. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9094 | 1220648-2025-28675 | 22077533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077533&pc=OZD | 05/01/2025 | 06/17/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637142 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 61-year-old female patient had an attempted impella cp implant and the implant was aborted due to patient anatomy. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella insertion aborted insertion due to patient anatomy. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9095 | 1220648-2025-28677 | 22077704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077704&pc=OZD | 05/02/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced reduced flow to the lower extremity. the impella was slowly weaned and explanted. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.
9096 | 1220648-2025-28678 | 22077743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077743&pc=OZD | 05/03/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559935 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient was on impella cp support and had oozing noted to dressing. one unit of blood was administered. the patient was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9097 | 1220648-2025-28679 | 22077856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077856&pc=OZD | 05/02/2025 | 06/18/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123079 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the soft buttons on an automated impella controller (aic) failed. the aic was removed and replaced to continue patient support. | [5 empty] | The investigation of automated impella controller (aic) - kbd/rotary knob issues has been completed. the console logs show alarm 420 (complete the procedure) at 20:11:39, during purge bag change, coinciding with high guard level (touch detected outside soft-key area), which is consistent with attempt to perform bag change and either reacting to, or inadvertently causing momentary interruptions in soft-key operation. some keypresses were intermittently detected (e.g. open 'display' menu). prior to pump disconnection (in order to transfer to backup aic) the performance level had not been turned down to p-0. console was tested, but the issue was not reproduced, and kbd was fully responsive. it is possible, however, that a liquid spill (e.g. purge fluid) on the front of the aic (glass kbd) caused interruptions in soft-key detection or erroneous soft-key detection, associated with ¿high guard level¿ (touch detected outside soft-key area). the root cause was not determined due to lack of sufficient evidence. d1 brand name was erroneously entered on the initial manufacturer device report number 1220648-2025-28679. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-28679. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9098 | 1220648-2025-28680 | 22077956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077956&pc=OZD | 05/01/2025 | 06/24/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576447 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during standard impella insertion on a patient, fluoroscopy showed that the impella position was abnormal. after troubleshooting the position, a decision was made to remove the device and restart the insertion process. once again impella orientation seemed abnormal and a conclusion was made that the cause was the patient's valve anatomy. wire was removed and impella was increased to p4. peel away sheath was removed and repo sheath was advanced. forward tension sutures were applied and the impella was locked in place. the patient had minimal deplorable flow to the lower extremity. during support, the patient received a total of three units of blood, two units of fresh frozen plasma, and two units of cryoprecipitate. the patient then developed a gastrointestinal bleed. the impella and extracorporeal membrane oxygenation sites were both oozing. the patient was currently not on any systemic anticoagulation. it was noted that the patient continued to have persistent bleeding from the impella site despite reapplying forward tension sutures, holding manual pressure, and redressing the site with angle matching. the patient received an additional six units of blood, four units of platelets, two units of cryoprecipitate, and five units of fresh frozen plasma in the last 16 hours. the patient was taken to the operating room for a washout at the impella site. left groin cutdown and primary repair of femoral artery was done. the impella site was still oozing and thromboelastography was abnormal. mattress suture was applied and new forward tension sutures were used which seemed to resolve the issue. the patient was escalated to an impella 5.5. | [5 empty] | The investigation of vascular damage - peripheral vessel, limb ischemia, positioning issues, and access site bleeding has been completed. the impella device was not returned for evaluation. vascular damage - peripheral 1: clinical details reported that at the same time the patient was bleeding from the impella and ecmo access sites on (b)(6) 2025, a gi bleed was reported. no further details were provided. since insufficient clinical details were provided regarding the complication and product wasn't returned for investigation, the root cause of the vascular damage could not be determined. positioning issues: clinical details report that during case start, the team noted that the position of the pump in the lv was abnormal. they were unable to troubleshoot, so they removed the pump and re-inserted, but the orientation of the pump continued to be abnormal. the team came to the conclusion that the patient's aortic valve was causing the complication, and left the pump in that position. data logs were reviewed and there was a single position in aorta alarm at the end of the case, presumably during pump explant. based on the clinical details provided that the patient's valve anatomy was causing the issue, the root cause of the positioning issue was determined to most likely be patient condition. access site bleeding: based on the clinical detail that the patient was bleeding from all access sites and had a low platelet count, the root cause was determined to most likely be patient condition. limb ischemia - reversible: clinical details report that on the first day of support, it was noted that the patient had minimal deplorable flow to the impella access site leg. of note, the access vessel (lcfa) was noted to be small during access, so an antegrade sheath was used. a patient update on (b)(6) 2025 noted that pulses had returned. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. vascular damage - peripheral 2: since insufficient clinical details were provided regarding when or how the vascular damage occurred and product wasn't returned for investigation, the root cause of the vascular damage could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9099 | 1220648-2025-28681 | 22077963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22077963&pc=OZD | 05/02/2025 | 09/25/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026636001 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 82 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Failure to Sense | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | PROTEK | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while an impella 5.5 was implanted in a patient the pump lost optical signal and could not be fully re-positioned to perfect left ventricle positioning. the pump was angled to posterior wall. the pump remained in support. | [5 empty] | The investigation of positioning issues has been completed. visually inspected and no major abnormalities identified to naked eye. without additional clinical details, the root cause of the reported issues could not be determined. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: impella 5.5 pump was returned. visually inspected and no major abnormalities identified to naked eye. optical 5s parameters were tested by connecting pump to optical bench(console) and 5s parameters were failing indicating hard failure with low signal not reliable (snr), light, and gain increasing, cavity length being maxed out. light continuity test was done and the light was losing out from optical fiber located above motor housing surface near cannula. the optical fiber was observed to be damaged during support with some ingress present below the damage fiber line zone. the cause of the damaged fiber line is unknown due to insufficient evidence from clinical. data logs were reviewed. lt log was showing placement signal not reliable (#1036 alarm) on 05/02/25 with ps maxing out and optical 5s signals being lost. rt log showing placement signal not reliable (#1036 alarm) with snr decreasing to almost zero, contrast decrease but increase in amplitude, while gain and lamp increase when ps is lost by maxing out to 3000. the identified failure pattern resembles a fiber break pattern. the root cause of the optical signal issue was a damaged optical fiber line at location above the motor housing surface where the fiber line connects to optical sensor. the cause of fiber damage/break is unknown. positioning issue: the root cause of the positioning issue was not determined due to insufficient clinical information. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "1559" and 114 based on the results of the investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9100 | 1220648-2025-28682 | 22078034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078034&pc=OZD | 05/02/2025 | 06/13/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding due to a complicated delivery. the patient was already noted to suffer chronic anemia and two units of packed red blood cells were transfused. the patient survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since the doctor use non-abiomed product which is not recommended per instructions for use which led to the bleed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9101 | 1220648-2025-28683 | 22078054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078054&pc=OZD | 04/29/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 04/29/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that multiple complications were encountered during impella cp support. two days into support, the patient experienced intermittent atrial fibrillation. an afterload reduction was initiated following the event and pump support was increased to p-2. three days later, the patient was found to have an intercostal arterial bleed. over two liters were drained from the site. heparin was being administered at the time of the noted issue. fourteen days later, the patient once again went into atrial fibrillation. weaning of the pump was initiated around this time. two days later, the patient had bloody secretions and required a bronchoscopy. the pump was explanted after twenty-three days of support. upon explant, a thrombus was noted across the entirety of the cannula and in the descending aorta. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smart assist system section: potential advhe impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9102 | 1220648-2025-28684 | 22078146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078146&pc=OZD | 05/03/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the pump was clotted which resulted in low pump flow. the pump was explanted and a new impella cp was placed. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9103 | 1220648-2025-28685 | 22078154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078154&pc=OZD | 04/30/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626840 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a hematoma at the access site after implantation of an impella 5.5. the pump was explanted to address the condition, and an impella cp pump was placed for ongoing support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9104 | 1220648-2025-28687 | 22078403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078403&pc=OZD | 05/02/2025 | 06/30/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 64 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp using a 14fr sheath was aborted due to patient anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the sheath was unable to pass through due to calcification and small size of the vessels and the insertion was aborted instructions for use: impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-28687.
9105 | 1220648-2025-28686 | 22078405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078405&pc=PYX | 05/03/2025 | 06/19/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026640593 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that upon ramping up support on an impella rp flex on a patient, the purge pressure increased to over 1000 mmhg. despite troubleshooting, the issue remained and the decision was made to replace the pump. the patient survived the procedure. | [5 empty] | Thrombus failure mode added based on investigation findings. high purge pressure: the returned product was inspected and biomaterial was found obstructing the purge gap and around the impeller. the pump was run in the lab after soaking and high purge pressure did not reproduce likely as a result of biomaterial shifting position. the data logs show a motor current spike about 3 minutes after pump start. there was a motor current spike with speed deviation and a drop in pump flows characteristic of biomaterial ingestion. this occurred outside a p-level change. after motor current spike there was a rise in purge pressure until alarms were triggered. the root cause of the high purge pressure was determined to be biomaterial as evidenced by the returned product analysis and data log showing ingestion pattern. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. h6 investigation findings and investigation conclusions were erroneously entered on the initial manufacturer device report number 1220648-2025-28686. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9106 | 1220648-2025-28688 | 22078533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078533&pc=OZD | 05/03/2025 | 06/24/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642695 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp implantation on a patient, the physician was pulling the wire back and knuckling at the valve. the physician told the staff that the wire was out, but it was not completely out of the body and ended up getting ingested in the motor housing on the impella cp. motor current spiked right away, and the pump completely stopped. the pump was removed, and a new impella cp was placed successfully. no patient harm was noted. | [5 empty] | The investigation of pump stop has been completed. the impella cp and guidewire were returned for evaluation. pump did not start: upon visual inspection, returned guidewire was kinked and partially fractured. pump did not have any visual abnormalities on the returned product. pump was connected to console and able to be run for approximately 5 minutes at p-9 without issue. data logs were reviewed and real time (rt) log of case showed immediate motor current spike and pump stop as soon as pump was turned on. clinical details noted that physician pulled the guidewire to the atrioventricular (av) valve but not all the way out of the cannula. when attempting to start the pump, an immediate pump stop occurred and a motor current spike had been noted on the automated impella controller (aic). as the pump was run for less that 5 minutes in the field and unable to be started, the failure mode of "pump stop" was changed to "pump did not start". the root cause of the pump did not start was due to a use related issue as the guidewire was still in the pump when pump was turned on. product damage (guidewire damage): based on analysis of returned product analysis, an additional failure mode of "product damage (guidewire damage)" was added. upon visual inspection, returned guidewire was kinked/partially fractured near tip of guidewire. clinical details noted that the guidewire was potentially still in the pump when the pump was turned on and an immediate pump stop occurred. the root cause of the guidewire damage was most likely a use related issue based on returned product and clinical details. | The pump and guidewire were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9107 | 1220648-2025-28689 | 22078627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22078627&pc=OZD | 05/03/2025 | 07/01/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610399 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was being placed with an impella cp over the .018 wire to the left ventricle. the team was removing the wire from the pump and it snapped. the wire and pump were removed completely and a new pump and wire were placed. the patient harm was noted. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | Reported the guidewire breaking when removing it from the pump. the pump was removed and replaced with a new pump. the pump and purge cassette were returned but the guidewire was not. there were no damages or abnormalities found on the returned pump. delivery issue - the root cause of the delivery issue was not determined due to insufficient clinical details. product damage - the root cause of the product damage was not determined since guidewire was not returned and insufficient clinical details were provided.
9108 | 1220648-2025-28690 | 22079096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079096&pc=OZD | 05/05/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/05/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649363 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient admitted in with st-segment elevation myocardial infarction. the pump was supporting despite ventricular fibrillation and defibrilate x 7. the impella cp pump was successfully explanted and the patient remained on impella rp flex support for nine more days. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9109 | 1220648-2025-28691 | 22079137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079137&pc=PYX | 05/05/2025 | 05/29/2025 |  | 05/23/2025 |  |  | 05/05/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026630672 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Atrial Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during support, the patient had atrial fibrillation (af). the patient was given diltiazem and support was continued. additionally, during the pump weaning, the pump ran at 1.8 liters per minute (lpm) for a few hours. a sudden spike in motor current and placement signal were noted. flows dropped to 0.5 lpm. the pump was explanted and biomaterial was observed. the procedural outcome was the patient survived the surgery. | [5 empty] | H6 health effect - impact code 4627 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28691. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
9110 | 1220648-2025-28692 | 22079220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079220&pc=OZD | 05/02/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636081 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 was implanted in a patient admitted with multiple cardiac comorbidities. the pump was supporting along with ecmo during transfer from community to tertiary medical center. during the support thrombocytopenia concerns were made as well as bleeding in the lungs. the patient had heparin removed from systemic infusion and sodium bicarb was added to the purge. the patient expired after seven days of support. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9111 | 1220648-2025-28694 | 22079380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079380&pc=OZD | 05/01/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654704 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient experienced ischemia. external bypass was inserted for good blood flow down the leg. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9112 | 1220648-2025-28693 | 22079411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079411&pc=OZD | 05/21/2025 | 08/08/2025 |  | 05/23/2025 |  |  | 05/21/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026638006 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support hooked the cord with the foot opposite of the impella causing the cp to come out of the left ventricle. as a result, the cp pump was replaced with an impella 5.5. | [5 empty] | The investigation of positioning issues has been completed. the impella device was discarded. based on the clinical details provided, patient movement (pulling on the pump) caused the positioning issue. the root cause of the positioning issue was determined to most likely be patient movement. e.1 facility name was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-28693. g.2 added a report source selection as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-28693. h.6 code 4114 was entered incorrectly on initial manufacturer device report number 1220648-2025-28693. the device was discarded so a new code was added. h.10 additional manufacturer narrative was incorrect on initial manufacturer device report number 1220648-2025-28693. the device was not returned because it was discarded. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9113 | 1220648-2025-28695 | 22079474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079474&pc=OZD | 05/01/2025 | 06/18/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582313 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the nurse reporting a plan to stop anticoagulation due to bleeding and oozing at sites. the nurse was reminded of recommendation to restart anticoagulation to maintain act 160-180 once bleeding is under control. the patient continued on support. when patient returned from computed tomography scan to assess neuro status, he went into cardiac arrest and return of spontaneous circulation was never achieved. the patient subsequently expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9114 | 1220648-2025-28696 | 22079563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079563&pc=OZD | 05/05/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/05/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596365 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed hemolysis during impella cp support. the patient was escalated to an impella 5.5 and the condition resolved. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9115 | 1220648-2025-28697 | 22079720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079720&pc=OZD | 05/01/2025 | 09/29/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | American Indian Or Alaskan Native | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The customer reported an optical sensor issue. | During investigation, a controller error was found due to an automated impella controller software issue. there were no adverse events caused by the reported console issue. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | The console was received from the customer and an investigation regarding the returned device has been completed. the logs showed that after 1 day and 18 hours, when the peak signal-to-noise ratio alarm was resolved, a ¿controller error (opm comm stopped) [optical pump connected] ¿ alarm,¿ occurred, which was automatically resolved in 17 seconds. the console was tested. the root cause of the controller error was due to an automated impella controller software issue.
9116 | 1220648-2025-28698 | 22079724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079724&pc=OZD | 05/01/2025 | 06/24/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED,INC. |  |  | US |  | 2026642709 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp started with suction after the patient was turned. the p-level was reduced, and suction was resolved. as the p-level was raised again, there was a sudden spiking in the motor current, which then went from biphasic to flat and impella in ventricle alarm occurred. echo ordered and impella noted to be 5.5cm in the left ventricle. continuous suction still noted at p2 with flows of 0.6 lpm. physician repositioned the device with no improvement in flow or resolution of suction alarm. the impella was explanted. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. data log reviewed: suctions occurred followed by dampened mc, mc spike and flow drop on 01-may. flows continuing to trend down. quick resolved position alarms seen after suctions. no placement signal issue seen. patient had mitra clip procedure one day prior reported flow issue. repositioning and blood administration didn't solve the flow issue and suspects clot stated. low pump flow: the cause of reported low pump flow most likely was biomaterial ingestion related. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-28698. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9117 | 1220648-2025-28699 | 22079987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22079987&pc=OZD | 04/30/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625384 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the left was accessed despite known hernia and a hematoma developed. the team gave blood and explanted the cp and placed a impella 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9118 | 1220648-2025-28700 | 22080057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080057&pc=OZD | 05/04/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/05/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609442 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a long-term outcome and quality indicator (loqi) impella registry reported stating a patient had ventricular fibrillation with relationship of 'definitely' related to the impella cp. loqi states: "an impella cp was placed but resulted in more ventricular ectopy which degenerated to vf s/p direct current cardioversion 1. the impella was removed with improvement in hemodynamics." the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9119 | 1220648-2025-28701 | 22080116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080116&pc=OZD | 05/04/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/04/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630350 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and during support, the patient developed renal failure with suspected hemolysis. the pump position was verified and was found to be at 6 centimeters. the doctor was reluctant to pull back and opted to leave as is for now and monitor. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9120 | 1220648-2025-28702 | 22080122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080122&pc=OZD | 05/01/2025 | 06/16/2025 |  | 05/23/2025 |  |  | 05/01/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251888 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller (aic) had a failed self test. staff attempted to restart the console and experienced the same issue upon turning it on. the aic was replaced. there was no patient involved. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9121 | 1220648-2025-28703 | 22080167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080167&pc=OZD | 05/03/2025 | 07/02/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622267 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be successfully delivered/positioned due to an emergency ascendes graft implant and anatomical issues. the pump was removed was explanted and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because no correct position was achieved due to emergency ascendes graft implant and anatomic issues.
9122 | 1220648-2025-28704 | 22080200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080200&pc=OZD | 05/02/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026647787 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5 the patient went into atrial fibrillation which resulted in tachycardia. the patient was cardioverted and cardiopulmonary resuscitation was provided. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9123 | 1220648-2025-28705 | 22080350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080350&pc=OZD | 05/02/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026639115 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and during support, there were several suction alarms. volume was given and support was continued. additionally, there was high purge pressure. lysis protocol was initiated. support was continued until successful weaning and explanation. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9124 | 1220648-2025-28706 | 22080466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080466&pc=OZD | 05/03/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/03/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608334 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2025 |  |  |  |  | The complainant had an impella cp pump placed to support an acute myocardial infarction/cardiogenic shock surgical patient. the pump was supporting when there was pump repositioning and suction due to pump placement along the heart structures. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9125 | 1220648-2025-28707 | 22080950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22080950&pc=OZD | 01/22/2025 | 05/23/2025 |  | 05/23/2025 |  |  | 05/02/2025 | 05/23/2025 |  |  |  |  | 05/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599684 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Kidney or Urinary Problem | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient experienced access site bleeding and worsening kidney injury during impella cp support. the bleed was believed to have a probably relationship to impella use while the kidney injury was thought to have a possible relationship to impella support. the patient received two units of packed red blood cells and 1 unit of cryo in response to the bleed. the patient's acidosis was managed with sodium bicarbonate pushed and high pressors. the patient unfortunately did not recover and expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9126 | 1220648-2025-28628 | 22068822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22068822&pc=OZD | 05/02/2025 | 08/13/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161463 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) failed to recognize the impella pump when it was plugged in. the aic was swapped and the issue resolved. no patient harm was noted. | [5 empty] | The investigation of pump not detected has been completed. the impella device was returned for evaluation. data analysis: a review of the logs shows that a case was started and the user was prompted to connect a pump to the automated impella controller (aic). however, even after the epm reset, no pump was ever detected. these symptoms align with the previously investigated epm crystal issue, which impacts pump detection. the user then initiated a shutdown through the graphical user interface, and the console carried out the shutdown. device analysis: the pump connector was inspected for any issues that might contribute to an incomplete or improper connector insertion, and nothing was found. the console was opened, and all cables and connections to the impellatronic pcb were inspected for integrity, with no issues identified. this process was repeated for the pump connector and the other pcbs in the console, and no issues were found. root cause: the pump detection issue is likely due to a rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9127 | 1220648-2025-28629 | 22068850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22068850&pc=OZD | 04/17/2025 | 07/10/2025 |  | 05/22/2025 |  |  | 04/17/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026617570 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Per the medical safety officer review the device was unable to be inserted; calcification lower limb. delay in/inability to delivery mcs very likely due to patient's underline anatomy. however, there remains not enough evidence to exclude the impella device used as an associated factor to the outcome of death. | It was reported that implantation of an impella cp was attempted but due to calcification of the lower limbs placement of the pump was aborted. |  |  |  |  | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the patient had calcification in the lower limbs was so severe that both insertions were difficult. | B1 selection of adverse event was inadvertently omitted on the initial manufacturer device report number. b2 selections death and date or death were inadvertently omitted on the initial manufacturer device report number. b5 revised to include mso statement and the patient outcome that was inadvertently omitted on the initial manufacturer device report number. h1 type of report was erroneously selected on the initial manufacturer device report number. h6 codes 4582 and 2199 were erroneously entered on the initial manufacturer device report number. new coded added to health effect - clinical code and health effect - impact code. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9128 | 1220648-2025-28630 | 22068876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22068876&pc=OZD | 05/02/2025 | 07/01/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1738105 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) had issues transitioning over to the wall outlet as a power source when plugged in. the aic continued to run on battery power and the battery drained when plugged into the wall outlet. the console was removed from service and another aic was used for patient support. no patient harm was noted. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. data analysis: the failure mode described was not reproduced during testing. additionally, following the investigation. during the case start on (b)(6) 2025, the ac power was shown to be connected, as no alarms were displayed. however, after two hours of pump operation, the "ac power disconnected alarm" (event 109) was triggered and could not be cleared until the console was shut down. the lt log indicates that the batteries were actively charging until the occurrence of event 109. device analysis: the console was returned for further investigation. upon inspecting the aic, an exposed live (black) wire was discovered inside the ac cord plug. further examination of the plug revealed that the wire was loose, and upon opening the plug, the wire completely came off. it appears that the wire covering or outer sheath was improperly cut, as per the vendor wiring instructions, as well as the internal black wire being shorter than the green and white wires. clamp marks are visible on the internal wires as well, which likely contributed to the wire becoming loose and detaching from the plug. see shortwire.png, shortwire2.png, and wiring instruction interpower.pdf. additionally, the batteries were checked, and no charging issues were found. the console was also tested with a known good pump and purge cassette, and no issues were observed. root cause: the reported complaint about the console not being able to charge even connected to the wall outlets was due to the loose live wire (black wire) inside the ac cord plug. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9129 | 1220648-2025-28631 | 22068922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22068922&pc=OZD | 04/30/2025 | 06/27/2025 |  | 05/22/2025 |  |  | 04/30/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614277 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58 year old female patient implanted with impella 5.5 for non-st elevated mi and sepsis. sepsis had resolved prior to insertion of the impella 5.5. during insertion it was noted that the axillary sheath was placed with some difficulty and surgeon noted inward collapse of sheath due to tension. heavy ties placed around graft and sheath. the surgeon attempted placing pump over steel core 018 wire through sheath and was not able to pass through due to collapse. surgeon was not concerned and decided to reposition the sheath with more heavy ties. ventricle reassessed with wire and catheter. impella 5.5 backloaded onto steelcore 018 wire and advanced into ventricle without difficulty. impella support was ramped up to p4 for rest of case. pci performed with single access companion sheath through axillary sheath. bleeding noted and more heavy ties applied to maintain hemostasis. after pci, impella 5.5 was explanted. patient survived the procedure. | [5 empty] | The investigation of difficulty inserting/removing pump through introducer, access site bleeding, and introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. difficulty inserting/removing pump through introducer: pump was not able to pass through sheath due to collapse. repositioned sheath with more heavy ties. ventricle re-accessed with wire and impella delivered without difficulty. the root cause of the difficulty inserting pump through introducer was most likely a sheath kink. cause of the sheath kink was unknown. access site bleeding - major: clinical details noted that patient was bleeding occurred at access site and was managed with more heavy ties around graft and sheath to maintain hemostasis. the root cause of the access site bleeding could not be definitively determined as limited clinical details were provided. introducer damage - mechanical: clinical details noted that the sheath collapsed inward after insertion of the sheath in the graft anastomosed to the carotid artery. physician noted inward collapse of sheath due to tension. heavy ties placed around graft and sheath. the root cause of the introducer damage - mechanical was most likely a sheath kink. cause of the sheath kink was unknown. instructions for use: impella 5.5 for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during impella section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during impella with automated impella controller ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28631. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28631. | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9130 | 1220648-2025-28632 | 22068952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22068952&pc=OZD | 05/02/2025 | 07/28/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116484 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when trying to prime an impella for in-service, an error appeared and then automated impella controller failure occurred. four different test pumps and five different cartridges were tried but they all resulted in the same error. there was no patient involved. another automated impella controller was used. | [5 empty] | Data analysis: the complaint was confirmed by the imc and rt logs. consistently on the complaint date, the console alerted the user that the priming was not completed and that a kink was detected "imcgui: casestart: msg c:90 e:3031 0x82918c8 screen curr:8 prev:8 (3009)" (see attached). this screen is triggered in the source code when purge pressure is greater than the difference between the fast prime max pressure and the change in pressure. the purge ticks would plateau each time and purge pressure would also flatten but would have some variability. there was also one instance during attempted priming of the aic getting purge pressure high into purge system blocked alarm which are related to the inability to prime. additionally, there were a multitude of impella stopped controller failure alarms received by the user with the pmd constantly resetting and 2 instances of a controller error from event #138 due to pump speed deviations. unrelated to the complaint, there were instances on the complaint date and 3/5 of the optical bench resetting with a pump connected triggering a log entry for optical bench communication lost while the bench resets. the rt logs show the optical sensor fast triggering to -1638.4 mmhg (sampled at 250 hz) but the placement signal is unaffected (25 hz). device analysis: the inability to prime was not reproduced during testing. additionally, the purge unit was a disassembled but there was no evidence of anything that could have prevented priming. purge pressure accuracy was confirmed to be passing step 12 of pm. the pump stops were also not reproduced as the console was able to constantly run pumps without issue. the optical bench failure was reproduced during testing. current into the optical bench was found to be around 430 ma with the original optical bench but dropped to 405 ma with a known good optical bench, and the fast-triggering raw optical sensor went away. root cause: the cause of the inability to prime was not determined since the failure mode was not reproduced. repair: as a precaution, please replace the purge assembly (0042-3026) and optical bench (0042-3015p) on ic1939 then perform sections 10-12 and 23 of pm (0042-7309) and functional check (0042-7316). | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9131 | 1220648-2025-28633 | 22069043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069043&pc=OZD | 05/01/2025 | 06/17/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  | 06/30/2025 |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024427204 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had oozing at the access site. systemic heparin was withheld and the dressing was changed. the patient survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9132 | 1220648-2025-28634 | 22069098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069098&pc=OZD | 05/02/2025 | 06/06/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1469546 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) had damage to the cover to the pump. the electrical connector outlet was broken. this device covers the location where the staff will plug in the electrical connector power source. no patient involvement was reported. | [5 empty] | The investigation of automated impella controller (aic) - damage during therapy has been completed. data analysis: not informative for this investigation. the logs did not show any anomalies. device analysis: the console was sent to field service for further investigation. it was determined that the external optical connector dust cover was broken. root cause: the broken dust cover was likely caused by user-related issues h6 component code 925 was erroneously entered on the initial manufacturer device report number 1220648-2025-28634. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9133 | 1220648-2025-28635 | 22069103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069103&pc=OZD | 04/23/2025 | 06/27/2025 |  | 05/22/2025 |  |  | 04/23/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609294 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hypovolemic Shock | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was implanted on a 64-year-old male patient in bailout situation after acute myocardial infarction emergency. support from impella was not enough, so an emergency ecmo implantation was performed. after ecmo, impella low flow occurred. however, the ecmo site was bleeding heavily which lead patient to hemorrhagic shock and the patient expired. | [5 empty] | The investigation of low pump flow has been completed. no product was returned. data logs show pump running on auto with good flows when started. an hour in, there were suction alarms and flow reduced in response to the alarms. pump is then turned down to p2 where flows and placement signal trend downwards. the root cause of the low pump flow was most likely patient condition related since there was a reduction in flows due to suction and the patient was placed on extracorporeal membrane oxygenation (ecmo). | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the patient's outcome was reviewed by abiomed's medical safety officer. it was determined that the hemorrhagic shock occurred on ecmo insertion. this is not an impella related death. even though the patient was anticoagulated for the use of the impella device, it was also said (per the reported information) the patient was dying due to massive blood loss; after ecmo was running a ¿big¿ hematoma on the ecmo insertion side was seen. regarding the low flow, if there was no flow and no other reason beyond why there was no flow, it could be likely because 1) the patient could have ¿bled out or 2) the impella could be malfunctioning at the same time, not providing enough flow. it is, however, more likely that the patient was bleeding out because of the volume issue. however, there is not enough information at hand to draw this conclusion. we need to be able to confirm the impella device was providing enough flow; it was not malfunctioning. this cannot be said when there is a line in the report stating low flow. there is no direct or indirect causal relationship between the demise outcome and the impella device use. thus, this is not an impella related death.
9134 | 1220648-2025-28636 | 22069279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069279&pc=OZD | 05/02/2025 | 06/24/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612823 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a purge blocked alarm appeared during impella 5.5 support. the purge cassette was replaced, but the issue remained. the decision was ultimately made to explant the pump. the patient survived. | [5 empty] | The investigation of high purge pressure has been completed. based on inconclusive information from data logs, the root cause of high purge pressure could not be determined. based on the trend in motor current immediately following high purge pressure in data logs and saturated flows reproducing on returned product with biomaterial residue, the root cause of the saturated flows was most likely biomaterial buildup. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9135 | 1220648-2025-28637 | 22069386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069386&pc=OZD | 05/02/2025 | 06/24/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642720 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella cp was implanted in a patient to treat cardiogenic shock. however, there was a ¿purge system open¿ alarm that would not clear despite troubleshooting measures. the pump was replaced with a second impella cp was used to continue treatment. no alarm was noted after the pump exchange. the patient survived. | [5 empty] | The investigation of purge leak ¿ pump has been completed. there were no data logs returned. the returned product was connected to a lab console and began to purge with the cassette returned. there were immediate "purge system open," alarms and it was noted that there was fluid was coming out from the sidearm filter at the location of the filter vents. attached to the complaint is communication with the rep stating it is unknown if any cleaning agents were used but it is a well-trained site. the root cause of the purge leak - pump is due to damaged sidearm filter but the cause of the damage is unknown. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9136 | 1220648-2025-28638 | 22069519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069519&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644947 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient developed signs of hemolysis during impella cp support. an echocardiogram showed the impella to be slightly shallow in the ventricle and repositioning was requested. the pump was explanted later that same day. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9137 | 1220648-2025-28639 | 22069624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069624&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646460 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support. a clot on the impella cannula was visualized on echocardiogram. the patient also had intramural thrombus and possible intra-aortic thrombus. a computed tomography scan and vascular consult regarding clots were pending. the patient continued on extracorporeal membrane oxygenation support. the patient's outcome was stable. | [5 empty] | The impella device was not returned, and the investigation has been completed. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9138 | 1220648-2025-28640 | 22069792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069792&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658429 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient lost blood flow to their lower left extremity (lle) during impella cp support. it was noted that an antegrade sheath had not been placed to assist with supplying flow to the leg. an occlusion of the lle occurred and the patient was escalated to an impella 5.5. an above knee amputation was subsequently required due to the ischemia. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9139 | 1220648-2025-28641 | 22069800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069800&pc=OZD | 05/01/2025 | 06/19/2025 |  | 05/22/2025 |  |  | 04/30/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656933 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Application Program Problem | [5 empty] | Hematoma | Hemolysis | Cardiogenic Shock | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [5 empty] | IMPELLA 5.5 SN (B)(6). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5 the automated impella controller (aic) generated a controller error. a replacement aic was used to complete implantation. the patient survived. | The 5.5 was placed and was supporting when after 2 days there was brown fluid in the purge sidearm. the purge is running with sodium bicarb and remains on for support. the patient is awaiting transplant. the access site developed a hematoma and this prompted the team to pause the anticoagulation. additionally months after the support was initiated the long-term outcome and quality indicator (loqi) impella registry (loqi) database reported upon 2 adverse events, cardiogenic shock and hemolysis both with "possible" relationship to the use of the pump. |  |  |  |  | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of purge leak ¿ pump, access site bleeding has been completed. no product was returned for evaluation. purge leak - pump: clinical details noted that the purge sidearm was leaking purge fluid and brown fluid from sidearm. purge pressure was not affected and pump was not replaced. the root cause of the purge leak - pump was most likely due to a damaged sidearm based on clinical details but the exact location of the leak was not able to be definitively determined as no product was returned for evaluation. access site bleeding - major: clinical details noted that a small hematoma was observed at access site while on support. the root cause of the access site bleeding could not be definitively determined as limited clinical details were provided as to how the hematoma occurred. b5 additional information received reporting cardiogenic shock and hemolysis, investigation results pending should any new information become available a supplemental report will be sent. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ the initial manufacturer device report number 1220648-2025-28641 was erroneously reported against the automated impella controller (aic). the following fields have been corrected to reflect the correct device; impella 5.5. the aic failure modes were previously reported in manufacturer device report number 1220648-2025-28620. b1 revised to reflect both adverse event and product problem. b2 revised to reflect outcome. b3 revised to reflect the correct event date. b5 revised to reflect correct description. d1 brand name. d4 model number, catalog, number, serial number, lot number, expiration date, and udi number. d6a and d6b d10 should have been left blank. g2 study was missing. g3 date should have been 5/1/2025. h1 type of event. h4 manufacture date. h5 single use. h6 health effect - clinical code, health effect - impact code, medical device problem code, and component code. h8 usage of device. h10 instructions for use.
9140 | 1220648-2025-28642 | 22069987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22069987&pc=OZD | 04/30/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 04/30/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653712 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had torsades followed by ventricular tachycardia (vt). the patient was shocked and support continued. reportable due to patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9141 | 1220648-2025-28643 | 22070341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22070341&pc=OZD | 04/29/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 04/29/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1748478 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure with the automated impella controller (aic). the console recognizes when it is plug into the outlet; however, charge was at 50%. battery failure displayed on screen. a check was made to ensure batteries were engaged with the underside buttons. consequently, the device was removed from service. it is unknown if the event occurred during a clinical procedure and, therefore, unknown if there was any adverse effect to a patient. | [5 empty] | Patient-specific information was unavailable and/or not provided at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller (aic) data logs and console were returned for evaluation and analysis. a product history review was completed, and no action was required as a result of this review. data review was completed, and logs were consistent with the reported complaint details. there were two instances of battery failure (transport connection blown/missing) alarms observed in the logs. device analysis was then completed; the console was tested after arriving in field service. the reported battery failure alarm issue was able to be reproduced. results of running the batttest utility revealed that cell 2 in battery 2 had a voltage that was significantly less than the voltages of the rest of the cell stacks in the battery. isolative testing was performed by swapping battery 2 with a known good one which resolved the alarm. in conclusion, this battery failure was identified more specifically as a battery cell failure (as opposed to the as reported failure mode "aic - battery failure." after investigating. the root cause of the battery failure was determined to be a cell imbalance in battery 2.
9142 | 1220648-2025-28644 | 22070499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22070499&pc=OZD | 05/01/2025 | 10/06/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while implanted in a patient an impella pc generated suction alarms, flattened motor current, and "impella in ventricle" alarm. the impella was repositioned and the ¿impella in ventricle¿ alarm resolved, but the suction alarm remained. as time passed, the suction alarm was cleared. patient care was delayed in addressing the alarms. the patient was successfully weaned from the pump and survived. | The customer reported that the device exhibited an optical signal issue. |  |  |  |  | The investigation of positioning issues has been completed. the data stick was received for investigation. positioning issues: the clinical details state that the ao waveform appeared ventricular and that it was less than 30, with pulsatile motor current. there was no placement signal low alarm and then "suction" and "impella position in ventricle" alarm. the impella was repositioned and the ao waveform became arterial and the left ventricle (lv) waveform was then showing but there was no numerical reading. the "impella position in ventricle" alarm resolved but there was still a suction alarm and no numerical values despite normal flows and motor current in the 900's. as time passed, the lv was not dashed out and the suction was cleared. due to lack of clinical details, the root cause of the positioning issues of the pump being too deep in the ventricle cannot be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. due to lack of clinical details, the root cause of the positioning issues of the pump being too deep in the ventricle cannot be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h4: corrected manufacture date. h6: added sensing code 1559.
9143 | 1220648-2025-28645 | 22070689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22070689&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656647 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was being placed for a high-risk cardiac procedure patient. a perforation of the left femoral artery (lfa) occurred. both the right femoral artery and the left femoral artery were scanned and the right showed significant disease. the lfa was accessed with micropuncture inserted and with first shot dissection noted. the decision was made to access the left axillary arter. the pump supported and was explanted after 2.5 hours. the patient survived the procedure. additional information was received. the site was being assessed for percutaneous coronary intervention. after the explant and case was complete, a stent was believed to be used. the pump was discarded as the physician did not state this was impella related. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9144 | 1220648-2025-28646 | 22070769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22070769&pc=OZD | 05/01/2025 | 06/19/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025483517 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while repositioning the device, the impella cp pump came back into the aorta. attempts to readvance the pump were unsuccessful and the decision was made to remove the device. the patient survived. | [5 empty] | The investigation of positioning issues has been completed. the data logs were not recovered. the pump was returned and inspected. the pump showed no physical abnormalities and ran in the lab for 10 minutes at nominal parameters. the cause of the positioning issue was most likely use related due to the pump being initially placed too deep in the ventricle and then pulled into the aorta during repositioning attempts. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9145 | 1220648-2025-28647 | 22070950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22070950&pc=OZD | 05/01/2025 | 06/30/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640526 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was moved by the physician and the impella cp got pulled into the aorta. the physician attempted to recross the aortic valve with the impella cp. however, attempts were unsuccessful. the impella cp was explanted. the patient survived the procedure. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement during repositioning of the pump. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9146 | 1220648-2025-28648 | 22070979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22070979&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656783 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 111 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support had a head ct showing multiple areas of infarctions. neurosurgeon was consulted for emergent craniotomy, but the family decided to withdraw care, and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9147 | 1220648-2025-28649 | 22071010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071010&pc=OZD | 04/29/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 04/29/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599194 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and during the approximately 14-hour mcs, the patient experienced bleeding issues at the puncture side. hemostasis was achieved with fem stop and the leg was noted well perfused. it is unknown if blood products were administered and/or if anticoagulation was suspended. further details around the bleeding event were unnoted. however, some time thereafter the bleeding event (same day), the patient was indicated to have expired while on support. no details surrounding the demise outcome were provided. | [5 empty] | Given the impella-related event with insufficient details and demise outcome, there is not enough evidence to exclude the impella cp device used as an associated factor in the overall outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9148 | 1220648-2025-28651 | 22071164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071164&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 58 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp the patient developed an apical thrombus. the impella was explanted and the patient survived. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9149 | 1220648-2025-28650 | 22071165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071165&pc=OZD | 05/03/2025 | 06/25/2025 |  | 05/22/2025 |  |  | 05/03/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619248 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following a successful implant, the physician attempted to reposition the pump and inadvertently pulled it back into the aorta. attempts to re-advance the pump were unsuccessful and the decision was made to replace the device. there was no noted patient injury. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was inadvertently pulled back into the aorta during repositioning.
9150 | 1220648-2025-28652 | 22071337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071337&pc=OZD | 05/03/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/03/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605615 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had lack of distal pulses on the impella extremity. the pump was removed and the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9151 | 1220648-2025-28653 | 22071398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071398&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654618 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed to support a patient who was in cardiogenic shock. the limb became ischemic and had a perfusion placed to restore flow. the pump remains on for support despite this. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9152 | 1220648-2025-28654 | 22071573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071573&pc=OZD | 05/03/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/03/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617728 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella 5.5 was placed on a st-segment elevation myocardial infarction induced ventricular septal defect patient. it was reported while on support, the team noticed uneven pupils and did emergent computed tomography scan of which showed global anoxia and the family decided to withdraw care. the patient expired and the relationship between the impella 5.5 and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9153 | 1220648-2025-28655 | 22071657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071657&pc=OZD | 05/03/2025 | 06/27/2025 |  | 05/22/2025 |  |  | 05/03/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634290 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella cp support and during support, the impella was repositioned due to patient movement. the patient later went into ventricular tachycardia (vt). the patient was given a bolus of amiodarone. it is unknown if the position of the pump contributed to the patient's vt. care was withdrawn and the patient expired. | [5 empty] | The investigation of positioning issues and vf/vt/af/at has been completed. the product and data were not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement since the issue occurred when turning the patient. the vf/vt/af/at could not be determined without additional clinical details. | H6 health effect - impact code 4644 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28655. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9154 | 1220648-2025-28656 | 22071788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071788&pc=OZD | 05/01/2025 | 06/26/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671060 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Gastrointestinal Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support of a 57-year-old male, the patient developed a gi bleed and heparin was discontinued. additionally, it was reported that the patient was not doing well, occasional intermittent vagal, bradycardia and alternating ventricular tachycardia. patient was on vasopressin. patient was made a do not resuscitate. ultimately case was withdrawn and the patient expired. | [5 empty] | The investigation of vascular damage - peripheral vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28656. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28656. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9155 | 1220648-2025-28657 | 22071807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071807&pc=OZD | 04/11/2025 | 06/26/2025 |  | 05/22/2025 |  |  | 04/11/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636291 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was later documented that a placement signal not reliable alarm appeared during support. the pump was replaced in response to the noted failure. | The user facility reported that a 66-year-old patient experienced some small runs of ventricular tachycardia following impella cp implant. the following day, the aic (console) alerted that the pump had shifted position. attempts to reposition the pump were unable to achieve optimal positioning away from the mitral valve. at this point in time, the anticontamination sleeve tore from the repositioning hub. the catheter was cleaned and tegaderm was placed over the tear. the staff expressed concerns if additional repositioning was necessary later in support. additional intervention was not required and support continued with the current setup. |  |  |  |  | The investigation of vf/vt/af/at, positioning issues, product damage, and optical signal issue has been completed. the root cause of the optical signal issue was due to sensor wear based on review of returned data logs. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided on how pump came out of position. the root cause of the product damage (sleeve damage) was a use related issue as the pump was attempted to be repositioned when the damage occurred. as the necessary information was not provided, the root cause of the vt/vf was not determined. d9 device returned to manufacturer date added, investigation of the actual device has not yet started. the above investigation results are based on data logs and clinical details. should any new information from the device analysis become available a supplemental report will be submitted. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28657. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28657. g2 report source selections were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28657. h6 medical device problem code 2917 and component code 4723 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28657. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
9156 | 1220648-2025-28658 | 22071820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071820&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617134 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support experienced ventricular tachycardia (vt). additionally, unsuccessful troubleshooting was done for impella in aorta placement alarm. the impella cp was explanted and discarded. the patients outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9157 | 1220648-2025-28659 | 22071914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22071914&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651730 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella site experienced oozing during patient support. the dressing was changed and heparin was placed on hold in response to the condition. the patient was additionally given two units of blood. support continued with the implanted pump following the intervention. the patient was eventually weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9158 | 1220648-2025-28660 | 22072003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072003&pc=PYX | 05/04/2025 | 06/23/2025 |  | 05/22/2025 |  |  | 05/04/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625116 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella rp flex was implanted in a patient for mechanical circulatory support and during the implant, a dissection of the left internal jugular occurred while advancing the sheath. access was changed to the right femoral vein. a mattress suture was placed and the impella sheath was advanced without an issue. the pigtail catheter was used to access the pulmonary artery, and the impella rp flex was placed without an issue and slowly ramped up to performance level p-8. the sheath was removed and extra mattress suture placed for hemostasis. pressors were started to be weaned down. no bleeding or hematoma was noted and the patient's hemodynamics remained unchanged. the patient was transferred to the unit in stable condition. three days later, it was noted that the patient went into atrial fibrillation with rapid ventricular response and was started on amiodarone. the patient continued on the impella mcs for an additional five days before being successfully weaned. the device was explanted with hemostasis achieved and outcome was survived. | [5 empty] | The investigation of vascular damage - peripheral vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. vt/vf : as the necessary information was not provided, the root cause of the vt/vf was not determined. vascular damage - peripheral vessel : the root cause of the vascular damage issue was most likely use related, the physician chose to advance sheath over standard j wire (non-abiomed product) and wire buckled when advancing. no blood returned after. the sheath was pulled back some and blood returned, but the physician chose to remove sheath and close the site. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
9159 | 1220648-2025-28661 | 22072008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072008&pc=OZD | 05/01/2025 | 07/29/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  | 08/31/2025 |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382989 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Approximately two months after implant, a placement signal low alarm appeared. the measured lv diastolic pressure became increasingly negative despite the position, motor current, and purge metrics remaining stable. the alarm was muted, however it returned shortly after. the alarm was subsequently disabled as a result of the persistent issue. the staff planned to monitor the flow, motor current and purge metrics to ensure ongoing functionality of the pump. | The user facility reported that a 24 year old male patient experienced a run of ventricular tachycardia during impella 5.5 support. the patient was defibrillated and cpr was performed to relieve the condition. it was unknown if a discovered bleed was the cause of the arrest, or if the bleed was a result of the chest compressions. the pump was repositioned in response to the vt. support continued uninterrupted with the implanted pump. |  |  |  |  | B1 revised to include product problem based on information received from the clinical site b5 revised to include additional information received from the clinical site h6 revised to reflect information received from the clinical site the impella pump was not returned for investigation of optical signal issue. should new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
9160 | 1220648-2025-28662 | 22072191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072191&pc=PYX | 05/02/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  | 05/31/2025 |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024346248 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella rp was aborted. the physician could not advance the impella rp due to the patient's anatomy. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
9161 | 1220648-2025-28663 | 22072241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072241&pc=OZD | 05/01/2025 | 10/02/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631226 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Death | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient was initially on impella cp and then was escalated to impella 5.5, when neurological deficits were noted. a computed tomography (ct) scan confirmed multiple old and new areas of infarct, possibly hemorrhagic and embolic. this case is related to (b)(4) which captures the impella 5.5 new cerebrovascular accident and patient¿s death. | The complainant also reported that the impella cp aortic placement signal reading was very dampened. the team reset aortic placement without any changes. |  |  |  |  | The investigation of stroke/cva has been completed. the returned pump showed no identifiable optical abnormalities were observed in the data logs, the returned pump passed all testing and limited information related to failure was provided. stroke/cva: strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this complaint is related to another complaint that's being reported impella 5.5 (serial number (b)(6)). as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the returned product was inspected and there were no physical abnormalities. there were no broken pins or kinks. the pump passed laser continuity test and had stable 5s parameters. the data logs from the case had just one placement signal low alarm which recovered within seconds. there were no other placement signal alarms or identifiable optical issue patterns. the root cause of the optical signal issue was not determined as no identifiable optical abnormalities were observed in the data logs, the returned pump passed all testing and limited information related to failure was provided a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem. b2 death date updated . b5 updated to include placement signal description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28663. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
9162 | 1220648-2025-28664 | 22072244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072244&pc=OZD | 05/02/2025 | 07/01/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  | 01/31/2026 |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455066 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that implantation of the impella cp was aborted due to the patient¿s vessel condition. the patient's vessels were heavily calcified in the aortic arch. the patient continued care on intra-aortic balloon pump. | [5 empty] | The root cause of the delivery issue was most likely patient condition related, cp aborted due to heavy calcification in the aortic arch after multiple attempts at insertion. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9163 | 1220648-2025-28665 | 22072428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072428&pc=OZD | 05/02/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614124 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon implantation of an impella cp, flows were suboptimal. the position and fluid status of the patient were both verified, but the flows remained low. the decision was ultimately made to replace the pump and the patient survived. | [5 empty] | The impella device is available to be returned but has not been received yet. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9164 | 1220648-2025-28666 | 22072538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072538&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/01/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615955 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient suffered a gastrointestinal bleed during impella 5.5 support. an esophagogastroduodenoscopy was negative but patient continues to have bloody stools. the patient received one unit of packed red blood cells to treat the condition. heparin administration was intermittent leading up to the event. support continued with the implanted pump following the intervention. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9165 | 1220648-2025-28667 | 22072589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072589&pc=OZD | 05/01/2025 | 05/22/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584077 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp tuohy-borst valve is suspected to be damaged. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9166 | 1220648-2025-28668 | 22072821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22072821&pc=OZD | 05/02/2025 | 07/31/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154857 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon start-up, an error message appeared on the automated impella controller instructing the staff to switch to the reserve console. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. there was no evidence in the logs related to power-on issues or a blank screen. most likely, the ¿aic - system self-check error¿ was misreported as power-on issues/blank screen. the root cause of the ¿system self check failure¿ was determined to be power battery manager (pbm) corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
9167 | 1220648-2025-28669 | 22073086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22073086&pc=OZD | 05/02/2025 | 06/24/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601468 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella 5.5 support when bleeding was noted at impella access site as the team could not get good seal with graft locks so a suture was around the graft/sheath. there was no patient harm and the patient remains on support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related since 10mm gelweave graft was used. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9168 | 1220648-2025-28670 | 22073368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22073368&pc=OZD | 05/02/2025 | 06/19/2025 |  | 05/22/2025 |  |  | 05/02/2025 | 05/22/2025 |  |  |  |  | 05/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631847 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support woke up agitated wrapped foot in wire and kicked it apart. the catheter and red plug came apart, only visible part was still connected, the purge line. a pump stop occurred. the pump was pulled back across the aortic valve. completely removed at bedside 30min later. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of pump stop has been completed. the pump was returned for evaluation. upon visual inspection of returned pump, catheter just past the red handle was broken and only purge line was still intact. electrical testing showed all three motor cable lines were open. no data logs were returned for analysis. clinical details noted that patient was agitated and wrapped pump wire around foot and kicked it apart and catheter and red plug came apart. the root cause of the pump stop was due to an electrical open of all three motor cable lines from excessive movement from patient based on returned product analysis. d6a and d6b was erroneously entered on the initial manufacturer device report number 1220648-2025-28670. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9169 | 1220648-2025-28591 | 22059414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059414&pc=OZD | 05/14/2025 | 07/23/2025 |  | 05/21/2025 |  |  | 05/14/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655974 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient expired while on support, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The user facility reported a patient was escalated from intra-aortic balloon pump (iabp) to impella 5.5 in anticipation of ventricular septal perforation (vsp) closure. during impella 5.5 management, the purge pressure rose, and the purge flow rate dropped with the pressure rise, causing purge-related alarms. there were no problems with the pump flow rate or motor current consumption. although the purge system was replaced and checked for bends, no improvement was found, so the impella 5.5 device was replaced. it was noted that when the existing impella 5.5 was removed, a thrombus was found to have adhered to the outflow site. the device was replaced without any issues and without affecting the patient's hemodynamics. the patient is in stable condition. |  |  |  |  | The investigation of high purge pressure and thrombus has been completed. the pump was not returned for investigation. the data log shows a 30-minute gradual rise in purge pressure, followed by a high purge pressure alarm without any observed motor current spike. purge pressure was seen to recover with a gradual trend after 12 hours. the cause of the high purge pressure issue was most likely biomaterial buildup, based on data log review. the root cause of the thrombus was unable to be determined due to insufficient clinical details. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28591. b5 patient outcome was erroneously entered on the initial manufacturer device report number 1220648-2025-28591. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28591. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28591. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9170 | 1220648-2025-28592 | 22059457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059457&pc=OZD | 05/12/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 05/13/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616983 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on ecpella (extracorporeal membrane oxygenation and impella cp) due to ventricular fibrillation storm. it was noted that during weaning, the motor current consumption increased and the pump stopped. after the pump stopped, the recommended procedure was carried out, however, there was no improvement so the decision was made to remove the impella cp. fortunately, since it was in the weaning stage, the pump was removed without the use of mechanical circulatory support (mcs) and the patient survived without any problems thereafter. it was noted that when the pump stopped, only an alarm related to the motor current consumption appeared. no alarms related to purge were observed. additionally, it was confirmed that no biological material was formed when the impella cp was removed. the patient is in stable condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9171 | 1220648-2025-28593 | 22059536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059536&pc=OZD | 05/02/2025 | 06/05/2025 |  | 05/21/2025 |  |  | 05/02/2025 | 05/21/2025 |  |  | 07/31/2025 |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399180 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 and developed some oozing at their impella insertion site. one unit of packed red blood cells was started. the patient's platelets were good. support continued uninterrupted and the patient's outcome is unknown. additional information was received. the pump was not leaking. the surgical pocket was oozing some serosang. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9172 | 1220648-2025-28594 | 22059661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059661&pc=OZD | 04/20/2025 | 05/30/2025 |  | 05/21/2025 |  |  | 04/29/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301537 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella support the aic battery would not charge. the aic was exchanged. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. the console logs were partially consistent with what was reported in an email relating to this complaint and aic. console was tested after arriving in field service. the reported battery issue could not be reproduced. however, further inspection of the aic revealed that the power cord that was connected to this console was not the original one we install in house console during manufacturing. the root cause of the reported battery issue was likely a defective foreign power cord. it is likely that this defective power cord caused intermittent charging of this console which resulted in battery level low alarms. | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
9173 | 1220648-2025-28595 | 22059723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059723&pc=OZD | 05/02/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 05/02/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654302 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced an episode of ventricular tachycardia. the pump was explanted and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9174 | 1220648-2025-28596 | 22059741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059741&pc=OZD | 01/04/2025 | 09/17/2025 |  | 05/21/2025 |  |  | 01/29/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548487 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. | The complainant reported the patient was on impella 5.5 support. while on support, high purge pressure was noted. as the purge pressure rose the team made choice to intervene by the addition of tissue plasminogen activator into the purge fluid. the pump remained on for support for 2 more days when then bridged to left ventricular assist device. there was no patient harm. |  |  |  |  | The investigation of high purge pressure has been completed. the cause of the high purge pressure issue was most likely biomaterial, based on returned product investigation and provided clinical details including data log. instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number: 1220648-2025-28596. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal, issue has been completed. the pump was returned for investigation. the returned pump was cut at the cannula near the motor housing. no investigation could be performed on the reported optical signal issue. data logs were reviewed and the placement signal not reliable alarm was confirmed. upon review of data logs, there is no device problem with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28596. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9175 | 1220648-2025-28597 | 22059819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059819&pc=OZD | 10/05/2023 | 06/25/2025 |  | 05/21/2025 |  |  | 05/02/2025 | 05/21/2025 |  |  | 07/31/2025 |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388049 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through abiomed's long-term outcome and quality indicator impella registry that a patient was implanted with an impella cp and developed hemolysis that was "possibly" related to the pump. the pump was explanted and the patient was implanted with an impella 5.5. the patient's outcome was stable. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. data log reviewed: many suction alarms observed on first 2 days of support. no motor current (mc) spike observed outside of p-level changes observed. no positioning issue alarms seen. around 10hrs of support drop in flow observed at p-3 and resolved on next day. a few purge pressure spikes observed during support. no placement signal issue observed. clinical account stated hemolysis observed but insufficient data provided if reported issue resolved or any attempts performed for resolving. hemolysis: the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. low pump flow: failure mode low pump flow added as a result of data log observation. the cause of the low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9176 | 1220648-2025-28598 | 22059849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059849&pc=OZD | 04/29/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/29/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633286 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient developed a small subdural hematoma during impella cp support. heparin was being given at the time of the bleed. blood was given to treat the condition and heparin was stopped. the patient also experienced runs of supraventricular tachycardia. support continued with the implanted pump and both the position and support level were being monitored. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9177 | 1220648-2025-28599 | 22059946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22059946&pc=OZD | 01/04/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 05/02/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581547 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient endured renal failure following impella cp support. it was noted that the condition was believed to be possibly related to the impella. it was additionally documented that the acute kidney injury was likely from the patient's tubular necrosis. continuous renal replacement therapy was initiated in response to the condition until it resolved roughly a month later. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9178 | 1220648-2025-28600 | 22060035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060035&pc=OZD | 04/30/2025 | 09/18/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026655276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp the patient developed a hematoma at the access site. while wiring through septals the patient had a run of ventricular tachycardia (vt) which was converted without defibrillation. furthermore, the pump had lost placement signal. there was two guide catheters in with the impella that may have had interaction with the sensor. impella flows remained the same throughout. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal, issue has been completed. no product was returned for analysis. the data logs show placement signal low alarm and contrast with a simultaneous increase in lamp. the pattern observed in the logs is most likely due to an interaction between the guide catheter and the sensor. the root cause of the optical signal issue was unable to be definitively determined as no product was returned to confirm possible sensor damage or effect of interaction. to determine the true root cause of the arrhythmia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the arrhythmia was not determined. the root cause of the access site bleeding - major was not determined as insufficient clinical information related to failure was provided a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem from the initial manufacturer device report number 1220648-2025-28600. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output - use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
9179 | 1220648-2025-28601 | 22060063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060063&pc=OZD | 11/30/2024 | 05/21/2025 |  | 05/21/2025 |  |  | 05/02/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573923 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported through abiomed's long-term outcome and quality indicator (loqi) impella registry an adverse event of acute blood loss anemia with a "possible" relationship to the use of the pump. the patient had a gastrointestinal bleed and two units of packed red blood cells were transfused. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9180 | 1220648-2025-28602 | 22060175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060175&pc=OZD | 05/01/2025 | 06/09/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644948 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient endured runs of ventricular tachycardia during impella cp support. the patient was defibrillated to a normal sinus rhythm. the patient remains on support. | [5 empty] | "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9181 | 1220648-2025-28603 | 22060235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060235&pc=OZD | 04/30/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635433 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 90 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 was implanted in a patient undergoing coronary artery bypass grafting and mitral valve replacement. during support the patient was intubated and transported to another facility with their chest open and significant bleeding/chest tube output. a stroke was called but a ct scan showed no blockage, only some narrowing of the middle cerebral artery. the patient continued to do well on support with no focal deficits. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: & impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9182 | 1220648-2025-28604 | 22060392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060392&pc=OZD | 04/30/2025 | 06/04/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1388031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during maintenance the automated impella controller (aic) had a system self-check failure. no patient was involved in the event. the aic was removed from service. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs from the reported day of event are consistent with complaint and show repeating communication issues with the pbm (power battery manager) during the initial bootup. due to the communication breakdown, the console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the emergency shutdown had to be performed. the reported failure mode was reproduced during testing - upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel of pbm1. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
9183 | 1220648-2025-28605 | 22060567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060567&pc=OZD | 05/15/2025 | 07/22/2025 |  | 05/21/2025 |  |  | 05/15/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2025595554 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76 year old female patient was placed on impella cp as a pre-op placement. a ventral septal perforation (vsp) operation was performed. after the operation, the assistance level was increased from surgical mode, and the purge system blocked alarm immediately sounded. after confirming that there were no kinks in the purge system, the purge system and the yellow connector of the cp pump were removed to attempt to replace the purge cassette, and the purge pressure returned to normal. therefore, it was determined that the problem was with the pump, and the cp was replaced with a new pump. the alarm was then resolved. | [5 empty] | The investigation of high purge pressure has been completed. the impella device and data logs were received for investigation. high purge pressure: data log review shows a 6-minute purge pressure increase starting once the pump is turned on to p-2 from surgical mode. purge pressure high and purge system blocked alarms then triggered. pump flow is immediately saturated upon startup from surgical mode and remains saturated the rest of the case. high purge pressure is sustained for 30 minutes but then resolves. pump flow was still saturated at this point. there were no motor current spikes besides the expected spike upon pump startup from surgical mode. the pump was in and out of surgical mode for ~3 hours. the product was returned and evaluated. there was biomaterial in the purge gap and wrapped around the impeller. no kinks were found on the catheter. high purge pressure and purge system blocked alarms reproduced, which resolved after cutting the catheter close to the motor housing. there was no blood ingress found in the purge line. the root cause of the high purge pressure was most likely biomaterial since biomaterial was found in the returned pump's purge gap and data logs show saturated flow begin at the same time as the 6-minute purge pressure (pp) rise; however, it is unknown whether the biomaterial was buildup or ingestion since there was no motor current spikes except the expected spike upon pump startup from surgical mode. saturated flow: the failure mode saturated flow was added since data log review showed saturated pump flow occurring at the same time as high purge pressure. upon pump startup from surgical mode, pump flow is immediately saturated and remains saturated for the remainder of the case. the returned product had biomaterial in the purge gap and wrapped around the impeller. there was no blood ingress seen in the purge line in the catheter. it is unknown if there was biomaterial ingestion since there were no motor current spikes prior to the saturated flow except the expected spike upon pump startup from surgical mode. the root cause of the saturated flow was most likely biomaterial since the saturated began at the same time as the pp increase in the data logs and the returned product had biomaterial in the purge gap. d9 device returned to manufacturer date added h6 medical device problem code 1250 was erroneously entered on the initial manufacturer device report number. h10 instructions for use were erroneously entered on the initial manufacturer device report number. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9184 | 1220648-2025-28606 | 22060604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060604&pc=OZD | 05/01/2025 | 06/24/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593594 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella cp pump was implanted in a patient undergoing a high-risk percutaneous coronary intervention. single access delivery was attempted but two of the abiomed 7 french companion sheaths failed despite modifying the location of the stick to the 14 french. the third 7 french sheath worked effectively and support proceeded. there was no patient harm. | [5 empty] | The investigation of single access issue has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9185 | 1220648-2025-28607 | 22060993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22060993&pc=OZD | 04/30/2025 | 07/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239545 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automatic impella controller (aic) displayed a controller failure when it was powered on. another aic was used for patient support and no adverse patient impact was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: on the complaint date, (b)(6) 2025, after the initial bootup, the console displayed a ¿controller failure (epm1 software corruption) ¿ alarm,¿ event 1029. this confirms the reported failure mode. device analysis: during testing in the lab, a "controller failure" alarm appeared when the console was powered on, confirming the ability to reproduce the issue as described. the connectors and cables attached to the impellatronic and other associated circuit boards were checked, and no loose connections were found. when the console was powered on, both leds on the impellatronic were illuminated. the epm 5v and pmd 20v rails were measured at 5.01v and 20.1v, respectively. the epm firmware was extracted and compared to a known good epm firmware, confirming the existence of firmware corruption. root cause: the root cause of the ¿controller failure¿ alarm was determined to be firmware corruption of the epm.
9186 | 1220648-2025-28608 | 22061462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22061462&pc=OZD | 04/30/2025 | 06/11/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571143 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella 5.5 pump placed to support the surgical patient who had three coronary artery bypass grafts. due to aggressive positioning of the pump a left ventricle perforation occurred. the patient came off the bypass pump and was medically managed with the impella 5.5 and pharmaceuticals. the patient survived. | [5 empty] | The investigation of ventricle perforation and positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the ventricle perforation is most likely due to use as the surgeon jammed the pump in farther resulting in a left ventricle perforation. the root cause of the positioning issue is most likely due to use as the pump was jammed into an improper position. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9187 | 1220648-2025-28609 | 22061948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22061948&pc=OZD | 04/26/2025 | 06/04/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1126568 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 74 YR |  | Male | 125 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the placement signal went out and a controller error alarm occurred on an automated impella controller (aic). the alarm resolved on its own. the console was not changed out due to concern for pump not restarting, as the impella pump was actively supporting a patient with non-st-segment elevation myocardial infarction. there was no harm to the patient reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the impella device was not returned for evaluation. the console logs from the reported day of event are consistent with complaint and show that several times during support with pump 585749, console presented with controller error alarms (1040 due to ambient pressure out of range, 1045 due to opm comm crc error.) real time (rt) logs confirmed that all signals received from optical bench (placement, signal not reliable (snr), contrast, lamp, gain) are represented as -16384 values due to the crc error. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28609. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28609. h6 type of investigation, investigation findings, and investigation conclusions were erroneously entered on the initial manufacturer device report number 1220648-2025-28609. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9188 | 1220648-2025-28610 | 22061979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22061979&pc=OZD | 04/30/2025 | 06/19/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025549780 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support for patient on an impella 5.5 pump, the access site had a hematoma develop. the team applied pressure dressing and reduced the systemic heparin. there was no purge heparin, as the patient had sodium bicarbonate. the pump remained on support and has survived. | [5 empty] | The root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9189 | 1220648-2025-28611 | 22062132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062132&pc=OZD | 04/30/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593532 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed to escalate care from an intra-aortic balloon pump (iabp) for the st-elevated mi patient with known ventral septal defect (vsd). the pump was placed but the pump had prolapsed and kinked with the delivery to the lv. the pump was explanted and replaced. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9190 | 1220648-2025-28612 | 22062366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062366&pc=OZD | 05/01/2025 | 06/23/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630021 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Tachycardia | Thrombocytopenia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported the patient was on impella 5.5 support and during support, there were suction events. the patient received four liters of lactated ringer¿s solution and 500 cubic centimeters of albumin. support was continued. additionally, the patient had atrial fibrillation and ventricular tachycardia (vt) during support. the patient had to be cardioverted due to the vt. amiodarone and lidocaine were given to the patient. furthermore, there was concern for heparin-induced thrombocytopenia. anticoagulation was switched to bival. the patient remains on support. |  |  |  |  | The investigation of thrombocytopenia, vf/vt/af/at, and low pump flow has been completed. the impella device was not returned for evaluation. data log reviewed: many suction alarms observed during impella support around days suctions issue reported. no mc spike outside of p-level changes. no positioning or placement signal issues seen. no drop in flow observed. low pump flow: the cause of the low pump flow most likely was patient condition related with fluid bolus administration/reduction in p-level resolving suction alarms. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. thrombocytopenia: the cause of the thrombocytopenia issue was not determined due to insufficient clinical details. b2 updated to include death and date of death. b5 updated with patient outcome. d6b explant date added. h1 type of event updated to death. h6 health effect - impact code 1802 added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9191 | 1220648-2025-28613 | 22062492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062492&pc=OZD | 04/30/2025 | 06/19/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637868 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, there was a pump stop that occurred. the pump was then removed. | [5 empty] | The product was returned for investigation. data logs were not provided, and they could not be retrieved from the remote server. no physical damage was observed on the returned pump. the pump was unable to start during testing and failed electrical testing on the motor current line. all three lines were open during electrical testing. further observation inside the red handle revealed that all three electrical motor lines were stretched at the monkey fist glue bond to the catheter. the cause of the pump stop issue was an open electrical line inside red handle at monkey fist glue bond. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
9192 | 1220648-2025-28614 | 22062595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062595&pc=OZD | 05/01/2025 | 05/29/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608801 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on impella cp for hemodynamic support and coronary perfusion. it was noted that during support heparin was being held due to hemoptysis and that the urine had cleared up since holding heparin. the patient was taken to the operating room for a coronary artery bypass graft (cabg), so the impella was successfully weaned and explanted. however, during the cabg, the patient decompensated and was found to have severe mitral regurgitation. intra-aortic balloon pump was placed. there was discussion about possible 5.5 placement. however, it was revealed that the patient had an ischemic leg in the former impella site which was requiring amputation. the procedural outcome was the patient survived the surgery. | [5 empty] | B2 was inadvertently left blank on the initial manufacturer device report number 1220648-2025-28614. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9193 | 1220648-2025-28615 | 22062659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062659&pc=OZD | 04/30/2025 | 06/18/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653505 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had high-risk cardiac surgery which the patient never recovered hemodynamically on va ecmo for one week, so impella 5.5 was attempted. the aortic valve (aov) was not open and difficult anatomy made crossing aov very challenging but was successful. impella placed over 035 super stiff amplatz and was unsuccessful at crossing aov despite several medical modalities to force ejection. during the case the patient had multiple episodes of ventricular fibrillation (vf), requiring over 20 shocks. the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of delivery issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. delivery issues: based on the clinical details the root cause of delivery issue is determined to be patient condition because pump was unable to pass through vasculature due to difficult anatomy of the patient. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist with automated impella controller: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with automated impella controller: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 incomplete instructions for use were provided on the initial manufacturer device report number 1220648-2025-28615. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b2 date of death unknown.
9194 | 1220648-2025-28616 | 22062685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062685&pc=OZD | 05/01/2025 | 05/29/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654596 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support and had an echocardiogram showing a left ventricle thrombus. the plan is to keep partial thromboplastin time level at 80-100 range due to the thrombus. additionally, an electrophysiology ablation was scheduled for the patients atrial fibrillation. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of thrombus and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9195 | 1220648-2025-28617 | 22062785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22062785&pc=OZD | 04/30/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648624 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 63-year-old patient developed ischemia in their impella leg during support. the cp pump was subsequently weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9196 | 1220648-2025-28618 | 22063084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22063084&pc=OZD | 04/30/2025 | 06/19/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635389 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient encountered multiple complications during impella cp support. following implant, flow down the leg was slow and a pulse was not dopplerable. the decision was made to explant the pump and perform an abdominal aorta run off. the common femoral artery (cfa) became completely occluded at this point in time. a thrombectomy was scheduled, during which a dissection was noted to the cfa. the patient was taken to the operating room for an endovascular cut down and surgical repair. the patient was eventually weaned and explanted. | [5 empty] | Limb ischemia - reversible: after 1 day on support, the patient hasn¿t had doppler pulses. the pump was returned and inspected. no visual abnormalities were noted of the pump, which ran at nominal parameters in the lab. the cause of the limb ischemia - reversible was not determined due to insufficient clinical details. vascular damage - peripheral vessel: during removal of the pump, a cfa dissection was noted when the pump was removed through a presumed clot. cfa was completely occluded. it was noted during pump implant procedure that the pt's femoral artery was diseased. the pump was returned and inspected. no visual abnormalities were noted of the pump, which ran at nominal parameters in the lab. the cause of the vascular damage - peripheral vessel was most likely patient condition related due to noted disease of the patient's artery and dissection occurring during pump removal through a suspected clot. thrombus: during removal of the pump, the patient began to clot everywhere requiring thrombectomy. the pump was returned and inspected. no visual abnormalities were noted of the pump, which ran at nominal parameters in the lab. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9197 | 1220648-2025-28619 | 22063322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22063322&pc=OZD | 04/30/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593493 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was having an impella cp implant and during the implant, staff was unable to deliver initial companion sheath despite obtaining access on 10 and 2oclock. the initial companion sheath was replaced and was able to deliver. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during impella. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
9198 | 1220648-2025-28620 | 22063495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22063495&pc=OZD | 04/30/2025 | 06/04/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127435 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant was placing a 5.5 pump for support with an automated impella controller (aic). the aic had a ¿controller error switch to backup controller.¿ and was replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the case associated with the reported complaint was started on 30th april 2025. at 16:23:55 on april 30th, the first log entries indicating "imc-kbd error: 0xa4 0x01 0x00" appeared, suggesting that the issue occurred during the reading of the chip id (read failure of chip id). after this, the controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) - alarm 24 issued. device analysis: the console ic2214 was sent back to field service for further investigation and was tested. after performing 100 power cycles, the failure could not be reproduced. no damage was found to the keyboard or the keyboard- 4in1 cable. the root cause for the controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) - alarm 24 could not be determined due to the inability to reproduce. d6a should have been left blank on the initial manufacturer device report number 1220648-2025-28620.
9199 | 1220648-2025-28621 | 22063529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22063529&pc=OZD | 04/30/2025 | 06/17/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during implant for an impella 5.5 on a patient for bridge to transplant, the impella 5.5 was too large to cross the anastomosis. the patient's vascular anatomy was too small to accommodate the device. the physician attempted multiple techniques and followed all recommendations with no success. the decision was made to abort the implant for patient safety reasons. the pump was discarded. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella 5.5 was too large to cross the anastomosis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9200 | 1220648-2025-28622 | 22063548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22063548&pc=OZD | 04/30/2025 | 06/19/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625302 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support ongoing for a patient admitted in with known disease and undergoing a high-risk percutaneous coronary intervention. the pump was supporting when there were signs of hemolysis, urine burgundy and decreased to 10/hr. the care was withdrawn and patient expired. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9201 | 1220648-2025-28623 | 22063884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22063884&pc=OZD | 04/30/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625427 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient admitted in with multiple cardiac and systemic comorbidities with new cardiogenic shock. the registered nurse reported that the patient had run of a flutter/wide complex ventricular tachycardia requiring cardioversion. the patients pressure dropped following and the nurse doubled vasopressin and levophed drips, adding amiodarone and dc¿d dobutamine. an echocardiogram repeated but not resulted. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9202 | 1220648-2025-28624 | 22064084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22064084&pc=OZD | 05/01/2025 | 06/20/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655587 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Sepsis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support. the patient was running a fever and was on contact isolation to rule out sepsis and clostridium difficile. additionally, the patient had some runs of ventricular tachycardia (vt) while on support which resolved on its own. the patient was treated for suspected sepsis. the patient continued to run a fever and experienced multiple rounds of vt on a couple days. the patient remains on support. | [5 empty] | Access site bleeding minor: drainage from the impella access site. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. infection/sepsis: the patient continued to run a fever and sepsis was suspected. "an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician physical examination findings. results of special investigations: laboratory test results & imaging studies microbiological culture and sensitivity results. response to previous treatments and therapies. any relevant epidemiological information or exposure history. concomitant devices in use. preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined." vt/vf: the patient had some runs of vt while on support. "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9203 | 1220648-2025-28625 | 22064129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22064129&pc=OZD | 02/09/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella 5.5 device used: ¿required cardioversion.¿. the patient recovered with no sequelae. | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia was not determined.
9204 | 1220648-2025-28626 | 22064274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22064274&pc=OZD | 04/30/2025 | 05/21/2025 |  | 05/21/2025 |  |  | 04/30/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622603 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella 5.5 was placed on a patient with ischemic cardiomyopathy. during support, it was reported that they tried to increase to p8 however continue to have suction on p6. additionally, amiodarone bolus and drip for atrial fibrillation given in addition to one unit packed red blood cells for hemoglobin 7.6. the impella was successfully weaned and explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9205 | 1220648-2025-28627 | 22064649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22064649&pc=OZD | 05/01/2025 | 07/14/2025 |  | 05/21/2025 |  |  | 05/01/2025 | 05/21/2025 |  |  |  |  | 05/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638396 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support, there were not positioning issues due to patient movement. the pump was re-positioned. the patient continued on support and the patients condition was noted to be stable, but poor. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issues most likely was due to patient movement. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28627. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9206 | 1220648-2025-28548 | 22051082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051082&pc=OZD | 05/15/2025 | 10/24/2025 |  | 05/20/2025 |  |  | 05/15/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536154 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a placement signal alarm appeared during impella 5.5 support. the device had been in use for more than thirty days at the time of the alarm. positioning was verified to be correct at the time of the alarm. the pump was replaced the following day in response to the noted issue. there was no patient harm. | [5 empty] | D6b "if explanted, give date" corrected. | E.1 first and last name are unknown of the physician. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | G3 ¿date received by manufacturer¿ corrected. | The investigation of optical signal issue has been completed. the returned pump showed biomaterial on the outflow cage and around the impellor blades. no catheter kinks noted. the pump passed laser continuity test. uson testing showed placement signal out of range when less than 150mmhg exterior pressure was applied. from 150-300mmhg exterior pressure applied, the pump showed approximately 150mmhg less than it was actually experiencing. the log review showed abnormal offset being applied to the pump however the g0 offset was accurate and matched its factory g0. the cause of the optical signal issue was not determined due to lack of clinical information and inconclusive data log review and product investigation. d.9 device returned to manufacturer date added. e.1 first and last name were added.
9207 | 1220648-2025-28549 | 22051192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051192&pc=OZD | 04/28/2025 | 06/12/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607805 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the implant of an impella 5.5 on a patient, repositioning the impella 5.5 away from the mitral valve was not successful. the physician elected to pull the pump out, rewire, and readvance the pump. after this the position was good, and the impella 5.5 started but immediately got saturated flows. the impella 5.5 was exchanged for a new one. the patient survived the procedure. | [5 empty] | Positioning issues. during implant of the pump, they were unable to maneuver it away from the mitral valve, so the pump was pulled out and rewired and advanced into a good position. due to limited clinical details, the root cause of the positioning issues of not being able to maneuver the pump away from the mitral valve cannot be determined. saturated flow. the clinical details state after readvancing the pump, there were immediate saturated flows. the data logs show that before the pump was removed there was motor current spikes. the pump was put back in and there was another motor current spike followed by saturated flows and variation in placement signal. the returned product showed biomaterial sitting on the top of the impeller. based on the data logs and returned product, the root cause of the saturated flow is due to ingested biomaterial. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9208 | 1220648-2025-28550 | 22051271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051271&pc=OZD | 04/30/2025 | 06/12/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654295 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was attempted to be implanted but the medical doctor was unable to advance due to patient anatomy and muscular contraction due to cerebral palsy. impella placement was aborted and the patient's outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump was unable to pass due to patient anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9209 | 1220648-2025-28551 | 22051399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051399&pc=OZD | 04/30/2025 | 06/24/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627414 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted in response to the noted difficulty. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, aborted insertion due to patient's anatomy.
9210 | 1220648-2025-28552 | 22051465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051465&pc=OZD | 04/28/2025 | 06/23/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622798 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during set up of an impella cp pump for support of a 62-year-old male patient, the pump was accidentally started outside the patient causing purge/priming alarms. the procedure was aborted as they could not troubleshoot the issue and the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation of pump did not start and priming issues has been completed. the pump was returned for investigation. no physical damage was present on the returned product. the red lumen was seen partially removed from the pump. the red lumen was removed before testing the pump in a bucket of water. no damage was noted on the red lumen. the pump was run according to specifications in the bucket of water, and the priming steps were successfully performed. only lt and imc logs were provided for the case run. the lt log shows the case start on console ic10749, and a pump stop-motor current high alarm was reported after several repeated case start complete notifications. the imc log was reviewed, and it was observed that there were several back keypress events to go back from step 7 (start impella screen) to step 6 (enter purge fluid information) during the case start, along with a purge pressure open alarm. according to the clinical details, the purge cassette tubing was loose and disconnected from the purge bag during priming. the cause of the priming issue was most likely use related to loose purge tubing connection. the cause of the pump did not start issue was a red lumen removal issue since pump was accidentally started outside of the body without removing the red lumen. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
9211 | 1220648-2025-28553 | 22051488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051488&pc=OZD | 04/28/2025 | 06/26/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  | 05/31/2025 |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355550 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support and approximately eight days after start of the support, the patient had low purge flows and high purge pressures. purge was exchanged for five percent dextrose in water with heparin and tissue plasminogen activator protocol was initiated. purge pressure and purge flows returned to normal ranges after switching over the heparin in the purge system. the patient continued on support. two days later there were suction alarms. with the first instance, the impella position was noted as great and volume looked fine. the next day, it was noted that suction was more volume related and there was no longer an issue. however, overnight the impella was repositioned due to suction alarms which continued throughout the next day morning. volume was added. the patient was eventually listed for a heart transplant and a couple of days later (on day 19 of support) the patient was planned for transplant. the device was explanted with a survived outcome. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. clinical details reported that replacing the purge cassette did not resolve the issue but using tpa and heparin in the purge did. additionally, the logs showed the gradual rise in purge pressure to be consistent with biomaterial build up. based on all this information, the root cause of high purge pressure was most likely biomaterial build up. though clinical details stated that tissue plasminogen activator (tpa) resolved the complication, the choppy spikes and drops in purge pressure as it trended up were most likely caused by a kinked purge line. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9212 | 1220648-2025-28448 | 22051613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051613&pc=OZD | 04/06/2025 | 06/03/2025 |  | 05/20/2025 |  |  | 04/08/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612352 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the pump had suction alarms requiring albumin; additional volume was given as well. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used: ¿(b)(6) 2025- (b)(6) 635 (286); iv fluids given, heparin gtt started (b)(6) 2025-suction alarms, given iv fluids (b)(6) 2025-(b)(6) cont to increase (b)(6) 2025-(b)(6) 1207, pfh 40, anti xa 0.4; cvp 2, started miv, remains on (b)(6) @5 (b)(6) 2025 - 5.5 removed in cvor upon oht¿. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and low pump flow has been completed. the impella device was not returned for evaluation. hemolysis: the data logs do not show any motor current spikes before reported increase in ldh levels. there were no positioning issues reported and it was noted there were suction alarms which required albumin; additional volume was also given. the pump was repositioned and ldh levels remained elevated. the root cause of the hemolysis was not determined as no product was returned, the data logs do not provide any definite analysis and limited clinical information related to failure was provided. low pump flow: the data logs ¿do not show any motor current spikes. suction alarms were confirmed. clinical mentioned additional volume was given after suction alarms which resolved the issue. no other suction alarms occurred for the remainder of the case. the root cause of the low pump flow was most likely patient condition related as evidenced by the issue resolution after volume administration.
9213 | 1220648-2025-28554 | 22051762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051762&pc=OZD | 04/28/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654685 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 70 year old female patient began to feel numbness in their impella leg during support. the pump was explanted that same day and the symptoms resolved. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9214 | 1220648-2025-28555 | 22051818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051818&pc=OZD | 04/26/2025 | 06/27/2025 |  | 05/20/2025 |  |  | 04/26/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647893 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported a patient was on impella cp support and during support, there were suction alarms. volume was given and the position was checked. however, further details stated that care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of low pump flow has been completed. the cause of the low pump flow was not determined since product was not returned, data logs were not recovered, and insufficient clinical details. only the cassette was returned for investigation, not the pump. | The purge cassette was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9215 | 1220648-2025-28556 | 22051918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051918&pc=OZD | 04/28/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613563 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that placement signal low alarms occurred with impella 5.5 while supporting a 54 year old male patient and the impella was repositioned. additionally, the patient experienced runs of ventricular tachycardia (vt) requiring defibrillation multiple times. the patient continued on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9216 | 1220648-2025-28557 | 22051948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051948&pc=OZD | 04/28/2025 | 06/20/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650836 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a surgical turndown patient was in for a high risk cardiac procedure with impella cp support. the impella cp was successfully placed prior to intervention to right coronary artery. however, the physician pulled the peel-away back and lost femoral access. the impella cp was removed and new access was obtained for the remainder of the case. the decision was made to leave the left anterior descending artery alone until another time when impella cp support could be re-established. the procedural outcome was the patient survived. | [5 empty] | Positioning issues: the cause of positioning issues most likely was use related since physician accidentally pulled peel away out and pump of artery during procedure to fix the rca. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9217 | 1220648-2025-28558 | 22051994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22051994&pc=OZD | 04/28/2025 | 09/17/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5-CA |  |  | OZD |  |  | IMPELLA 5.5 -CA |  |  | ABIOMED, INC. |  |  | US |  | 2025603895 |  | 0550-0004 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Sepsis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. | The user facility reported that the impella 5.5 had numerous suction alarms occurring over p-7. a decision was made to use tee to determine position of impella given cvp was 11. on tee, the position was noted to be 6.4cm deep. the physician elected to pull the device back about 1.5 cm under tee guidance, measured at 4.7 cm. it was further reported that faulty positioning metric was causing the aic to sense that the impella 5.5 was in the aorta. confirmation of position under tee revealed the impella was in the correct spot and was not abutted against other structures in the heart. repositioning was done to optimal position did not have an effect on the alarm. the alarm was resolved by disabling position metrics. it was further reported that care was withdrawn and the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue, infection, and positioning issue has been completed. data logs were reviewed and the positioning alarms were confirmed and concentration of alarms. returned product was investigated, and all 5s parameters were within specification when connected to a lab console and passed laser continuity testing. there were no abnormalities with returned product, the root cause of the optical signal issue could not be determined. since insufficient clinical details were provided regarding how the pump came out of position and returned product had no abnormalities, the root cause of the positioning issues could not be determined. because this evidence has not been provided, the root cause of the infection cannot be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-28558. d9 device returned to manufacturer date added. h3 device evaluated by manufacturer updated as the initial manufacture report stated no product return which is incorrect. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-28558. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ the patient outcome was reviewed by abiomed's medical safety officer. there is no enough evidence to exclude the impella device as an associated factor to the outcome. there was report of sepsis, and the impella device cannot be ruled out.
9218 | 1220648-2025-28559 | 22052119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052119&pc=OZD | 04/29/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539674 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support for non-st segment elevation myocardial infarction. during support, the patient went into atrial fibrillation. it was also noted that anticoagulant was off due to bleeding. the bleeding slowed. however, heparin was also off. the patient was also septic and vasopressor was off. the patient was on subcutaneous heparin rather than intravenous infusion. there was a concern about whether the pocket had a hematoma or if the pocket was the patients body habitus, but the site soft. the impella 5.5 was successfully weaned and explanted. the patient outcome is survived. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9219 | 1220648-2025-28560 | 22052174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052174&pc=OZD | 04/29/2025 | 06/10/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1754606 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller battery failure (red alarm) triggered after unpacking from its box. several attempts with abiomed's customer service center to resolve the alarm were unsuccessful. | [5 empty] | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. data analysis: the aic logs confirmed the reported failure. there was a battery failure alarm, battery level low, and battery failure alarms was issued just after the aic was turned on. further investigation into the aic logs revealed that the unit had been unused for four months prior to the reported issue. device analysis: the console was tested on a lab bench, where the failure mode was successfully reproduced: the aic would not turn on using only battery power. upon arrival, the batteries were found to be dead, and the logs confirmed lack of use during the four-month period. upon replacing the original batteries with known good ones, there was no battery failure alarm, and the batteries were able to charge and the power battery manager (pbm) was working as per the specification. root cause: the battery failure alarms was due to the depleted batteries. the underlying cause of the battery depletion was likely the extended period of inactivity, as the aic was kept unplugged for four months. | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
9220 | 1220648-2025-28561 | 22052305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052305&pc=OZD | 04/29/2025 | 06/12/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  | 09/30/2025 |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445455 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella cp on a patient suddenly began to experience suction in response to a loss of patient pulsatility and the motor current became flat. despite a return of pulsatility and support level manipulation, the motor current on the pump remained flat and the decision was made to explant the pump. the purge had been changed to bival because it was thought that the patient might have heparin-induced thrombocytopenia and platelets were not better. the pump was explanted and the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of low pump flow and thrombocytopenia has been completed. low pump flow: on the same day as implantation, impella experienced suction events and motor current (mc) became flat. the data logs were recovered and showed suction events throughout case and purge flow trending down with purge pressure ticking upwards. no apparent mc spike outside of p-level changes. the pump was returned and inspected. inspection revelaed. biomaterial wrapped around the impellor blades. low flow was able to be reproduced when pump was run in the lab, where flows were at 1.13l/min for p-6 and expected values for impella cp at p-6 are 2.5-2.9l/min. the cause of the low pump flow was biomaterial due to the biomaterial found wrapped around the impellor blades and low flow being reproduced in the lab. thrombocytopenia: after 2 days on support, anticoagulant changed to bilval because care team thought patient might have hit- platelets not better. the cause of the thrombocytopenia was not determined due to insufficient clinical details.
9221 | 1220648-2025-28562 | 22052320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052320&pc=OZD | 04/29/2025 | 05/30/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1404610 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella automated impella controller (aic) purge fluid driver jammed and was unable to deliver the required flow rate from purge solution. also, the spring-loaded door was not opening despite side button depressed multiple times. it is unclear if the aic was exchanged. the physician commented a delayed start to procedure, despite patient cannulated. it was noted there was no harm to the patient as a result of the event. | [5 empty] | The investigation of automated impella controller (aic) - purge issue - other has been completed. the console logs were consistent with what was reported in the complaint. a single instance of purge system failure (piston blocked) alarm was observed in the logs. inspection of the console revealed that there was dried glucose ingress around the purge assembly and the purge motor sprocket. the root cause of the reported issues was identified as dried glucose ingress, which temporarily caused both the purge motor sprocket and the purge door to become stuck. | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
9222 | 1220648-2025-28563 | 22052490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052490&pc=OZD | 04/29/2025 | 06/12/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  | 04/30/2025 |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338710 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 support and developed a large right axillary hematoma. surgery will evaluate and likely evacuate. the patient continued on support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9223 | 1220648-2025-28564 | 22052519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052519&pc=OZD | 04/22/2025 | 06/20/2025 |  | 05/20/2025 |  |  | 04/22/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632563 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was escalated from a impella cp to a impella 5.5 for escalation of care. during the 5.5 support it was reported that the pump was losing left ventricle signal and waveform in the intensive care unit and eventually motor current was dampened. physician felt uncomfortable with keeping the remaining pump in place and decided to exchange pump to another impella 5.5. upon pump removal, both inlet and outlet had several large clots in place. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of saturated flow and thrombus has been completed. saturated flows: data logs were reviewed, and there were several of the signatures of a biomaterial ingestion event on the reported date. there was a motor current (mc) spike to 1000 ma that correlated with a speed deviation. this was followed by dampened motor current and saturated flows. additionally, a false position in aorta alarm was confirmed to be caused by the ingestion since mc became dampened while the ps remained pulsatile. this is supported by the patient update reporting that position was good at the time of the alarm. after 16 hours, saturated flows resolved. returned product was investigated, and there was some biomaterial residue in the outflow cage. the pump was connected to a console and run in a bucket. initially, pump flows were variable, likely due to the biomaterial, but then normalized to values within specification after it got kicked out. the biomaterial noted on returned product may not have been the ingested biomaterial seen on data logs since saturated flows resolved in the field. based on the signatures of an ingestion in data logs, the root cause of the saturated flows was determined to most likely be biomaterial. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. d9 device returned to manufacturer date added | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9224 | 1220648-2025-28565 | 22052581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052581&pc=OZD | 04/29/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655351 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was used to support a patient with non-st segment elevation myocardial infarction and during a high risk cardiac procedure. the patient had decompensated prior pump implantation requiring resuscitation and a balloon pump placed. during support, the patient experienced runs of ventricular tachycardia (vt). blood pressure and flows remained stable. on a different date it was again noted of episodes of vt, that required shock and that patient was on lidocaine and amiodarone. oxygen was increased. the patient was unstable and required reintubation. the patient was taken down to the catheterization lab to attempt to open the left anterior descending artery with no success. ultimately, care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9225 | 1220648-2025-28566 | 22052610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052610&pc=OZD | 04/25/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/25/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623944 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was escalated from a impella cp to the impella 5.5 as advanced therapies. during support the patient went into ventricular tachycardia and shocked 5 times. it was noted that the impella was close to septum so was repositioned. sitting closer to mitral valve after reposition. the 5.5 was successfully weaned and explanted with bridge to left ventricular assist device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9226 | 1035166-2025-00033 | 22052760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052760&pc=DYB | 05/05/2025 | 06/05/2025 |  | 05/20/2025 |  |  | 05/06/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DOR01986 |  | 0052-3021 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We have been notified of a complaint involving an oscor 23 fr introducer (material #0052-3021, lot # dor01986). the complainant reported that during attempted insertion, the 23 fr introducer sheath cracked. the device was removed and a 23 fr medtronic dilator was utilized for the planned impella rp implant. the customer stated the patient was stable. the device was used in an impella rp implant procedure. no other devices used in the procedure. no medical intervention/additional surgical procedure required. the procedure was completed using another device brand. the device was removed without issue and a medtronic sheath was utilized for the implant. | [5 empty] | The following sections were updated in follow-up 1: b4, g3, g6, h2, h6 and h11. no product was returned to oscor inc. for evaluation device not returned; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. based on the information provided by the supporting documentation, during attempted insertion, the 23 fr introducer sheath cracked. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. based on the investigation, a capa is not required. in addition, there was no new failure mode identified, and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
9227 | 1220648-2025-28567 | 22052768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052768&pc=OZD | 04/30/2025 | 06/24/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  | 11/30/2025 |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024424174 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was being escalated from an impella cp to an impella 5.5 pump. the impella cp had supported for percutaneous coronary intervention of multivessel coronary artery disease. the impella 5.5 access site had developed a large hematoma that was surgically treated and the patient was given two units of blood. the pump additionally was malpositioned along the septum but remained supporting the patient in the intentive care unit for six days. the patient was successfully weaned. | [5 empty] | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding - reported large hematoma forming at the axillary cut down site. activated clotting time (act) was 290. patient received two units of blood and bleeding was surgically resolved. the root cause of the access site bleeding was most likely high patient act values since act was at 290 at the time of the bleed. positioning issue - reported impella towards septum and unable to rotate it away from the septum after delivery. right ventricle (rv) was contracting and extremely small. the root cause of the positioning issue was most likely patient condition related since the patient had an extremely small right ventricle. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9228 | 1220648-2025-28568 | 22052848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22052848&pc=OZD | 04/28/2025 | 07/07/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582558 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately four days after start of the support, constant suction alarms were encountered. the impella cp device was removed, and biomaterial was identified at the inlet (noted as the likely caused the suction alarms). the patient survived the explant and was noted to be in stable condition. | [5 empty] | Additional details noted, aside from the suction alarms, there were no further alarms, and the pump did not stop. as well, the patient had a mitral valve replacement prior to the impella cp device implant. extracorporeal membrane oxygenation was running in parallel with impella mcs, and the activating clotting time was always over 180 seconds. patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | Low pump flow: clinical states that the pump had constant suction alarms and biomaterial chunk was found on the outflow cage at explant. pump was retuned cut from catheter. biomaterial chunk was found on the inflow cage/pigtail. the pump was soaked in warm water to clean the blood. biomaterial traces were observed on the impeller after the pump was soaked. data log were returned. the data logs showed no suction alarms at the end of the case. few suction alarms were present on (b)(6) 2025 and (b)(6) 2025. no motor current spikes or purge changes observed. the flow was in specification throughout the case. therefore, the root cause of the low pump flow issue was most likely biomaterial found on the impeller. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
9229 | 1220648-2025-28569 | 22053021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053021&pc=OZD | 04/30/2025 | 06/26/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Tamponade | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, there were persistent suction alarms. the p level was dropped to p2 with no improvement. the doctor believed there to be clot entrapment in the impella. the pump was removed and replaced. a computed tomography scan was performed and a pericardial thrombus was noticed and was causing tamponade of the heart. the patient's outcome is unknown. | [5 empty] | The investigation of low pump flow, thrombus, and vascular damage - great vessel has been completed. the pump was not returned for investigation. low pump flow: the data log shows low pump flow from the start of the case, followed by a motor current spike and the explant of the pump. the cause of the low pump flow issue was most likely biomaterial, based on data log review and provided clinical details. vascular damage - great vessel: the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between impella and vascular damage was unknown. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9230 | 1220648-2025-28570 | 22053306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053306&pc=OZD | 04/30/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502212 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) had a battery failure. the aic would not hold a charge or charge. troubleshooting by turning the aic off and on from the base unit was sttempted but the console continued to lose the battery and failure alarm remained. the aic was exchanged and no patient harm was reported. | [5 empty] | The investigation for the automated impella controller battery charging issue has been completed. the console logs from the event were consistent with what was reported in the complaint. multiple instances of battery failure (transport connection blown/missing) alarms were observed in the logs. the console was returned for testing. the reported battery failure alarm issue was able to be reproduced. the results of running the batttest utility revealed that cell 1 in battery 2 had a voltage that was significantly less than the voltages of the rest of the cell stacks in the battery. isolative testing was performed by swapping battery 2 with a known good one which resolved the alarm. the root cause of the battery failure was determined to be a cell imbalance in battery 2.
9231 | 1220648-2025-28571 | 22053460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053460&pc=OZD | 04/29/2025 | 06/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the impella began alarming, low purge pressure. purge cassette was changed, d 10 substituted as purge fluid, and console was switched in an effort to alleviate alarm. upon further inspection, once the field was cleared, it was discovered that the purge side arm at the red impella plug was leaking profusely. decision was quickly made to remove pump and replace with new device. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the root cause of the purge leak - pump was most likely due to a damaged sidearm, however, the exact location of the leak could not be definitively determined as no product was returned for evaluation.
9232 | 1220648-2025-28572 | 22053563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053563&pc=OZD | 04/29/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650207 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was placed to support a patient admitted in with known end stage renal disease on dialysis, biventricular failure and deep vein thrombosis. the patient was transferred from the community to tertiary center. the pump was causing ischemia so the team placed an external bypass catheter. support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9233 | 1220648-2025-28573 | 22053619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053619&pc=OZD | 04/29/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885741 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the implant, the impella began alarming, low purge pressure. purge cassette was changed, d 10 substituted as purge fluid, and console was switched in an effort to alleviate alarm. upon further inspection, once the field was cleared, it was discovered that the purge side arm at the red impella plug was leaking profusely. decision was quickly made to remove pump and replace with new device. this complaint was created to document the automated impella controller (aic) exchange. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9234 | 1220648-2025-28574 | 22053730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053730&pc=OZD | 04/30/2025 | 05/29/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593601 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the arterial dupplex on the impella cp left distal extremity showed monophasic flow. a distal perfusion catheter (dpc) to the left superficial femoral artery (sfa) also connected on the impella re-access port. however, the foot remained mottled. lsfa dpc was connected to extracorporeal membrane oxygenation dpc and left foot color and temperature improved. the next day, a fasciotomy was done to the left leg. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 health effect - impact code 4625 was was erroneously entered on the initial manufacturer device report number 1220648-2025-28574. new health effect - impact code entered. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9235 | 1220648-2025-28575 | 22053748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22053748&pc=OZD | 04/29/2025 | 06/09/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643471 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following insertion of the 14 fr sheath, the patient experienced right hip and leg pain. a short run-off confirmed a distal occlusion. an external fem-fem bypass was performed and the symptoms were alleviated. the impella cp pump was then inserted, but decreased flows were encountered. a fluoro revealed a kink in the impella cannula and the decision was made to replace the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of product damage (cannula kink), limb ischemia, and low pump flow has been completed. limb ischemia: as all the necessary information was not provided, the root cause of the limb ischemia was not determined. product damage - cannula kink: the returned product confirmed cannula kink. clinical did not provide any information related to origin of kink. the root cause of the product damage - cannula kink was most likely a cannula kink though origin of kink is unknown due to limited clinical information. low pump flow: the returned product was inspected and there was no biomaterial observed. the pump was run in the lab and low pump flow did not reproduce. the logs do not show any motor current spikes, positioning issues, placement signal trends or continuous suction alarms. the root cause of the low pump flow issue was not determined as no log abnormalities were observed and issue did not reproduce on returned product. limited clinical information related to low pump flow was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9236 | 1220648-2025-28576 | 22054107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054107&pc=OZD | 04/29/2025 | 06/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578426 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 4 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient's extracorporeal membrane oxygenation arterial cannula popped off during impella 5.5 support. the resulting bleed was managed, but the drop in hemoglobin levels prompted blood to be given. anticoagulants were being given at the time of the bleed. eight days into support, the patient stopped moving their left extremity. a computed tomography scan revealed an infarction in the middle cerebral artery. support remained ongoing despite the condition. | [5 empty] | Vascular damage: the root cause of vascular damage is determined to be use issue because of improper ecmo cannula attachment causing the bleed. stroke/cva: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. -timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient. -neurological examination findings and any focal deficits observed. -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). -laboratory test results, including coagulation profiles and cardiac markers. -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿ on (b)(6) 2025: patient¿s arterial cannula popped off due to ¿bad connection ¿. patient bled out for a bit but connection issue was resolved. hgb dropped from 7 to 5. 1 rbc is being given. patient now stable and impella flowing without issue. echo/cxr was done, everything now ok. on (b)(6) 2025: patient bled out for a bit but connection issue was resolved. hgb dropped from 7 to 5. 1 rbc is being given. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9237 | 1220648-2025-28577 | 22054314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054314&pc=OZD | 04/30/2025 | 07/31/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618866 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that seven days after implantation of an impella cp, a pump stop occurred. the pump was restarted with high purge pressure and remained on supporting the patient. | [5 empty] | The investigation of temporary pump stop has been completed. pump was not returned for investigation. data log was not provided for this case run. the cause of the temporary pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | H6 medical device problem code 1491 was erroneously entered on the initial manufacturer device report number.
9238 | 1220648-2025-28578 | 22054410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054410&pc=OZD | 04/30/2025 | 06/09/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026647759 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during impella 5.5 support, the patient had ventricular tachycardia (vt). the patient was cardioverted. the pump position was verified. amiodarone and lidocaine were administered and support was continued. the patient was in stable condition. | [5 empty] | The investigation of vt/vf/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9239 | 1220648-2025-28579 | 22054637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054637&pc=OZD | 04/30/2025 | 06/27/2025 |  | 05/20/2025 |  |  | 04/30/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618819 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Defective Device | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was unsuccessful due to unsuitable anatomy. severe kinking and an aneurysma on both the right and left sides made it impossible to advance the impella. the long peel-away was used to bridge the kinky artery. the force from the implanter lead to demolition of the distal end of the peel-away sheath. the implant was aborted. there was no report of adverse impact to the patient. | [5 empty] | The introducer was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and introducer damage ¿ mechanical has been completed. a data log review was not required for this case. the introducer had a kink and tip split damage. the cause of the introducer damage issue was sheath tip split, which is most likely related to patient's diseased vessels. the cause of the delivery issue was most likely patient condition related to diseased vessels.
9240 | 1220648-2025-28580 | 22054779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054779&pc=OZD | 05/01/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 05/01/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653613 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | An impella 5.5 was implanted in a patient presenting with coronary artery disease and reduced ejection fraction. the patient exhibited chest tube drainage from a pleural tube and the bleeding appeared old. the patient was taken back to the operating room for a chest washout. vasopressors were administered. the patient maintained pressure on minimal support. the decision was made to explant the impella 5.5. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9241 | 1220648-2025-28581 | 22054851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054851&pc=OZD | 04/29/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600633 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with triple vessel disease and needs impella cp support. while on support the patient was having diffuse bleeding from multiple access sites, including hematoma to impella site. impella was sutured to skin and patient transported back to the intensive care unit (icu). while transporting to icu, patient lost native heart pulsatility and impella flows dropped. upon arrival to icu, unable to find palpable pulse and cardiopulmonary resuscitation initiated. patient coded in icu room for approximately 30 minutes. return of spontaneous circulation obtained. during support it was also noted of bleeding continued, not just from impella site, but all access sites. in addition, to unable to doppler bilateral lower extremity pulses. surgical team notified that if patient stabilized, escalation to impella 5.5 may be necessary. however care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9242 | 1220648-2025-28582 | 22054925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22054925&pc=OZD | 05/01/2025 | 07/28/2025 |  | 05/20/2025 |  |  | 05/01/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129496 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during prep the automated impella controller (aic) displayed incorrect/old prompts on the screen. the aic was removed for service. no adverse impact to patient management was reported. | [5 empty] | The investigation of pump not detected has been completed. the pump and data logs were received for investigation. the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9243 | 1220648-2025-28583 | 22055026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055026&pc=OZD | 04/29/2025 | 06/23/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648106 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had a tortuous vasculature resulting in some kinking of the inserted introducer. upon attempting to advance an impella cp pump for implant, the physician was unable to move the impella through the sheath. due to the noted resistance, the decision was made to abort the implant. there was no resulting patient harm. | [5 empty] | The investigation of difficulty inserting/removing pump through introducer and introducer damage ¿ mechanical has been completed. difficulty inserting pump through introducer: the returned sheath was inspected and sheath kink was confirmed. the root cause of the difficulty inserting pump through introducer was determined to be due to a sheath kink as evidenced by returned product analysis and clinical detail mentioning kinking of sheath due to patient anatomy with strong vessel kinking and tortuosity. introducer damage - mechanical: the returned sheath was inspected and sheath kink was confirmed. clinical mentioned kinking due to tortuous patient anatomy. the root cause of the introducer damage - mechanical was determined to be due to a sheath kink as evidenced by returned product analysis and clinical detail mentioning kinking of sheath due to patient anatomy with strong vessel kinking and tortuosity. d9 device returned to manufacturer date added. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28583. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9244 | 1220648-2025-28584 | 22055084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055084&pc=OZD | 05/01/2025 | 06/24/2025 |  | 05/20/2025 |  |  | 05/01/2025 | 05/20/2025 |  |  | 01/31/2026 |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025459856 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with st-segment elevation myocardial infarction was scheduled for implant with an impella cp. while trying to load the abiomed .018 guidewire, the wire was unable to be inserted through the red easy guide lumen. the red flag was pulled and the wire would not pass through the inlet of the cannula. multiple attempts to resolve were performed and eventually the device was aborted and a new impella cp was retrieved. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. clinical states that the 0.018 wire was unable to be inserted through the red easy guide lumen. red flag pulled and wire would still then not pass through the inlet of the cannula. the pump was returned and no abnormalities or damages were observed. a sample lab guidewire was able to be inserted into the pump without issue. the cause of the delivery issue was not determined due to lack of clinical information, red lumen and guidewire not returned for investigation, and sample guidewire able to be inserted into returned cp pump with no issues. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9245 | 1220648-2025-28585 | 22055265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055265&pc=OZD | 04/29/2025 | 06/02/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486654 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiac Arrest | Hemolysis | Ischemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received that the patient could have likely limb ischemia. the team was working on determining a plan of action when the patient arrested, so no interventions had a chance to be performed. | The complainant reported that the patient on impella cp support arrived at another facility with red urine and the cp was repositioned from 6.5cm to 4.5cm. 30 min post repositioning the patient was asystolic. despite multiple rounds of advanced cardiac life support, pulse was resumed briefly and impella repositioned under echo despite optical sensor failure after initial compressions. the patient went asystolic again, and healthcare team unable to regain perfusable rhythm. the patient expired. |  |  |  |  | The investigation of positioning issues, vf/vt/af/at, hemolysis, limb ischemia, and optical signal issue has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis was most likely due to improper positioning based on the clinical details provided. the clinical stated that the impella was measuring 6.5cm. it is unknown the root cause of the positioning issues due to lack of clinical details. the clinical stated that there was optical sensor failure after compressions. however, there was no product or data logs returned so the root cause could not be determined. as the necessary information was not provided, the root cause of limb ischemia and vt/vf was not determined. b2 require intervention was inadvertently omitted on the initial manufacturer device report number. b5 information added that was inadvertently omitted on the initial manufacturer device report number. h6 health effect - clinical code 1942, health effect - impact code 4647 and component code 3038 were inadvertently omitted on the initial manufacturer device report number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9246 | 1220648-2025-28586 | 22055334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055334&pc=OZD | 04/29/2025 | 06/05/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401384 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | Noise, Audible | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs). approximately eight days after start of the support, the automated impella controller (aic) started to have a "loud" fan noise which continued, and consequently, the aic was exchanged without noted incident. the latest update indicates the patient continues on impella mcs and patient status is unknown. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. although console logs from the reported day of event did not contain any obvious indications of fan or power battery manager (pbm) failure, the aic battery and power data embedded in the log for pump 583225 showed that pbm2 charger status was changing frequently between 4 different states: cc11 ac present, battery present, thermistor not underrange, thermistor not hot, charge inhibited. 4c10 battery present, thermistor hot, thermistor underrange. cc10 ac present, battery present, thermistor not underrange, thermistor not hot. cc91 ac present, battery present, thermistor hot, thermistor underrange, charging voltage value invalid. when console was tested in danvers, the issue was not reproduced. there were no observable malfunctions of fans or the pbm. monitoring of communication between batteries and pbm also did not reveal any anomalies. however, knowing that the pbm controls the fan speed and adjusts it to maintain adequate temperature of console components, it is most likely that battery temperature or condition might have been misreported, which sped up the fans to cool batteries down. ultimately, the issue of ¿loud fan noise¿ was not reproduced during testing, but it was most likely caused by miscommunication between batteries and pbm, erroneously indicating overheating condition, which increased fan speed.
9247 | 1220648-2025-28587 | 22055349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055349&pc=PYX | 04/29/2025 | 06/09/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  | 01/31/2026 |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025462237 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella rp flex support and during support, the patient had oozing. an additional suture was placed and the oozing resolved. support continued. it was further reported that the patient expired on support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since tightening the suture and adding additional suture resolved the bleed. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
9248 | 1220648-2025-28588 | 22055569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055569&pc=OZD | 04/28/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/28/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653770 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) was suspected of having an infection. the patient was on intra-aortic balloon pump (iabp) preoperatively and underwent coronary artery bypass graft surgery x 4 in the morning and while weaning cardiopulmonary bypass (cpb), the patient became hypotensive, and the decision was made to place an impella 5.5 device. the impella 5.5 device was successfully implanted under fluoroscopy and transesophageal echocardiogram via the right axillary artery. the position was optimized and the pump turned on to performance level p-6 with flows of 3.5 liters per minute. cpb was weaned and the iabp was removed. the patient was transported to the unit on impella mcs in stable condition. approximately seven days after start of the support, an infection was suspected. blood cultures were drawn for the suspicion, and the plan was to continue to wean as the patient tolerated. the following day, the automated impella controller (aic) started to have a "loud" fan noise which continued, and consequently, the aic was exchanged without noted incident. the latest update indicates the patient continues on impella mcs and the patient's status is unknown. additional information was received that stated it is unknown id the infection was confirmed but the patient was being treated prophylactically with antibiotics. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9249 | 1220648-2025-28589 | 22055674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22055674&pc=OZD | 04/29/2025 | 06/11/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623467 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the 67 year old male patient experienced a gastrointestinal bleed during impella cp support. heparin was put on hold and the patient was taken to the operating room for surgical closure. as support continued, the impella began to experience suction issues. despite attempts to reposition the pump, plans were made to replace the device. | [5 empty] | Vascular damage - peripheral vessel. the clinical stated that on april 29th the patient had a gastric bleed. due to lack of clinical details provided, the root cause of the vascular damage - peripheral vessel cannot be determined. low pump flow. there were lower flows throughout the case as well as lower ecmo flows. suction alarms were seen up until 4/30. the patient updates state that the patient was to have received a bolus of volume and there were no suction alarms after, however lower flows were still noted. flows were only slightly lower then the expected range and there was a trend in the ps. based on the clinical information and data log review, the root cause of the low pump flow is most likely due to the patients condition as they received volume, had lower ecmo flows and there was a trend seen in the ps. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9250 | 1220648-2025-28590 | 22056037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22056037&pc=OZD | 04/29/2025 | 05/20/2025 |  | 05/20/2025 |  |  | 04/29/2025 | 05/20/2025 |  |  |  |  | 05/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658078 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 66 year old male patient on impella support who experienced bleeding at access site. a hematoma was present. vascular was notified and heparin was stopped. in addition, hematuria was present. also reported hemolysis and dark urine. the impella was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9251 | 1220648-2025-28508 | 22043123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043123&pc=OZD | 04/26/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/26/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622599 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient on impella 5.5 pump support. the support was ongoing when there was an observation made of atrial fibrillation on day 5 of the support. the team adjusted the pacer and support continued. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9252 | 1220648-2025-28509 | 22043207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043207&pc=OZD | 04/28/2025 | 06/19/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552315 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed symptoms of hemolysis during impella cp support. the patient's urine become bloody and lactate dehydrogenase rose. hemoglobin and hematocrit were stable. the patient is on dual antiplatelet therapy and systemic heparin. the impella cp was lowered from p7 to p4 and the pump was later explanted in response to the condition. the patient survived. | [5 empty] | The investigation of hemolysis has been completed. data logs showed pump ran without issue during support. visual inspection found delaminated epo-tek on the sleeve bearing of the pump. no other issues were found on the rest of the pump. pump failed hemolysis test. the root cause of hemolysis was most likely due to delaminated epo-tek on the sleeve bearing of the pump. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9253 | 1220648-2025-28510 | 22043254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043254&pc=OZD | 04/26/2025 | 09/24/2025 |  | 05/19/2025 |  |  | 04/26/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654917 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that the engineer indicated that the aorta placement signal went out, however, continued to have a pulsatile motor current. the console was not exchanged due to concerns of the pump not restarting. | It was reported that a patient presented for primary percutaneous coronary intervention with impella 5.5 placement. during support it was reported that the patient experienced some ectopy and ventricular tachycardia post weaning sedation. the patient was successfully weaned and the pump was explanted. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal and cardiac arrythmia. the pump was not returned for investigation. the data log shows that the optical signal 5s parameters temporarily lost to 16384 during the case run, with placement signal (ps) lost to 3000 and controller failure alarm. upon review of data logs, it is apparent that the console is the potential cause of the issue. cardiac arrythmia can be reported during impella support. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28510. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9254 | 1220648-2025-28511 | 22043354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043354&pc=OZD | 04/26/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/26/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had ventricular fibrillation. milrinone and norepinephrine were given. support continued and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9255 | 1220648-2025-28512 | 22043384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043384&pc=PYX | 04/28/2025 | 06/20/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026636065 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced bleeding during impella rp flex support. the bleeding was observed around the catheter. an extra suture was placed. the patient recovered with no sequelae. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of access site bleeding has been completed. the returned pump was visually inspected and no abnormalities were observed. the root cause of the access site bleeding was not determined since insufficient clinical details were provided. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
9256 | 1220648-2025-28502 | 22043436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043436&pc=OZD | 04/27/2025 | 06/20/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655549 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Thrombosis/Thrombus | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and had intermittent episodes of atrial fibrillation. the patient also had micro clots which had gone to extremities. the patient had bilateral dusky fingertips and toes but there was no concern for compromised flow on the impella arm. additionally, the patient had a stroke and could not move their left side. the patient was successfully weaned from the impella. | [5 empty] | Vt/vf- intermittent episodes of afib. "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." stroke/cva: . patient still not moving left side after confirmed stroke. "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿ thrombus: patient showered micro clots. "thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined". | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9257 | 1220648-2025-28513 | 22043591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043591&pc=OZD | 04/30/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654601 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient went into atrial fibrillation. actions taken are unknown. the patient continued on support until the device was successfully weaned. the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9258 | 1220648-2025-28514 | 22043661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043661&pc=OZD | 04/26/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/26/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617599 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 developed gastrointestinal bleeding. systemic heparin was held. it was further reported that the family decided to make patient comfort care and the patient expired. per medical safety officer review there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9259 | 1220648-2025-28515 | 22043844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043844&pc=OZD | 04/26/2025 | 09/24/2025 |  | 05/19/2025 |  |  | 04/26/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654667 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Arrhythmia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that the placement signal reading became inaccurate. support was able to continue despite inconsistencies with the optical sensor. | The complainant reported that an implanted impella 5.5 pump began to experience some suction during support resulting in bouts of ventricular tachycardia. the pump was repositioned in response to the condition. suction alarms continued intermittently, for which the device was repositioned, volume was given, and support levels were lowered. the patient remains on support. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for low pump flow, positioning issue, optical signal issue, and cardiac arrythmia. data logs were reviewed and lt logs showed suction alarms occurring when pump was running above p-4. additionally, persistent "impella position unknown" alarms noted in first half of logs. no motor current (mc) spikes outside of p level changes. placement signal (ps) is slightly trending up. the root cause of the low pump flow could not be definitively determined as no product was returned for evaluation and lack of clinical details. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided on how the pump came out of position. pump had been repositioned multiple times throughout case but ao discrepancy continued. the root cause of the optical signal issue could not be definitively determined as limited clinical details were provided, and no product was returned for evaluation. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28515. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9260 | 1220648-2025-28516 | 22043983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22043983&pc=OZD | 04/30/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633363 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, the patient had a left ventricle thrombus. the patient did not have heparin because of low platelets. it was noted that the device would likely to be explanted, but the patient was noted to be stable. subsequently, heparin was restarted, the device was explanted, and the patient continued on extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves, ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9261 | 1220648-2025-28517 | 22044017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044017&pc=OZD | 05/06/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 05/14/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623424 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the lower limb on the impella side showed a tendency for ischemia. the pump was explanted and replaced with intra-aortic balloon pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9262 | 1220648-2025-28518 | 22044167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044167&pc=OZD | 04/26/2025 | 06/11/2025 |  | 05/19/2025 |  |  | 04/26/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631845 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient awaiting transplant and on the list for a heart was on impella 5.5 support. on day 11 of the support, there was an episoded of ventricular fibrillation that required cardioversion back into sinus, etomidate and amiodarone bolus. the impella 5.5 position was checked and unchanged pre & post cardioversion. no further arrhythmia since has been reported. additionally, the sterile sleeve was found with a tear at the proximal end of the sleeve on the side towards the red impella plug. the team as a result placed tegaderm over tear. the pump remained on despite this. | [5 empty] | H6: no known explant date. the investigation of vf/vt/af/at and product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9263 | 1220648-2025-28519 | 22044277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044277&pc=OZD | 04/27/2025 | 06/25/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501622 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the axillary sheath for 5.5 was damaged out of the package. once placed into the 10mm platinum graft and two graft locks placed there was significant bleeding from perforation points down the sides of the hub. new axillary sheath placed, issue resolved. the patient received 3 units prbcs. additionally, the patients' outputs dropped overnight and suction alarm occurred on p8. lv seemed to sucked down. down to p7 and flolan started to help right to left flow. patient recovered and was eventually weaned off levo. the patient continues on support. additional information received indicating that the patient had a right mca stroke. care was ultimatly withdrawn and the patient expired. | [5 empty] | The investigation of introducer damage ¿ mechanical, access site bleeding and low pump flow has been completed. introducer damage - mechanical: during implantation procedure, axillary sheath for 5.5 was damaged out of the package. once placed into the 10mm platinum graft and two graft locks placed there was significant bleeding from perforation points down the sides of the hub. new axillary sheath placed, issue resolved. the axillary sheath was returned and inspected. visual inspection revealed dried blood at the hemostatic valve of the sheath where the two halves would peel away. the sheath was leak tested where fluid would flow out of the valve at 105.6mmhg, far below the 180mmhg required of the valve. the cause of the introducer damage - mechanical was a damaged valve due to bleeding from the side of the introducer hub and leaking of the introducer reproducing under 180mmhg of pressure. access site bleeding - major: after there was significant bleeding from perforation points down the sides of the hub. new axillary sheath placed, issue resolved. the axillary sheath was returned and inspected. visual inspection revealed dried blood at the hemostatic valve of the sheath where the two halves would peel away. the sheath was leak tested where fluid would flow out of the valve at 105.6mmhg, far below the 180mmhg required of the valve. the cause of the access site bleeding - major was a damaged valve due to bleeding from the side of the introducer hub and leaking of the introducer reproducing under 180mmhg of pressure. low pump flow: after 1 day on support, pt's output dropped causing suction alarms on p-8. down to p7 and flow started to help right to left flow. ef around 20-25%. pt on vent. pump was not returned. data log review showed to separated sets of suction events. both suction events resolved by reducing the p-level from p-8 to p-7, indicating the suction events were triggered by a lack of blood volume in the lv. the cause of the low pump flow was most likely patient condition related due to the suction events resolving when the p-level was reduced. h10 narrative on initial manufacturer device report number 1220648-2025-28519 stated the device was not going to be received, however the device had been received. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9264 | 1220648-2025-28520 | 22044500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044500&pc=OZD | 04/27/2025 | 06/26/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560107 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was in need of impella cp support post intervention. the medical doctor attempted to place the impella sheath from the percutaneous coronary intervention site (right femoral artery). while attempting to dilate for the impella sheath, the wire got bent and as the doctor was trying to remove the dilator, the wire was accidentally pulled, losing access to the vessel. the impella cp placement was aborted. the patient survived the procedure. | [5 empty] | The investigation of delivery issues and product damage (guidewire damage ¿ peripheral vessel) has been completed. the product was not returned for investigation. a data log review was not required for this case run. according to the clinical details, pump delivery was aborted after the doctor accidentally pulled the wire while removing the dilator. additionally, the wire was damaged during the dilation of the impella sheath. no clinical image was provided for wire damage. the reported device part number on delivery issue failure mode was changed from the guidewire to the pump, as the pump's delivery was aborted during the procedure. the cause of the delivery issue was most likely use related since guide wire was damaged and the doctor accidently loose the access during procedure. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9265 | 1220648-2025-28521 | 22044529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044529&pc=OZD | 04/27/2025 | 06/16/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644262 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that multiple complications were encountered involving impella support on a 74 year old male. during setup, the staff followed the appropriate steps, but consistently encountered a "purge fluid not detected" alert. troubleshooting was performed, but ultimately the purge cassette had to be replaced to resolve the issue. following placement of the repositioning sheath, bleeding was noted from around the access site. manual pressure was held and femo-stop was applied, but the issue persisted. upon removal of the dressing and femo-stop, it appeared that the blue suture hub had been pushed into the arteriotomy. the bleeding resolved temporarily with minimal manual pressure, but returned the following day. one unit of blood was given in response to the bleed and both femostop and a mattress suture were applied. support continued with the implanted pump following the intervention. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical information, the root cause of the access site bleeding is most likely due to use as the blue suture hub was pushed too far into the skin. h6 health effect - impact code 4643 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28521.
9266 | 1220648-2025-28522 | 22044610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044610&pc=OZD | 04/27/2025 | 06/22/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644952 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the pump was shallow and was repositioned. additionally, the patient had a hematoma with bleeding at the access site. integrilin was withheld. the patient was successfully weaned from the pump and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft. to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of access site bleeding - major issue most likely was anticoagulation management because bleeding stopped since stopping integrilin. positioning issues: the cause of positioning issues cannot be determined due to lack of clinical details.
9267 | 1220648-2025-28523 | 22044666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044666&pc=OZD | 04/27/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657982 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 93 year old male patient on impella cp support developed gi bleeding. heparin was held. patient outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9268 | 1220648-2025-28524 | 22044905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22044905&pc=OZD | 04/27/2025 | 06/20/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Cardiac Arrest | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the 66-year-old male patient on impella cp support coded. initial rhythm was pea; pulse lost with the patient at bedside with immediate recognition. cpr started and the patient was intubated with return of spontaneous circulation. echo performed after code and the impella was repositioned. the patient was successfully weaned. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. vt/ vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. positioning issue: ¿as per the clinical information provided, the root cause of the positioning issue was use related to cardiopulmonary resuscitation (cpr) given to the patient as a response to patient coding. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9269 | 1220648-2025-28525 | 22045061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045061&pc=OZD | 04/27/2025 | 06/22/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654373 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 67-year-old female patient was on impella 5.5 support and during support, the patient had ectopy/runs of ventricular tachycardia (vt). position was verified and a vasodilator was given to the patient. additionally, suction alarms occurred during support. albumin bolus was given and two units of blood products were given to the patient. furthermore, the patient was heparin-induced thrombotic thrombocytopenia (hitt) positive. the patient was switched from heparin drip to argatroban drip. support continued. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of vf/vt/af/at, low pump flow, and thrombocytopenia has been completed. low pump flow: pump was not returned for investigation. partial data logs were available only for the first start of case up to 2025-apr-24, but the logs around the time of suction event were not provided. there are a few intermittent suction alarms prior to the reported event. the root cause of the low pump flow issue was most likely patient condition related since giving volume to the patient resolved the suction alarms as stated in the clinical. thrombocytopenia and vf/vt: the root cause of the thrombocytopenia and vt/vf were not determined due to insufficient clinical details.
9270 | 1220648-2025-28526 | 22045080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045080&pc=OZD | 04/28/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637056 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella cp explant, the patient experienced ischemia in both their legs. both legs were cool and the patient was taken to the catheterization lab for a thrombectomy on both legs. the patient is in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9271 | 1220648-2025-28527 | 22045262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045262&pc=OZD | 04/28/2025 | 06/23/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642525 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Deformation Due to Compressive Stress | Detachment of Device or Device Component |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that multiple complications were encountered during impella cp support on a patient. upon insertion of an ar mod guide for an angiogram of the right coronary artery, the guide went into the left ventricle and was pulled into the inlet of the impella. both the guide and impella were pulled out of the left ventricle, at which point the cannula on the impella folded on itself. the guide was successfully removed and the staff attempted to also explant the pump. when the impella reached the peel-away sheath, the motor and outlet broke away from the rest of the cannula, leaving the remaining part of the impella in the peel away sheath. vascular was consulted and both the catheter and cannula were removed successfully. the patient is in stable condition. | [5 empty] | Product damage (cannula kink): clinical details report that after a guide catheter got pulled into the impella cannula, both devices were pulled out of the lv. the impella cannula folded on itself, and it was reportedly caused by the interaction with the guide catheter. returned product was investigated and there was confirmation of kinks along the cannula. based on returned product having signs of kinking, the root cause of the product damage was determined to be a cannula kink. based on the clinical details provided, the cannula kink occurred due to an interaction with another device when trying to remove both. difficulty removing pump through introducer: clinical details report that once the ar mod was pulled out of the pump inflow cage, they attempted to remove the pump through the introducer and met resistance. the motor housing ended up separating from the cannula, and vascular had to be consulted to remove the cannula. returned product was investigated and the cannula detachment was confirmed. additionally, there was tip damage to the introducer sheath, with some collapsing inward. it is unclear if the damage to the introducer occurred when removing the ar mod or when trying to pull out a kinked pump. either way, the most likely cause of the difficulty was an interaction between the introducer and the guide catheter causing the introducer tip damage that made it difficult to remove the kinked pump through; the interaction between the pump and guide catheter that resulted in a kinked cannula which was difficult to remove through the introducer; or a combination of both. based on the clinical details provided, the root cause of the difficulty removing the pump through the introducer was most likely due to the interaction with another device. it is not clear which of the device interactions was at the root of the issue. cannula detachment: clinical details reported that initially, the cannula was kinked due to an interaction with another device. then, when trying to remove the pump through the introducer, the cannula then detached from the motor housing. returned product was investigated and there was clear stretching/tearing of the cannula and evidence of glue residue on the motor housing. based on the clinical details provided that the cannula kinked due to an interaction with a guide catheter prior to the detachment while trying to remove the pump through the introducer sheath, the root cause of the cannula detachment was determined to most likely be an interaction with another device. introducer damage - mechanical: returned product was investigated and there was damage to the tip of the sheath that was likely related to the reported complications. for this reason, a failure mode was added. there were no reports of diseased vessels and there were no reports of difficulties during access. based on the clinical details provided, this damage could have occurred due to the interaction with the ar mod that had been pulled into the pump inlet during removal or during the removal of the kinked pump cannula. either way, it was determined that the introducer was most likely damaged during an interaction with another device based on the clinical details provided and the state of returned products. vascular damage - peripheral: clinical details reported that vascular repair had to be done, though the patient remained stable without any major vascular concerns. no further details were provided. there were several different potential complications that could have caused the vascular damage. there was damage to the introducer sheath tip, which could have been caused by the removal of the guide catheter or the kinked pump cannula. the pump cannula also detached from the motor housing during removal. all of these device interactions could play a role in vascular damage. there was no report of diseased vasculature or difficulty during access. based on the complications that occurred during the removal procedure as a result of device interactions and no report of diseased vasculature or access complications at case start, the most likely root cause of the vascular damage was interactions with another device. based on investigation findings after reviewing the clinical data the following were revised b1 revised to both product problem and adverse event h1 type of event revised h6 health effect - clinical code and health effect - impact code revised. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves . ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9272 | 1220648-2025-28528 | 22045514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045514&pc=OZD | 04/28/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603449 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that following successful implant on a patient, the pump was found to be positioned incorrectly in the left ventricle. multiple attempts to reposition the pump were unsuccessful and the staff opted to proceed with support despite the suboptimal position. as the scheduled procedure proceeded, the patient experienced ventricular tachycardia and required defibrillation. the supported procedure was completed,. however, a hematoma was discovered at the impella access site following placement of the repositioning sheath. manual pressure and femostop were applied and anticoagulation was placed on hold to manage the condition. support was continued following the intervention and the patient is in stable condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9273 | 1220648-2025-28529 | 22045527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045527&pc=OZD | 04/27/2025 | 06/22/2025 |  | 05/19/2025 |  |  | 04/27/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 65-year-old male patient was on impella cp support and during support, the patient had hematoma/bleeding at the access site. manual pressure was held, a femstop was applied and heparin was withheld. support continued. additionally, during support, there were suction alarms. cyanokit and albumin were given to the patient. support continued. | [5 empty] | The investigation of low pump flow and access site bleeding has been completed. the pump was not returned for investigation. the data log shows several suction alarms in the initial days of pump start, with noted low pump flow. there was no motor current spike, nor was any positioning issue recorded in the data log. the placement signal was also observed trendy while the flow was dropping intermittently at steady p-levels. according to the clinical notes, cyanokit and albumin were given, resulting in the resolution of the suction alarms. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. the cause of the access site bleeding issue was most likely high patient activated clotting time (act) values. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9274 | 1220648-2025-28530 | 22045679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045679&pc=OZD | 04/28/2025 | 07/07/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1206664 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that shortly after initiating support, a purge system failure alarm appeared on the automated impella controller (aic). the purge cassette was replaced but the issue remained. the aic was subsequently swapped to resolve the failure. there was no patient harm. | [5 empty] | Data analysis: on the complaint date, (b)(6) 2025, one hour after the pump (sn: (b)(6)) was started, the console displayed ¿purge system failure (piston blocked) ¿ alarm,¿ event #417. following this, there was a ¿controller error (pbm voltage out-of-spec) ¿ alarm,¿ event #501, caused by abnormal vpurge voltage from the pbm. the cassette was then changed from s/n (b)(6) to s/n (b)(6), however, event #417 persisted (imc_logs_ic3937.pdf). this confirms the reported failure mode. device analysis: this console was tested upon arrival at fs. the reported failure mode, ¿purge system failure¿ alarm, was reproduced by running the optical test pump and the cassette. although there was no dextrose residue on the purge motor gasket, the console displayed event #417 and event #501. it is likely that the pud drew more current, resulting in a drop in the vpurge voltage and triggering event #501. the vpurge voltage was measured at tp43, fluctuating between 23.7v and 2.4v under ac power. note: when the console was turned on and idle without the cassette and pump, the vpurge voltage remained stable at 23.8v. after replacing the original pud with a known good one, and upon running the test cassette and pump, there were no alarms, and the vpurge voltage remained stable at 23.8 v under ac power. root cause: the root cause of the purge system failure alarm was a pud malfunction. ad hoc task: before returning this console (ic3937) back to the customer, fs was instructed to perform the following, replace the pud pca (0042-1105) due to a malfunction. perform sections 10, 11, 12, 18, 19, and 20 of the pm procedure (0042-7309). perform a full functional check (0042-7316). return the original pud pca back to the engineering for further analysis. | A4 and a5 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9275 | 1220648-2025-28531 | 22045731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045731&pc=OZD | 04/28/2025 | 06/23/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654986 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was implanted to support a patient, along with extracorporeal membrane oxygenation (ecmo), for non-st segment elevation myocardial infarction. the pump experienced a placement signal loss and this prompted the premature explant. the team made the decision to remain on ecmo. no patient harm was reported. | [5 empty] | The investigation of optical signal issue has been completed. optical signal issue: the root cause of the optical signal issue was not determined since the issue could not be reproduced in the lab and data logs were not returned. positioning issues was added as failure mode: clinical details state that the surgeon had difficulty positioning pump and had torqued it multiple times before finally removing the pump and re-wiring. the pump was then re-inserted and positioned without issue. no patient anatomy issues were noted. data logs were not returned. the root cause of the positioning issues was not determined since insufficient clinical details were provided. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9276 | 1220648-2025-28532 | 22045817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045817&pc=OZD | 04/28/2025 | 07/07/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026646377 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female | 90 KG |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp placement, when the guidewire was removed, the impella cannula kinked under the outlet. a suction alarm occurred, and the impella was exchanged. the patient survived. | [5 empty] | Product damage (cannula kink): clinical states that after the guidewire was taken out, the cannula appeared to be kinked near the motor housing of the pump. pump was not returned for investigation. therefore, the root cause of the product damage issue was a cannula kink but the reason for the kink is unknown since pump was not returned for investigation. low pump flow: clinical states that after the guidewire was taken out, the cannula appeared to be kinked near the motor housing of the pump. initially no flow, then suction at p2 with 1l flow. pump was not returned for investigation. data logs were not returned as well. therefore, the root cause of the low pump flow issue was most likely a kinked cannula as per the clinical information provided but rest of the details are unknown. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9277 | 1220648-2025-28533 | 22045829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045829&pc=OZD | 04/30/2025 | 06/24/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618602 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during repositioning, the implanted impella cp pump came back across the aortic valve. attempt to re-advance the pump were unsuccessful and the decision as made to remove the device. there was no patient harm reported. | [5 empty] | The investigation of positioning issues has been completed. the pump was not returned. clinical reports upon repositioning pump it popped out of the left ventricle (lv) and was unable to get it back across valve. the cause of the positioning issue was most likely use related as the pump was accidentally pulled out of the lv during repositioning. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9278 | 1220648-2025-28534 | 22045952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22045952&pc=OZD | 04/30/2025 | 06/12/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648135 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the team was unable to deliver an impella cp pump due to anatomy. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump was unable to pass due to patient anatomy.
9279 | 1220648-2025-28535 | 22046173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046173&pc=OZD | 04/28/2025 | 06/10/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1630913 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the automated impella controller (aic) had a controller error. there was no patient involvement reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench are represented with negative values due to the crc error. device analysis: the console was returned for further investigation. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. additionally, impella connect module flight button had a mechanical issue, needed hard push to make it operate. lastly, the purge system was examined along with the pump connected to the console, and no issues were found. root cause: the cause of the aic controller error alarm and controller error was due to an optical bench fw communication protocol anomaly. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9280 | 1220648-2025-28536 | 22046241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046241&pc=OZD | 04/30/2025 | 06/24/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639417 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Low Blood Pressure/ Hypotension | Diminished Pulse Pressure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during an impella cp procedure post percutaneous coronary intervention to the left anterior descending artery and left circumflex artery, the patient became acutely hypotensive and loss of pulsatile. cardiopulmonary resuscitation (cpr) was performed and the patient was intubated. post cpr the patient was stable but was noted on placement signal conflicting reading versus arterial line and placement signal low not correlating to patient condition or pump position. the decision was made to exchange the pump in concerns regarding pump metrics post cpr. the peel away sheath remained, so the patient was explanted and a new device is to be placed. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of optical signal issue and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9281 | 1220648-2025-28537 | 22046501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046501&pc=OZD | 04/28/2025 | 06/30/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025603446 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 51 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | After case, the physician expressed concern that possible dissection occurred with initial wire insertion versus from device itself. | The complainant reported that following startup, the implanted impella 5.5 pump encountered immediate suction alarms. proper positioning was verified via echo; however, an unknown echogenic area was noted for concern. the pump was subsequently removed and tissue was found on the inlet of the device. the patient survived. |  |  |  |  | The investigation of vascular damage - peripheral vessel, low pump flow, and thrombus has been completed. product damage (cannula kink) was added based on the investigation. product damage (cannula kink): there was a kink found in the cannula and low pump flows seen. due to lack of clinical details, the root cause of the product damage (cannula kink) cannot be determined. low pump flow: after insertion and starting the pump, there were immediate suction alarms and low flow. there was concern of obstruction of the inlet. the data logs do not show an ingestion signature but suction and varying flows with positioning alarms. the returned product showed a significant kink in the cannula. the root cause of the low pump flow is most likely due to a cannula kink based on the data logs and returned product, however it is unknown how the kink occurred. vascular damage - peripheral vessel: due to limited clinical details, the root cause of vascular damage - peripheral vessels cannot be determined. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined d9 device returned to manufacturer date added b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28537. h6 health effect - clinical code 3160 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28537. h6 medical device problem code 1491 was erroneously entered on the initial manufacturer device report number 1220648-2025-28537. new code added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9282 | 1220648-2025-28538 | 22046566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046566&pc=OZD | 04/30/2025 | 06/19/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629972 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient has three liters negative fluid balance for 24 hours. suction alarms occurred at p-8. the pump was repositioned with no improvement. support was continued. it is unknown if the suction alarms resolved. additionally, the patient had an embolic cerebrovascular accident (cva). the patient had been on and off anticoagulation for a gastrointestinal bleed. support was continued and the patient is in stable condition. | [5 empty] | The investigation of stroke/cva, low pump flow, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. stroke: as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. low pump flow: clinical states that the patient had suction alarms at p-8, pump was repositioned but no improvement was observed. doctor didn't give volume. pump was not returned for investigation. data logs were investigated. there was a motor current spike 1200 ma corresponding to the suction alarms. slight spike of purge pressure and low flows were observed. pump had low flows at p-7 as well. therefore, the root cause of the low pump flow issue was most likely biomaterial as per the data log review. vascular damage - peripheral vessels: clinical states that there was a gi bleed but no other information was provided. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9283 | 1220648-2025-28539 | 22046637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046637&pc=OZD | 04/30/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636009 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was on impella cp support and experienced heavy bleeding. the impella cp was weaned and explanted. the patient's care was withdrawn and the patient expired. abiomed's medical safety officer reviewed the event and determined there is not enough evidence to exclude the impella as an associated factor to the patient's outcome. additional information was received. the patient came into the clinic spontaneously breathing. after first examination, it was decided to do a cardiac catheterization. a pronounced occlusion of the left main stem was noted. the patient presented with an ejection fraction of 25% and a systolic blood pressure well below 90. it was decided to intervene under impella support. during the application, the patient was in need of resuscitation for atrial fibrillation. after about two minutes, the patient had a return of spontaneous circulation. intervention went well under impella protection. stenting of the left main was successful, after which the patient was transferred to the intensive care unit. the patient was very unstable and developed respiratory insufficiency and had to be intubated and ventilated. the patient had an increasing need for vasopressors/inotropes. diffuse bleeding was noted as there was a drop in hemoglobin due to liver failure. there was also bleeding from the insertion site of the impella (despite best practice) so the decision was made to explant the impella cp. the patient died in the course of the afternoon from severity of the shock. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9284 | 1220648-2025-28540 | 22046640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046640&pc=OZD | 04/28/2025 | 06/04/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1701427 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller had a broken dust cover. there was no patient involved in this event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. data analysis: not informative for this investigation. the imc logs did not show any anomalies. device analysis: the console was sent to field service (fs) for further investigation. it was determined that the external optical connector dust cover was broken. root cause: the broken dust cover was likely caused by user-related issues
9285 | 1220648-2025-28541 | 22046852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22046852&pc=OZD | 04/29/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/29/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654480 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had impella cp support being placed for a high-risk percutaneous coronary intervention patient admitted in with cardiomyopathy and other cardiac risk factors. it was reported that the catheter shaft of the impella kept buckling in the abdominal aortic aneurysm and was not able to advance. support was aborted and there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system s: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9286 | 1220648-2025-28542 | 22047053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22047053&pc=OZD | 04/30/2025 | 05/19/2025 |  | 05/19/2025 |  |  | 04/30/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653487 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient became septic and developed thrombocytopenia during impella 5.5 support. the patient's blood cultures came back positive for serratia and one unit of platelets was given in response to the thrombocytopenia. the pump was subsequently explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9287 | 1220648-2025-28543 | 22047104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22047104&pc=OZD | 04/28/2025 | 06/19/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026657553 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Detachment of Device or Device Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had oozing at the access site despite angle matching and a purse string suture. a femstop was placed and blood products were transfused to obtain hemostasis. additionally, the pump was repositioned due to alarms and the anticontamination sleeve was torn. care was withdrawn and the patient expired. | [5 empty] | The investigation of access site bleeding, product damage ( sleeve damage), and positioning issues has been completed. the impella device was not returned for evaluation. as per the clinical information provided, the root cause of the access site bleeding issue was most likely anticoagulation management related due to high patient activated clotting time (act) values. positioning issue: clinical states that there were "impella in ventricle" alarms overnight and the pump was repositioned which helped troubleshoot the alarms. pump was not returned. the reason for the pump position change is not clear. therefore, the root cause of the positioning issue was most likely use related as position issue resolved with troubleshooting during support. product damage (sleeve): clinical states that the sterile sleeve was torn on (b)(6) 2025 but no other supporting information is provided. pump was not returned. therefore, the root cause of the product damage - sleeve issue was not determined since product was not returned for investigation and insufficient clinical information was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9288 | 1220648-2025-28545 | 22047549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22047549&pc=OZD | 04/28/2025 | 06/22/2025 |  | 05/19/2025 |  |  | 04/28/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650906 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 77 year old female patient had an attempted impella cp implant and it was aborted. the impella sheath unable to be placed in right. attempted access in left femoral, but patient found to have iliac stent. stent ballooned but impella still unable to pass and the implant was aborted. | [5 empty] | The investigation of delivery issues and difficulty inserting/removing introducer has been completed. the product was not returned for investigation. data log review was not required for this case run. as per the clinical, impella sheath could not be placed in the right femoral artery, and access to the left femoral artery was unsuccessful due to an iliac stent, which, despite ballooning, still prevented the passage of the impella. the cause of the difficulty inserting introducer issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the delivery issue was most likely patient condition related to diseased vessels. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high-risk cp: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9289 | 1220648-2025-28546 | 22047809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22047809&pc=OZD | 04/22/2025 | 11/30/2025 |  | 05/19/2025 |  |  | 04/22/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Pump | Pumping Stopped | [5 empty] | Hemodynamic instability | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and would experience pump failure due to an accident with the automated impella controller. the patient would eventually expire. the impella cp was successfully placed, and percutaneous coronary intervention was completed via a drug-eluting stent to the left anterior descending artery. that evening an echocardiogram was completed with pump at "good" position, 3.3cm from aortic valve. the next morning discussion was made with cardiologist and cardiovascular and thoracic surgeon, and the plan was noted to escalate to 5.5. the request was made and accepted to transfer the patient to another hospital later in the day. the transfer was completed by flight crew. the aic was attached to side of stretcher and when entering the elevator, the blue portion of the impella plug hit the elevator door and broke off. pump failure occurred; support level diminished to p-0. the patient was in critical condition. consequently, the aic was exchanged and the impella cp device pump 1 was explanted and replaced with another impella cp device pump 2 for continuation of therapy. on the replacement impella cp device pump 2, the patient would develop limb ischemia undergoing a fem-fem bypass and eventually expire three days thereafter. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted there is not enough information to exclude the automated impella controller (aic) as an associated factor to the outcome (patient's demise). the pump stop/failure for the impella cp pump 1 is a serious injury as well as impella cp pump 2 (replacement pump) where the patient developed limb ischemia. limb ischemia can lead to more complications down the line, but the immediate issue is the aic pump failure (and this is what led to or contributed more to the demise outcome). the aic is one of three devices associated with the event. note the following: refer to initial medical device report for the impella cp pump 1 serial number (b)(6) with date of event 22-apr-2025 and patient identifier ending in (b)(6). refer to manufacturing report number 1220648-2025-28324 for the impella cp pump 2 (replacement pump) report. the automated impella controller device was not received from the customer; therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. log review: the console logs show a sudden red "impella stopped (mechanical/electrical issue" alarm in the logs, confirming the complaint. this is followed immediately by a red "impella disconnected (running pump disconnected)" alarm. the console is shut down ~28 minutes later. the console was not returned for the investigation. the cause of the issue was determined to be use related. h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation.
9290 | 1220648-2025-28547 | 22048117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22048117&pc=OZD | 04/22/2025 | 06/10/2025 |  | 05/19/2025 |  |  | 04/22/2025 | 05/19/2025 |  |  |  |  | 05/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645889 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Pump | Pumping Stopped | [5 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and would experience pump failure due to an accident with the automated impella controller. the patient would eventually expire. the impella cp was successfully placed, and percutaneous coronary intervention was completed via a drug-eluting stent to the left anterior descending artery. that evening an echocardiogram was completed with pump at "good" position, 3.3cm from aortic valve. the next morning discussion was made with cardiologist and cardiovascular and thoracic surgeon, and the plan was noted to escalate to 5.5. the request was made and accepted to transfer the patient to another hospital later in the day. the transfer was completed by flight crew. the aic was attached to side of stretcher and when entering the elevator, the blue portion of the impella plug hit the elevator door and broke off. pump failure occurred; support level diminished to p-0. the patient was in critical condition. consequently, the aic was exchanged and the impella cp device pump 1 was explanted and replaced with another impella cp device pump 2 for continuation of therapy. on the replacement impella cp device pump 2, the patient would develop limb ischemia undergoing a fem-fem bypass and eventually expire three days thereafter. | [5 empty] | Pump stop: clinical details note that during patient transfer the patient cable plug hit the elevator door and snapped off, leading to pump failure. data log review did not show any alarms associated with a pump stop at the end of the log. it is unknown if the pump had been transferred to another aic before the pump stop. the pump was returned and evaluated. the piece of the patient cable plug with the connector pins was separated from the rest of the patient cable plug and showed signs of excessive force. all wires soldered to the backside of the pins inside the plug were broken off near the solder joints. the root cause of the pump stop was most likely an electrical open since the wires soldered to the backside of the pins inside the plug were broken off on the returned product and clinical details note pump failure when the plug was snapped in the field. | Medical safety reviewed the event and noted there is not enough information to exclude the automated impella controller (aic) as an associated factor to the outcome (patient's demise). the pump stop/failure for the impella cp pump 1 is a serious injury as well as impella cp pump 2 (replacement pump) where the patient developed limb ischemia. limb ischemia can lead to more complications down the line, but the immediate issue is the aic pump failure (and this is what led to or contributed more to the demise outcome). the impella cp pump 1 is one of three devices associated with the event. note the following: -refer to manufacturing report number 1220648-2025-28546 for the aic report. -refer to manufacturing report number 1220648-2025-28324 for the impella cp pump 2 (replacement pump) report. the impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
9291 | 1220648-2025-28464 | 22034441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034441&pc=OZD | 04/25/2025 | 10/24/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  | 11/30/2025 |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024433064 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | Device Sensing Problem |  |  |  |  | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that a placement signal unreliable alarm sounded which resolved after shifting patient position. | The complainant reported that a patient was implanted with an impella rp flex and exhibited ventricular tachycardia, and amiodarone was given. additionally, the patient had pink/red urine when attempting to increase p-levels. the pump was noted to be positioned deep, with the inlet crossing the tricuspid valve, and the pump was kinked at the outflow cage. the pump was repositioned. the patient survived. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella rp flex returned for evaluation. optical signal issue: clinical details noted that the patient had transferred hospitals and pump position was deep and "placement signal not reliable" alarms were noted for a few minutes but resolved after patient position was adjusted. upon visual inspection, no abnormalities noted on returned product. pump 5s optical parameters were within normal range. data logs were reviewed and cvp was trending down with decreased (snr) before triggering psnr alarms. this lasts about 8 minutes before snr and cvp were able to recover and remain stable for remainder of case. raw optical signal was affected at time of psnr alarm but raw placement signal and mc were stable. the root cause of the optical signal issue could not be determined per returned product, data log review, and limited clinical details. positioning issues: clinical details noted that patient transferred hospitals and pump position was deep on x-ray but position was stable. the root cause of the positioning issues was most likely due to patient movement as mal-positioning of pump was found after patient transferred hospitals. hemolysis: returned product showed no damage to impeller blades. biomaterial was present on bottom of impeller blade. data logs were reviewed and a motor current spike was present on the morning of reported hemolysis on 26-apr-2025, with a speed deviation, drop in pump flows, and step up in placement signal at time of ingestion signature. additionally, clinical details noted that patient was experiencing pink/red urine when pump level was attempted to be increased. the root cause of hemolysis was most likely due to biomaterial per data log review and biomaterial on returned product. product damage (cannula kink): clinical details noted that pump was deep and cannula was kinked near outflow cage. no cannula kink was able to be identified on returned product. the root cause of the product damage - cannula kink could not be definitively determined due to lack of clinical information. vf/vt: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | D6b "if explanted, give date" corrected. | Impella rp flex returned for evaluation. optical signal issue: clinical details noted that the patient had transferred hospitals and pump position was deep and "placement signal not reliable" alarms were noted for a few minutes but resolved after patient position was adjusted. upon visual inspection, no abnormalities noted on returned product. pump 5s optical parameters were within normal range. data logs were reviewed and cvp was trending down with decreased snr before triggering psnr alarms. this lasts about 8 minutes before snr and cvp were able to recover and remain stable for remainder of case. raw optical signal was affected at time of psnr alarm but raw placement signal and mc were stable. the root cause of the optical signal issue could not be determined per returned product, data log review, and limited clinical details. positioning issues: clinical details noted that patient transferred hospitals and pump position was deep on x-ray but position was stable. the root cause of the positioning issues was most likely due to patient movement as mal-positioning of pump was found after patient transferred hospitals. hemolysis: returned product showed no damage to impeller blades. biomaterial was present on bottom of impeller blade. data logs were reviewed and a motor current spike was present on the morning of reported hemolysis on (b)(6) 2025, with a speed deviation, drop in pump flows, and step up in placement signal at time of ingestion signature. additionally, clinical details noted that patient was experiencing pink/red urine when pump level was attempted to be increased. the root cause of hemolysis was most likely due to biomaterial per data log review and biomaterial on returned product. product damage (cannula kink): clinical details noted that pump was deep and cannula was kinked near outflow cage. no cannula kink was able to be identified on returned product. the root cause of the product damage - cannula kink could not be definitively determined due to lack of clinical information. vf/vt: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual, section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
9292 | 1220648-2025-28465 | 22034476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034476&pc=PYX | 04/25/2025 | 09/18/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146396 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller was exchanged due to placement signal unreliable alarm. the alarm displayed for 10 minutes and resolved after shifting patient position. the patient survived. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. the associated impella rp flex was reported in manufacturer device report number 1220648-2025-28464. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. there was no log record indicating the ¿placement signal unreliable¿ alarm. this console was tested. no optical signal-related issues were found. the analog output showed no issues and matched the known good analog output from the test fixture. the 5s parameters and bulb power were within specifications as per the preventative maintenance procedure. the root cause of the "placement signal unreliable" alarm was not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
9293 | 1220648-2025-28466 | 22034545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034545&pc=OZD | 04/25/2025 | 06/17/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614236 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was being placed for a patient admitted in with non-st elevated myocardial infarction and other cardiac risk factors. the impella 5.5 loaded over the 0.018 wire but was unable to pass through the axillary artery. the surgeon removed the impella 5.5 and used an impella cp through the graft attached to the axillary artery. the patient's outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through the vasculature.
9294 | 1220648-2025-28467 | 22034621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034621&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026627618 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp experienced a dissection to the right axillary. a peripheral balloon tamponade was performed with a covered stent placed to resolve the dissection, and no further intervention was required. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9295 | 1220648-2025-28468 | 22034627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034627&pc=OZD | 04/25/2025 | 06/12/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639887 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the impella cp was pulled into the aorta after the peel away was taken out. the impella cp was explanted and replaced. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | No product was returned for evaluation. data logs were reviewed and "impella in aorta" alarms were triggered. clinical details noted that when removing peel-away sheath, pump was pulled into aorta. the root cause of the positioning issues was most likely use related as the pump was pulled into the aorta when removing the peel-away sheath.
9296 | 1220648-2025-28469 | 22034666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034666&pc=OZD | 04/25/2025 | 07/28/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1663415 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA PUMP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller generated a purge disc failure. the team replaced the controller and no harm or injury noted from any support delay. patient data was not provided by the complainant. | [5 empty] | A1, a2, a3, a4, and a5 are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: console logs from the reported day of event do not contain any information relevant to this failure mode. no issue were found regarding the purge cassette and pump being not detected. additionally, no logs were available for the other device (b)(4) for the pump sn: (b)(6). note: console's serial number was confirmed with the coordinator and pump used for the case. device analysis: the console (b)(6) was sent back to fs for further investigation. the reported issue was not reproduced; the console successfully detected the pump and purge cassette, as well as the purge disc. furthermore, the console was tested with a known good purge cassette and pump, and no other issues were observed. the purge pressure accuracy was measured and was found to be within specification limits. root cause: the root cause of the problems detecting purge cassette and pump was not confirmed (no problem found). corrective actions: before returning the console to the customer, the following actions are instructed to perform: perform sections 10 - 16 of the pm procedure (0042-7309). perform a full functional check on the console and optical system (0042-7316).
9297 | 1220648-2025-28470 | 22034710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034710&pc=OZD | 04/18/2025 | 07/16/2025 |  | 05/16/2025 |  |  | 05/14/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623420 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp placed along with extracorporeal membrane oxygenation for cardiomyopathy and ventricular tachycardia (vt) storm. the patient was seen to be thrombosed at the atrioventricular valve and aorta. however, the patient expired. the cause of death is not yet determined. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source selection foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28470. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9298 | 1220648-2025-28471 | 22034800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034800&pc=OZD | 04/25/2025 | 06/11/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026631217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 53 KG |  | Asian | Malfunction | No | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was placed on bipella support after experiencing ventricular tachycardia. it was noted that the impella cp was placed too deep into the ventricle. it was pulled back across the valve and kinked. therefore, the pump was removed and the case was aborted. | [5 empty] | The clinical stated that the patient had a small arch. however, it was discovered the pump was placed too deep into ventricle. pulled back and pulled across valve and kinked. unable to move back into ventricle during positioning after placement in cath lab. the root cause of the product damage (cannula kink) is most likely due to use as the pump was placed too deep and it was pulled back and kinked on itself. the root cause of the positioning issue is most likely due to use as it was stated the pump was placed too deep. | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9299 | 1220648-2025-28473 | 22034970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034970&pc=OZD | 04/25/2025 | 06/09/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615947 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient suffered a massive stroke during impella 5.5 support. the decision was made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | Stroke/cva: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. -timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient. -neurological examination findings and any focal deficits observed. -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). -laboratory test results, including coagulation profiles and cardiac markers. -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿ | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9300 | 1220648-2025-28472 | 22034975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22034975&pc=OZD | 04/25/2025 | 06/23/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026643527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that oozing was noticed after impella cp implantation. the blue suture hub was advanced securely to the patient's skin and the perclose sutures were tightened to resolve the issue. additionally, the patient had a seizure episode which resulted in the impella becoming malpositioned, and impella alarms occurred. the impella was repositioned to resolve the alarms. the patient survived. | [5 empty] | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use related since the bleeding was resolved after perclose sutures were tightened. the root cause of positioning issue was most likely patient movement related since the issue occurred while the patient was having seizure. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28472. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9301 | 1220648-2025-28474 | 22035134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22035134&pc=OZD | 04/25/2025 | 06/19/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245459 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) will not turn on. when it is turned on, it appears to be loading but does not actually load. troubleshooting was unsuccessful. the aic was removed from service and there was no patient harm. | [5 empty] | Data analysis: unable to download the logs from the console, as it couldn¿t boot a hard drive or enter service mode. this aligns with the reported failure mode. device analysis: this console was tested after arriving at fs. the reported failure mode was reproduced - bootup issue (fm_reproduced_(b)(6)). upon moving the original flash cards to the known good test fixture, the bootup issue was reproduced, indicating corrupted flash cards. upon mounting the known good flash cards into the (b)(6), the bootup issue did not occur. additionally, the console was power cycled 100 times with the known good flash cards, and it was able to boot up as intended. there was no flash card corruption. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). root cause: the root cause of the bootup issue was corrupted flash cards. ad hoc task: before returning this console ((b)(6)) back to the customer, fs was instructed to perform the following, replace the compact flash cards (2025-0019) due to corruption. replace the front panel optical connector (0042-2736) due to debris. perform sections 13, 14, 15, 16 and 23 of the pm procedure (0042-7309). perform a full functional check (0042-7316). | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9302 | 1220648-2025-28475 | 22035220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22035220&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632540 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 51-year-old male patient with known aortic aneurysm was transferred for advanced cardiac therapies and taken to the or for impella 5.5 placement, ecmo decannulation, intra-aortic balloon pump (iabp) removal, and tracheostomy. it was noted that the impella 5.5 placement was successfully placed. va ecmo was decannulated surgically repaired and a wound vac inserted.the left groin iabpwas turned off, removed, and site preclosed. the patient was then prepped for tracheostomy. upon trach exchange there was an absence of ventilation for 10 seconds. patient went into ventricular tachycardia (vt), hypotensive and consequently defibrillated. vasopressin 2 units were administered with compressions and patient recovered. impella turned back up to p5. during support it was reported that the patient was in/out of supraventricular tachycardia (svt). in addition, the patient was recannulated after having arrhythmias and drop in mixed venous. the patient had paradoxical breathing and was not following commands. a head ct revealed diffuse encephalopathy. the patient's neuro status was not changing and eventually care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)".
9303 | 1220648-2025-28476 | 22035264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22035264&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553806 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp for a high risk cardiac procedure. the impella sheath was placed after the percutaneous coronary intervention sheath had kinked and caused bleeding. the impella cp was supporting when there was an observation made of new ventricular tachycardia which was treated by medical therapy and pump repositioning. the patient was eventually weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9304 | 1220648-2025-28477 | 22035291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22035291&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654475 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 79-year-old male patient had an attempted impella 5.5 implant and the pump would not pass through the patient's anatomy. the pump was removed and debris was noted on the pump after removal. the pump was flushed and became contaminated. the implant was aborted a new pump was used. the patient's outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9305 | 1220648-2025-28478 | 22035965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22035965&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Migration or Expulsion of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump and ecmo supporting a 69 year old male patient admitted in with a st-elevated myocardial infarction and multiple cardiac and systemic comorbidities. the pump came out of position and was removed/replaced. support continued after replacement. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. d.4 lot number, expiration date and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown.
9306 | 1220648-2025-28480 | 22036010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036010&pc=OZD | 04/26/2025 | 05/28/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and had multiple arrhythmias and rapidly decompensated requiring an increase of pressor support. the patient needed escalation of care. however, the family chose to transition to comfort measures and the patient expired. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9307 | 1220648-2025-28479 | 22036026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036026&pc=OZD | 04/25/2025 | 06/09/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639796 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump could not be implanted on a patient due to the patient's extremely tortuous iliacs. the implant was aborted. there was no patient harm. | [5 empty] | Procedure aborted due to extremely tortuous iliacs. the root cause of the delivery issue was most likely patient condition related, impella procedure aborted due to extremely tortuous iliac. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9308 | 1220648-2025-28481 | 22036151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036151&pc=OZD | 03/18/2025 | 09/22/2025 |  | 05/16/2025 |  |  | 03/18/2025 | 05/16/2025 |  |  | 03/31/2025 |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024320157 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered an optical signal issue with an impella 5.5. the optical sensor was unreliable and was reading falsely elevated readings. however, the motor current was stable and flows were accurate. the pump was also in good position. additionally, it was noted that the device was not properly unloading. low flow issues were encountered prompting pump removal. upon removal, a thrombus was confirmed in the inlet of the device. the patient survived the event and a new 5.5 pump was implanted. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. neither data logs nor the pump were returned for investigation. due to limited clinical details and product not being returned for investigation, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28481 b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
9309 | 1220648-2025-28482 | 22036334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036334&pc=OZD | 04/24/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured cardiac arrest with a "probable" relationship to the impella device used: ¿due to cardiac arrest, both impellas were removed and the sheaths removed. the 14fr impella sheath advanced into the right femoral artery and the pigtail catheter advanced into left ventricle. cpr was initiated and the impella wire inserted. the impella cp advanced into left ventricle. impella was turned on. impella rp was removed. the patient had return of spontaneous circulation achieved. the impella cp removed¿. the patient recovered with no sequelae. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1-a.5 are unknown. d.4 model number, catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
9310 | 1220648-2025-28483 | 22036555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036555&pc=OZD | 04/26/2025 | 06/09/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515227 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Thrombocytopenia | Ectopic Heartbeat | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella 5.5. during support, it was reported that system heparin was on hold due to chest tube bleeding. the next day, it was noted that chest tube output increased and bleeding was from the dialysis line. electrolytes were received. one unit of platelets, one unit of cryoprecipitate, one unit of red blood cell, and one unit of fresh frozen plasma were given. additionally, the patient developed left upper extremity (lue) ischemia that had progressively worsened. the registered nurse noted that the lue was done the day before, and there was no evidence of thrombus. the impella graft was on the ascending aorta and right supraclavicular exit. no lines were currently in the lue. it is unknown if there was an art line in place in lue at any time. there was a call regarding the purge flow blocked. they had been watching closely since the purge flow had been around 3 milliters per hour. the team had switched; the purge cassette and flow improved slightly to 1.8-2 milliters per hour. the patient was also diagnosed heparin-induced thrombocytopenia positive. they also reported "purge pressure high" alarm. tissue plasminogen activator was started in the purge. purge flow/purge pressure was back in the normal range. the patient became tachycardic into the 120's and now 140's. the patient was weaned off epiphrine. however, amiodarone was started due to some occasional ectopy competing with paced rhythm. the impella was successfully weaned and explanted. | [5 empty] | The investigation of vascular damage - peripheral vessel, limb ischemia, high purge pressure, thrombocytopenia, and vf/vt/af/at has been completed. high purge pressure: the returned product was inspected and a block of calcium was observed blocking the purge gap. there were no other physical abnormalities with the pump. the pump was run in the lab at p9 and the high purge pressure did not reproduce. the purge pressure was in the 800s (mmhg) and there was a gradual rise in purge pressure until high purge pressure alarms and purge system blocked alarms were triggered. there was high purge pressure for about 2 hours until it resolved to normal ranges. clinical mentioned tissue plasminogen activator (tpa) protocol was initiated. the root cause of the high purge pressure was determined to be biomaterial as evidenced by the gradual rise in the data logs and the lack of high purge pressure reproduction on the returned product. vascular damage - peripheral vessel, limb ischemia, thrombocytopenia and vt/vf/at/af: the cause of these issues were not determined due to insufficient clinical details. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9311 | 1220648-2025-28484 | 22036575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036575&pc=OZD | 04/25/2025 | 06/17/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620922 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported 51 year old male patient on impella cp support and during support, the impella access site began to ooze overnight during patient support. one unit of blood was transfused and gauze was placed under the cp site to maintain angle matching position. the patient was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is most likely use issue because gauze is placed under cp insertion site to maintain angle matching to manage bleeding.
9312 | 1220648-2025-28485 | 22036622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036622&pc=OZD | 04/26/2025 | 06/19/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492219 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Ventricular Fibrillation | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during the implant and during support, the patient had ectopy and ventricular tachycardia/ventricular fibrillation. amiodarone was given to the patient. additionally, the pump appeared deep during support. it is unknown if the pump was repositioned. the procedural outcome was the patient survived the surgery. | [5 empty] | Vt/ vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. positioning issue: clinical states that the pump was found to be seating deep but clear of the mitral valve after vt/ vf runs. no other information related to the positioning of the pump is provided. pump was not returned for investigation. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9313 | 1220648-2025-28486 | 22036819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036819&pc=OZD | 04/19/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/19/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630513 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was escalated from a impella cp to the impella 5.5 due to worsening condition. during support on the 5.5 it was noted that heparin was discontinued due to concerns of bleeding. the 5.5 was repositioned. additionally, a computed tomography (ct) scan revealed a small subdural hematoma. neurosurgery consulted with no intervention needed. a repeat ct showed stable subdural and multiple scattered small infarcts throughout the brain. bival was put on hold. ultimately, the patient expired on support. a medical safety officer reviewed this case and determined the demise outcome is associated with the impella cp device accordingly under manufacture report 1220648-2025-28270. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a medical safety officer reviewed this case and determined the demise outcome is associated with the impella cp device accordingly under manufacture report 1220648-2025-28270. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9314 | 1220648-2025-28487 | 22036829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22036829&pc=OZD | 04/26/2025 | 06/19/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632642 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, there were suction alarms. albumin bolus was given to the patient. support was continued. additionally, the patient had a head bleed. heparin was withheld. computed tomography showed a cerebrovascular accident (cva). the patient was recovering. | [5 empty] | The investigation has been completed. clinical states that there were suction alarms on (b)(6) 2025 and albumin was given. pump was not returned for investigation. data logs confirm suction alarms on (b)(6) 2025 which are resolved shortly after. no motor current spikes or placement signal trend observed. there are no other suction alarms after this event. the root cause of the low pump flow issue was most likely patient condition related as the suction alarms were resolved after albumin was given per data log and clinical review. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9315 | 1220648-2025-28488 | 22037007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037007&pc=OZD | 04/29/2025 | 06/05/2025 |  | 05/16/2025 |  |  | 04/29/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1668177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported by the biomed engineer that the automated impella controller screen froze on the "please wait" screen when attempting to change the cassette. the controller was sent in to field service for investigation and there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. console logs from the reported day of event are mostly consistent with complaint and show that on 2025-04-25 error messages ¿imc-ahp error: 0xff 0x8f 0xdc¿ were recorded in the log, indicating possible rfid issues, while purge cassette was presumably present. at the onset of purge-related error messages, purge motor shaft was no longer advancing the piston, purge pressure started decreasing and calculated purge flow started showing unrealistic values of 300-400ml/hr, as evidenced in real time (rt) log data. six minutes later, a purge change wizard was launched via on-screen menu, however the purge would not start, and ¿complete the procedure¿ alarm (420, due to user interaction stopped) was presented and would not resolve. eventually pump (b)(6) and purge cassette sn (b)(6) were transferred to a backup console (ic1141), where purge issues did not repeat. console was tested, but the issue was not reproduced. purge system was able to correctly detect a known-good purge cassette, and no rfid reading errors were observed. purge motor would start without problems, and correct values of purge pressure were measured. monitoring of rfid-pud serial communication did not reveal any anomalies. it is most likely that reported purge issue was caused by malfunction of the purge system (communication failure of either pud or rfid). exact root cause was not identified due to inability to reproduce. (aic) - display issue: there was no issue found during the case. the device functioned/operated to specification. aic - purge issue - other: the cause of the aic issue was a defective purge assembly. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9316 | 1220648-2025-28489 | 22037110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037110&pc=OZD | 04/29/2025 | 07/28/2025 |  | 05/16/2025 |  |  | 04/29/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1736233 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was skipping prompts at the start up of the process. the aic was exchanged and there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9317 | 1220648-2025-28490 | 22037229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037229&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527813 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella was placed on a 60 year old male patient and during the cardiac procedure, the patient had lots of ectopy and went into ventricular fibrillation (vf) multiple times and was shocked. during support, the patient experienced suction alarm and was given bolus, however events of suction continued through out support. additionally, it was noted that due to compartment syndrome in left lower extremity post cannulation and a fasciotomy was performed. the cp was explanted and a 5.5 was implanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9318 | 1220648-2025-28491 | 22037455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037455&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/25/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654965 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support waiting for an left ventricular assist device, had runs of ventricular tachycardia (vt) overnight and the impella was repositioned. on day 16 of impella support the patient had vt ablation. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9319 | 1220648-2025-28492 | 22037458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037458&pc=OZD | 04/29/2025 | 06/16/2025 |  | 05/16/2025 |  |  | 04/29/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625147 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was not advancing through the peel away sheath due to vessel tortuosity. the physician changed to a competitor 14 french sheath, but the impella catheter kinked. the impella insertion was aborted. there was no patient harm. | [5 empty] | Impella cp and 14fr x 25cm introducer were returned for evaluation. product damage (cannula kink): upon visual inspection, impella cp had a cannula kink present just distal of the motor housing. clinical details noted that a kink occurred when attempting to insert pump through a gore sheath. the root cause of the product damage was due to a cannula kink as the kink occurred when attempting to insert the pump through a gore sheath which is not an abiomed-approved introducer per ifu 0048-9007 rr. difficulty inserting/removing pump through introducer: clinical details noted that when using 14fr x 25cm impella introducer, the pump was unable to be advanced through the introducer due to patient's tortuosity. returned product showed a kink in the introducer near the distal tip of the introducer. the root cause of the difficulty inserting the pump through the introducer was most likely due to a sheath kink as a result of the patient's tortuous anatomy. delivery issues: clinical details noted that the 14frx25cm abiomed introducer was replaced by a 14fr x 33cm gore sheath for pump insertion but pump was unable to pass through gore sheath hemostatic valve. the root cause of the delivery issues was a use related issue as the introducer used was not an abiomed approved introducer per ifu 0048-9007 rr. introducer damage - mechanical: analysis of returned product showed a kink present in the 14fr x 25cm introducer close to the distal tip of the introducer. a failure mode of "introducer damage - mechanical" was added to the investigation based on damage found on returned product. the root cause of the introducer damage - mechanical was due to a sheath kink based on returned product analysis. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
9320 | 1220648-2025-28493 | 22037723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037723&pc=OZD | 04/29/2025 | 06/19/2025 |  | 05/16/2025 |  |  | 04/29/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557264 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, distal perfusion to the impella leg was absent. a distal perfusion sheath was placed to restore flow to the limb. as support continued, ongoing suction alarms appeared. an ultrasound showed a large pericardial effusion and a perforation was suspected. upon pulling the impella back, the pigtail appeared to be caught on a structure and straightened. when it came free, the effusion continued to reaccumulate at a faster rate. the family made the decision to withdraw care and the patient passed. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | Ventricle perforation: clinical details note that a large pericardial effusion was found after the patient had gotten more unstable with very low pulsatility, decreased flows, and suction alarms. it is unknown when the effusion began. the effusion kept reaccumulating even after drainage, signaling a rv or lv perforation. the impella pigtail appeared to be caught on a structure, which became free after pulling the pump back. the timing of the repositioning is unknown. the effusion reaccumulated at a faster rate after this and flows began decreasing again. clinical details noted that it was difficult to rule out the impella as the cause of the perforation since impella insertion was the only procedure performed this day; however, clinical detail also notes the infiltrative myocarditis may have made the lv tissue more friable. data log review showed no positioning alarms. the returned product did not have any cannula kinks. the root cause of the ventricle perforation was not determined since insufficient clinical details were provided. low pump flow: clinical details note the patient had gotten more unstable with very low pulsatility, decreased flows, and suction alarms before a pericardial effusion was found. flows continued decreasing with more suction alarms as the effusion kept reaccumulating. the impella pigtail appeared to be caught on a structure, which became free after pulling the pump back. data log review showed low pump flow during the entire case at different p-levels. it also showed downward trending placement signal and decreasing motor current with stable p-level. there were several suction alarms but no positioning alarms or motor current spikes. the returned product did not have any cannula kinks. there was small biomaterial in the purge gap. low pump flow was not reproduced with the returned product. the root cause of the low pump flow was most likely patient condition related as the patient was unstable during support with very low pulsatility and pericardial effusion based on clinical details and data log review showed downward trending in placement signal. limb ischemia - reversible: clinical details state distal perfusion to impella side leg was absent confirmed by ultrasound. the treatment team decided to place distal perfusion sheaths, which restored flow. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9321 | 1220648-2025-28494 | 22037982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037982&pc=OZD | 04/29/2025 | 06/19/2025 |  | 05/16/2025 |  |  | 04/29/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026638692 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with a cardiac history and transfer from community to tertiary center for escalation of care. the team placed an impella 5.5 in the operating room. the pump position had to be re-adjusted after discharge from operating room to intensive care unit. additionally the patient went into ventricular tachycardia/ventricular fibrillation (vt/vf) and was cardioverted. the patient had an automatic implantable cardioverter-defibrillator already placed but the vt/vf occurred when the impella was sucked down in the left ventricle. the patient remains on support. additional informatin was received from the area clinical consultant that the vt/vf was attributed to running the impella at too high of a speed while the patient was under volume resuscitated. the position was thought to be deep when bedside point-of-care ultra-sonography was done with the fellow but was confirmed to be in good position under formal transthoracic echocardiogram. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. product was not returned for review. the root cause of positioning issue and vt/vf was unable to be determined as limited clinical details were provided and no product was returned for review. d6b explant date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9322 | 1220648-2025-28495 | 22037984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22037984&pc=OZD | 04/26/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656761 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and after the implant, motor current was flat while there was pulsatile placement signal. transesophageal echocardiogram showed proper position. the staff suspected the pump ingested a thrombus. the pump was explanted. the thrombus was not visualized or confirmed. the patient was put on a new impella 5.5 and survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9323 | 1220648-2025-28496 | 22038033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038033&pc=OZD | 04/27/2025 | 06/02/2025 |  | 05/16/2025 |  |  | 04/27/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1754606 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) with a battery issue. the aic was plugged in and still the battery was still reading "low". the team swapped the aic and support was initiated. battery issues were noted prior to support initiation. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. during the data analysis the console immediately after booting up, the console displayed ¿battery failure (transport connection blown/missing) ¿ alarm,¿ this confirms the reported failure mode. the console was tested and though the battery switch was engaged, there was a ¿battery failure¿ alarm immediately after booting up. under ac power, battery 1 is not being charged, as the batttest shows i: 0 ma. additionally, there is a cell imbalance in cell 4 of battery 1. the root cause of the battery failure alarm was cell imbalance in cell 4 of battery 1 and cell 1 of battery 2.
9324 | 1220648-2025-28497 | 22038174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038174&pc=OZD | 04/26/2025 | 07/02/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647891 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp purge luer lock was leaking. the cassette was already replaced but there was no improvement. purge filter bypass was successfully created, and pump was explanted the next day. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | Purge leak - pump: clinical details state that there was a leak from the sidearm yellow luer and a purge filter bypass was performed. product was returned and evaluated. the leak was reproduced at the sidearm luer, where a small crack was found. data log review showed purge pressure drop suddenly that was eventually resolved. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc., per supplier evaluation as seen under complaint (b)(4). investigation reports attached to complaint (b)(4).
9325 | 1220648-2025-28498 | 22038251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038251&pc=OZD | 04/26/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/27/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650459 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient lost distal pulses to impella limb and the pump was removed. the patient was taken to the operating room for thrombectomy and fasciotomy. the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9326 | 1220648-2025-28499 | 22038268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038268&pc=OZD | 04/26/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623963 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient suffered a stroke during impella 5.5 support. no intervention was required to manage the condition. three days later, the family opted to pursue comfort care and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions and precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
9327 | 1220648-2025-28500 | 22038482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038482&pc=OZD | 04/25/2025 | 06/17/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644137 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had hemolysis with worsening acute kidney injury. the patient had elevated lactate dehydrogenase. the pump was repositioned due to being in contact with wall of left ventricle. care was withdrawn and the patient expired. | [5 empty] | The investigation of hemolysis, renal failure, and positioning issues has been completed. the impella device was not returned for evaluation. hemolysis: the root cause of the hemolysis was most likely improper positioning since the hemolysis was successfully managed by repositioning the pump per clinical details. renal failure: the cause of the renal failure was not determined due to insufficient clinical details. positioning issues: the root cause of the positioning issues was most likely patient condition related as the patient had an incredibly small ventricle per clinical details. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9328 | 1220648-2025-28501 | 22038614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038614&pc=OZD | 04/27/2025 | 07/02/2025 |  | 05/16/2025 |  |  | 04/27/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657752 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient with an impella cp had an impella in ventricle alarm. the p-level was reduced to p2. echocardiogram was done and the impella was pulled back 2.5 centimeters. additionally, the patient had absent distal pulse via doppler. the physician ordered an arterial us study. the patient survived the procedure. | Additionally the patient experienced ventricular fibrillation/pulseless electrical activity (pea) arrest requiring reintubation and emergent venoarterial extracorporeal membrane oxygenation cannulation. escalated to impella 5.5. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Positioning issue occurred after about 5 hours of implantation and the pump was repositioned. days later into support report more positioning alarms, but unknown other details such as if it was repositioned. ischemia happened on the 2nd day & vf/pea happened a day before escalating to 5.5. positioning: data review: data logs confirmed positioning alarms corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. limb ischemia: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition, any concomitant medication, vascular size or variations thereof, detailed description of the symptoms experienced by the patient, physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. the cause of the vt was unable to be determined due to insufficient clinical details. the failure mode vt/vf/at/af is used as a category for injuries including arrhythmia, asystole, bradycardia, cardiac arrest, heart block, paroxysmal atrial fibrillation, tachycardia, ventricular fibrillation, and ventricular premature complexes. | B5 additional information was inadvertently omitted on the initial manufacturer device report 1220648-2025-28501. h6 health effect - clinical code and health effect - clinical codes were inadvertently omitted on the initial manufacturer device report 1220648-2025-28501.
9329 | 1220648-2025-28503 | 22038881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038881&pc=OZD | 04/25/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637956 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following impella 5.5 implant, an echocardiogram revealed a thrombus in the left ventricle. the pump was pointed away from the mass and the decision was made to keep the pump in the patient for continued support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9330 | 1220648-2025-28504 | 22038891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038891&pc=OZD | 04/27/2025 | 06/25/2025 |  | 05/16/2025 |  |  | 04/27/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618526 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant of an impella 5.5, the pump ingested a clot and experienced saturated flows. a purge blocked alarm appeared shortly after, at which point anesthesia gave another bolus of heparin in an effort to dissolve the possible thrombus. the alarm eventually resolved and tissue plasminogen activator continued to be run through the purge. support continued with the implanted pump following the intervention. | [5 empty] | The investigation of thrombus and high purge pressure has been completed. the pump was not returned for investigation. after pump start, data log shows a small motor current spike while running on steady p-level p8, followed by a 5-minute gradual rise in purge pressure, which resulted in a purge system block alarm. purge pressure gradually returned to normal limits in 2 hours of case run. pump flow was seen saturated during the start of the case, and this resolved as the high purge pressure was resolves. according to the clinical information, purge flow returned to normal after tissue plasminogen activator (tpa) was initiated in the purge line. saturated flow: failure mode was added as an investigated failure mode since the data log confirmed the pump flow saturation during the high purge pressure event. the cause of the saturated pump flow issue was most likely biomaterial ingestion, based on data log review. high purge pressure: the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider.
9331 | 1220648-2025-28505 | 22038950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22038950&pc=OZD | 04/24/2025 | 06/13/2025 |  | 05/16/2025 |  |  | 04/24/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095648 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN: (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during a procedure the soft keys on the automated impella controller (aic) were not working. the aic was successfully exchanged. there was no harm to the patient. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that after step 7 of case start wizard, keyboard detected high guard level due to touch detected outside soft-key area, evidenced by multiple consecutive log entries ¿imc-kbd error: 0xa1 0×01 ¿. during and after these entries there¿s no indication of any soft-key presses recorded in log. console¿s ac cord was disconnected and reconnected (presumably in order to troubleshoot the issue), then pump was disconnected from the console without prior reducing p-level to p-0. console was successfully shut down (soft keys are not involved in this step). console was returned to danvers, and during troubleshooting it¿s been observed that the aic would not power-up when operated on batteries. it¿s been verified that internal dc power supply was producing nominal 24v and pbm seemed operational, as other components - pud, impellatronic, optical bench - were receiving power. monitoring of voltages on pbm responsible for switching on and powering the epc via 4-in-1 carrier revealed some anomalies in pc_on signal. after replacing pbm with a known-good one, issue was resolved. reported issue with non-responsive soft keys was not reproduced during testing, however the ¿high guard level¿ condition could be triggered for example by liquid spill on the console¿s glass kbd assembly, which is known to interfere with proper soft-key detection, temporarily disabling it. no defect of kbd or its communication cable was observed, and root cause of malfunction was not determined. precautionary replacement is recommended. during console testing it¿s been also observed that optical system failed ¿5s¿ parameter testing, due to defect observed in ccd detector output, which was affecting (lowering) snr for certain sensor cavity lengths (pressures). power-on issues and optical bench issue were unrelated to reported complaint. repair: replace pbm assembly, and optical bench. replace glass kbd assembly as a precaution. perform functional check of the console. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9332 | 1220648-2025-28506 | 22039083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22039083&pc=OZD | 04/26/2025 | 09/24/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599810 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that there were placement signal not reliable alarms. | The complainant reported the patient was on impella 5.5 support and during support, the patient had blood in their stool post colonoscopy, care team is waiting for another bowel movement before increasing heparin. the procedural outcome was the patient survived surgery. |  |  |  |  | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: after 11 days on support, patient had blood in stool post colonoscopy. the pump was not returned. there was no issue found during the case. the device functioned/operated to specification. blood in stool is likely unrelated to impella function. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. after 13 days on support, a placement signal not reliable (psnr) alarm occurred. the pump was not returned. data logs recovered showed psnr alarms and controller error alarms during a 19 hour placement signal (ps) blackout period starting on 28apr25. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial. manufacturer device report number 1220648-2025-28506. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9333 | 1220648-2025-28507 | 22039469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22039469&pc=OZD | 04/26/2025 | 05/16/2025 |  | 05/16/2025 |  |  | 04/26/2025 | 05/16/2025 |  |  |  |  | 05/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629959 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had pulseless ventricular tachycardia followed by atrial fibrillation with rapid ventricular response. support continued until successful explant and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9334 | 1220648-2025-28419 | 22024506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22024506&pc=OZD | 05/02/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 05/02/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632468 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support (mcs) and bridge to transplant. approximately 13 days after start of the support, placement signal low alarms with suction were encountered. the alarms were noted to have self-resolved. an ultrasound was completed at bedside and revealed the impella 5.5 was malpositioned. an echocardiogram measured the cannula at 6.7 centimeters from the aortic valve annulus. the surgeon repositioned with echocardiogram at bedside with cannula measurement now 5 centimeters from the aortic valve annulus. the patient continued on impella mcs and was noted for possible durable ventricular assist device placement work up. approximately 10 days later, placement alarms were triggered. an echocardiogram was performed, and the device was measured at 6.58 centimeters. the patient was taken to interventional radiology for a procedure. the impella support was on p-2 and awaiting repositioning. it was later reported that the patient experienced a retroperitoneal bleed and had gone to interventional radiology for cauterization of the bleed noted to be unrelated to the impella 5.5 device. the patient was lethargic and placed on a ventilator over that weekend. the patient was noted on several drips. lactic had a downward trend and the latest echocardiogram measured 5.4 centimeters from the aortic valve annulus. a couple of days later it was noted the patient's condition was noted to be "actively" declining. a decision was made with the heart failure team and family to withdraw care, and the patient would expire after approximately 27 total days on impella mcs. | [5 empty] | As it appears heparin was in use for the impella device (and noted subsequently paused), bleeding is being reported accordingly (even though it was noted the retroperitoneal bleed is unrelated to the impella) as it unknown if the use of the heparin adversely impacted and/or exacerbated the situation. additionally, given the patient¿s demised outcome, there is not enough information to exclude the impella device used as an associated factor given the event. the impella device was not returned, and the investigation has been completed. the root cause of vascular damage peripheral vessel could not be determined since the product was not returned and insufficient clinical details were provided.
9335 | 1220648-2025-28428 | 22025373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22025373&pc=OZD | 05/09/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 05/11/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611240 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient had bleeding at the access site. blood products were given and the pump was replaced using a 16 french sheath. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9336 | 1220648-2025-28429 | 22025908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22025908&pc=OZD | 05/04/2025 | 07/11/2025 |  | 05/15/2025 |  |  | 05/04/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630489 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that a patient was being supported with an impella cp. however, the impella cp was escalated to an impella 5.5. during impella 5.5 support, it was reported that the flow rate suddenly dropped. in addition, at that time, the pump position in the aorta alarm was also frequent. however, purge pressures, purge flow rates, and motor current consumption values were all normal. echocardiography and contrast-enhanced computed tomography scans were used to check the amount of fluid and pump position. however, there was no particular problem. a blood clot in the cannula was suspected and the policy was removed. subsequently, the impella 5.5 was removed and a large amount of blood clots were found in the discharge site. the patient's hemodynamics were maintained with increased catecholamines, but it was determined that continued treatment with the impella was required. an additional graft was anastomosed and it was replaced with an impella 5.5. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Optical signal issue/low pump flow: no product was returned. the data logs show that there are a few small spikes in motor current from (b)(6) to (b)(6). on (b)(6) into (b)(6) the motor current and flow dampen and there are positioning alarms seen. there is a motor current spike around 7pm on (b)(6) and flows become slightly saturated. shortly after, at around 9pm there is another motor current spike but then the flows began to decrease and there are multiple suction alarm and impella position in aorta alarms. flows remain low until the pump is explanted. the clinical stated that there was a large thrombus seen in the outlet cage upon removal. based on the data log analysis the root cause of the low pump flow is most likely due to ingested biomaterial. based on the data log analysis, the root cause of the optical signal issue is most likely due to ingested biomaterial that causes a false triggering of the alarms. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined.
9337 | 1220648-2025-28425 | 22025968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22025968&pc=OZD | 04/24/2025 | 06/26/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561423 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the 62-year-old male patient on impella 5.5 support had ventricular tachycardia (vt) without suction event. no known intervention. on day four of 5.5 support, there was suction after 2l of urine output which resolved by dropping from p-5 to p-4. the patient continued on support. | Code stroke called yesterday evening. left cerebellar cerebrovascular accident (cva), believed to be caused by left vertebral artery, no neurological changes. |  |  |  |  | H6: no explant date ... patient still on support as of (b)(6) 2025. no product was returned as patient still on support. access site bleeding - minor: the root cause of the access site bleeding - minor was most likely patient condition related as clinical mentioned hematoma resolved without known intervention vt: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. low pump flow: no product returned as of 5/15. the data logs confirm suction alarms. there were also impella position unknown alarms but clinical did not suggest any positioning issues. there were no motor current spikes and no placement signal trend indicating ingestion or other abnormality. the root cause of the low pump flow was not determined as no product was returned for analysis and no log abnormalities were observed. stroke/cva: strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. priming issue: no product was returned for analysis. the data logs confirm 'air in purge' alarm. during a cassette change, there was a 'purge line click-on not detected' alarm which was resolved. the 'air in purge' alarm triggered after this cassette change procedure likely as a result of a use-related issue. the use-related issue is unable to be confirmed as imc logs of relevant time frame were not provided but could be due to a clamped purge line or the purge bag not being spiked - this alarm has been reproduced on a working purge cassette in lab testing when these events occurred. clinical mentions the purge cassette was changed and there were no following issues. the root cause of the priming issue was most likely due to a use-related issue as evidenced by the data log observation of the clinical events. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28425. h6 health effect - clinical code 1770 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28425. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28425. h6 type of investigation codes, 4111 and 4114, investigation findings codes 3221 and 114, and investigation conclusions codes 50 and 19 were erroneously entered on manufacturer device report number 1220648-2025-28425-1. h10 priming issues and access site bleeding minor investigation results were erroneously included in the on manufacturer device report number 1220648-2025-28425-1.
9338 | 1220648-2025-28420 | 22026023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026023&pc=OZD | 04/22/2025 | 06/18/2025 |  | 05/15/2025 |  |  | 04/22/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638926 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was inserted with an impella cp without any issues. while on support, labs were hemolyzed and the patient's right lower extremity was cold with no distal pulses. the patient was brought to the operating room for escalation to an impella 5.5. right axillary was cutdown and the graft was sewn. the impella cp was pulled into descending aorta on p1. unsuccessful placement of the impella 5.5 over 0.018 wire was attempted and left ventricle access was lost. after looking at transesophageal echocardiogram again prior to obtain left ventricle access again, there was concern for a new left ventricle thrombus and impella 5.5 placement was aborted. the impella cp was then removed by vascular surgery via femoral cutdown and the patient remained on peripheral venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of hemolysis, limb ischemia, and thrombus has been completed. complaint device not returned for analyzing. no more details provided regarding attempts to resolve the reported hemolysis issue since they planned to escalate to 5.5. data log reviewed: motor current (mc) spikes observed during impella support. first spike occurred after 4hrs of support followed by resolved impella stopped- mc high alarm and couple spikes at last 3hrs of support followed by drop in flow and placement signal trending up toward the end of support. also purge pressure spike seen prior first mc spike. failure mode temporary pump stop added due to data log observation (resolved impella stopped- mc high alarm). hemolysis: the cause of hemolysis issue most likely was biomaterial ingestion. temporary pump stop: the cause of the temporary pump stop issue most likely was due to biomaterial ingestion. limb ischemia and thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined.
9339 | 1220648-2025-28421 | 22026068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026068&pc=OZD | 04/22/2025 | 09/23/2025 |  | 05/15/2025 |  |  | 04/22/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026647333 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that there were placement signal not reliable alarms. position was verified, and audio was disabled. support continued. | The user facility reported a patient was on impella 5.5 support and during support, the patient had internal bleed at the axillary site. heparin was withheld. support was continued. |  |  |  |  | B1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28421. d10 concomitant medical products and therapy dates updated. | The pump was not returned for investigation. on the 5th day of use, the data log showed a gradual decreasing trend in snr, followed by the psnr alarm. gain was noted to be gradually increasing, and contrast changed its amplitude 10-30. the light was also seen to gradually decrease to 1.2. the placement signal was never seen to recover until the end of the case. there is no definitive failure optical system failure pattern that matches observed trend on the 5s parameters. therefore, the cause of the optical signal issue could not be determined without a returned product investigation. the cause of the access site bleeding issue was not determined since the product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9340 | 1220648-2025-28422 | 22026148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026148&pc=OZD | 04/24/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644049 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 73 year old male patient developed a large hematoma following placement of the repositioning sheath. manual pressure was applied, but the condition persisted. four units of blood were given and the patient was taken to the operating room for a cutdown, hematoma evacuation, and pump explant. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9341 | 1220648-2025-28423 | 22026161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026161&pc=OZD | 04/23/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613052 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5. support and developed a hematoma around their right axillary jp drain following physical therapy. there was no output in the drain itself. oozing was noted the next day around the jp drain and the tubing was stripped which cleared some clots and the drain put out 30 milliliters per hour. an ultrasound was ordered and hemoglobin continued to trend up. heparin drip was put on and bicarbonate was running through the purge. support was continued. | [5 empty] | The impella device was not returned, and the investigation has been completed. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9342 | 1220648-2025-28424 | 22026200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026200&pc=OZD | 04/23/2025 | 05/30/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114581 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) had an optical sensor failure during patient support. the aic was swapped and support was continued with the same impella cp pump without issue. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. based on the device and data analysis the optical system sensor errors were caused by a low signal not reliable (snr) optical bench.
9343 | 1220648-2025-28426 | 22026289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026289&pc=OZD | 04/23/2025 | 06/20/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635336 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Puncture/Hole | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that upon insertion of the impella catheter through the 14 french sheath blood was seen spewing through the catheter outside of the body as if the catheter was pierced. the catheter was replaced and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Product damage (cannula kink): the root cause of the product damage(cannula kink issue was a catheter kink as the product was not returned and insufficient clinical details provided. access site bleeding: the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. | The root cause of the product damage hole in catheter was not determined as the product was not returned and insufficient clinical details provided. h10 narrative inadvertently cannula kink on manufacturer device report 1220648-2025-28426-1. h6 medical device problem code 2889 was erroneously entered on the initial manufacturer device report number 1220648-2025-28426. h6 type of investigation 4111 and 4114 and investigation findings 4248 were erroneously entered on the initial manufacturer device report number 1220648-2025-28426-1.
9344 | 1220648-2025-28427 | 22026300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026300&pc=OZD | 02/23/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598908 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Delirium | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received an impact eu study where the patient was on impella 5.5 support and during support, the patient suffered from "delir". the study stated, "this is not directly an impella-related event. a connection to the entire procedure is theoretically possible." no further information was provided. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9345 | 1220648-2025-28430 | 22026723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026723&pc=OZD | 04/28/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/28/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648196 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced hemolysis approximately ten days after start of support. the staff was advised to reposition the pump and add volume to resolve the condition. the patient was reported to be in stable condition following the event. impella mcs continued with plans to explant the device the following day. the patient was not a candidate for left ventricular assist device and heart transplant. the patient was successfully weaned and, approximately three days later, the device was explanted with a survived outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9346 | 1220648-2025-28431 | 22026970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22026970&pc=OZD | 04/28/2025 | 06/11/2025 |  | 05/15/2025 |  |  | 04/28/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609474 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after implant, same day, a small hematoma and slight oozing were noted at the impella access site. one unit of pack red blood cells was administered, and dressing was changed with angle match technique to manage the condition. support continued with explant of the impella cp device planned for the next day which appears to have been completed as the latest update on this following day noted the patient was no longer on support. the patient's status after event and explant of the device was unknown. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since hematoma was stable after adjusting the angle matching. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9347 | 1220648-2025-28432 | 22027115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027115&pc=OZD | 04/23/2025 | 06/06/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620847 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and went into ventricular tachycardia and atrial fibrillation. the physician planned to cardiovert the patient. additionally, there were suction alarms. one bolus was transfused to resolve the suction alarms. the patient remained on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed and suction alarms were present at end of case at time of reported low flows. pump flows were within p-level range at time of suction alarms. clinical details noted that patient received 250cc bolus and suction alarms resolved. the root cause of the low pump flow was most likely patient condition related as a fluid bolus resolved the issue. as the necessary information was not provided, the root cause of the vt/vf was not determined.
9348 | 1220648-2025-28433 | 22027170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027170&pc=OZD | 03/04/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 03/04/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605547 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the impella cp that was implanted on a 76-year-old male, was taken for escalation to an impella 5.5 as there were less than adequate flows. the pump was removed via surgical cutdown and repair. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9349 | 1220648-2025-28434 | 22027370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027370&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573961 |  | 1000100 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted in a patient diagnosed with cardiomyopathy. impella in aorta alarm was active despite the pump being too deep, repositioned and correct position was verified. motor current was narrow and placement monitoring was disabled. the pump was supporting for a month when it was explanted and replaced due to persistent position alarms. there was no lasting harm or injury and the patient survived. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Hemolysis was added as an email from the rep claimed of rising ldh. saturated flow was added from the investigation. the product was returned and it was cut. no further testing was able to be conducted. there was no biomaterial around the impeller, it looked very clean. data logs indicate an ingestion event though. hemolysis; the pump was removed due to a rise in ldh on the 24th. there was an ingestion noted on the 23rd and no other information related to the patient. based on the data log analysis, the root cause of the hemolysis is most likely due to ingested biomaterial. optical signal issue; it was stated that there was persistent impella position in aorta alarms despite the position being too deep, so positioning was adjusted but the alarms persisted. the data log analysis shows that there was a motor current spike around 1:45pm on 4/23 followed be changes in the ps and saturated flow. the motor current remains high for the rest of the case. about 3 hours later there are impella position in aorta alarms which are due to the discrepancy between the motor current and placement signal. based on the data log analysis, the root cause of the optical signal issue is most likely due to ingested biomaterial. saturated flow: the data log analysis shows that there was a motor current spike around 1:45pm on 4/23 followed be changes in the ps and saturated flow. the motor current remains high for the rest of the case. based on the data log analysis, the root cause of the saturated flow is most likely due to ingested biomaterial. access site bleeding: it was stated that there was a hematoma and minimal drainage from the jp drain. due to lack of clinical details on why there was bleeding or how it resolved, the root cause of the access site bleeding cannot be determined. positioning issues; it was stated that on april 23rd there were impella position in aorta alarms despite the pump being too deep and it was repositioned and position was verified. there are no other clinical details indicating how the pump became too deep, so the root cause of the positioning issues cannot be determined. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined
9350 | 1220648-2025-28436 | 22027586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027586&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632694 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant of an impella cp, the device was repositioned under fluoroscopy and the flow rate dropped. suction alarms appeared and the device appeared to be shallow in the ventricle. it was additionally noted that the cannula appeared to be pinched under the motor. due to the noted failures, the decision was made to replace the pump and the patient is stable. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Returned complaint cp pump visually inspected. cannula kink near outlet observed. no biomaterial, no broken blade found. physician was unsure of impella location and moved the device frequently. echo showed device shallow. cannula kink occurred during attempted repositioning. product damage (cannula kink): the cause of the product damage (cannula kink) issue was cannula kink due to improper technique. positioning issues: the cause of the positioning issues most likely was use related dur to improper technique while attempted repositioning. low pump flow: the cause of the low pump flow issue was cannula kink due to improper technique.
9351 | 1220648-2025-28435 | 22027592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027592&pc=OZD | 04/10/2025 | 05/27/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp patient transferred to the tertiary center for more care and escalation to the impella 5.5 pump. weeks later the loqi reported upon 2 adverse events. there was thrombocytopenia that required treatment with transfusion due to platelet drop. additionally there was anemia/blood loss after the sheath/pump exchange. the patient had more transfusion and the patient expired. | [5 empty] | The investigation of thrombocytopenia and access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since the perclose failed. the root cause of thrombocytopenia cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9352 | 1220648-2025-28437 | 22027666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027666&pc=OZD | 04/24/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637872 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Arrhythmia | Stroke/CVA | Ischemia | Vascular Dissection | Thrombocytopenia | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced multiple complications during impella 5.5 support. upon the initial attempted delivery, it was noted that the j-wire and glide catheter were unable to advance passed the innominate artery. an angiogram revealed a dissection was present and the devices were removed. a sternotomy was performed and a graft to the subclavian artery was placed to bypass the dissection. impella implant was then successful, however, the patient lost distal flow down their right arm from the subclavian down. support continued despite the noted condition and three days later, the patient suffered a couple runs of ventricular tachycardia. cardioversion was required to treat the condition. the following day, the patient was determined to be hit positive and heparin was paused due to oozing and gastrointestinal bleed concerns. later that same day, the patient's computed tomography scan results showed that they had suffered an embolic stroke. the family opted to withdraw care and the patient passed away. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: cautions - ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9353 | 1220648-2025-28438 | 22027674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027674&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580806 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance past the innominate artery during attempted delivery. an angiogram of the right axillary artery showed that the innominate artery was occluded with evidence of severe disease to the axillary artery. the decision was ultimately made to abort the implant. no patient harm. | [5 empty] | The root cause of delivery issue is determined to be patient condition because the pump was unable to advance pass innominate artery due to occlusion and severe disease to axillary artery. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9354 | 1220648-2025-28439 | 22027749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027749&pc=OZD | 04/24/2025 | 09/24/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622693 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support placed to support a patient admitted in with cardiogenic shock. during the 11 day support there were multiple issues, the first being loss of optical signal which did not prompt pump replacement. additionally the patient had oozing at the access site which prompted dressing change and placement of quickclot. also, when conducting echocardiogram imaging there was an observation made of right atrial clot/thrombus. the procedural outcome was the patient survived surgery. | The complainant also reported that there was a loss of optical signal which did not prompt pump replacement. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue, access site bleeding, and thrombus. the data logs show that the placement signal and 5s parameters were stable throughout case, however at each trigger of the impella position unknown alarms, the signal-to-noise ratio (snr) became lower. these alarms were triggering correctly as there was lacking motor current modulation and placement signal pulsatility. flows also slightly fluctuated. however, the clinical confirmed the pump was in good position. due to lack of clinical details related to the patient¿s condition and no product returned, the root cause of the optical signal issue could not be determined. it was stated that there was oozing from the dressing site so quickclot was placed and it resolved. due to lack of clinical details provided, the root cause of the access site bleeding cannot be determined. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28439 b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
9355 | 1220648-2025-28440 | 22027819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027819&pc=OZD | 04/30/2025 | 07/10/2025 |  | 05/15/2025 |  |  | 05/01/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316955 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. during support, there was a controller failure alarm that occurred while waiting for insertion. the alarm could not be cancelled so a spare automated impella controller was prepared and used. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) and aic - controller failure (yellow alarm) has been completed. based on the data analysis, device analysis, and clinic details, the following conclusions were determined. the root cause of the controller failure alarm (pmd communication loss) was an impellatronic malfunction. the root cause of the controller error alarm (epm communication lass) was not determined due to the inability to reproduce.
9356 | 1220648-2025-28441 | 22027931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22027931&pc=OZD | 01/28/2025 | 06/17/2025 |  | 05/15/2025 |  |  | 01/28/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567224 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Infusion or Flow Problem | Material Split, Cut or Torn | [5 empty] | Arrhythmia | Bacterial Infection | Cardiac Arrest | Fungal Infection | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced cardiac arrest during impella 5.5 support. cardiopulmonary resuscitation was performed and both epi and sodium bicarb were given to manage the condition. as support continued, the patient developed an infection of unknown origin. both bacteremia and candidemia were present at the time of the infection. the patient was given iv meropenem, iv micafungin, and serial blood cultures to monitor/treat the infection. in addition, the central lines and arterial lines were removed. despite the noted intervention, infections of various degrees remained intermittent throughout support. it was additionally reported that the anti-contamination sleeve was torn. the sleeve was secured via tape and support remained uninterrupted. forty-six days into support the patient had run of ventricular tachycardia (vt). the patient received one shock in response to the vt. a week later, the patient again had a short run of vt treated with a shock. roughly 108 days into support, a spike in motor current was encountered and saturated flows were identified. in response to the malfunction, the decision was made to withdraw care and the patient passed away. | [5 empty] | The investigation of vf/vt/af/at, infection/sepsis, product damage (sleeve damage), and saturated flow has been completed. the impella device was not returned for evaluation. product damage (sleeve damage): due to limited clinical details and lack of product returned, the root cause of the product damage (sleeve damage) cannot be determined. saturated flow: there was a shift in the ps as well as the purge pressure. the root cause of the saturated flow is most likely due to ingested biomaterial based on data log analysis and the clinical details provided. thrombus: failure mode added based on data analysis. as the necessary information was not provided, the root cause of the thrombus was not determined. infection/sepsis: as the necessary information was not provided, the root cause of the thrombus was not determined. vt/vf: as the necessary information was not provided, the root cause of the thrombus was not determined. b2 required intervention was inadvertently not selected on the initial manufacturer device report number 1220648-2025-28441. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9357 | 1220648-2025-28442 | 22028067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22028067&pc=OZD | 04/28/2025 | 06/09/2025 |  | 05/15/2025 |  |  | 04/28/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637910 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy and an ejection fraction of 20% was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced runs of ventricular tachycardia and as well as device malpositioning. the patient would subsequent expire noted to be ineligible for any advance therapies. the patient, with multiple hospitalizations, was being evaluated for heart failure and the decision was made for an impella 5.5 device placement and left ventricular assist device workup. the impella 5.5 device was implanted without incident, and the patient was transferred to the unit on impella mcs. two days after start of the mcs, the medical team became aware of pink urine post bivalirudin initiation. the need for echocardiogram was discussed to confirm impella 5.5 device placement, and the results revealed the pump was slightly towards mitral structure but not interacting. further noted was the patient experienced runs of ventricular tachycardia requiring five shocks of the implantable cardioverter defibrillator. the following day, the patient became delirious from suspected ethyl alcohol withdrawal. labs and hemodynamics were stable, and the pink tinged urine self-resolved the day prior. the next days, the patient continued to be delirious and confused and pulled out the swan catheter. additionally, there was occasional suction when the patient was thrashing in the bed. a nosebleed was present and bivalirudin was placed on hold. a computed tomography scan completed one week prior revealed nothing remarkable. the psychiatrist noted the patient may be having some type of psychosis or delirium. anticoagulation remained on hold due to nose bleeding. the patient was noted as stable and the physician wanted additional time to manage. however, the family withdrew care due to no change in the patient's neurological status and ineligibility for advanced therapies. the patient would subsequently expire after a total of approximately 16 days on impella mcs. | [5 empty] | The investigation of vt/vf/af/at and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: starting three days after implantation on continuing on and off throughout case, suction events occur. blood volume was given for some but not all suction reports. the pump was not returned. data logs were recovered and review confirmed suction events occur without case. suction events resolve by reducing p-level. the cause of the low pump flow was most likely patient condition related due to suction events occurring around runs of low flow that seem to resolve the problem along with blood volume given. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9358 | 1220648-2025-28443 | 22028311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22028311&pc=OZD | 04/04/2025 | 05/27/2025 |  | 05/15/2025 |  |  | 04/04/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631403 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Inaccurate Information | Wrong Label | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon being plugged into an automated impella controller, the serial number of the pump did not match the box labeled serial number of the delivered set. support was completed without issue despite the labeling issue. | [5 empty] | The investigation of packaging issues has been completed. impella cp pump set box labeled as sn (b)(6) contained a different pump serial number (sn (b)(6)) when the pump was connected to the console. pump and images were not returned for investigation. the cause of the packaging issue was an incorrect label. h6 medical device problem code 4051 was erroneously entered on the initial manufacturer device report number 1220648-2025-28443. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9359 | 1220648-2025-28444 | 22028363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22028363&pc=OZD | 04/24/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620635 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a the patient developed thrombocytopenia during impella 5.5 support. two units of platelets were given to treat the condition. the patient remained stable and support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9360 | 1220648-2025-28445 | 22028589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22028589&pc=OZD | 04/23/2025 | 06/16/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655069 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was unable to advance past the innominate artery during attempted delivery. after multiple failed attempts, the decision was made to abort the implant and an impella cp pump was placed for patient support instead. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on clinical details the root cause of the delivery issue was most likely patient condition related, impella unable to be advance past the innominate artery after multiple attempts.
9361 | 1220648-2025-28446 | 22028868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22028868&pc=OZD | 04/24/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654410 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, there were concerns for the patient's hemolysis. additionally, the patient had oozing at the access site with a hematoma. manual pressure was held and it is unknown if blood products were given. reportable due to patient's possible hemolysis and access site issues. furthemore, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9362 | 1220648-2025-28447 | 22029060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029060&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/15/2025 |  |  | 04/23/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653702 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Detachment of Device or Device Component | Material Split, Cut or Torn | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 procedure, the sterile sleeve ripped out of the back lock while advancing the blue suture hub. sterility was maintained because of the sterile drape. additionally, the patient had ventricular tachycardia episode while leaving the operating room which required defibrillation and the impella was pulled back one centimeter. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | No product was returned and logs are not applicable. product damage (sterile sleeve damage). clinical details state that the sterile sleeve was pulled out of back lock while advancing the blue suture hub. the root cause of the sterile sleeve damage was not determined since product/images were not returned. vf/vt/af/at. clinical details note there was a vt arrest while leaving operation room. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
9363 | 1220648-2025-28449 | 22029139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029139&pc=OZD | 04/24/2025 | 06/27/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618588 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed hemolysis during impella cp support. the power level was lowered and support was continued uninterrupted. | [5 empty] | The investigation of hemolysis has been completed. hemolysis: clinical details state that hemolysis was present on the first day of support and the physician felt like it was caused by the impella. repositioning did not resolve the issue. the patient's condition worsened overnight and p-level was reduced to p-6 due to hemolysis. the patient was improving by the next day. on day 4 of support, the pump was dislocated due to mobilization and was quickly successfully repositioned. the hemolysis was not present anymore the next day. the product was returned and evaluated. there was no biomaterial found on or around the impeller. no cannula kinks were observed. there was no damage to the impeller or abnormalities with the pump. the pump passed hemolysis testing with an mih of 2.18. data log review did not show any motor current spikes indicating biomaterial ingestion or suction alarms on the day the hemolysis started. there was no significant trend in placement signal. there was one impella in aorta alarm on the first day that was resolved within a minute. the root cause of the hemolysis was most likely patient condition since the patient's condition was worsening the night of the hemolysis report, the hemolysis had resolved by the last day of support, and the returned pump passed hemolysis testing. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9364 | 1220648-2025-28450 | 22029150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029150&pc=OZD | 04/24/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/24/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637888 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemolysis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had ectopy and frequent runs of ventricular tachycardia usually related to positioning. echo was performed and position was confirmed. additionally, the patient had a large clot extracted from the posterior side of the heart. it is unknown if the clot was related to the impella. furthermore, the p level was decreased due to concern for hemolysis. the patient had an increase in pfhgb. later the patient went into atrial fibrillation (afib) yesterday, amino bolus given and rate went back up to 1mg. attempted impella wean but unsuccessful yesterday due to cardiac rhythm being afib. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9365 | 1220648-2025-28451 | 22029285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029285&pc=OZD | 04/27/2025 | 06/06/2025 |  | 05/15/2025 |  |  | 04/27/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp support developed a gastrointestinal bleed. heparin was discontinued. the patient was weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9366 | 1220648-2025-28452 | 22029342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029342&pc=OZD | 04/25/2025 | 06/17/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025544904 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was attempted but could not cross the aortic valve with the wire. implantation was aborted. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella aborted due cardiac and vascular anatomy. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller. section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-28452. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
9367 | 1220648-2025-28453 | 22029553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029553&pc=OZD | 04/25/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026633548 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 112 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was dislodged while the patient was awaiting transport. upon successful transfer, the pump was removed. the patient was subsequently assessed for impella 5.5 implant in response to the pump removal. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9368 | 1220648-2025-28454 | 22029585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029585&pc=OZD | 04/24/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630528 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that blood was noticed in the impella 5.5 purge sidearm. bicarbonate replaced the heparin in the purge. implantation was successful. there was no harm to the patient reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9369 | 1220648-2025-28455 | 22029604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029604&pc=OZD | 04/25/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026619288 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female | 168 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp suffered from ventricular tachycardia. despite being defibrillated, the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9370 | 1220648-2025-28456 | 22029719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029719&pc=OZD | 04/25/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636791 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the patient vasculature during attempted delivery. after multiple failed attempts, the decision was made to abort the implant. a second pump was subsequently implanted for patient support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9371 | 1220648-2025-28457 | 22029913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029913&pc=PYX | 04/25/2025 | 06/10/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026622790 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella rp flex was removed and replaced with another pump after less than 24 hours after implantation due to red alarms of purge system blocked and purge pressure high with flows dropping below 0.5 liters per minute. the patient survived. | [5 empty] | The investigation of high purge pressure has been completed. high purge pressure: no product was returned for analysis. the data logs confirm high purge pressure alarms and show a motor current spike with speed deviation. immediately there was a drop in purge flow, pump flow and a shift in placement signal. this pattern is characteristic of biomaterial ingestion. purge pressure remained elevated until explant. the root cause of the high purge pressure was most likely due to biomaterial ingestion as evidenced by motor current spike with placement signal shift and speed deviation. as the necessary information was not provided, the root cause of the thrombus was not determined g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2025-28457. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9372 | 1220648-2025-28458 | 22029921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029921&pc=OZD | 04/25/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025547445 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp experienced internal bleeding in their ieg leading to development of a hematoma. the leg was swollen and a ct scan verified an active bleed. two units of blood were given and the pump was explanted in response to the condition. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9373 | 1220648-2025-28459 | 22029989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22029989&pc=OZD | 04/25/2025 | 06/17/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630006 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during attempted implant of an impella 5.5 pump, the wire was backloaded onto the impella and placement into the left ventricle was unsuccessful due to being unable to pass innominate. a decision was made to exchange the guidewire. the impella 5.5 was removed from the body and when attempting to exchange the guidewire, the impella 5.5 fell off the table and onto the floor. the implant was aborted and a new impella was implanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump was unable to pass through innominate.
9374 | 1220648-2025-28460 | 22030136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22030136&pc=OZD | 04/25/2025 | 06/16/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026654083 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 107 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Embolism/Embolus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 began complaining of chest pain and profound shortness of breath, then went into ventricular fibrillation. unsuccessful attempts were made to convert the sinus rhythm and cpr was initiated. the decision was eventually made to halt resuscitation efforts and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." | The investigation of vf/vt/af/at and embolism has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28460. h6 health effect - clinical code 4438 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28460.
9375 | 1220648-2025-28461 | 22030385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22030385&pc=OZD | 04/25/2025 | 06/22/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026655072 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 72 KG | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon attempting to implant an impella 5.5 pump, a dissection was discovered in the axillary artery. it was noted that the physician declined imaging before the procedure. the implant was aborted, the vessel was repaired, and the access site was sutured to resolve the issue. the patient survived. | [5 empty] | The investigation of vascular damage - peripheral vessel and delivery issues has been completed. the impella device was not returned for evaluation. during advancement of the 5.5 into the axillary artery a dissection was noted in the axillary artery, which was noted to have calcification. md suggested the vessel was too small. resistance was met trying to pass through vasculature and excessive force was used. vascular damage - peripheral vessel: the pump was not returned. the cause of the vascular damage - peripheral vessel was most likely patient condition related due to the noted calcification and smaller vessel size leading to a challenging delivery requiring excessive force to be used. based on the clinical data failure mode delivery issues was added. the pump was not returned. the cause of the delivery issue was most likely patient condition related due to the noted calcification and smaller vessel size. instructions for use. impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28461. b2 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28461. h6 medical device problem code 2682 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28461. h10 instructions for use was incomplete on the initial manufacturer device report number 1220648-2025-28461. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9376 | 1220648-2025-28462 | 22030551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22030551&pc=OZD | 04/25/2025 | 05/15/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026626864 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Injury | Yes | No | Infusion or Flow Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with st-elevation myocardial infarction and was implanted with an impella cp. during support, suction events started and the patient was put on levosimendan and intubated. the patient also lost distal pulses and developed hemolysis. the decision was made to escalate the patient to an impella 5.5. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9377 | 1220648-2025-28463 | 22030861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22030861&pc=OZD | 04/25/2025 | 06/11/2025 |  | 05/15/2025 |  |  | 04/25/2025 | 05/15/2025 |  |  |  |  | 05/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025610405 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump experienced saturated flow. the pump was removed and replaced to resolve the issue. there was no noted patient harm. | [5 empty] | Impella 5.5 was returned for evaluation. upon visual inspection, a large biomaterial was present within outflow cage and covering impeller blade. data logs were reviewed and pump flow saturation occurring immediately upon start-up with full saturation occurring approximately 10 seconds after pump start. a small motor current spike and motor speed spike were present just before full pump saturation. clinical details noted that pump was inserted without issue and flow saturation was occurring immediately after pump was started and decision was made to replace the pump. the root cause of the saturated flow was most likely due to biomaterial based on returned product and data log analysis. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9378 | 1220648-2025-28377 | 22017361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017361&pc=OZD | 04/22/2025 | 07/25/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637514 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 16 YR |  | Male |  |  |  | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that multiple complications were encountered during impella 5.5 support on a patient. seven days into support, the patient showed signs of hemolysis including dark urine and low haptoglobin levels. the condition was being monitored and the staff was advised to check on the positioning of the pump. one day later, the patient suffered from recurrent tachycardia, for which cardioversion was required. support continued despite the noted conditions. roughly fourteen days into support, the anti-contamination sleeve was damaged. support continued despite the noted damage. the patient was eventually weaned and bridged to left ventricular-assist device. the patient is stable. | [5 empty] | The pump was not returned for investigation. hemolysis: hemolysis was reported on 22nd april and confirmed on 24th april, and it was seen to resolve on 2nd march. the data log shows several suction alarms during the case run. there were no motor current spikes, positioning issues, or any major placement signal trends noted that would suggest a concern related to hemolysis. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was provided. the cause of the access site bleeding issue was not determined without sufficient clinical information. the cause of the purge leak cassette issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the purge leak cassette issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the sterile sleeve damage issue was not determined since product was not returned and sufficient clinical information was not provided. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | G1 reporting contact address and country were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28377. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-28377.
9379 | 1220648-2025-28378 | 22017472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017472&pc=OZD | 04/22/2025 | 06/27/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a guidewire was defective and could not be used for impella implant. another guidewire from a third party was utilized for the implant. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. reported the guidewire was unraveling and was not used. resistance was during guidewire placement. no excessive force was applied and no vessel conditions were noted. guidewire was returned and the tip was unraveled. the root cause of the delivery issue was not determined due to insufficient clinical details. | The guidewire was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. a.1, a.4 and a.5 are unknown. d.1 exact brand name is unknown. d.4 exact model number, catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
9380 | 1220648-2025-28379 | 22017577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017577&pc=OZD | 04/23/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194345 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that a console had a battery failure error during setup. there was no patient involvement. | [5 empty] | The investigation for the automated impella controller (aic) battery failure - red alarm has been completed. console logs were were reviewed and were consistent with what was reported in the complaint. multiple instances of battery failure (transport connection blown/missing) alarms were observed in the logs. the console was returned for evaluation and the reported battery failure was able to be reproduced. batttest results revealed that the cells in both batteries had voltages of below 2400 mv. the battery failure alarm resolved after testing this console with two known good batteries. the root cause of the battery failure was depleted batteries resulting in battery lockout.
9381 | 1220648-2025-28380 | 22017611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017611&pc=OZD | 04/07/2025 | 10/02/2025 |  | 05/14/2025 |  |  | 04/07/2025 | 05/14/2025 |  |  | 11/30/2025 |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428296 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 117 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63 year old male patient on an impella 5.5 that had runs of ectopy on monitor during support. also, the patient was treated with antibiotics due to sepsis. additionally, the patient experienced retroperitoneal bleed and two units of blood were administered. the patient continued on support. | The complainant also reported that a placement signal not reliable alarm was noted. |  |  |  |  | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for topical signal issue. clinical reports placement signal not reliable with absent placement signal. no events or clinical details occur during placement signal (ps) outage. the pump was not returned. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28380. a4 updated. d6b updated. d10 concomitant medical products and therapy dates updated.
9382 | 1220648-2025-28381 | 22017695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017695&pc=OZD | 04/23/2025 | 09/19/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026618608 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 64 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was scheduled for implantation with an impella 5.5 pump for support while awaiting surgery. the team repeatedly tried multiple techniques to place the pump but could not due to patient's anatomy. the patient survived and was evaluated for high risk percutaneous coronary intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical review, it was found that the patient had tortuous vascular anatomy which caused insertion issues. it was determined that the root cause for the delivery issue was patient condition due to tortuous vascular anatomy.
9383 | 1220648-2025-28382 | 22017760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017760&pc=OZD | 04/22/2025 | 06/09/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594280 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and was being prepped for transfer to the intensive care unit. (icu) tegaderm was placed to protect the pre close strings, and this tegaderm changed the angle of the impella. upon arrival in the icu, the impella site was noted to be bleeding. dressing and manual pressure was held for an hour and a powder hemostatic agent was used for hemostasis. a decision was made to leave the impella cp in overnight and explant the next day. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be anticoagulation management because the activated clotting time (act) were high and not brought to the therapeutic range when the patient was transferred to intensive care unit (icu) from catheterization lab.
9384 | 1220648-2025-28383 | 22017778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017778&pc=OZD | 04/21/2025 | 06/26/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026627026 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after implantation of an impella 5.5 a spike in the motor current curve appeared and the pump stopped. the automated impella controller (aic) was exchanged but the impella did not restart. the impella was explanted and support continued with va ecmo. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. after 5 days on support, high motor current (mc) occurred and sudden pump stop followed. the console was exchanged with no success. impella was then removed and it was reported that the impella had sucked some of these particles into the inflow basket. data logs were not recovered. the pump was returned and inspected. imaging of the impella blade and inlet cage showed no abnormalities with no biomaterial found. electrical testing showed the electrical lines were open which was confirmed by imaging showing broken electric lines in the red handle with no bond issues noted. the cause of the pump stop was an open electrical line most likely caused by excessive force due to the broken wiring found in the red pump handle. d9 device returned to manufacturer date added
9385 | 1220648-2025-28384 | 22017816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017816&pc=OZD | 04/22/2025 | 05/30/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616030 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that implantation of an impella 5.5 could not be completed because the patient's vessels were too small. the patient survived. | [5 empty] | A.5 ethnicity, is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the vessel size was small and the insertion was aborted.
9386 | 1220648-2025-28385 | 22017938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017938&pc=OZD | 04/22/2025 | 06/13/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643253 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp could not be complete. the pump could not cross the abdominal aorta lesion. the support was aborted without any harm or injury. | [5 empty] | A.4 is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella would not advance in the abdominal aorta.
9387 | 1220648-2025-28386 | 22017942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22017942&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026654425 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 65 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced blood was oozing out of the red plug area. the pump was replaced with another impella 5.5. the patient survived. | [5 empty] | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Reported blood oozing from the red plug area. impella 5.5 s2 was returned. multiple holes in the catheter and other damages were found around the 30cm mark, as well as blood in the red handle. the root cause of the product damage was most likely use related since there was blood in the red handle and a hole in the catheter on the returned pump.
9388 | 1220648-2025-28387 | 22018100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22018100&pc=OZD | 04/22/2025 | 06/23/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026641248 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombocytopenia | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on impella 5.5 for hemodynamic support. during support, it was reported that they had negative fluid balance due to fursosemid and no suction alarms after volume. in addition, a call from the physician was received to report purge pressure with red alarm "purge system blocked". there were no kinks and the purge system was changed. motor current was stable. additionally, they switched to argatroban due to heparin-induced thrombocytopenia. the pump was exchanged and the patient survived. | [5 empty] | A.4 is unknown. the impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of high purge pressure, thrombocytopenia, and low pump flow has been completed. insufficient product was returned for evaluation. a partial pump, cut at the catheter leaving only the cannula was returned hence no reproduction testing could be carried out. biomaterial was observed in the purge gap and in the outlet cage. no data logs were returned for analysis. high purge pressure: the root cause of high purge pressure was not determined as no data logs were returned for analysis and partial product was returned which could not be tested. no strong evidence to determine definitive root cause. low pump flow: the root cause of low pump flow was not determined as no data logs were returned for analysis and partial product was returned which could not be tested. no strong evidence to determine definitive root cause. thrombocytopenia: the cause of the issue was not determined due to insufficient clinical details. thrombus: as the necessary information was not provided. the root cause of the thrombus was not determined.
9389 | 1220648-2025-28388 | 22018127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22018127&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026633149 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 63 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella cp had flow saturation, so it was exchanged for a new impella cp. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Pump saturation happened as soon as the pump went in and was replaced by a second pump. data log review confirmed saturated flow where mean flow at p-6 was at 3.75 l/min, where expected flows for cp are 2.5-2.9l/min. the pump was returned and inspected. inspection revealed biomaterial wrapped around the impellor blade. the saturated flow reproduced in lab where at p-6, flows are 3.7l/min where expected range is 2.5-2.9l/min. the cause of the saturated flow was biomaterial due to the biomaterial found wrapped around the impellor blade and the saturated flow reproducing during investigation.
9390 | 1220648-2025-28389 | 22018358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22018358&pc=OZD | 04/23/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025522924 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited bleeding from the femoral site. the care team applied pressure and stopped heparin. the family made choice to withdraw care and the patient expired. | [5 empty] | A4 and a5, race, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9391 | 1220648-2025-28390 | 22018499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22018499&pc=OZD | 04/23/2025 | 06/06/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026638713 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 108 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp support exhibited hemolysis. attempts were made to reposition the pump and reduce the flow to p-1. additionally, suction alarms occurred. the impella cp was explanted and replaced with an impella 5.5. | [5 empty] | The investigation of hemolysis, positioning issues, and low pump flow has been completed. the data logs were not returned for review. hemolysis: visual inspection found biomaterial inside the pump cannula and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of hemolysis was most likely biomaterial ingestion based on return product analysis and clinical review. positioning issues: as per clinical the repositioning reported seems to be a troubleshooting done likely in response to suction event. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no problem was found on the pump. low pump flow: visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other pump cannula kinks or damaged impeller issues were found on the pump. as per clinical, 0.1l/min flows at p-2 with suction and left ventricle (lv) was simply decompressed with some spikes in motor current (mc) observed. the root cause of low flow was most likely biomaterial ingestion based on return product analysis and clinical review. d9 device returned to manufacturer date added b1 product problem was inadvertently not selected on the initial manufacturer device report number 1220648-2025-28390. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28390. | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9392 | 1220648-2025-28391 | 22018717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22018717&pc=OZD | 04/23/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025612254 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 83 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed thrombocytopenia during impella 5.5 support despite continued platelet transfusions. support was able to continue with the implanted pump despite the condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9393 | 1220648-2025-28392 | 22018972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22018972&pc=OZD | 04/22/2025 | 06/17/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025609975 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 88 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a hematoma during impella cp support. the pump was removed and the groin was repaired. a new pump was placed in the right groin for ongoing support. the patient survived. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. vascular damage - peripheral vessel: vascular damage failure mode was added due to vascular cut during the repair in left groin. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-28392. h6 health effect - impact code 4624 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28392.
9394 | 1220648-2025-28393 | 22019053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22019053&pc=OZD | 04/22/2025 | 05/19/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025605501 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. the pump was placed and support ongoing when there was venticular tachycardia (vt). the position of the pump was assessed when the vt was present. the patient also had to have all systemic heparin stopped due to the patient going for a colonoscopy. the patient received one unit of packed red blood cells due to low hemoglobin and hematocrit. stool sample tested positive but colonscopy was negative for active bleeding. the pump remains on for support. | [5 empty] | A.4 is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and vascular damage, peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9395 | 1220648-2025-28394 | 22019265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22019265&pc=OZD | 04/08/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026616051 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 104 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during removal of an impella 5.5 placement of a wound vac to the right groin was required as access was obtained through a lymph note and lymphatic fluid was draining. the right axillary incision was opened and a pocket wash out was performed. the impella 5.5 was successfully explanted, the incision was closed, and doppler pulse were confirmed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9396 | 1220648-2025-28395 | 22019603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22019603&pc=OZD | 05/04/2025 | 06/12/2025 |  | 05/14/2025 |  |  | 05/04/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607174 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed heparin induced thrombocytopenia during impella 5.5 support. the patient received plasmapheresis in response to the condition. as support continued, it was also documented that high purge pressure alarms were encountered. the purge fluid was changed to tissue plasminogen activator and the staff was advised to monitor the motor current for any upward trends. support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | Thrombocytopenia: after 15 days on support, patient developed hit during impella 5.5 support. the patient received plasmapheresis in response to the condition. the pump was not returned. the cause of the thrombocytopenia was not determined due to insufficient clinical details. high purge pressure: after 19 days on support, a high purge pressure event occurred with purge system blocked. rn stated they observed purge flow decreased to <3 ml/hr and pp >900 mmhg. purge fluid d5w with bicarb. reported that no kinks in tubing. the pump was not returned. data log review showed the lt log ended at 27apr25 and the last rt log recovered ended at 28apr25 at 1206. both of these times are prior to the complaint occurrence of 07may25. the cause of the high purge pressure was not determined since product was not returned, insufficient clinical details, and logs recovered did not include logs from the day the complaint occurred.
9397 | 1220648-2025-28396 | 22019867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22019867&pc=OZD | 04/24/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/24/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026638732 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 YR |  | Female | 51 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while an impella cp was implanted in a patient, there was a pump stop with a motor current spike. the pump was restarted and support continued. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
9398 | 1220648-2025-28397 | 22019970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22019970&pc=OZD | 04/22/2025 | 09/25/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635342 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | Hematoma | Unspecified Infection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the device exhibited placement signal not reliable. | It was reported that a patient was implanted with an impella 5.5. during support, a deep vein thrombus was found in venous system from the swan. it was noted that it was not impella related. additionally, it was reported that the patient had another hematoma at the access site and heparin was being held. the impella was also repositioned and the patient was being treated for clostridioses difficile. the patients outcome is unknown. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917." | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Access site bleed - major: it was reported that the patient had another access site hematoma on (b)(6) 2025. it is unclear when it was first noted. the decision was made to hold heparin. it was not reported if the complication resolved with this troubleshooting method, and no other details were provided. product was not returned for investigation. since product wasn't returned for investigation and clinical details did not report if the complication resolved by reducing heparin, the root cause of the access site bleed could not be determined. pain at insertion site: clinical details report that the patient had pain at the impella insertion site throughout support. the team troubleshot by ambulating the patient. no further details were provided. product was not returned for investigation. the cause of the pain at insertion site was not determined due to insufficient clinical details. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. positioning issues: clinical details report that the pump was repositioned on (b)(6) 2025. it was not reported how the pump became out of position. data logs were reviewed and there were only two positioning alarms during case start on (b)(6) 2025 that each resolved within a minute. product was not returned for investigation. since insufficient clinical detail was provided and product wasn't returned for investigation, the root cause of the positioning issue could not be determined. infection/sepsis: an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations: laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. optical signal issue: clinical details report that the psnr alarm triggered on (b)(6) 2025. the pump nor console were replaced due to the issue. data logs were reviewed and the alarm was confirmed. at this time, contrast began to oscillate between +/-3200. the placement signal did not recover through the rest of the case. however, there was a previous instance of psnr in the case, on (b)(6) 2025, with the same signature of oscillating contrast that resolved after 8 minutes. product was not returned for investigation. based on the intermittent nature of oscillating contrast during the case noted in data log review, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console.
9399 | 1220648-2025-28398 | 22020155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22020155&pc=OZD | 04/23/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/24/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026626825 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella 5.5 exhibited stroke-like symptoms. neurology intervention was done for the embolic stroke and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9400 | 1220648-2025-28399 | 22020191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22020191&pc=OZD | 04/23/2025 | 06/12/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the console was exchanged after the purge cassette change failed twice. the procedure outcome was successful after the exchange of the console. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the cause of the purge disc not detected alarm was most likely contamination of the purge pressure sensor.
9401 | 1220648-2025-28400 | 22020345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22020345&pc=PYX | 04/23/2025 | 05/28/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600719 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella rp flex placed to support a patient admitted in with right ventricle fibrillation and newly placed heartmate 3. the pump was placed but explanted when after low flows made the team question thrombus. there was an access site bleed/ooze that prompted the team to tighten the suture and place a new extra purse string suture. the pump was explanted after the flows dropped. the new pump was successfully placed and support proceeded. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Based on investigation, the failure mode 'optical signal issue' was changed to 'placement signal issue'. the investigation of access site bleeding, low pump flow, and thrombus has been completed. the returned product was inspected and thrombus was observed on the impeller. the returned pump was received clamped down with forceps along the cannula. the pump passed the light continuity test and had 5s parameters within range. the pump passed electrical testing. placement signal issue: the data logs show placement signal drifting up and pump flows drifting down. there were no motor current spikes outside p-level changes and motor current was stable. cvp also stable and no psnr alarms were triggered. the pattern observed is characterized as a possible sensor drift. the pump was flow loop tested and dp sensor drift was reproduced in the lab. the root cause of the placement signal issue was determined to be sensor drift as evidenced by log pattern and issue reproduction on returned product. low pump flow: no problem was found as the root cause of low pump flow. the data logs show motor current stability with no spikes outside p-level changes. low pump flow concern was due to the effect of the placement signal issue (sensor drift) on the calculated flow causing user to suspect pump was not flowing at expected rate. access site bleeding - major: clinical reported an access site bleed that needed tightening of suture and an extra purse suture. it is unknown if there was lack of forward suturing. the root cause of the access site bleeding - major was not determined as insufficient clinical information related to failure was provided h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-28400.
9402 | 1220648-2025-28401 | 22020463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22020463&pc=OZD | 04/24/2025 | 07/21/2025 |  | 05/14/2025 |  |  | 04/24/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025514311 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on impella 5.5 support exhibited bleeding from the surgical insertion site and had to undergo cardiopulmonary resuscitation briefly. the patient survived and support continued. | Additionally, a pump stop occurred with impella stopped motor current high alarm. staff tried to restart the pump but this was not possible due to the impella stopped motor current high alarm. the pump was explanted and a small thrombus was noted on on motor housing and catheter. |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, pump stop, and thrombus has been completed. pump stop - data logs show pump stop with alarm 115 and motor current spikes to 30000. real time (rt) logs for that time period were not returned. impella 5.5 was returned. there was a catheter kink noted below the motor housing and the catheter was stretched between 60cm to 70cm. broken motor cables were found in the red handle. the root cause of the issue was most likely due to open electrical lines based on data logs and returned product analysis. access site bleeding - the root cause of the access site bleeding was most likely use related since the anastomosis had an insufficiency and a surgical revision resolved the bleeding. thrombus - as the necessary information was not provided, the root cause of the thrombus was not determined. d9 device returned to manufacturer date added instructions for use: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock . section: direct aortic insertion . ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28401. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28401. h6 health effect - clinical code 4440, health effect - impact code 4627, and health effect - impact code 1503 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28401. h10 incomplete instructions for use were provided on the initial manufacturer device report number 1220648-2025-28401.
9403 | 1220648-2025-28402 | 22020585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22020585&pc=OZD | 04/23/2025 | 06/30/2025 |  | 05/14/2025 |  |  | 04/23/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639382 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was implanted for a patient admitted in cardiogenic shock. after five days of support, a purge cassette leaked occurred. there had been no purge pressure or purge flow alarms but, the purge cassette leakage prompted the pump to be explanted. the patient survived the procedure. | [5 empty] | The impella pump was not received from the customer and the impella catheter/purge cassette were discarded. therefore, the investigation of the device was not possible. should the pump or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of purge leak ¿ cassette has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9404 | 1220648-2025-28404 | 22020789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22020789&pc=OZD | 04/26/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/26/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was being supported by impella cp. during support, the patient developed hemolysis and the impella cp was explanted. the patient was put on venoarterial extracorporeal membrane oxygenation and survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9405 | 1220648-2025-28405 | 22021072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021072&pc=OZD | 04/25/2025 | 06/20/2025 |  | 05/14/2025 |  |  | 04/25/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634511 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a cardiogenic shock patient with cardiac history/prior coronary artery bypass graft. it was reported that during support the impella cp was weaned from p9 to p6 overnight due to hemolysis. later, the impella in aorta alarmed overnight after episodes of violent coughing from the patient. the doctor repositioned the pump with echocardiogram guidance. additionally overnight, the patient had gastrointestinal bleeding and was given two units of packed red blood cells. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Hemolysis: product & data were not returned for review. the root cause of the hemolysis issue was not determined since product and data logs were not returned and due to inconclusive evidence from clinical. positioning: the root cause of positioning issue was most likely patient movement related since the issue occurred after patient had violent cough. vascular damage: the root cause of gi bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
9406 | 1220648-2025-28406 | 22021091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021091&pc=OZD | 04/22/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026628992 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support. during support the patient had ventricular tachycardia (vt) and the impella position was assessed. additionally, the patient had an ischemic stroke that turned hemorrhagic. impella support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9407 | 1220648-2025-28408 | 22021246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021246&pc=OZD | 04/22/2025 | 06/10/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025612355 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with cardiomyopathy had an impella 5.5 pump implanted for 16 days while they were preparing for transplant. the pump had repeated high purge pressure alarms that prompted a pump explant and replacement. the patient survived. | [5 empty] | Clinical details reported that several high purge pressure alarms triggered between (b)(6) 2025. tpa was added to the purge after the third alarm, but the pump was removed due to the complication. data logs were reviewed and the alarms were confirmed. there were three instances of high purge pressure alarms in the final 48 hours of the case. the first rise was a gradual 3.5 hour rise that started at the time of a purge cassette change. purge pressure had returned to normal levels after 2.5 hours, and the start of the return correlated with another cassette change. the second event was another gradual rise over 40 minutes that resolved after 30 minutes. there were no events to correlate the onset and resolution with. the final event was a relatively gradual 7 hour rise that did not resolve for the remainder of the case. there were no abnormalities to correlate the onset with. returned product came back with a cut in the catheter, so limited reproduction testing could be done. there was some biomaterial residue in the outflow cage, and after tearing down the motor housing, more signs of biomaterial were found on the distal bearing and the underside of the impeller. a window was cut in the catheter just proximal to the motor housing and no blood ingress was noted. additionally, the pump was connected to a purge cassette and primed without issue, indicating there were no blockages in the purge line proximal to where the cut in the catheter was made either. based on the gradual nature of the rise in purge pressure seen on data logs and biomaterial residue being found within the motor housing, the root cause of the high purge pressure was most likely biomaterial buildup. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9408 | 1220648-2025-28409 | 22021496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021496&pc=OZD | 04/29/2025 | 06/09/2025 |  | 05/14/2025 |  |  | 04/29/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654288 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella 5.5 support. during support, it was noted that the patient had left leg thrombectomy/fasciotomy and the patient was hemodynamically stable and neurologically intact. additionally, the patient was taken to operating room for extracorporeal membrane oxygenation decannulation. intra-operative transesophageal echocardiogram showed probable left ventricle (lv) thrombus near the mitral valve. heparin was switched to bival due to possible lv clot near the mitral valve. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Lv thrombus near mitral valve. "thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined."
9409 | 1220648-2025-28410 | 22021526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021526&pc=PYX | 04/22/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/22/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2026635501 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 78 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with bipella support of rp flex/5.5 pump had low flows. the patient was treated with albumin and the p level was adjusted. the patient was successfully weaned and the pump explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9410 | 1220648-2025-28411 | 22021569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021569&pc=OZD | 03/01/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 03/01/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609820 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp to support a patient needing mechanical circulatory support. the pump was placed for the hrpci but despite all pump repositioning the pump had persistent low flow and suction. the pump was explanted and patient survived. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. no biomaterial observed. no cannula kink or broken blade found. pump connected to aic to reproduce the low pump flow issue and run for more than 10 minutes. no low pump flow reproduced. no other abnormality seen. data log reviewed: suctions occurred during impella support. flows are low from the start of the case. quick resolved positioning alarms observed. the cause of the low pump flow issue most likely cannot be determine due to lack of clinical information and issue not reproduced in lab.
9411 | 1220648-2025-28412 | 22021594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021594&pc=OZD | 04/26/2025 | 06/20/2025 |  | 05/14/2025 |  |  | 04/26/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593575 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had ventricular fibrillation and needed to be defibrillation with amiodarone /lidocaine started. the pump was also repositioned during support. furthermore, there were suction alarms that were silenced and the pump was set to p2. later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Pump was not returned for investigation. data log was not provided for this case run. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was provided. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
9412 | 1220648-2025-28413 | 22021697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021697&pc=OZD | 04/26/2025 | 09/25/2025 |  | 05/14/2025 |  |  | 04/26/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026635559 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | During support, it was reported that the device exhibited an optical sensor issue. the customer was advised that the sensor was most likely up against a wall and therefore miscalculating flows. hemodynamics were stable. there was also placement signal not reliable (psnr) alarm. the impella was successfully weaned and explanted. | It was reported that a patient was on impella rp flex. during support, it was reported that there was oozing from spinal fusion site and they were holding heparin. the impella rp flex was successfully weaned and explanted. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. placement signal issue: clinical details reported that the psnr alarm triggered on (b)(6) 2025. the pump was not replaced due to the issue. data logs were reviewed, and the report was confirmed. on this date, there were two instances of the dp ps drifting down to -50 and then returning over the course of 4 hours, which aligns with the csc call noted in the clinical. at the end of the second instance, alarm 1052 triggered (psnr) and did not resolve for the rest of the case. during the instances of drift in the dp ps as well as after psnr triggered, the raw optical signal and motor current both remained stable. product was not returned for investigation. based on the signatures noted in data log review, the most likely root cause of the placement signal issue is sensor drift. it cannot be confirmed without product return. access site bleed: clinical details report that the patient had oozing coming from a spinal fusion access site. the team troubleshot the complication by holding heparin. product was not returned for investigation. since the clinical details provided report that there was only bleeding coming from a non-impella access site, the root cause of the access site bleed was determined to most likely be patient condition. optical signal issue: an optical signal issue failure mode was added after data log review pertaining to the report that the team was unable to transduce the cvp and it was being calculated as 0 mmhg on the automated impella controller (aic) on (b)(6) 2025. a 500cc bolus of fluid was administered with no change. data logs were reviewed, and the optical signal 5s were all within specification the entire case, indicating a strong signal. that being said, there was an isolated instance around 5:00 pm on (b)(6) 2025 where the cvp calculation was measuring values below -100 mmhg. when this occurs, the cvp calculation will become dashed out on the console screen. this is typically indicative that the sensor is in a suction condition. at the time, pump flows were noted to be variable while running at p-7, however, there were no suction alarms and clinical details reported that a 500 cc fluid bolus didn't resolve the complication. product was not returned for investigation. since product wasn't returned for investigation, data logs did not show definitive evidence to explain a negative cvp calculation, and insufficient clinical details were provided, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Placement signal issue: clinical details reported that the psnr alarm triggered on (b)(6) 2025. the pump was not replaced due to the issue. data logs were reviewed, and the report was confirmed. on this date, there were two instances of the dp ps drifting down to -50 and then returning over the course of 4 hours, which aligns with the csc call noted in the clinical. at the end of the second instance, alarm 1052 triggered (psnr) and did not resolve for the rest of the case. during the instances of drift in the dp ps as well as after psnr triggered, the raw optical signal and motor current both remained stable. product was not returned for investigation. based on the signatures noted in data log review, the most likely root cause of the placement signal issue is sensor drift. it cannot be confirmed without product return. access site bleed: clinical details report that the patient had oozing coming from a spinal fusion access site. the team troubleshot the complication by holding heparin. product was not returned for investigation. since the clinical details provided report that there was only bleeding coming from a non-impella access site, the root cause of the access site bleed was determined to most likely be patient condition. optical signal issue: an optical signal issue failure mode was added after data log review pertaining to the report that the team was unable to transduce the cvp and it was being calculated as 0 mmhg on the aic on (b)(6) 2025. a 500cc bolus of fluid was administered with no change. data logs were reviewed, and the optical signal 5s were all within specification the entire case, indicating a strong signal. that being said, there was an isolated instance around 5:00 pm on (b)(6) 2025 where the cvp calculation was measuring values below -100 mmhg. when this occurs, the cvp calculation will become dashed out on the console screen. this is typically indicative that the sensor is in a suction condition. at the time, pump flows were noted to be variable while running at p-7, however, there were no suction alarms and clinical details reported that a 500 cc fluid bolus didn't resolve the complication. product was not returned for investigation. since product wasn't returned for investigation, data logs did not show definitive evidence to explain a negative cvp calculation, and insufficient clinical details were provided, the root cause of the optical signal issue could not be determined.
9413 | 1220648-2025-28415 | 22021765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021765&pc=OZD | 04/25/2025 | 05/28/2025 |  | 05/14/2025 |  |  | 04/25/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647498 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had oozing at the access site. dressing was changed, one unit of blood products were given and anticoagulation was withheld. support was continued. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely patient movement since patient moved right leg overnight causing some oozing.
9414 | 1220648-2025-28416 | 22021946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021946&pc=OZD | 04/24/2025 | 07/07/2025 |  | 05/14/2025 |  |  | 04/24/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | OZD |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026619153 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [5 empty] | NON-ABIOMED V18 WIRE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that while attempting to place a rp for the support of a 79 year old male patient for support during a high risk procedure, the pump delivery was complicated and caused v18 wire to perforate vasculature. the v18 is a non-abiomed wire. the rp delivery was aborted in favor of ecmo when the ascending aorta, superior vena cava, and inominate arteries all were damaged. the patient outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | Priming issue: clinical mention introducer was dropped out of sterile field. the root cause of the priming issue was most likely a use-related issue when handling the device as evidenced by clinical detail. delivery issue: clinical mentions md felt resistance when pulling back 0.018 wire and guidewire was kinked. the root cause of the delivery issue was most likely due to patient condition as evidenced by resistance met. vascular damage - great vessel: the root cause of the vascular damage - great vessel was not determined as limited clinical information related to how perforation occurred was provided and no product was returned for analysis.
9415 | 1220648-2025-28417 | 22021974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22021974&pc=OZD | 04/08/2025 | 05/14/2025 |  | 05/14/2025 |  |  | 04/21/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639770 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | Thrombocytopenia | Mitral Valve Insufficiency/ Regurgitation | [7 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported through abiomed's long-term outcome and ouality impella registry database multiple adverse events with possible/probable/definite relationship to the use of the impella cp. the patient experienced thrombocytopenia. therefore, the heparin drip was changed to argatroban. the patient experienced limb ischemia, access site bleeding, and hemolysis. the pump sheath was replaced and was transfused six units of packed red blood cells. the patient experienced mitral valve regurgitation and was treated with norepiphrine, vasopressin, nitroprusside, and milrinone. the patient became septic and was treated with vancomycin and cefepime. the family decided to withdraw life support and the patient expired. | [5 empty] | B.2 exact date of death is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ impella cp with smartassist system section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist system section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9416 | 1220648-2025-28418 | 22022208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22022208&pc=OZD | 04/24/2025 | 07/07/2025 |  | 05/14/2025 |  |  | 04/24/2025 | 05/14/2025 |  |  |  |  | 05/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1197231 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during prepping, the impella aic (console) had a system self-check failed alarm. the aic was exchanged. no patient involvement. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. e1-initial reporter name is unknown. | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: after the reboot, the console displayed a "system self check failed" screen due to the persistent cpld communication issue. this confirms the reported failure mode. device analysis: this console was tested. when powered on, the console fails in its initial attempt to boot. it automatically reboots once in an attempt to boot successfully, but again fails to do so, finally displaying the "system self check failed" screen. the console was then forcefully shut down. the piezo buzzer on the epm board sounded for the duration that the console was powered on. root cause: as evidenced by the power motor driver (pmd) communication errors found in the device logs, and the inability of the software debugger to establish a connection to the pmd, the root cause of the ¿system self check failure¿ was determined to be an impellatronic malfunction.
9417 | 1220648-2025-28337 | 22008410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008410&pc=OZD | 04/25/2025 | 06/19/2025 |  | 05/13/2025 |  |  | 04/25/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616458 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Inaccurate Flow Rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump began to develop low pump flow during patient support. the pump was subsequently removed and a thrombus was noted on the inlet of the device. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. as per the clinical information and the data log review, the root cause of low pump flow issue was most likely biomaterial ingestion. as per the clinical information and data log review, the root cause of the saturated flows was also biomaterial ingestion.
9418 | 1220648-2025-28338 | 22008494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008494&pc=OZD | 04/21/2025 | 06/18/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628398 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67 year old female patient with an impella cp for support pre surgery. it was noted that the patient had calcified short ascending aorta to aortic arch, that impeded the impella flow causing suction alarm and low flow due to bending of the cannula. decision was made to replace the pump. | [5 empty] | The investigation of low pump flow and product damage (cannula kink) has been completed. reported impella cp having suction and low flow after implant. a kink was noticed in the cannula due to calcification of the aorta. patient had calcified short ascending aorta to aortic arch. imaging confirmed good pump position. no product was returned and no data logs were returned. low flow - the root cause of the low pump flow was most likely a kinked cannula since a cannula kink was observed at the time of low flow. cannula kink - the root cause of the product damage cannula kink was most likely patient condition related to the patient's calcified short ascending aorta to aortic arch. h6 medical device problem code 1158 was erroneously entered on the initial manufacturer device report number 1220648-2025-28338. new medical device problem code entered. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9419 | 1220648-2025-28340 | 22008622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008622&pc=OZD | 04/21/2025 | 06/13/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that several attempts were made to place the impella cp into the left ventricle but unsuccessful. the impella insertion was aborted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through the vasculature.
9420 | 1220648-2025-28341 | 22008663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008663&pc=OZD | 04/30/2025 | 07/11/2025 |  | 05/13/2025 |  |  | 05/01/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP- JP |  |  | ABIOMED, INC. |  |  | US |  | 2026623381 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that ecpella was used for ventricular storm (vt) in hypertrophic cardiomyopathy (hcm) patient. the patient expired despite treatment. the cause of death was fatal arrhythmia and multiorgan failure. the doctor commented that there was no causal relationship with the impella. there is not enough information to exclude the impella used as an associated factor with the event and for this reason this event will be reported. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28341. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-28341 as it is unknown if the initial reporter also sent the report to the food and drug administration.
9421 | 1220648-2025-28342 | 22008695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008695&pc=OZD | 04/20/2025 | 06/17/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566800 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and during the implant, the device appeared to kink as inserted with advanced cardiac life support in progress, and proper placement could not be achieved. device appeared to be damaged therefore was removed and replaced. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issues: the root cause of the delivery issues was most likely use-related as the pump was attempted to be delivered during cardiopulmonary resuscitation (cpr). product damage: the root cause of the product damage catheter kink was most likely use related to the pump having a challenging delivery as it was attempted to be delivered during cpr.
9422 | 1220648-2025-28343 | 22008725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008725&pc=OZD | 04/25/2025 | 06/17/2025 |  | 05/13/2025 |  |  | 04/25/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621472 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and suffered an embolic cerebrovascular accident with limb weakness and expressive aphasia. the patient also exhibited bleeding from the access site. as a result, anticoagulation was stopped. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding and stroke/cva has been completed. the impella device was not returned for evaluation. access site bleeding: the patient additionally had bleeding from all sites. the root cause of the access site bleeding issue was most likely patient condition related, the patient had bleeding from all sites. stroke/cva: as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿
9423 | 1220648-2025-28344 | 22008868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008868&pc=OZD | 04/21/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634533 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp and the p-level was dropped as there were suction alarms. the medical doctor left the lower p-level and explanted the pump the following day. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9424 | 1220648-2025-28345 | 22008877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22008877&pc=OZD | 04/24/2025 | 05/29/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643500 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon removal of the impella cp, the patient went into ventricular fibrillation and arrested. the patient was shocked twice and compressions were initiated to stabilize the condition. the decision was ultimately made to place another impella cp. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9425 | 1220648-2025-28346 | 22009076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009076&pc=OZD | 04/24/2025 | 06/13/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648831 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. during support, the patient required fem stop for groin oozing. forward tension sutures with angle matching were used. urine was dark red and awaiting plasma free hemoglobin results. the impella cp was repositioned under echocardiogram guidance. 500 cubic centimeters bolus was give twice for suction alarms. urinary output was less than 30 and was still dark red. the intensivist was aware. purge fluid was switched to draw with bicarbonate and heparin drip was started. the groin was still oozy, however, improved. the patient was transferred to another facility and reported to be stable and the impella cp with good pulsatility and maps in the 90's. touhy confirmed locked, angle of entry matched, however, groin appeared oozy. the plan was to stabilize and decide escalation pathway. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Access site bleeding - major: immediately following implant procedure, oozing from the groin noted with act >400. forward tension sutures applied with angle matching. fem stop to groin for oozing. the pump was not returned. the cause of the bleeding was high (400) patient act values. positioning issues: after pt was transferred to icu, pump was repositioned at bedside due to suction alarms. the pump was not returned. data log review showed elevated ps with suction alarms that had corresponding low flows. the cause of the positioning issue was not determined due to insufficient clinical details. low pump flow: after pt was transferred to icu, suction alarms occurred and ump was repositioned at bedside. the pump was not returned. data log review showed flows for p-7 at 2.5l.min, below the expected 2.9-3.3l/min, which occurred only during the suction events. no mc spikes. flows returned to nominal for the remainder of the case. the cause of the issue was most likely patient condition related since volume resolved the suction alarms. hemolysis: immediately following implant procedure, urine was noted as dark red. however, the lab results for pfhgb were never reported. the pump was not returned. the cause of the hemolysis was not determined since product was not returned and insufficient clinical details as to the occurrence and results for pfhgb were never reported from the red urine event described in the patient updates.
9426 | 1220648-2025-28347 | 22009153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009153&pc=OZD | 04/21/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568639 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was suction at p6. the pump was lowered to p5. additionally, the patient had a computed tomography and electroencephalogram and it showed no brain activity and brain herniation. it is unknown if the impella caused or contributed to the brain injuries. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9427 | 1220648-2025-28348 | 22009260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009260&pc=OZD | 04/23/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125048 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was supporting a patient when prior to explant, the battery dropped to 68% after ambulation. the aic had been kept plugged in overnight but never charged higher than 68%. when it was going to disable the battery for transfer, the battery was completely dead and was unable to charge or turn it back on. there was no patient harm. | [5 empty] | The investigation has been completed. the console was returned for investigation. data review of the console logs were consistent with what was reported in the complaint. an instance of battery level low (50%) alarm was observed in the imc logs. the console was tested and the reported charging issue was able to be reproduced. the battery level on the screen of the automated impella controller would not charge past 70%. results of running the batttest revealed that the battery in position 1 was not receiving a charge. troubleshot by swapping the battery terminals but the battery in position 1 still would not receive a charge.in a second attempt at troubleshooting, the power battery manager (pbm) was swapped out for a known good one which resolved the issue. both batteries were able to charge with the known good pbm. the battery charging issue was isolated to the pbm. the root cause of this issue was a defective pbm.
9428 | 1220648-2025-28349 | 22009269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009269&pc=OZD | 04/21/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639561 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support developed hematuria. medical doctor was aware however intervention is unknown. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9429 | 1220648-2025-28350 | 22009508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009508&pc=OZD | 04/24/2025 | 06/04/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025483513 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. while on support, lactate continues to rise from 4.8 to 16. the physician repositioned the impella cp under echocardiogram. later it was reported that pulses were lost and the leg turned blue on the leg with the impella cp. the pump was removed at bedside. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9430 | 1220648-2025-28351 | 22009638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009638&pc=OZD | 04/21/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551286 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after initial access and difficulty engaging the left coronary artery the patient became bradycardic and hypotensive requiring dopamine and impella cp access. impella on in p-auto, immediate hemodynamic stabilization and the patient's rhythm went from a junctional rhythm to atrial fibrillation (afib) rapid ventricular response (rvr). the patient was given 2 150mg boluses of amiodarone with unsuccessful cardioversion. the patient was sedated again and a single synchronized shock at 100j was delivered with successful conversion form afib rvr to sinus brady and a rate of 58 bpm. the patient was noted to have several unifocal premature ventricular contractions but remained awake and alert to commands. the peel-away sheath was removed, position optimized under fluro. the groin was dressed and the patient is sent to intensive care unit for hemodynamic monitoring. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9431 | 1220648-2025-28352 | 22009712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009712&pc=OZD | 04/24/2025 | 06/12/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612007 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Material Split, Cut or Torn |  |  |  |  | Atrial Fibrillation | Atrial Flutter | Tachycardia | Ectopic Heartbeat | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was needing mechanical circulatory support. during impella 5.5 support, the patient was in atrial fibrillation/atrial flutter. it was also noted of ventricular tachycardia runs and ectopy on electrocardiogram. echocardiogram (echo) showed the impella 5.5 measuring 6.1 centimeters. the impella 5.5 was repositioned under echo. it was noted that the impella 5.5 was still slightly touching the septum, small left ventricle (lv) cavity. additionally, a newly noted torn anticontamination sleeve at the distal end, tegadermed to seal and the team was all aware and unconcerned at this time. there was also a report of episode of suction that resolved by dropping from p-6 to p-4. however, acute change in aorta (ao) and lv placement signal values at that time which have persisted despite resolving suction. ao and lv were still peak to peak in systole; however, ao and lv were reading in 40s systolic and are both displaying negative diastolic values. motor current pulsatile and flows were maintained. no hemodynamic changes were noted. again, it was noted of occasional impella position unknown with deranged placement signals. echo showed device shallow at 3.6 centimeters, however, the team did not want to pursue repositioning. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Product damage (sterile sleeve): clinical states that the sterile sleeve of the pump was torn while changing the dressing today. pump was not returned for investigation. therefore, the root cause of the product damage(sterile sleeve) issue was not determined since product was not returned. vt/ vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. access site bleeding: clinical states that a slight ooze was noted on (b)(6) 2025 and dressing was changed. patient thinks that its due to lying on back. no other information available. pump was not returned for investigation. therefore, the root cause of the access site bleeding issue was most likely patient movement related as patient was lying on the pump leading to oozing event which resolved with dressing change. low pump flow: clinical states that on (b)(6) 2025, suction alarms were present and p-level reduction resolved the alarms. pump was not returned for investigation. data logs were investigated. lt log image shows suction alarms at p6 and p-level to decreased to p-4 to resolve the suction. placement signal is trending up. no other abnormalities with motor current or position of the pump observed. therefore, based on the clinical information provided and data log review, the root cause of the low pump flow issue was not determined since logs and clinical information were inconclusive. positioning issue: clinical states that the pump was measuring 6.1cm and was repositioned to 5.1cm. patient had vt runs and ectopy overnight but the reason for the pump out of position is unknown. pump was not returned. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient information about the improper positioning was provided.
9432 | 1220648-2025-28353 | 22009798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009798&pc=OZD | 04/21/2025 | 06/12/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604731 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during the implant, the patient had ventricular tachycardia (vt) and was defibrillated. the patient also had vt on support and was cardioverted. additionally, the patient was found to have a deep vein thrombus. furthermore, the pump was repositioned during support. the patient was escalated to impella 5.5 and survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta & suction alarms. logs also showed ps spiked up to 3000 mmhg during the case but resolved quickly. no other issues were found on the logs. positioning issues: product was not returned for review. the root cause of positioning issue was most likely patient movement related since the issue occurred during turning the patient. vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. the cause of the vt storm was unable to be determined due to insufficient clinical details. the failure mode vt/vf/at/af is used as a category for injuries including arrhythmia, asystole, bradycardia, cardiac arrest, heart block, paroxysmal atrial fibrillation, tachycardia, ventricular fibrillation, and ventricular premature complexes. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
9433 | 1220648-2025-28354 | 22009812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22009812&pc=OZD | 04/24/2025 | 06/18/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642495 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there was a concern for distal perfusion. the decision was made to explant the pump. the procedural outcome was the patient survived the surgery. | [5 empty] | Limb ischemia: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication vascular size or variations thereof detailed description of the symptoms experienced by the patient physical examination findings, including pulse assessment and visual examination of the affected limb diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography laboratory test results, including blood tests for markers of inflammation or clotting disorders any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9434 | 1220648-2025-28334 | 22010073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22010073&pc=OZD | 03/26/2025 | 05/23/2025 |  | 05/13/2025 |  |  | 03/26/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026628988 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Stroke/CVA | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 72 year old male patient was on impella 5.5 and during support, the patient had atrial fibrillation (af). support continued. additionally, the pump was found to be fixated on pap muscle. the team did not reposition at the time, but eventually the pump was repositioned. furthermore, multiple acute infracts were found on the ct. later in support, new frontal lobe subacute infarct were found in an additional ct scan. the patient also had a sepsis workup. it is unknown if the sepsis was related to the impella. support continued. | [5 empty] | The investigation of vf/vt/af/at, stroke/cva, and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. based on the clinical details the cause of the positioning issue most likely was patient condition related due to small left ventricle. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
9435 | 1220648-2025-28355 | 22010109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22010109&pc=OZD | 04/24/2025 | 09/24/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633626 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that prior to initiating impella cp support, there was concern with arterial access and distal flow to the access leg. the leg post percutaneous cardiac intervention procedure did not have a return of arterial flow, and the patient was taken to the operating room for a vascular revision. it was further reported that there was a placement signal not reliable alarm during shockwave therapy. there was no intervention for the placement signal issue. the impella cp was explanted due to ischemia. | It was reported that there was a placement signal not reliable alarm during shockwave therapy. there was no intervention for the placement signal issue. |  |  |  |  | Impella cp was returned for evaluation. optical signal issue: returned pump was evaluated and 5s parameters were consistent with a damaged sensor. pump passed laser continuity testing and imaging of the sensor face was performed and showed damage of the sensor face. clinical details noted that the pump lost ao placement signal after shockwave therapy was initiated. proximity of shockwave device to impella was not noted in clinical details. data logs were reviewed and pump optical parameters had sudden drop of placement signal to 0 with all 5s optical parameters consistent with a damaged sensor face. motor current and pump flows remained stable at time of loss of placement signal. the root cause of the optical signal issue was due to a damaged sensor as a result of use of shockwave while on impella support, however, exact distance between shockwave device and impella optical sensor was not noted. ifu 0048-9007 rr notes that use of the shockwave device within 20mm can cause damage to the impella optical sensor. limb ischemia - reversible: clinical details noted that patient had pad and preexisting stents in iliac artery that required ballooning to accommodate the impella access sheath. when pump was inserted, patient loss distal pulse in right leg and required vascular revision after pump was removed to restore flow. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. impella cp was returned for evaluation. optical signal issue: returned pump was evaluated and 5s parameters were consistent with a damaged sensor. pump passed laser continuity testing and imaging of the sensor face was performed and showed damage of the sensor face. clinical details noted that the pump lost ao placement signal after shockwave therapy was initiated. proximity of shockwave device to impella was not noted in clinical details. data logs were reviewed and pump optical parameters had sudden drop of placement signal to 0 with all 5s optical parameters consistent with a damaged sensor face. motor current and pump flows remained stable at time of loss of placement signal. the root cause of the optical signal issue was due to a damaged sensor as a result of use of shockwave while on impella support, however, exact distance between shockwave device and impella optical sensor was not noted. ifu 0048-9007 rr notes that use of the shockwave device within 20mm can cause damage to the impella optical sensor. limb ischemia - reversible: clinical details noted that patient had pad and preexisting stents in iliac artery that required ballooning to accommodate the impella access sheath. when pump was inserted, patient loss distal pulse in right leg and required vascular revision after pump was removed to restore flow. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication vascular size or variations thereof detailed description of the symptoms experienced by the patient physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
9436 | 1220648-2025-28336 | 22010299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22010299&pc=OZD | 04/16/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/16/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629007 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Delivered as Unsterile Product | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was admitted to er with abdomen pain and chest pain with an inferior myocardial infarction. the patient was noted to be confused on arrival with hypoxia and intubated. given profound acidosis and comorbities with 3 pressers and not a surgical or extracorporeal membrane oxygenation candidate. patient proceeded to cath lab. right heart cath was completed confirming cardiogenic shock and left ventricle gram revealed vsd. patient transferred for further medical treatment. the complainant reported that the packaging of an introducer opened irregularly resulting in the sidearm getting contaminated. a second introducer kit was opened from a different 5.5 kit for use. following a successful implant, a wound vac was placed at the access site for bleeding. upon arrival to icu patient continued to decline. lactate continued to rise, greater than 20. she was anuric and would not tolerate continuous rental replacement therapy. patient was also not a candidate for extracorporeal membrane oxygenation due to body mass index and groin access. after further discussion with family, decision was made to withdraw care once family was able to arrive at bedside. withdraw of care occurred and patient expired quickly after. | [5 empty] | The medical safety officer reviewed and determined there is not enough information to exclude the impella as an association factor to the patient's outcome. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9437 | 1220648-2025-28335 | 22010536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22010536&pc=OZD | 04/25/2025 | 06/12/2025 |  | 05/13/2025 |  |  | 04/25/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639576 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and developed hemolysis with pink/light red tinged urine. the pump was removed and the patient survived. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the data was not returned for review. the pump was returned for analysis. visual inspection found no issues on the pump. pump passed hemolysis test. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no problem was found on the pump.
9438 | 1220648-2025-28356 | 22010553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22010553&pc=OZD | 04/24/2025 | 06/23/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624671 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was on impella 5.5 support and during preparation for impella 5.5 support, axillary vessel loops were tied distal and proximal for bleeding control. after the implant, the patient sat up abruptly and began kicking and swinging. when this happened, blood started to pour out into the jp drain and began to ooze through the incision. the patient became hypotensive requiring more levophed and epi. sedation was given. protamine was given. the patient was urgently taken back to the operating room where it was discovered that the vein clip had come undone when they woke up and vagaled causing a significant venous bleed. several clips were placed on the vein and then it was ligated with suture. the pocket was flushed and then closed again. support was continued and the patient is stable. | It was further reported that the patient experienced a stroke 24 hours post implant. the patient was successfully explanted after 598.83 hours of support and bridged to left ventricular assist device (lvad). |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. b5 updated to include additional information received from the clinical site d6b explant date added. h6 updated to reflect the additional failure mode the investigation of stroke/cva has not been completed should any new information be received a supplemental report will be submitted instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Based upon further investigation vascular damage - peripheral vessel investigation was updated: the root cause of the vascular damage was most likely patient movement as the bleeding occurred after the patient sat up abruptly and began kicking and swinging. stroke/cva: as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. h6 revised based on investigation results.
9439 | 1220648-2025-28357 | 22010986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22010986&pc=OZD | 04/24/2025 | 06/06/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644059 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's anatomy during attempted delivery due to the patient's narrowing vasculature. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings, ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to narrowing and the insertion was aborted.
9440 | 1220648-2025-28333 | 22011293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22011293&pc=OZD | 03/03/2025 | 06/18/2025 |  | 05/13/2025 |  |  | 04/25/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526703 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation, an impella cp was found to be kinked at the pump cannula and gave low flows. the pump was replaced and the patient survived. | [5 empty] | The investigation of low pump flow, product damage (cannula kink), and thrombus has been completed. clinical reports inability to flow appropriately. kink in cannula noted under angiography. pump removed due to possible thrombus ingestion. repositioning didn't resolve the issue. blood volume not given. patient was short and kink was observed near the motor housing area. decision made to remove cp and the replacement pump also had kinking issues. low pump flow: pump was returned for investigation. there was a kink confirmed near the motor housing of the pump and some biomaterial was present in the cannula near the impeller and outflow cage of the pump. data logs were investigated and there were suction alarms present. no motor current spike was observed. the pump was auto attempting to run on higher p-levels but consistent low flows observed. the root cause of the low pump flow issue was most likely a kinked cannula as per the product investigation, data log review and clinical information available. product damage (cannula kink): pump was returned and kink was confirmed near the motor housing of the pump. patient was a short women with other pump also having a kinked cannula. the root cause of the product damage cannula kink issue was most likely patient condition related to the patients smaller size as per clinical information and issue occurring on replacement device. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. b1 adverse event was inadvertently omitted on the initial manufacturer device report. b2 was inadvertently omitted on the initial manufacturer device report. h1 type of report was erroneously selected on the initial manufacturer device report number. h6 health effect - clinical code 4582 was erroneously selected on the initial manufacturer device report number. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9441 | 1220648-2025-28358 | 22011338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22011338&pc=OZD | 04/21/2025 | 05/28/2025 |  | 05/13/2025 |  |  | 04/21/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639818 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and the following day there were frequent suction alarms. volume was provided to the patient and impella mcs continued. approximately three days later, the patient experienced recurrent ventricular tachycardia, went asystole, brady, which was resolve with eight shocks. amino and lidocaine were administered. the patient as given packed red blood cells and platelets after labs. the following day the patient's overall condition was noted as unchanged but stable. the patient was a full code; however, the care team capped care with no plans for further interventions. two days later attempted update was unsuccessful as the patient was not in the designated room on the intensive care unit, and the impella pump was not streaming. status of the patient is unknown. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information received regarding the patient outcome; death. b2 revised changed life threatening to death and date of death added. b5 updated with additional information received from the clinical site d6b explant date added. h1 type of event revised to death. h6 health effect - impact code 1802 added.
9442 | 1220648-2025-28332 | 22011493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22011493&pc=OZD | 05/07/2025 | 07/15/2025 |  | 05/13/2025 |  |  | 05/08/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637996 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Defective Device | [6 empty] | Cardiac Tamponade | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp smart assist support and during the implant, the patient's iliac was found to be tortuous, and a 16 french long sheath was used, but a kink was found in the 16 french sheath. after placement, the patient's hemodynamics deteriorated, cardiac tamponade was confirmed, and surgical hemostasis was performed. when the thoracotomy was performed, several centimeters of the tip of the impella cp smartassist were found outside the ventricle. the perforation was repaired and the pump was explanted. the patient's outcome is serious. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system a section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ | Ventricle perforation: no product was returned for analysis. clinical details mention patient had very tortuous anatomy making it difficult to insert impella. insertion was done while pushing the pump quite strongly. the root cause of the ventricle perforation was most likely patient condition related as evidenced by the detail of very tortuous anatomy making it difficult to insert impella. introducer damage - mechanical: no product was returned for analysis. clinical details mention the patient's iliac was found to be tortuous, and a 16fr long sheath was used, but a kink was found in the 16fr sheath. the root cause of the introducer damage - mechanical was not determined as no product or images were returned for analysis.
9443 | 1220648-2025-28359 | 22011566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22011566&pc=OZD | 04/24/2025 | 05/23/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp for support during a high risk cardiac procedure. it was reported that after being transferred to the unit, the white repo sheath was significantly outside of the body and all four eyelets of the blue suture hub were sutured flat causing slow arterial oozing. the implanting physician was notified and came up to reposition. during reposition, the physician tented the repo sheath upwards as he advanced and a hematoma quickly developed. oozing had stopped, proper suture was placed and direct pressure was being placed on hematoma. it was noted that due to the circumstances of history and age and culprit vessel the physician decided to leave the patient on impella cp support. it was noted that the patient received a stent for culprit left ascending descending artery. oozing and hematoma were addressed and the intensivist team was holding heparin due to the hematoma. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. pump was not returned for investigation. data log review was not required for this case run. the cause of the access site bleeding issue was most likely use-related since the reposition sheath was pulled out and the blue suture hub was sutured flat down. the cause of the positioning issue was most likely use-related since the secured portion of the sheath was compromised during use.
9444 | 1220648-2025-28360 | 22011811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22011811&pc=OZD | 04/24/2025 | 06/30/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1156155 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) alarmed ¿controller failure: the purge system has stopped, switch to back-up controller.¿ upon starting up the console. the console was then replaced. no patient harm was reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other and aic - controller failure (red alarm) has been completed. apart from the controller failure (red alarm) due to a malfunction of the purge pressure sensor, no other purge-related issues were observed in either the logs or during the console testing. hence, 'aic-purge issue-other' is a duplicate failure mode of 'aic-controller failure (red alarm). the root cause of the controller failure alarm was a malfunction of the purge pressure sensor.
9445 | 1220648-2025-28361 | 22011865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22011865&pc=OZD | 04/22/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618545 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and approximately two days after start of the support, the nurse reported to the call center a purge fluid leak with purge fluid and purge pressure within normal limits. the nurse was advised to monitor the purge system and notify with any changes. the patient was planned for left ventricular assist device the following day, which was successfully completed. the impella 5.5 device was successfully wean and explanted. there was no report of harm to the patient as a result of the event. | [5 empty] | Explant date is unclear/unknown at the time of this mdr writing and was left bank accordingly. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: cleaning, ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9446 | 1220648-2025-28362 | 22012113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012113&pc=OZD | 04/24/2025 | 07/02/2025 |  | 05/13/2025 |  |  | 04/24/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026651060 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp-extracorporeal membrane oxygenation support was being escalated to an impella 5.5 for more support when the team observed a left atrium clot. the impella 5.5 was still placed but, when the flows were not as expected the left ventricle and pump were assessed and were seen to be full of thrombus burden. the impella 5.5 was removed and flushed. the patient went into ventricular fibrillation and was shocked and the clotting situation prompted the team to withdraw/stop support and care. the patient expired. the european team communicated the pump was not to be the cause. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus, low pump flow, and vt/vf has been completed. low flow: reported pump unable to generate flow when started. pump was removed and no thrombus was noted on the pump and ran well in water. impella 5.5 was returned and no damages or abnormalities were found. data logs reviewed and no motor current (mc) spike, placement signal trends or suction alarm observed. pump was unable to run with flows at almost zero during initiation of pump run. the root cause of the low pump flow was most likely patient condition related since pump was unable to run in the patient but ran without issue in water after it was explanted and also patient had volume issues with huge thrombi blocking and completely filled in left ventricle (lv). thrombus: reported large thrombus in left ventricle. as the necessary information was not provided, the root cause of the thrombus was not determined vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28362. h6 medical device problem code 1248 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28362.
9447 | 1220648-2025-28363 | 22012126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012126&pc=OZD | 04/22/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595728 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. approximately two hours after start of the support limb ischemia concern was noted and the impella cp device was explanted. the physician, on the unit, assessed distal pulses and decided the patient no longer needed the impella cp device. the impella cp device performance level was lowered to p-2, device pulled across the valve and then lowered to p-0 followed by removal. manual pressure was applied until hemostasis was achieved. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9448 | 1220648-2025-28364 | 22012326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012326&pc=OZD | 04/25/2025 | 06/23/2025 |  | 05/13/2025 |  |  | 04/25/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653605 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a surgical patient on impella 5.5 support. during the support, there was suction issues remedied by blood products. the access site was oozing as well and thus kept the team from adding heparin for anticoagulation. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. low flow: data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for 1 hour and 15 minutes, low flow occurred that triggered auto suction response & suction alarms. this event remained for about 1.5 hours then resolved and pump ran without issue to the rest of the case. based on clinical description and data review, the root cause of low flow was most likely patient condition related since cvp has drop when the issue happening.
9449 | 1220648-2025-28365 | 22012457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012457&pc=OZD | 04/22/2025 | 06/16/2025 |  | 05/13/2025 |  |  | 04/28/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026648736 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 85 KG | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after implantation of an impella cp intermittent suction alarms occurred. the pump was turned to p1 to resolve the alarms. additionally, the impella stopped multiple times. the pump was able to be restarted but due to signs of biomaterial ingestion the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of pump stop, thrombus, and low pump flow has been completed. the impella device was not returned for evaluation. pump stop: the data log shows several pump stops with alarm 13 - motor current high on the 6th day of pump support. a motor current spike was reported at 1400 during the pump stop. there were no abnormalities reported on the purge parameters. the cause of the pump stop issue was not determined without returned product investigation. low pump flow: the data log shows low pump flow at p8 during the start of the case, accompanied by a suction alarm. the pump was turned to p2, and the p level was never increased throughout the case run. several suction alarms were reported, and the pump flow fluctuated between 0 to 1 l/min. the patient was on va ecmo throughout the case. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. h6 medical device problem code 1248 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28365.
9450 | 1220648-2025-28366 | 22012537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012537&pc=PYX | 04/23/2025 | 05/22/2025 |  | 05/13/2025 |  |  | 04/23/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561690 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella rp flex and exhibited oozing. hemostasis was obtained with a purse string and mattress sutures. the patient had one episode of ventricular tachycardia. additionally, there was concern for hemolysis. the impella rp flex was decreased to p6 from p8. the patient was successfully weaned and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of access site bleeding, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9451 | 1220648-2025-28367 | 22012548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012548&pc=OZD | 04/23/2025 | 06/05/2025 |  | 05/13/2025 |  |  | 04/23/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026638673 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Arrhythmia | Atrial Fibrillation | Hemolysis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5. during support, the patient experienced ventricular tachycardia and atrial fibrillation, and was started on amiodarone infusion. the patient also had loss of pulsatility, so the p level was adjusted. additionally, the impella was turned down to p6 due to red/pink urine. the patient survived. | [5 empty] | The investigation of hemolysis, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 product problem selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28367. d6a was erroneously entered on the initial manufacturer device report number 1220648-2025-28367. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28367. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
9452 | 1220648-2025-28368 | 22012795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012795&pc=OZD | 04/09/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025599641 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient endured hemolysis, worsening anemia, and ischemia during impella cp support. the ischemia was believed to have a definite relationship to the impella device. the patient's leg was noted to be cold with no dopplerable pulse. similarly, the hemolysis was also believed to have a definite relationship to the impella device. the pump received two units of blood to treat the condition. finally, the worsening anemia was believed to have a possible relationship to the device used. two units of blood were given to the patient to treat the condition. the pump was removed in response to the conditions. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
9453 | 1220648-2025-28369 | 22012899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22012899&pc=OZD | 04/27/2025 | 05/30/2025 |  | 05/13/2025 |  |  | 04/27/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643662 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Unknown |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | Hemolysis | Renal Failure | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and exhibited a drop in p levels during pump repositioning. the patient became non-pulsitile and a code blue was initiated. the patient was defibrillated and return of spontaneous circulation was achieved. repositioning of the implella was attempted to retract the catheter as an echocardiogram showed shallow placement but the decision was made to leave the impella cp at five centimeters within a dilated left ventricle. the patient exhibited dark colored urine and was placed on continuous dialysis. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues, vf/vt/af/at, hemolysis, and renal failure has been completed. the impella device was not returned for evaluation. based on the clinical and the data log analysis, the root cause of the hemolysis is most likely due to improper positioning. due to lack of clinical details provided, it is unknown of why the pump came out of position. the cause of the renal failure and vt/vf were not determined due to insufficient clinical details. b1 product problem selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28369.
9454 | 1220648-2025-28371 | 22013034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22013034&pc=OZD | 04/21/2025 | 06/12/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025589765 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 57 KG | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impella 5.5 pump was implanted in a patient when the care team observed purge alarms, in addition to pump saturation. the team troubleshot by the addition of tissue plasminogen activator. when the pump was explanted biomaterial was seen at the cannula. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of high purge pressure, thrombus, and saturated flow has been completed. returned product was investigated and after soaking the pump, there was biomaterial noted to be in the outflow cage and purge gap. however, there is potential that this was not the ingested biomaterial, since saturated flows resolved in the field. the pump was connected to a console and run. for the first 10 seconds, pump flows were saturated, but quickly became low, likely due to movement of the biomaterial in the cannula. high purge pressure did not reproduce after running the pump two separate times. this could potentially be explained by tissue plasminogen activator (tpa) being added to the purge or other factors after the pump was explanted. high purge pressure: data logs showed the purge pressure rise suddenly over 2 minutes immediately after a biomaterial ingestion signature. returned product had biomaterial in the outflow cage and around the purge gap. however, high purge pressure did not reproduce when connecting the pump to the console. this suggests that biomaterial on returned product may not have been the biomaterial that was ingested. based on the signature noted on data logs, the root cause of the high purge pressure was determined to most likely be biomaterial ingestion. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. saturated flow: data logs showed the signatures of a biomaterial ingestion event by means of a motor current spike and motor speed deviation followed by a drop in diastolic ps. returned product was investigated and there was biomaterial in the outflow cage and around the purge gap. that being said, there is potential that this was not the biomaterial that was ingested in the case, since saturated flows resolved in the field. saturated flows only reproduced when running the pump in a bucket for a few seconds, and then quickly turned to low pump flows. this can likely be explained by the biomaterial on returned product shifting location in the cannula. again, it was not expected that saturated flows would reproduce since they resolved in the field. based on the signature seen in data logs, the root cause of the saturated flow was most likely biomaterial ingestion.
9455 | 1220648-2025-28372 | 22013187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22013187&pc=OZD | 04/22/2025 | 06/18/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after five days of support, an impella cp was explanted due to concerns of hemolysis and low flows/continuous suction alarms. the team could not troubleshoot them effectively with either fluid bolus or pump repositioning. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of low pump flow and hemolysis has been completed. the impella device was not returned for evaluation. low pump flow: clinical states that the patient had continuous suction alarm and patient's left ventricle (lv) was severely underfilled, had volume issues and 1.2l volume was given for troubleshooting but no change. repositioning was unsuccessful as well. pump was not returned for investigation. data logs were investigated and suction alarms were confirmed. the pump flow was observed to be gradually decreasing at p2 going below the lower range 1.1l/ min. there was a motor current spike at the start of the decreased flow when the pump p-level was being reduced. the placement signal was slightly drifting down corresponding to the low flow. therefore, as per the clinical information and data log review, the root cause of the low pump flow issue was most likely lv severely underfilled patient condition related. hemolysis: clinical states that the patient was experiencing frank hemolysis during the low flow event. pump was repositioned but repositioning did not help troubleshoot the situation. pump was not returned for investigation. data logs confirm the pump was out of position with multiple "impella position in aorta" alarm. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device per echo showing pump shallow which was confirmed in the data logs.
9456 | 1220648-2025-28373 | 22013370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22013370&pc=OZD | 03/11/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 03/11/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025587344 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during removal of an intra-aortic balloon pump, a patient implanted with an impella 5.5 went in and out of ventricular tachycardia (vt). the patient received one shock before returning to a normal sinus rhythm. arrhythmia and vt episodes returned intermittently over the next three days, for which defibrillation and cardioversion were performed. it was additionally reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry, that the patient endured hemolysis during support. the condition was believed to have a possible relationship to the impella device used. one unit of blood was given to the patient and support levels were manipulated to meet the needs of the patient. roughly twenty days into support, the connection between the cassette yellow luer and the purge sidearm became stuck. despite attempts to soak the connection, it would not loosen and plans to change the cassette were temporarily abandoned. support continued until planned explant the following day. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
9457 | 1220648-2025-28374 | 22013553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22013553&pc=OZD | 04/27/2025 | 10/09/2025 |  | 05/13/2025 |  |  | 04/27/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636085 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while implanted with an impella 5.5, the patient had nasal bleeding and heparin was stopped. systemic anticoagulants were withheld due to concerns of a gastrointestinal bleed. additionally, there was a purge pressure high alarm. the purge cassette was changed and the patient continued on support. | The complainant also reported that a placement signal not reliable alarm occurred. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | Optical signal issue: after 10 days on support, psnr alarms occurred. the pump was not returned. data log review showed intermittent ps failure where snr dropped to zero and contrast barcoded to +/-3000. psnr alarm occurred. review of prior pumps used by ic5204 showed similar psnr issues. upon review of the data logs, it is apparent that the console is the potential cause of the optical signal issue. please refer to complaint # (b)(4) for an investigation into the console high purge pressure: after 12 days on support, purge pressure high occurred with blocked purge. the hpp did not resolve before explant. troubleshooting efforts included 2 bags of tpa. the pump was not returned. data log review showed gradual purge pressure rise with purge flows decreasing. no mc spike was noted and mc did not trend during hpp trend and purge pressure blocked alarms. logs also show troubleshooting occur with bag and cassette changes, none of which were long changes. the cause of the high purge pressure was not determined since product was not returned and insufficient clinical details. vascular damage - peripheral vessel: after 4 days on support, pt experienced a nose bleed. pt taken off heparin. concerns for gi bleed. the pump was not returned. the cause of the vascular damage - peripheral vessel was not determined since product was not returned and insufficient clinical details. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated.
9458 | 1220648-2025-28375 | 22013733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22013733&pc=PYX | 04/22/2025 | 10/09/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2025571328 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Sense | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an implanted impella rp flex generated placement signal spikes despite the motor current on the pump remaining stable. it was additionally noted that low pump flows were encountered at this time. support continued with the noted pump despite the reported issue. there was no patient harm. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. placement signal issues: the pump passed electrical testing and was flow loop tested in the lab. the dp sensor drift reproduced on the returned product. the data logs from the case had no placement signal not reliable alarms but show a gradual drift in placement signal concurrent with a drift down in calculated pump flow. motor current remained stable and central venous pressure as well. the pattern observed is characteristic of a sensor drift of the dp sensor. the root cause of the placement signal issue was determined to be sensor drift as evidenced by log pattern and issue reproduction on returned product a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. | The investigation of low pump flow has been completed. the returned pump had dried blood by the impeller. no problem was found as the root cause of the low pump flow. the data logs show motor current stability with no spikes outside p-level changes. low pump flow concern was due to the effect of the placement signal issue (sensor drift) on the calculated flow. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-28375.
9459 | 1220648-2025-28376 | 22013937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22013937&pc=OZD | 04/22/2025 | 05/13/2025 |  | 05/13/2025 |  |  | 04/22/2025 | 05/13/2025 |  |  |  |  | 05/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that implantation of an impella cp was aborted due to incompatibility with the patient's anatomy. no other information was provided. device identifiers were requested but not provided. the patient survived. | [5 empty] | A4 and d4 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9460 | 1220648-2025-28288 | 22001342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001342&pc=OZD | 05/01/2025 | 07/16/2025 |  | 05/12/2025 |  |  | 05/07/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601391 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was being supported with an impella 5.5 pump for high-risk surgery. the pump supported for over 19 days when it was explanted due to low purge flow. there were concerns for a pump stop from pump thrombosis, which may have been causing the purge flow issues. the patient survived the procedure. | [5 empty] | Impella 5.5 was returned for evaluation. high purge pressure: upon visual inspection, no abnormalities noted. minimal dried blood around edge of impeller blades. purge gap was clear. pump was purged in the lab for approximately 18 minutes, then run at p-9 for approximately 5 minutes and high purge pressure did not reproduce. data logs were reviewed and rt logs at time of reported issue showed a motor current spike with speed deviation and step up in pump flow. 10 minutes later, multiple additional motor current spikes with small speed deviations occurred with a slight increase in purge pressure and drop in purge flow. no purge alarms were triggered. pump flow also became fully saturated at that time. clinical details noted that the motor waveform suddenly became flat and purge flow rates dropped to about half of what they were previously reading. the root cause of the high purge pressure was most likely due to biomaterial ingestion based on returned data log analysis. thrombus: clinical details noted thrombus was observed on explant. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9461 | 1220648-2025-28289 | 22001441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001441&pc=OZD | 04/14/2025 | 06/13/2025 |  | 05/12/2025 |  |  | 04/14/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) for recovery or bridge to decision. approximately 11 days after start of the impella mcs, high motor current presented. following, the pump stopped and was unable to restarted with the automated impella controller after multiple attempts. the patient was cannulated with extracorporeal membrane oxygenation for continued support and was taken to the operating room for explant of the impella cp. the patient was noted to have survived the impella cp explant and was noted to be in stable condition following the event. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-28082 for report on the impella cp pump. | Data analysis: at 7:18 on 4/3, cp pump 576601 was connected to (b)(6). at 1:38 on 4/14, the console receives an event #13 which indicates a pump stop due to high motor current ((b)(6) imclog pump stop.pdf). the console tried to restart the pump seven times including resetting the pmd but the pump stopped every time. this pump was not used on any other consoles. the rt logs confirm that directly preceding the first motor current high pump stop, the motor current spikes which would have triggered the alarm and pump stop ((b)(6) rtlog pump stop.png). device analysis: the console was returned for investigation but the failure mode was not reproduced during testing. root cause: there was a pump stop during the case but the root cause was found to not be with the durable product (see (b)(4)). (b)(6) performed as expected. repair: there was no problem found with (b)(6), the pump stop was unrelated to the durable product. please perform a functional check on the console.
9462 | 1220648-2025-28290 | 22001497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001497&pc=OZD | 04/21/2025 | 06/09/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647459 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon insertion of an impella cp on patient, the flows were noted to be .8 in auto. looking at the x-ray, the cannula was kinked. attempts were made to straighten out the cannula and reposition but were unable to straighten it. a decision was made to place a new impella device for the high-risk percutaneous coronary intervention. | [5 empty] | Low flow - reported pump unable to flow higher than 0.8 and noted a kink in the cannula. impella cp was returned and a cannula kink was found. data logs show pump was set on auto and was running at p2 with low flows from case start then explanted. the root cause of the low pump flow was a most likely a kinked cannula since a cannula kink was reported and found on the returned product. cannula kink - reported some difficulty tracking at the arch but placed properly and started in auto with cannula kink. product was returned and a cannula kink was found. lack of clinical details such as any vessel conditions or any excessive force used. the root cause of the product damage was a kinked cannula but the reason for the kink is unknown due to lack of clinical information. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9463 | 1220648-2025-28291 | 22001513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001513&pc=OZD | 11/27/2024 | 06/20/2025 |  | 05/12/2025 |  |  | 05/02/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025573935 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematoma | Gastrointestinal Hemorrhage | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through abiomed's loqi study database that a patient implanted with an impella 5.5 exhibited anemia. the relationship is noted as "possibly" related to the 5.5 pump support. loqi noted: that hemoglobin dropped from 14.7 to 11.1. it was found to have hematoma and likely initial cause of acute anemia. anemia did not improve post hematoma evacuations. patient also had recurrent gi bleeding secondary to perforated duodenum. patient had multiple blood products throughout hospital stay. the patient expired. | [5 empty] | Access site bleeding: the root cause of access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. vascular damage peripheral vessel: the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9464 | 1220648-2025-28292 | 22001605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001605&pc=OZD | 04/21/2025 | 06/06/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1719164 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomedical engineer reported the automated impella controller touch screen was not working. this issue occurred during use. there was no patient harm reported. | [5 empty] | Data analysis: the case associated with the reported complaint was started on (b)(6) 2025. on (b)(6), the first log entries indicating "imc-kbd error: 0xa1 0x00 0x02" appeared, suggesting that the issue occurred as touch detected outside soft-key area just before the purge cassette was changed. following this incident, several additional error messages were logged, highlighting ongoing issues with multiple touches detected outside the soft-key area. notably, these entries included the "imc-ahp error: 0x91 0x83 0x24," which corresponds to an rfid read block error (ahp: error rfid). (b)(6). device analysis: the console (b)(6) was sent back to field service for further investigation and was tested after arriving in fs. the reported issue was even after performing 100 power cycles. the console was run with pump and purge cassette for couple of hours and no issues were observed. additionally, the internal circuitry and cables of the console were examined, and everything appeared to be in good condition. root cause: the root cause for the reported the touch screen isn't working could not be determined due to the inability to reproduce. corrective actions: before returning the console to the customer, the following actions are instructed to perform: replace the glass keyboard as a precaution (0042-1090). clean the purge system from inside out, as well as the front panel as a precaution. perform section 8 - 16 from preventive maintenance procedure (0042-7309). perform a full functional check on the console and optical system (0042-7316). | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9465 | 1220648-2025-28293 | 22001608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001608&pc=OZD | 12/08/2024 | 05/12/2025 |  | 05/12/2025 |  |  | 12/08/2024 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501327 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 35-year-old male patient was on impella 5.5 support and during support, the patient had several runs of ventricular tachycardia (vt). the patient was given amio and lido. support continued. additionally, the long-term outcome and quality indicator (loqi) database reported upon an adverse event for acute kidney injury with relationship to pump being listed as "possibly" related. loqi states: "creatinine peak at 2.52 in the setting of rv dysfunction and acute tubular necrosis, post lvad. resolved with effective diuresis under ionotropic support. the patient bridged to lvad. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
9466 | 1220648-2025-28294 | 22001758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001758&pc=OZD | 04/20/2025 | 06/11/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627468 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp and had a pulmonary hemorrhage during a bronchoscopy. heparin drip was held to manage the condition. two days later, suction alarms were encountered at higher p-levels. heparin administration was increased in response to the low flow, however, the patient had a do not resuscitate order and passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of vascular damage - peripheral vessel and low pump flow has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: it was stated that there was a pulmonary hemorrhage found during a bronchoscopy. due to lack of clinical details, the root cause of the vascular damage - peripheral vessel cannot be determined. low pump flow: on 4/22 there was suction alarms and lower flows. there was a trend in the motor current and placement signal. the clinical stated that heparin was added to increase the response to the low pump flow but the patient was made do not resuscitate (dnr) and passed away the following day. based on the clinical details and data log analysis, the root cause of the low pump flow is most likely due to the patients condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9467 | 1220648-2025-28295 | 22001864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22001864&pc=OZD | 02/04/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 05/01/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025534709 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured recurrent ventricular tachycardia with a "possible" relationship to the impella cp device used: ¿multiple episodes of ventricular tachycardia. required amiodarone and cardioversion.¿. the patient recovered with no sequelae. it was also noted that the patient was cardioverted from atrial fibrillation to sinus tachycardia. in addition to the patient had a fever so antibiotics were given. the impella cp was explanted and replaced with an impella 5.5. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9468 | 1220648-2025-28296 | 22002068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002068&pc=OZD | 02/16/2025 | 06/04/2025 |  | 05/12/2025 |  |  | 02/16/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511288 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old female on a impella 5.5 for cardiomyopathy. during support the patient had hemolyzed labs and a suction event which they gave volume. it was noted that the patient was not a candidate for transplant. care was ultimately withdrawn and the patient expired. the patient was on impella support for a total of 2,288.40 hours. | [5 empty] | The investigation of hemolysis and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9469 | 1220648-2025-28297 | 22002225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002225&pc=OZD | 04/20/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582088 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hemorrhagic Stroke | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted in a patient post extracorporeal membrane oxygenation cannulation. during support, the patient exhibited bleeding, suspected to be from cardiopulmonary resuscitation and massive transfusion protocol was given. in addition, doppler pulses could not be found on the impella right side. the patient received one unit of packed red blood cells. the patient then suffered a hemorrhagic stroke. care was withdrawn and the patient expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9470 | 1220648-2025-28298 | 22002294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002294&pc=OZD | 04/13/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/13/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589981 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella cp was placed via axillary artery after 5.5 attempt was unsuccessful. in icu, bleeding was noted from axillary surgical site. the dressing was changed and one unit of blood was given. on day six of impella support, the patient went into atrial fibrillation with rvr and the patient seemed very septic and not a surgical candidate. it was further reported that the family decided to withdrew care and the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the patient's outcome was reviewed by medical safety officer. although care was withdraw, there is not enough evidence to exclude the impella device (related complications: access site bleeding/bleeding requiring transfusion, arrhythmic event and sepsis) as an associated factor to the overall outcome.
9471 | 1220648-2025-28299 | 22002436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002436&pc=OZD | 04/21/2025 | 06/12/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1191152 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomedical engineer reported that the automated impella controller had an error on the screen: controller error. no other information was provided. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs are consistent with complaint and show controller error alarm (1039 due to defective optical sensor lamp) present shortly after console boot-up. it is preceded by out-of-spec ¿5s¿ parameters detected, not resolved after optical bench self-reset. console was tested and the issue was reproduced, as controller error alarm 1039 occurred, preceded by log entries reporting out-of-spec parameters: inspection of the optical pump connector revealed its heavy contamination, which resulted in light ¿leaking¿ or coupling into the contaminant, instead of being reflected back from the fiber facet, into the optical bench. after cleaning the connector, the issue was resolved, and controller error alarm no longer presented. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9472 | 1220648-2025-28300 | 22002515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002515&pc=OZD | 04/18/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589256 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female | 85 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 exhibited persistent bleeding. the pump remained on for support for weeks. the choice was then made to withdraw care and patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
9473 | 1220648-2025-28301 | 22002652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002652&pc=OZD | 04/20/2025 | 06/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644253 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention implanted with an impella cp device for mechanical circulatory support experienced limb ischemia, positioning issue, and intermittent low pump flow and would eventually expire. the patient was consulted for coronary artery bypass graft surgery and was promptly turned down. after the physicians collaborated it was determined that protected (pci) with impella was the best course of action for the patient. the impella was successfully placed, and extensive intervention of the left anterior descending (4 drug-eluting stents) and left main (1 drug-eluting stent) arteries was completed with "semi" good result. there were many episodes of no coronary re-flow that resulted in ventricular fibrillation x3 with successful return of spontaneous circulation (no cardiopulmonary resuscitation was noted). patient was very hypotensive post intervention. fluids and levophed were started while on support. the groin was warm, dry and not bleeding. the patient was transported back to the unit in "very" critical but stable condition. the following day, the patient was transferred to a different facility. the patient¿s right leg (impella leg) was severely mottled and cold to the touch, no pulses. the left leg was cool to touch and had a femoral pulse. upon arrival to the receiving facility's unit, there was a femstop on the right leg that was swiftly removed. bedside echocardiogram showed the impella at 5.1cm from the aortic valve. the fellow did not feel comfortable adjusting to pull back even though impella was intermittently having suction against the septal wall. reevaluation was planned for the next morning. the echogram also showed new bi-ventricular failure. the right ventricle was hardly moving, and the left ventricle was not moving at all. the patient had some asystolic episodes for 30-45 seconds but maintained a mean arterial pressure with the impella. discussion regarding escalation to an impella 5.5 was pending with cardiothoracic surgery. however, the patient would eventually expire this same day. care was withdrawn. | [5 empty] | Product was not returned for review. data review: data logs only available for first day of support. suction alarms were reported on second day of support. data logs showed about 3 hours into support, flow dropped slightly from 3.5 l/min to 3.1 l/min that triggered auto suction response & suction alarms. no positioning alarm was triggered on the logs. clinical states bedside echo showed impella at 5.1cm from the aortic valve. fellow did not feel comfortable adjusting to pull back even though impella was intermittently having suction against the septal wall. echo also showed new bi-v failure with the rv hardly moving and the lv is not moving at all. low flows: the root cause of low flow was most likely patient condition since echo showed bi-ventricle failure with the rv hardly moving and the lv not moving at all. positioning issue: the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. limb ischemia: ischemia occurred after patient was transferred to new hospital on the 2nd day of impella support. echo found pump was in improper position with suction alarms & lack of moving on bi-ventricles. pump then was explanted as the family withdrew the care. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition; any concomitant medication; vascular size or variations thereof; detailed description of the symptoms experienced by the patient; physical examination findings, including pulse assessment and visual examination of the affected limb; diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography; laboratory test results, including blood tests for markers of inflammation or clotting disorders; any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension; patient's response to previous treatments or interventions for vascular conditions; with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | Care was withdrawn and the patient expired. however, given the severity of the adverse event, suboptimal positioning and intermittent low pump flow potentially compromising hemodynamic support, there is not enough information to exclude the impella cp device as an associated factor. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9474 | 1220648-2025-28303 | 22002811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002811&pc=OZD | 04/21/2025 | 06/09/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621030 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Malfunction | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and after the implant, the impella cannula kinked just below the outlet and was unable to flow higher than 0.8. the pump was explanted and was replaced. the patient expired on support of the replacement pump. | [5 empty] | Impella cp was returned for evaluation. low pump flow: upon visual inspection, a minor kink was noted on pump cannula just distal of the outflow cage. data logs were reviewed and rt log of case showed multiple instances of low flow with lack of motor current pulsatility during 4 minutes of pump run. clinical details noted that a cannula kink was observed during the pump run and pump was not able to run higher than 0.8l/min. the root cause of the low pump flow was most likely due to a cannula kink based on returned product and data log analysis, however, the cause of the cannula kink was not noted. product damage (cannula kink): returned product analysis showed a minor kink just distal of the outflow cage along the cannula. clinical details noted that a kink was observed during the case and pump was not able to flow higher than 0.8l/min. the root cause of the product damage was most likely due to a cannula kink, however, the cause of the cannula kink could not be definitively determined as limited clinical details were provided. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. b.2 exact date of death is unknown.
9475 | 1220648-2025-28304 | 22002856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002856&pc=OZD | 04/20/2025 | 06/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654952 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella 5.5 device to a patient with cardiomyopathy was unsuccessful due to the cardiac anatomy. the impella 5.5 device had been backloaded over the guidewire and resistance was met in the right axillary, past the graft anastomosis. the decision made to stent right axillary with an 8mm omni link stent and post dilated with 8mm armada balloon. the stent was fully opposed. the physician attempted to pass the impella and was again unsuccessful. the decision made to abort the impella 5.5 implant and place an impella cp instead. there was no report or indication of harm to the patient. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella 5.5 met resistance met in right axillary, past the graft anastomosis. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28304. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9476 | 1220648-2025-28305 | 22002953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22002953&pc=OZD | 04/12/2025 | 06/05/2025 |  | 05/12/2025 |  |  | 04/12/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628500 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support was taken to the operating room for attempted impella cp swap to intra-aortic balloon pump. the patient decompensated and the exchange was aborted. since returning from the operating room, there has been intermittent position alarms and suction at p7. the pump was repositioned by the medical team without imaging. cm marking now was 84 and was previously 91. echocardiogram was done now to assess placement and the pump appeared slightly deep but away from ventricular anatomy. the decision was made not to reposition. the next day, there was an impella in aorta alarm in the morning during nasoduodenal tube placement. the patient was bearing down that resolved when placement was completed. impella position unknown advisory alarm was also present and the patient on this day was able to be escalated to impella 5.5. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely patient movement related as the pump had position alarms after transfer to/from the operating room and also during naso-duodenal (nd) tube placement with patient bearing down that resolved after the procedure.
9477 | 1220648-2025-28306 | 22003000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003000&pc=OZD | 04/20/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026655482 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 88 KG | Non Hispanic | White | Injury | Yes | No | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella cp the patient had oozing at the access site. heparin was withheld and the dressing was changed. additionally, the pump was unable to get above p7 without suction alarms. the pump pigtail appeared to be caught under papillary muscle. support continued and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9478 | 1220648-2025-28307 | 22003158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003158&pc=OZD | 04/12/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/12/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627316 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 patient needing mechanical circulatory support. there was a leak in the purge filter. the leak appeared to be occurring inside the filter. a pocket of purge solution was collecting inside the chamber and leaking down the cable. no obvious cracks were observed internally or externally. the patient remained stable. | [5 empty] | The impella device was not returned, and the investigation has been completed. the data logs showed that purge pressure was stable and there were no purge alarms seen. without sufficient product returned, the root cause of the purge leak - pump could not be determined. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9479 | 1220648-2025-28308 | 22003238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003238&pc=OZD | 04/20/2025 | 06/16/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026653679 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 95 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 pulled the pump across the left ventricle. the patient was intubated in the icu and decompensated quickly. va ecmo was placed at the bedside. the surgeon wanted to attempt getting the impella across the valve a few more times before going to surgery. during those attempts the anticontamination sleeve was damaged. the patient was taken to surgery and the impella was exchanged. the patient survived. | [5 empty] | The investigation of positioning issues and product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. positioning issue: clinical states that patient pulled his impella and pump was noted to be fully across the left ventricle (lv). pump was tried to be repositioned but was unsuccessful. pump was not returned for investigation. as per the clinical information provided, the root cause of the positioning issue was patient movement related to patient pulling back the catheter of the pump. product damage (sterile sleeve): clinical states after patient pulled the catheter, the pump was attempted multiple times to reposition it but was unsuccessful and during the attempts, the sterile sleeve of the pump was damaged. pump was not returned for investigation. as per the clinical information provided, the root cause of the product damage issue was use related to sleeve being damaged wile reinserting the pump. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9480 | 1220648-2025-28309 | 22003350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003350&pc=OZD | 04/21/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644274 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp was implanted to support a percutaneous coronary intervention procedure. when closing the access site, two wires were placed and double angioseal was attempted with significant bleeding post deployment. manual pressure was held for 40 minutes and femstop was placed before leaving the catheterization lab to intensive care unit (icu). femstop was removed on arrival to icu, 30 minutes of manual pressure was held and hemostasis achieved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9481 | 1220648-2025-28310 | 22003538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003538&pc=OZD | 04/21/2025 | 06/27/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654838 |  | 0048-0044 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Detachment of Device or Device Component |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the impella cp required repositioning. when the physician moved to do this, they noted that the catheter required a significant amount of force to move, .to a point that it was considered dangerous given the patient¿s anatomy and repeated suction alarms and hemolysis. .the anti-contamination sleeve broke due to stiffness with the repositioning sheath and inability to move the impella cp without significant force.. the pump was removed and the patient outcome is unknown. | [5 empty] | The investigation of positioning issues, product damage (sleeve damage), low pump flow, and hemolysis has been completed. positioning issues: data logs were not available for review. returned product was investigated and there were no abnormalities related to the issue. the touhy-borst valves were able to tighten/loosen without issue. since clinical details reported that pump flows were being limited due to a patient condition that would allow the pump to move position easily and there were no abnormalities with returned product, the root cause of the positioning issue was most likely patient condition. low pump flow: data logs were not available for investigation. returned product was investigated, and there were no visual abnormalities. the pump was connected to a console in the lab and left to run in a bucket of water for several hours. pump flows were within specification. since clinical details reported that pump flows were being limited due to a patient condition that would allow the pump to move position easily, the pump was being repositioned at the time that suction alarms were reported, and low flows did not reproduce on returned product, the root cause of the low pump flow was most likely improper position. product damage (sterile sleeve): returned product was investigated and there was no damage noted to the sterile sleeve. the touhy-borst valves were both loosened and the catheter was able to slide through without resistance. the valves were then tightened, and the catheter was secured in place. since the physician and an abiomed representative both reported that the catheter would not move through the touhy-borst valves without significant force and the issue did not reproduce, the root cause of the product damage could not be determined. hemolysis: data logs were not available for investigation. returned product was investigated, and there were no visual abnormalities. the pump was hemolysis tested, and the resulting average mih was 9.15, which is a passing result. since clinical details reported that pump flows were being limited due to a patient condition that would allow the pump to move position easily, the pump was being repositioned at the time that hemolysis symptoms were reported, and returned product passed hemolysis testing, the root cause of the hemolysis was most likely improper position. d9 device returned to manufacturer date added h6 component code 4723 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28310. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9482 | 1220648-2025-28311 | 22003705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003705&pc=PYX | 04/21/2025 | 09/26/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025611299 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella rp flex was placed for acute myocardial infarction/cardiogenic shock and with good position on fluoroscopy but could not get the impella flow above one liter per minute. as the impella rp flex was pulled back slightly, flow improved to less than two liters per minute on p5. the patient continued to deteriorate, and the family made the patient do not resuscitate and decided to withdraw care. the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of low pump flow has been completed. low pump flow: clinical states that the pump was getting low flows upon activation. pump was repositioned and console was changed and adequate flows were observed at p5. pump was returned for investigation. the pump had no biomaterial, kink or any physical or mechanical defects. low pump flow was not reproduced. the pump worked as per the specification. data logs were investigated. low flows were confirmed. the pump was attempting to run at higher p-levels but was getting flow of between 1l/min to 2l/min. the expected flow at p7 is 2.7l/min to 3.4l/min. no suction alarms. console was replaced and showed no low flow. the pump had flows within the range on p5. no suction alarms present. placement signal was seen slightly drifting up. therefore, based on the product investigation, data log review and clinical information, the root cause of the low pump flow issue was most likely improper positioning of the device since after changing the console and repositioning the pump, flows were within the range. d9 device returned to manufacturer date added g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number: 1220648-2025-28311. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. low pump flow: clinical states that the pump was getting low flows upon activation. pump was repositioned and console was changed and adequate flows were observed at p5. pump was returned for investigation. the pump had no biomaterial, kink or any physical or mechanical defects. low pump flow was not reproduced. the pump worked as per the specification. data logs were investigated. low flows were confirmed on console ic1503. the pump was attempting to run at higher p-levels but was getting flow of between 1l/min to 2l/min. the expected flow at p7 is 2.7l/min to 3.4l/min. no suction alarms. console was replaced. ic2218 shows no low flow. the pump had flows within the range on p5. no suction alarms present. placement signal was seen slightly drifting up. therefore based on the product investigation, data log review and clinical information, the root cause of the low pump flow issue was most likely improper positioning of the device since after changing the console and repositioning the pump, flows were within the range. placement signal issue: clinical states that the placement signal was improving on the console 2 overnight. pump was investigated and no placement signal alarms were observed. data logs were investigated and a placement signal rise was observed corresponding to a dip in flow. the flow started drifting down and placement signal started drifting up. cvp was trending down as well. no related alarms were observed. therefore, the root cause of the placement signal issue was most likely patient condition related as per the clinical information and data log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. additional information: h6. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9483 | 1220648-2025-28312 | 22003770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003770&pc=OZD | 04/20/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631389 |  | 1000080 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp device was successfully implanted, and the patient was transported to the unit on impella mcs. upon arriving to the unit, purge system block alarmed was triggered and did not resolve despite confirming there were proper connections, no kinks, tuohy was tight, attempted de-air, and changing of the purge cassette. the automated impella controller (aic) would not complete the steps of the cassette change. therefore, the decision made to exchange the aic and immediately thereafter the purge pressure alarm resolved. the patient continued on support. however, a bedside echocardiogram was completed and showed a shallow pump placement. the impella cp device was therefore positioned. the patient continued on impella mcs with noted plan to wean pressers and monitor status with the goal for possible impella cp explant the following day. later that night the patient had a decrease in urine output but remained greater than 30cc. the urine was pink in color; impella cp device was repositioned. overnight the patient became febrile with increasing white blood cells and still requiring moderate pressor support despite hyperdynamic left ventricular function. sepsis was suspected. plan to culture, empiric intravenous antibiotics started, hydrate, with attempt to wean pressors as tolerated. the decision was made to maintain the patient on impella mcs. the following day, the decision made to explant the impella cp device despite still on levophed due to adequate left ventricular function and suspected sepsis. the explant was completed with a survived outcome. however, that status of the patient post explant is unknown. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two devices associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-28286 for the automated impella controller report. instructions for use for the related event are as follows: impella cp with smartassist system. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9484 | 1220648-2025-28313 | 22003945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003945&pc=OZD | 04/21/2025 | 05/29/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654664 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support had a run of ventricular tachycardia (vt) and was shocked to get back into normal sinus rhythm. the following day, the patient was reported to continue to go into vt when moved. lidocaine was started for vt. it was further reported that a decision was made to withdrew care and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28313. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9485 | 1220648-2025-28314 | 22003982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22003982&pc=OZD | 04/20/2025 | 05/20/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025607718 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Malposition of Device |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the pump had on going suction alarms. support continued. furthermore, the pump also had high purge pressure. tissue plasma activator was added to the purge. | It was reported that a patient implanted with an impella 5.5 required repositioning to advance the pump into proper position. the pump had on-going suction alarms. additionally, the patient had a right internal jugular thrombus noted. the patient survived and remained on support. |  |  |  |  | D6b explant date was added. analysis summary "thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28314. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28314. h6 medical device problem codes 1491 and 1248 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28314. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9486 | 1220648-2025-28315 | 22004136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22004136&pc=OZD | 04/20/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025610314 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 107 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp had a low placement signal issue. the cp was repositioned two days later. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9487 | 1220648-2025-28316 | 22004328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22004328&pc=OZD | 04/20/2025 | 08/29/2025 |  | 05/12/2025 |  |  | 04/20/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025574898 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 102 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced cardiac arrhythmias and multiple suction alarms. the pump position was verified via echo, the p level was reduced, and multiple blood products were transfused. however, the suction alarms did not resolve. the decision was made to exchange the pump for another impella 5.5. | [5 empty] | The product was returned in damaged condition as the catheter was cut off. visual inspection found no issues on the pump. pump run test was unable to be conducted since the pump was damaged. | The investigation of vf/vt/af/at and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. data review: data logs confirmed the failure mode & clinical narrative. logs showed alarm 13 (impella stopped-motor current (mc) high) and 115 (impella stopped (rpm deviations)) were triggered after pump started but resolved quickly. after that, pump ran with fluctuated flows but still be in normal range and multiple suction alarms for 927 hours then low flow occurred and remained to the rest of the case. failure mode temporary pump stop added as failure mode based on data review. the root cause of temporary pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. entering cardiac output (impella cp with smart assist). section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28316. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9488 | 1220648-2025-28317 | 22004680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22004680&pc=OZD | 04/21/2025 | 06/06/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026653458 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 developed bleeding from the impella access site. an additional suture was added to control a hematoma, and two units of packed red blood cells and one unit of platelets were transfused. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of access site bleeding is determined to be use issue because hematoma was controlled was and the bleeding was resolved by adding an additional suture at the site
9489 | 1220648-2025-28318 | 22004753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22004753&pc=OZD | 04/21/2025 | 06/10/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  | 07/31/2025 |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374152 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump supporting a patient. while on support, the patient experienced a ventricular tachycardia storm which required the impella 5.5 to be repositioned and slightly pulled back. ectopy resolved after repositioning and the patient remained stable. | [5 empty] | Positioning issue: clinical states that the pump had to be repositioned due to vt storm over the weekend. pump was not returned. clinical information provided is very limited to investigate the repositioning issue. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. vt/ vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9490 | 1220648-2025-28319 | 22004889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22004889&pc=OZD | 04/21/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026625095 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male | 94 KG |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while a patient was implanted with an impella cp they experienced suction alarms and volume was administered. the pump was explanted due to poor limb perfusion. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9491 | 1220648-2025-28320 | 22004931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22004931&pc=OZD | 04/21/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1708923 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) failed to detect in inserted purge disc during preparation. the purge cassette was replaced. however, the issue persisted. consequently, the aic was exchanged to successfully resolve the issue. there was no patient harm are a result of the event. | [5 empty] | The investigation has been completed. the automated impella controller and logs were received for evaluation/analysis and noted the following: data logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically is caused by fluid ingress into the purge pressure sensor assembly.
9492 | 1220648-2025-28321 | 22005106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005106&pc=OZD | 04/20/2025 | 05/23/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629128 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 75-year-old female patient was on impella cp support and during support, there was access site bleeding. site was redressed; better angle match and forward tension was applied. heparin was withheld due to the bleeding and femstop was applied. additionally, the patient had a sustained run of atrial fibrillation (af). the patient was treated with metoprolol. support continued. additionally, the patient developed hemolysis on support. the hemolysis led to organ damage and continuous renal replacement therapy (crrt) was started. blood products were given. care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding, hemolysis, vt/vf/af/at, and renal failure has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major was most likely due to patient movement as evidenced by clinical detail of hematoma developing after moving patient leg the root cause of the hemolysis, renal failure, and af was not determined as no product or data logs were returned for analysis and limited clinical information related to failure was provided. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9493 | 1220648-2025-28322 | 22005143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005143&pc=OZD | 04/21/2025 | 06/12/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637918 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted, pre-operative placement, with an impella 5.5 device for mechanical circulatory support (mcs) and experienced low pump flow and thrombocytopenia requiring intervention. the patient was a workup for a coronary artery bypass graft and the decision was made for pre-operative place of an impella 5.5 device due to multi- vessel disease, mitral valve and tricuspid valves replacement. while inducing, the patient coded for two minutes. transesophageal echocardiogram showed biventricular failure. the impella 5.5 device was placed and a transesophageal echocardiogram showed 4.8 centimeters secured at 49 centimeters and re-evaluated right ventricle showed right ventricle failure. the impella 5.5 device was unable to flow above p-6 without suction alarms. pulmonary artery pulsatility index was 0.5 and a decision was made to implant an impella rp flex, which was successfully completed, for bipella mcs. after the implant procedure, the patient was transported to the unit on support. overnight, one dose of albumin was given for suction events. the impella rp flex support level was decreased to p-7 and noted with surgery planned. three days later, the patient received a platelet infusion. the following day, there was a slight increase in platelet count post infusion. the patient remained off heparin due to thrombocytopenia and hemoptysis. abiomed clinical support discussed with the nurse the importance of therapeutic activated clotting time while the impella 5.5 and impella rp flex are in use. bipella mcs continued. a couple of days later, the protected percutaneous coronary intervention was completed and the impellas were weaned and successfully explanted approximately seven days later. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | Thrombocytopenia: the root cause of the thrombocytopenia issue was not determined since insufficient clinical information was provided. low pump flow: clinical states that on (b)(6) 2025, patient had suction alarms and albumin was given for troubleshooting. no heparin in purge due to thrombocytopenia. pump was not returned for investigation. data logs were investigated and it was observed that the pump had suction alarms and lower flows at p7. no motor current spikes outside p-level changes observed. placement signal is seen drifting slightly down a little. the suction alarms resolve after the clinical mention of albumin given. flows stated to be slightly low at p-6 on may 5, however data logs show pump at p-4 in this time frame with flows in range. the root cause of the low pump flow was most likely patient condition related as the suction alarms resolved with volume.
9494 | 1220648-2025-28323 | 22005274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005274&pc=OZD | 04/21/2025 | 06/09/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653687 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after impella 5.5 was inserted on a 58 year old female patient for support for cardiogenic shock, the repositioning placement signal was lost and eventually pump stopped. the impella 5.5 was exchanged with a new one. there was no patient harm reported. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | Optical signal issue: clinical details report that towards the end of the case, the pump was being repositioned and the placement signal went out. the pump was ultimately replaced, due - at least in part - to the issue. data logs were not available for review. returned product (pump sn (b)(6)) was investigated and there were no visual abnormalities. all 5s parameters were within specification and the pump passed laser continuity. the pump was connected to a console and run in a bucket, and the psnr alarm triggered. imc logs of the lab console were reviewed, and it showed that calibrated g0 = 0. this does not align with the clinical narrative provided, because this is an issue that only should occur during initial connection/calibration of the pump in the field, which would result in a psnr alarm within minutes of case start. the field rep was contacted regarding the timeline of events, and it is unclear if the returned pump is the pump that experienced the complication or not. the rep also confirmed that the report was correct, and psnr triggered during repositioning late in the case and there was no point where the pump became unplugged. confounding clinical details were provided when compared to the state of returned product, and field data logs were not available for analysis to compare to the provided timeline. for these reasons, the root cause of the optical signal issue could no be determined. pump stop: clinical details report that shortly after the placement signal went out, the pump stopped. the pump was replaced due to the complication. data logs were not available for investigation. returned product was investigated and there were no visual abnormalities. the pump passed electrical testing. it was connected to a console and run in a bucket without reproducing a pump stop. since limited clinical details were provided, data logs were not available, and the failure did not reproduce, the root cause of the pump stop could not be determined.
9495 | 1220648-2025-28324 | 22005323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005323&pc=OZD | 04/22/2025 | 05/19/2025 |  | 05/12/2025 |  |  | 04/22/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639744 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was initially implanted with an impella cp device (pump 1) for mechanical circulatory support and would experience pump failure due to an accident with the automated impella controller in which another impella cp device pump 2 (for this report) was implanted for continuation in therapy. the patient would eventually expire. the impella cp pump 1 was successfully placed, and percutaneous coronary intervention was completed via a drug-eluting stent to the left anterior descending artery. that evening an echocardiogram was completed with pump at "good" position, 3.3cm from aortic valve. the next morning discussion was made with cardiologist and cardiovascular and thoracic surgeon, and the plan was noted to escalate to 5.5. the request was made and accepted to transfer the patient to another hospital later in the day. the transfer was completed by flight crew. the aic was attached to side of stretcher and when entering the elevator, the blue portion of the impella plug hit the elevator door and broke off. impella cp pump 1 failure occurred; support level diminished to p-0. the patient was in critical condition. consequently, the aic was exchanged and the impella cp device pump 1 was explanted and replaced with another impella cp device pump 2 for continuation of therapy. on this replacement impella cp device pump 2, the patient would develop limb ischemia. there was no distal flow. the underwent a fem-fem bypass and was transferred to the unit to recover and plans for the upgrade/escalation in care. however, the patient would expire three days thereafter. care was withdrawn. | It was reported that a patient was implanted with an impella cp and had a limb that became ischemic. a femorofemoral bypass was placed. after 68 hours, the patient had care withdrawn and the patient expired. |  |  |  |  | Medical safety reviewed the event and noted there is not enough information to exclude the automated impella controller (aic) as an associated factor to the outcome (patient's demise). the pump stop/failure for the impella cp pump 1 is a serious injury as well as impella cp pump 2 (replacement pump) where the patient developed limb ischemia. limb ischemia can lead to more complications down the line, but the immediate issue is the aic pump failure (and this is what led to or contributed more to the demise outcome). section h6 evaluation codes (health impact/medical device problem) were updated accordingly based on medical safety review. the impella cp pump 2 is one of three devices associated with the event. note the following: -refer to manufacturing report number 1220648-2025-28546 for the aic report. -refer to initial medical device report for the impella cp pump 1 serial number (b)(6) with date of event 22-apr-2025 and patient identifier ending in (b)(6). update was made to section b5 event description to include relevant details surrounding the patient's implant experience. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9496 | 1220648-2025-28325 | 22005407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005407&pc=OZD | 04/21/2025 | 09/24/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613030 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the placement signal was significantly lower than arterial line. | The user facility reported the 63 year old male patient was on impella 5.5 support and during support, the placement signal was significantly lower than arterial line. the flows and motor current with in normal limits. additionally, the pump was up against the wall and the papillary. the pump will need to pulled back. the pump was weaned and explanted. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. positioning issues: on (b)(6) 2025, pump noted to be measuring deep and near papillary. pump was successfully repositioned under tte. data log review showed several "impella position unknown" alarms on this day. the root cause of the positioning issue was not determined since insufficient clinical details were provided. optical signal issue: on (b)(6) 2025 after the repo unit was placed, the placement signal was noted to be significantly lower than arterial line. arterial line values are unknown. flows and mc were noted to be within normal limits at this time. patient was being weaned off cpb after impella implant. on (b)(6) 2025, it was noted again that the placement signal was unreliable with no other details. on (b)(6) 2025, a wide variation in placement signal and patient hemodynamics was noted. echo was also not consistent with ps and lv waveforms. patient could not tolerate flows higher than p-5 during the case and there was noted rv dysfunction. the pump was not removed due to the optical signal issue. data log review showed only "placement signal low" (ps) alarms on (b)(6) 2025 and (b)(6) 2025, no psnr alarms. within the first hour of starting the pump, the ao ps read between ~20 and 60 mmhg. there were no related issues observed with the 5s optical parameters nor motor current spikes indicating biomaterial ingestion during the case. there was also a 20 mmhg lv offset applied during the case. the product was returned and evaluated. the outlet cage was filled with biomaterial, which covered the optical sensor. the 5s parameters were within spec before and after removing the biomaterial and the pump passed laser continuity testing. when plugged into the lab console, the ao ps read -20/-20. the pump was pressurized with the uson and the ps gradually drifted upward and downward in no specific pattern while at a stable pressure. the root cause of the optical signal issue was not determined since insufficient clinical details of the exact issue were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | Positioning issues. on (b)(6) 2025, pump noted to be measuring deep and near papillary. pump was successfully repositioned under tte. data log review showed several "impella position unknown" alarms on this day. the root cause of the positioning issue was not determined since insufficient clinical details were provided. optical signal issue: on (b)(6) 2025, after the repo unit was placed, the placement signal was noted to be significantly lower than arterial line. arterial line values are unknown. flows and mc were noted to be within normal limits at this time. patient was being weaned off cpb after impella implant. on (b)(6) 2025, it was noted again that the placement signal was unreliable with no other details. on (b)(6) 2025, a wide variation in placement signal and patient hemodynamics was noted. echo was also not consistent with ps and lv waveforms. patient could not tolerate flows higher than p-5 during the case and there was noted rv dysfunction. the pump was not removed due to the optical signal issue. data log review showed only "placement signal low" alarms on (b)(6) 2025 and (b)(6) 2025, no psnr alarms. within the first hour of starting the pump, the ao ps read between ~20 and 60 mmhg. there were no related issues observed with the 5s optical parameters nor motor current spikes indicating biomaterial ingestion during the case. there was also a 20 mmhg lv offset applied during the case. the product was returned and evaluated. the outlet cage was filled with biomaterial, which covered the optical sensor. the 5s parameters were within spec before and after removing the biomaterial and the pump passed laser continuity testing. when plugged into the lab console, the ao ps read -20/-20. the pump was pressurized with the uson and the ps gradually drifted upward and downward in no specific pattern while at a stable pressure. the root cause of the optical signal issue was not determined since insufficient clinical details of the exact issue were provided. | The impella device was received. upon completion of the investigation, a supplemental mdr will be filed.
9497 | 1220648-2025-28326 | 22005517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005517&pc=OZD | 04/21/2025 | 05/28/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637952 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was implanted in a patient being assessed for transplant. the patient experienced a stroke. the patient survived and later the team observed a pump stop that was able to be restarted. the pump remains on for support. the associated automatic impella controller (aic) was replaced. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | D6: explant date unk. the investigation of pump stop and stroke/cva has been completed. the impella device was not returned for evaluation. data logs were returned for review. there was no problem found during the case with pump relative to temporary pump stop event. the device functioned/operated to specification. please refer manufacturer device report number 1220648-2025-28328 for an investigation on console. there was no problem found during the case with pump relative to temporary pump stop event. the device functioned/operated to specification. as clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿
9498 | 1220648-2025-28327 | 22005660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005660&pc=OZD | 04/21/2025 | 09/24/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551260 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that after impella cp was placed on a 73-year-old female patient, a small hematoma was noted on the access site. the site was redressed with angle matching. around an hour later, it was noted that hematoma had grown. two units of blood products were administered. the cardiologist ordered manual pressure, and bleeding was resolved the next day. it was further reported that during the night the impella in aorta alarm was present after the patient was coughing. echo was done and impella was in an appropriate position. then in the morning, when impella was being weaned in preparation for explant, once it was decreased to p3, began having multiple position in aorta alarms that would intermittently resolve and then reappear. since the impella was going to be explanted no repeat echo was done. the pump was weaned and explanted. | The complainant reported that the device exhibited an optical sensor issue. |  |  |  |  | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. clinical mentions small hematoma observed before access obtained but after access obtained, it was inappropriately dressed without angle matching. hematoma grew and hemostasis achieved with manual pressure. the root cause of the access site bleeding - major was most likely due to a use issue as evidenced by improper angle matching stated in clinical details. based on the investigation, the failure mode 'positioning issue' was changed to optical signal issue as complaint described was in relation to false 'impella in aorta' alarms which does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally. based on the investigation findings b1 was revised to adverse event only. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. based on the investigation, the failure mode 'positioning issue' was changed to optical signal issue as complaint described was in relation to false 'impella in aorta' alarms access site bleeding - major: clinical mentions small hematoma observed before access obtained but after access obtained, it was inappropriately dressed without angle matching. hematoma grew and hemostasis achieved with manual pressure. the root cause of the access site bleeding - major was most likely due to a use issue as evidenced by improper angle matching stated in clinical details optical signal issue: no product was returned for analysis. the data logs from the case confirm 'impella position in aorta' alarms when pump turned down to p3. the 5s parameters were stable. there were no motor current spikes or other log abnormalities. clinical mentioned alarm would intermittently resolve with no changes and that position was confirmed via echo. the root cause of the optical signal issue was most likely patient condition related as evidenced by no log abnormalities and confirmed position via clinical imaging. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
9499 | 1220648-2025-28328 | 22005753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005753&pc=OZD | 04/21/2025 | 06/12/2025 |  | 05/12/2025 |  |  | 04/21/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1883673 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was implanted in a patient being assessed for transplant. the patient experienced a stroke. the patient survived and later the team observed a pump stop that was able to be restarted. the pump remains on for support. the associated automatic impella controller (aic) was replaced. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. the impella 5.5 associated with this automated impella controller was reported in manufacturer device report number 1220648-2025-28326. | Data analysis: on 3/27 5.5 pump 524733 was connected to (b)(6) and started running at 16:02. 77 minutes later, there is the log entry "pmd ok staterec818 0 2000 257 0 0 0 0xa313 0xa313 cpld 3" directly preceding sau_motor communication reinitializing and an event #115 pump stop ((b)(6) 524733 25_03_27 pump stop 1.png). there was another pump stop on 3/31 with similar log entries but no reinitialization of sau_motor communication ((b)(6) 524733 25_03_31 pump stop 2.png). the pump was disconnected and reconnected on 4/2 by the user for no discernible reason in the logs. on 4/6, there was another pump stop that occurred twice in quick succession ((b)(6) 524733 25_04_06 pump stop 3.png). the pump was disconnected on 4/7. on 4/15, a new 5.5 pump, 581850, was connected to (b)(6). on 4/20 there is another pump stop ((b)(6) 581850 25_04_20 pump stop 4.png) this time with a new pump. there were two more back to back pump stops on 4/21 at 8:57 and then again at 9:16 ((b)(6) 581850 25_04_21 pump stop 5.png). the pump was disconnected on 4/22. it was also noticed for every pump stop, that ac power was disconnected. the rt logs confirm all of the pump stops but do not edify the cause. time stamp pump ac power pmd staterec alarms pmd reset after sau motor entries sau motor reset (b)(6) 2025 17:19:15, 524733. yes ok. 115. yes. yes. yes. (b)(6) 2025 17:19:43, 524733. yes ok. 115+18. no. no no. (b)(6) 2025 13:09:06, 524733. yes ok. 115. yes. no. no. (b)(6) 2025 14:35:03, 524733. yes ok. 115. yes. no. no. (b)(6) 2025 14:35:06, 524733. yes ok. 115. no. yes. no. (b)(6) 2025 14:37:00, 524733. yes ok. 115. yes. yes. no. (b)(6) 2025 08:38:50, 581850. yes ok. 115. yes. yes. no. (b)(6) 2025 08:57:04, 581850. yes ok. 115. yes. yes. no. (b)(6) 2025 09:15:58, 581850. yes ok. 115. yes. yes. yes. device analysis: the console was run overnight twice with a pump on p-8 (all pump stops were with the pump on p-8) but the failure mode was not reproduced. additionally a pump was run for multiple hours on battery power but still the failure mode did not reproduce. root cause: the cause of the multiple pump stops across multiple pumps was not determined since the failure mode was not reproduced. root cause: as a precaution, please replace the impellatronic board (0042-1115) and the pbm (0042-1125), plus connectors 0042-2709 and 0042-2711. then run a test pump on battery power for 30 minutes followed by sections 13-15 of the pm (0042-7309) and functional check (0042-7316).
9500 | 1220648-2025-28329 | 22005991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22005991&pc=OZD | 04/23/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/23/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported on a patient was on impella 5.5 support. during use, they stated that they were changing the cassette. however, they were unable to complete the procedure due to purge disc not detected alarm. they tried two cassettes before calling and could not complete the cassette change. all connections were secure and the blue purge disc securement appeared intact. they switched to a new console with resolution. | [5 empty] | The investigation for the aic (console) purge disc/flag issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was returned and examination found glucose residue observed around the purge pressure sensor during inspection of the console. the root cause of the purge disc not detected alarm was likely the contamination of the purge pressure sensor.
9501 | 1220648-2025-28330 | 22006074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22006074&pc=OZD | 04/22/2025 | 05/12/2025 |  | 05/12/2025 |  |  | 04/22/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Asystole | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced ventricular fibrillation and asystole events. the patient was escalated to an impella 5.5 the following day. the patients outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9502 | 1220648-2025-28331 | 22006386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22006386&pc=OZD | 04/22/2025 | 06/06/2025 |  | 05/12/2025 |  |  | 04/22/2025 | 05/12/2025 |  |  |  |  | 05/12/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653840 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon implanting an impella 5.5, the guidewire was kinking at the tip and it could not be advanced. a steel-core 0.18 wire was used to successfully implant the pump. additionally, the patient was in and out of atrial fibrillation while on support. support was continued and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The product was not returned for investigation. according to the clinical details, a kink was noted on the guidewire, and the impella was unable to advance on it. the doctor used a new steelcore 0.18 wire, and the advancement was successful. clinical images were not provided for the damaged guidewire. the cause of the guidewire damage issue was not determined since product was not returned. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
9503 | 1220648-2025-28240 | 21992663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21992663&pc=OZD | 04/17/2025 | 05/23/2025 |  | 05/09/2025 |  |  | 04/17/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635316 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced positioning issues and the pump needed to be pulled back to optimize flows. the medical doctor is very knowledgeable with the impella and how to operate it during surgical cases and did not feel it necessary to pull back fully during coronary artery bypass graft x 3 and aortic valve replacement/mitral valve replacement with impella in place. an injury to the left ventricle was discovered and was suspected to have occurred when physically manipulating the heart during surgery. the pump was explanted and venoarterial extracorporeal oxygenation central cannulation was inserted. the patients outcome was stable. | [5 empty] | The investigation of positioning issues and vascular damage - great vessel has been completed. the impella device was not returned for evaluation. positioning issues: clinical mention pump was pulled back in the icu to optimize flows. the root cause of the positioning issues was not determined as no product was returned for analysis and limited clinical information to how pump became mispositioned was provided. vascular damage - great vessel: the root cause of the vascular damage - great vessel was most likely use related as decision to leave the pump in the lv resulting in a perforation indicates a use-related issue indicating user did not have sufficient space needed and unintended manipulation resulted in the injury. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9504 | 1220648-2025-28241 | 21992787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21992787&pc=PYX | 04/18/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  | 12/31/2025 |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025438949 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was in for a right ventricular assist device. however, an alternative decision was made to implant an impella rp flex. it was noted that during insertion due to patient's right ventricle size, it was difficult to remove unnecessary length in right ventricle. however, repositioning was required after the physicians completed tightening of the purse string suture as the device was advanced forward with inflow in the right ventricle and appeared to cause arrhythmias/impella rp flex suction alarms. arrhythmias and alarms subsided after the device was retracted. catheter fixated on the patient with stat locks and sterile dressing were placed on the impella rp flex insertion site. during support, it was noted that the patient was having sustained ventricular tachycardia (vt). vt ablation was planned. however, aborted due to instability. the impella rp flex was successfully weaned and explanted with patient outcome of no harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
9505 | 1220648-2025-28242 | 21992854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21992854&pc=OZD | 04/18/2025 | 06/16/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631946 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 pump was supporting a patient admitted in and awaiting multi-organ transplant. during the month-long support, there were positional alarms and the pump was being shallow. the pump was repositioned. the team also noted suction during the support. it was later reported that care would be withdrawn due to a hemorrhagic cerebrovascular accident. once care was withdrawn, the patient expired. | [5 empty] | The investigation of positioning issues, low pump flow, and stroke/cva has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issues was most likely patient condition related to the patient's small lv. low pump flow: clinical reports suction alarms above p-6, patient had small lv and unable to repositioning and optimize position. pump left shallow and suction reported with no resolve. data log reviewed: no motor current (mc) spike observed outside of p-level changes. few resolved position unknown alarms. many suctions alarms occurred during the case. the cause of the low flow was most likely improper positioning of the device as suction alarms were reported and the pump was stated to be shallow. stroke/cva: as necessary clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9506 | 1220648-2025-28243 | 21992931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21992931&pc=OZD | 04/19/2025 | 05/22/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204293 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) had a console error alarm that was triggered and disappeared multiple times, with loss of placement signal. no kinks of the cable and shaft were seen. the aic was exchanged and patient outcome was no harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | D6: explant date unknown. the investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the root cause of the low signal not reliable (snr) was due to optical bench malfunction (this is an additional analysis finding that is unreported and unrelated to the reported failure mode, but is related to the clinical procedure).
9507 | 1220648-2025-28244 | 21992964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21992964&pc=OZD | 04/21/2025 | 07/11/2025 |  | 05/09/2025 |  |  | 05/02/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025612130 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | Muscle/Tendon Damage | Hemodynamic instability | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after insertion of the impella cp, a new papillary muscle rupture occurred. although there was originally a rupture of the posterior wall papillary muscle due to myocardial infarction (mi), the papillary muscle of normal myocardium not related to mi was also ruptured. this led to a hemodynamic collapse and emergency mitral valve regurgitation. the pump was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. e1 initial reporter name unknown. | The investigation of the reported failure mode has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 event date was erroneously entered on the initial manufacturer device report number 1220648-2025-28244. h6 health effect: clinical code 4532 was erroneously entered on the initial manufacturer device report number 1220648-2025-28244. h6 health effect: clinical code 4451, 4868, and health effect: impact code 4624 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28244.
9508 | 1220648-2025-28245 | 21992987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21992987&pc=OZD | 04/17/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/17/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627419 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced limb ischemia, hemolysis, and renal failure requiring intervention. the impella cp device was placed without issues and the surgeon turned the patient down for surgery. therefore, the physician proceeded with protected percutaneous coronary intervention (pci) of left main-left anterior descending artery and left circumflex stenting. the patient's pulse pressure went away during the left main-left anterior descending artery intervention and resolved after stenting was complete. the decision was made to leave impella in over the weekend and let patient rest. the patient was transferred to the unit on impella mcs. same day, post implant, the patient was transferred back to the catheterization lab for fem-fem external bypass. the patient as noted as stable and support continued. the next day, mcs was planned for the weekend, and the physician discussed recommendation to withdraw impella 0.5cm due to noted hemolysis. over the next couple of days, the patient was noted as doing well. distal pulses were found with doppler bilaterally. groin site was noted as stable. however, it was noted the patient was transferred to an outside facility in which renal failure was noted and possible start of continuous renal replacement therapy. pulse remained confirmed via doppler. groin site clean dry and intact. update noted hemolysis was still present and the impella mcs support level was turn down from p-9 to p-6. additional information noted crrt was initiated due to lack of urine output which could have been due to impella and hemolysis. a couple of days thereafter, the patient was successfully weaned, and the device was explanted with a survived outcome. using double wire technique, the physician perclose successfully and the angiogram of the right leg was taken with no noted issues. it was further noted that the patient remained on crrt after the impella was removed. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9509 | 1220648-2025-28246 | 21993028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993028&pc=OZD | 04/19/2025 | 06/06/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649555 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after a patient on impella cp support arrived in intensive care unit, bleeding profusely occurred from the endotracheal tube and pressure dropped. suction alarms started and the impella cp was turned down to p2. over two liters of blood was suctioned out of the patient's lungs. the patient was taken emergently to operating room (or) for extracorporeal membrane oxygenation (ecmo) cannulation. the patient went into ventricular fibrillation (vf) arrest in or and cardiopulmonary resuscitation was done. the ecmo was placed and the patient was stabilized. the impella cp was placed on p3 to vent the left ventricle. the patient continued to bleed heavily from the lungs and received multiple blood products. the patient was transferred to another facility and a left groin hematoma was noted at the impella cp site. the patient received massive transfusion protocol and extensive or intervention for the pulmonary hemorrhage. anticoagulation was held. it was further reported that on day two of impella support, the impella cp was noted to be shallow and was advanced at bedside. the patient was eventually successfully weaned and explanted. the patients outcome was stable. | [5 empty] | The investigation of vascular damage - peripheral vessel, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9510 | 1220648-2025-28247 | 21993078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993078&pc=OZD | 04/18/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558794 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 62 year old male patient who had a cp pump placement for support due to myocardial infarction. the patient had hemolysis observed after insertion of the foley catheter. after over 16 hours of support the patient was made a do not resuscitation (dnr) and expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9511 | 1220648-2025-28248 | 21993109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993109&pc=OZD | 04/20/2025 | 07/02/2025 |  | 05/09/2025 |  |  | 04/20/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644610 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was unable to be implanted due to the patient¿s vascular anatomy. several attempts without success were made. the physician decided to complete the percutaneous coronary intervention procedure without impella cp support. the impella cp placement was aborted and the patient's outcome was no harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella was not able to overcome the stenosis.
9512 | 1220648-2025-28249 | 21993173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993173&pc=OZD | 04/18/2025 | 06/11/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  | 07/31/2025 |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390821 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62 year old male patient placed on impella cp support due to st-elevated mi. the impella was prepped and inserted without incident. flows started on auto achieving 3.4 lpm. pci successfully completed to rca. peel away sheath was removed and reposition sheath advanced. echo done to verify position. the impella position measured at 1.93 cm from aortic annulus, however when the impella was advanced it went into lv. under fluoroscopy the physician pulled back. it was reported that the impella jumped across valve. the impella was removed and a new pump was placed. the patient survived the procedure. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely use issue since under fluoroscopy the physician pulled back and impella jumped across valve. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
9513 | 1220648-2025-28250 | 21993332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993332&pc=OZD | 04/18/2025 | 05/23/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643359 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the second perclose failed following impella explant. angioseal was not effective in manage the resulting bleed, at which point vascular was consulted. vascular imaging noted a dissection at the bifurcation. the patient was subsequently taken to the operating room for an open procedure. the patient remained stable and was explanted in the procedural area. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of vascular damage peripheral vessel was most likely use issue since the second perclose failed. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9514 | 1220648-2025-28251 | 21993358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993358&pc=OZD | 05/01/2025 | 07/11/2025 |  | 05/09/2025 |  |  | 05/02/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026623730 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old female with an impella 5.5 placed as escalation of therapy. it was reported that during ventricle septal perforation (vsp) patch closure, the impella 5.5 caused wall perforation during dislocation. the ascending aorta was short, the catheter was inserted deeper than 5cm when the ascending aorta was clamped. it is believed that the catheter dislocated from there and penetrated the left ventricle. the impella was removed, and the ventricle was repaired. the pump was explanted but the patient continued on ecmo support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation of ventricle perforation and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: it was stated that the patient's aorta was short and the device was inserted deeper than 5cm when the ascending aorta was blocked. it became dislocated during vsp patch closure. based on the clinical information provided, the root cause of the positioning issues is most likely due to use as the pump became dislocated during the patch closure. ventricle perforation: the device was inserted deeper than 5cm, and during the patch closure, the 5.5 became dislocated and it was believed to have perforated the left ventricle. based on the clinical information, the root cause of the ventricle perforation is most likely due to improper positioning of the device. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28251. h6 health effect - clinical code 2135 was erroneously entered on the initial manufacturer device report number 1220648-2025-28251. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-28251.
9515 | 1220648-2025-28252 | 21993484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993484&pc=OZD | 04/20/2025 | 06/11/2025 |  | 05/09/2025 |  |  | 04/20/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632526 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support, and during support, the pump was experiencing lower than expected flows. the pump was exchanged and biomaterial was observed. the patient survived. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9516 | 1220648-2025-28253 | 21993486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993486&pc=OZD | 04/18/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517644 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 65 year old female patient was on impella cp support and after the explant, the sheath was reinserted due to the perclose sutures failed. manual pressure was held and a femstop was applied. patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9517 | 1220648-2025-28254 | 21993721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993721&pc=OZD | 04/18/2025 | 06/26/2025 |  | 05/09/2025 |  |  | 04/18/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617983 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the during placement of an impella cp, there was a single access technique failure with the other manufacturer, arrow, sheath. per the documentation they "could not pass a sheath next to the impella, so second femoral access was used in the other leg. access complication was managed by two-way access twice in the femoral artery". no further intervention was needed. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a5- unknown. d6a/d6b is unknown. | The investigation of single access issue has been completed. the impella device was not returned for evaluation. the cause of the single access issue was most likely use-related as the use of the teleflex-arrow braided sheath for single-access is not recommended by abiomed. per cp training guide, sheaths up to 7 fr can be utilized for single-access procedures, but abiomed recommends use of the 7fr companion sheath. d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28254.
9518 | 1220648-2025-28255 | 21993744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993744&pc=OZD | 05/02/2025 | 07/18/2025 |  | 05/09/2025 |  |  | 05/02/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621994 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | Aortic Valve Insufficiency/ Regurgitation | Hemodynamic instability | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the patient expired on support. the cause of death was determined to be circulatory failure due to abdominal aortic injury and abdominal hemorrhage. however, it is unknown if the impella was related to the cause of death. | The complainant reported a patient was on impella cp and during the implant, the placement wire could not be removed and formed a loop in the descending aorta. the wire was removed pulling it from the radial artery using a snare catheter. however, the blood vessel was damaged by the 18-inch wire, and the patient experienced intra-abdominal bleeding, so a stent graft was placed. additionally, the patient had severe aortic regurgitation (ar) and hemodynamic collapse occurred. the pump was repositioned; however, this did not improve the ar. the patient's outcome is unknown. |  |  |  |  | B2: death and date of death added based on additional information received regarding the patient outcome. b5: updated with additional information regarding the patient outcome. h1: type of report revised based on additional information received regarding the patient outcome. h6: health effect - impact code 1802 death added. heart valve damage: after the introduction of impella, severe aortic regurgitation occurred and the patient's hemodynamics collapsed. the device was switched to va-ecmo and repositioned, but the patient's condition did not improve. cp explanted and ar was alleviated. the pump was not returned. the cause of the aortic regurgitation was not determined since product was not returned and insufficient clinical details. product damage (guidewire damage - great vessel): following initial placement of the pump, placement wire could not be removed and formed a loop in the descending aorta. the 18-inch wire was removed by pulling it from the radial artery using a snare catheter. however, the blood vessel was damaged by the 18-inch wire, and the patient experienced intraabdominal bleeding, so a stent graft was performed. no product was returned. the cause of the product damage (guidewire damage - great vessel) was not determined since product was not returned and insufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter. impella catheter inlet to the aorta, ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms, ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
9519 | 1220648-2025-28256 | 21993897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21993897&pc=OZD | 04/22/2025 | 07/10/2025 |  | 05/09/2025 |  |  | 04/28/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599625 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 12 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for the support of a pediatric patient. it was reported that upon explant of the pump it was seen to be kinked. the patient remained stable and no patient harm was noted. | [5 empty] | The investigation of product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the root cause of the product damage cannula kink issue was not determined since product was not returned for investigation and lack of clinical details. b3 date of event was erroneously entered on the initial manufacturer device report number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9520 | 1220648-2025-28257 | 21994021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994021&pc=OZD | 04/21/2025 | 06/17/2025 |  | 05/09/2025 |  |  | 04/21/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646474 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient developed bleeding from their impella cp access site during support. multiple measures were taken in an attempt to manage the condition, but the decision was ultimately made to escalate to impella 5.5 support. the patient's outcome is stable. | [5 empty] | Reported bleeding from access site on (b)(6) 2025. standard troubleshooting attempted with no relief from bleeding, impella cp was explanted and escalated to impella 5.5. impella cp was returned and no damages or abnormalities were found. the root cause of the access site bleeding was not determined due to insufficient clinical details. | G3 date was inadvertently omitted on manufacturer device report number 1220648-2025-28257-1. h2 device evaluation selection was inadvertently omitted on manufacturer device report number 1220648-2025-28257-1. h3 evaluation summary attached was erroneously selected on manufacturer device report number 1220648-2025-28257-1. h6 type of investigation 4112 and 4110 were inadvertently omitted on manufacturer device report number 1220648-2025-28257-1. h6 investigation findings 4248 was erroneously entered on manufacturer device report number 1220648-2025-28257-1. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
9521 | 1220648-2025-28258 | 21994030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994030&pc=OZD | 04/19/2025 | 11/07/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1731799 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | Unintended Application Program Shut Down |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during the impella case a controller error alarm was present. automated impella controller (aic) to aic transfer was performed. once the impella catheter was disconnected from the aic, the aic was not able to get turned off. a new aic was used and there was no patient harm. | The complainant also reported that during the impella case the placement signal was displayed, however, after less than one minute, the placement signal disappeared. |  |  |  |  | The investigation of automated impella controller (aic) - shutdown or restart issue and aic - controller error (yellow alarm) has been completed. console logs were consistent with what was reported in the complaint. the console was tested on an engineering lab bench. the reported issues were unable to be reproduced while running the console with a known good pump and purge cassette kit. there was no issue with the console hardware. the root cause of the reported issues was an aic software issue. the root cause of the inability to shutdown was an aic software issue. the "as reported" code associated with the unexpected reboot was not confirmed in the data or device analysis. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data review: console logs were consistent with what was reported in the complaint. the aic and rt logs align with a transient loss of optical data. controller error (opm comm stopped) alarm [event set #1043] was observed in the logs which confirm the reported loss of placement signal and controller error. the logs also reported processsamplingdatafromopticalbench: sentstopacquisitioncmd ack, indicating ossiopsens receive thread gets a data sampling packet with an unexpected length. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are lastly reported as +16384 values and then no more samples were recorded. device analysis: the console was tested on an engineering lab bench. the reported issues were unable to be reproduced while running the console with a known good pump and purge cassette kit. there was no issue with the console hardware. conclusion: the absence of snr samples impacted the console¿s ability to recognize the pump disconnection which resulted in the user not being able to shutdown the console. the root cause of the reported issues was an aic software issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal description. h6 updated.
9522 | 1220648-2025-28259 | 21994145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994145&pc=OZD | 04/19/2025 | 06/06/2025 |  | 05/09/2025 |  |  | 04/21/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632715 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was admitted in with acute myocardial infarction. the case was done independent of any impella support on site. the team observed the sheath to kink at insertion and the pump could not be delivered. upon removal of the impella cp, the pump separated. the entire unit was able to be removed and the site was closed. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of introducer damage ¿ mechanical, delivery issues, and product damage (detached inflow/outflow peripheral vessel) has been completed. no product returned for analysis. delivery issues: the root cause of the delivery issue was most likely patient condition related as evidenced by the clinical detail of patient having tortuous iliac artery. introducer damage - mechanical: no product or images were returned for analysis. the root cause of the introducer damage - mechanical was not determined as no product or images were returned for analysis product damage (detached inflow/outflow - peripheral vessel): no product or images were returned for analysis. the root cause of the product damage (detached inflow/outflow - peripheral vessel) was not determined as no product or images were returned for analysis.
9523 | 1220648-2025-28260 | 21994286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994286&pc=OZD | 12/22/2024 | 05/09/2025 |  | 05/09/2025 |  |  | 12/22/2024 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580212 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella 5.5 support had ventricular tachycardia with deep suctioning and turning. later loqi database reported upon an acute thrombotic stroke with relationship to impella listed as "possibly" related. loqi states: patient had impella 5.5 insertion for cardiogenic shock. patient began to experience left sided weakness and stroke code activated. a computed tomography (ct) scan revealed an acute right middle cerebral artery bifurcation occlusion. the patient underwent emergent cerebral thrombectomy. ct showed hemorrhagic conversion (subarachnoid hemorrhage). it was noted that the ct was stable, resolving small bleed. persistent left sided hemiplegia through discharge. the patient remained stable and survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9524 | 1220648-2025-28261 | 21994325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994325&pc=OZD | 04/22/2025 | 07/07/2025 |  | 05/09/2025 |  |  | 04/22/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139057 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) error appeared during impella rp flex support to a patient. the staff was walked through an aic-to-aic transfer in response to the failure. there was no harm to the patient as a result of the event. | [5 empty] | Data analysis: the case associated with the reported complaint was initiated on april 20, 2025, with pump 521361. during the support on the complaint date, april 22, 2025, the console displayed ¿controller error (loss of comm (pbm1)) ¿ alarm,¿ event #1023, and the log recorded a continuous powermod_1 communication issue. this confirms the reported failure mode. device analysis: this console was tested after arriving at fs. visual inspection confirms corrosion on the pbm. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). note: failures of this type will be monitored for trends. (b)(4) was opened in response and was escalated to (b)(4). (b)(4) was opened in response. root cause: the root cause of the controller error was pbm corrosion due to the leakage of electrolyte from the battery failure alarm supercapacitors. ad hoc task: before returning this console (ic2572) back to the customer, fs was instructed to perform the following, replace the front panel optical connector (0042-2750) due to debris. replace the pbm (0042-1125) due to the corrosion. perform sections 18, 19, 20 and 23 of the pm procedure (0042-7309). perform a full functional check (0042-7316). return the original pbm back to the engineering for further analysis. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9525 | 1220648-2025-28262 | 21994468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994468&pc=OZD | 04/19/2025 | 06/16/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613577 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  | MW5169498 | Yes | Initial submission | Followup | Followup | [7 empty] | The complainant had placed a impella 5.5 pump for a patient needing mechanical circulatory support. the impella 5.5 was placed and started. when device was about to get secured with 3 point external fixation, the nurse noticed blood coming out of the red plug. the care team decided to replace 5.5 due to concern for pump integrity. there was no patient harm. | [5 empty] | Medwatch mw5169498 was received and the following fields were updated on manufacturer device report number 1220648-2025-28262. e4 revised to yes. h9 revised to include medwatchnumber. | The investigation of product damage (hole in catheter) has been completed. product damage (hole in catheter): when device was about to get secured w/ 3 point external fixation, rn noticed blood coming out of the red plug. the device was returned and inspected. visual inspection revealed blood in the red pump handle and a hole in the catheter near the 30cm mark the cause of the product damage (hole in catheter) was most likely use related due to a hole found in the catheter and blood in the red handle. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9526 | 1220648-2025-28263 | 21994511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994511&pc=PYX | 04/22/2025 | 06/16/2025 |  | 05/09/2025 |  |  | 04/22/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025567920 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Embolism/Embolus | Insufficient Information | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella rp flex pump experienced a temporary stop. upon restart, the pump performed at lower flows. it was recommended to replace the pump, however, the doctor continued support. care was withdrawn and the patient expired. | [5 empty] | The investigation of temporary pump stop and low pump flow has been completed. the impella device was not returned for evaluation. temporary pump stop: data log review showed some motor current spikes in the time leading up to the pump stop. after the pump was restarted, there were lower pump flows and motor current disturbances. purge pressure also began to rise a few minutes before the pump stop. the root cause of the temporary pump stop was most likely biomaterial ingestion since there were motor current spikes leading up to the pump stop with motor speed deviations and reduced flows afterward. low pump flow: as mentioned above, data log review showed the pump had lower flows after being restarted. there were motor current spikes leading up to the pump stop and a slight rise in purge pressure. purge fluid was changed from bicarb to heparin a few hours after the pump stop. the low pump flow abruptly resolved about 6 hours after the pump stop. the root cause of the low pump flow was most likely biomaterial ingestion since there were motor current spikes and a pump stop prior to the drop in flows. thrombus: failure mode added based on investigation results. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9527 | 1220648-2025-28264 | 21994674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994674&pc=OZD | 04/19/2025 | 09/23/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635350 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 118 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that, while the nurse was rounding on the patient, the impella began to alarm ¿impella position in aorta.¿ the patient was already scheduled to go to the operating room in the next hour to replace the existing impella. since the current impella was going to be removed, the team decided against repositioning. the removal and replacement were successful. | The complainant reported that a patient was escalate to impella 5.5 for advanced therapies evaluation. the 5.5 was successfully placed. during support it was noted of suspected hemolysis, no recent alarms, however decision was made to replace existing 5.5 with a new 5.5. it was further reported that a clot was visible on impella motor housing. the procedural outcome was the patient survived surgery. |  |  |  |  | B1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28264. a4 patient weight updated. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | There was no product returned. it was stated that on the 19th, there was suspected hemolysis and the pump was confirmed to be in the correct position. additionally, there was "impella position in aorta" alarms but repositioning was not attempted since the impella was going to be replaced with another 5.5 in an hour. the data logs show that during the time of the reported hemolysis and "impella position in aorta," alarms, there was a spike in the motor current followed by a change in placement signal and flows. there were "impella position in aorta," alarms as well as suction. the root cause of the hemolysis is most likely due to ingested biomaterial. the root cause of the optical signal issue is most likely due to ingested biomaterial. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis". "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure."
9528 | 1220648-2025-28265 | 21994696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994696&pc=OZD | 04/22/2025 | 05/20/2025 |  | 05/09/2025 |  |  | 04/23/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616533 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on impella 5.5 for hemodynamic support. during support, it was noted that placement signal low alarm routinely alarming due to positioning. an echocardiogram was performed multiple times to confirm the pump was in good position. once the patient was repositioned in bed, the alarms resolved. again, intermittent placement signal low alarms occurred when the patient turned on their right side. the care team was aware that the pump was positioned deep and was considering repositioning the pump. the patient survived. | [5 empty] | D6: explant date unknown. the investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 medical device problem code 1559 was erroneously entered on the initial manufacturer device report number 1220648-2025-28265. new medical device problem code entered. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9529 | 1220648-2025-28266 | 21994971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21994971&pc=OZD | 04/23/2025 | 06/06/2025 |  | 05/09/2025 |  |  | 04/23/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478138 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during preventive maintenance, the automated impella controller experienced a console error on a restart. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the console logs from the reported day of the event align with the complaint received. it was noted that the issue associated with the complaint began on (b)(6) 2025. the console triggered a controller error (defective optical sensor lamp ¿ event set 1039) alarm immediately following the bootup. notably, the same error persisted even after rebooting the console twice. device analysis: the console ic8089 was returned for further investigation. while testing the console "controller error (optical lamp defective)" alarm was reproduced. voltage measurements taken from the power battery manager (pbm) to the optical bench were within the required specifications, and no loose cable connections were found. further the aic was examined for the 5s parameter values; however, the results indicated that these values did not meet the specified criteria as mentioned in the aic pm procedure. additionally, a visual inspection of the console's optical bench showed that the red fiber cable was broken, with light leaking out. root cause: the root cause the controller error (defective optical sensor lamp) - alarm issued was due to optical bench malfunction (broken fiber). there was no issue found related to the controller failure (red alarm) as reported.
9530 | 1220648-2025-28267 | 21995053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21995053&pc=OZD | 04/14/2025 | 06/10/2025 |  | 05/09/2025 |  |  | 04/14/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620780 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a hematoma developed at the impella cp access site following removal of the peel away sheath. steady bleeding occurred at all access sites which the physician attributed to disseminated intravascular coagulation. coagulation factors and one unit of blood were given to the patient to treat the condition. support was continued with the implanted pump following the intervention. however, when bleeding was continued from other site, the decision was ultimately made to withdraw care and the patient passed away. the medical safety officer reviewed and determined there is not enough evidence to exclude the impella device as an associated factor to the patient's outcome. | [5 empty] | The bleeding was not resolved and the hemostasis was not achieved as per the clinical details. the root cause of the access site bleeding was determined to be patient condition because of the bleeding from all access sites. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
9531 | 1220648-2025-28268 | 21995248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21995248&pc=PYX | 04/16/2025 | 06/06/2025 |  | 05/09/2025 |  |  | 04/16/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502629 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support, post-operative aortic and mitral valve replacement surgery five days prior and was implanted with an impella rp flex device via the right femoral vein for bipella mechanical circulatory support. prior to bipella support, heparin was stopped due to a nosebleed. thereafter, approximately two days after start of the bipella support, the patient's urine was noted to be port wine color. haptoglobin was low and labs were hemolyzing. the following day, an assessment was made for further patient bleeding (outcome of the assessment unnoted). hemolysis was still present, however, with low haptoglobin, elevated lactate dehydrogenase, and red urine. additionally, the impella rp flex flows were not within normal limits for given the performance level, and there was an abrupt change in motor current/flows noted at 06:00 in the morning. previous posterial nasal bleeding led to clots, respiratory failure, and reintubation, and the team was, therefore, hesitant to restart anticoagulation. the next day, the patient vomited and aspirated overnight requiring reintubation. an orogastric tube was placed and set to suction with 500 milliliters of dark red output. three units of packed red blood cells and two units of fresh frozen plasma were given. the patient lost pulses in right lower extremity, systemic heparin drip was initiated. the impella rp flex flows increased from 1.8 liters per minute (l/min) to 2.4 l/min. the patient had been started on continuous renal replacement therapy. on chest x-ray, the impella rp flex outlet appeared deep in left pulmonary artery. the physician was notified and bipella support was continued. the following day, the patient was taken to the operating room for veno-arterial extracorporeal membrane oxygenation (va ecmo) cannulation and impella rp flex explant. the impella 5.5 remained implanted and in place as verified by echocardiogram. however, three days later the patient would expire. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted the patient experienced adverse events with the impella 5.5 and is a serious injury. however, there is not enough evidence to exclude the impella rp flex as an associated factor to the patient's outcome (demise) with the report of inadequate pump flow and suboptimal positioning in addition to an adverse event. the impella rp flex is one of two devices associated with the event. note the following: refer to initial medical device report for the impella 5.5-serial number (b)(6) with date of event on 11-apr-2025 and patient identifier ending in (b)(6). the impella rp flex pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of hemolysis, positioning issues, and low pump flow has been completed. data review: data logs confirmed the low flow and clinical narrative. logs showed after pump started and run for 49 hours, motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. no positioning alarm was triggered during support. product analysis: pump was returned in damaged condition (catheter was cut off and& only cannula was returned for analysis). visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. conclusion: the root cause of hemolysis was biomaterial ingestion. the root cause of the low flow was biomaterial ingestion. the cause of the positioning issue was not determined due to insufficient clinical information.
9532 | 1220648-2025-28269 | 21995418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21995418&pc=OZD | 04/19/2025 | 06/05/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629881 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella 5.5 insertion on a patient was aborted after numerous attempts were unsuccessful in crossing the transcatheter aortic valve replacement valve. the impella cp was placed instead. the patient was eventually explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the 5.5 was unable to pass through valve and cp was inserted successfully.
9533 | 1220648-2025-28270 | 21995784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21995784&pc=OZD | 04/19/2025 | 06/25/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629135 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was supporting a patient with myocardial infarction. no prior cardiac history was noted. during support, it was noted that the patient had little to no pulsatility. in addition, the patient's condition continued to worsen. multiple drips were added. the patient experienced continuous suction alarms and the pump was repositioned. urine output dropped off. labs were hemolyzing so there was incomplete lab data. the decision was made to escalate to an impella 5.5. the impella cp was explanted and an impella 5.5 was successfully implanted. however, the patient coded and received several shocks. the patient expired on impella 5.5 support. | [5 empty] | The investigation of low pump flow, positioning issues, and hemolysis has been completed. the impella device was not returned for evaluation. positioning issues: the morning after device implantation the night before, the device was repositioned at bedside due to suction alarms, but no clinical information was provided as to the cause of the positioning. the pump was not returned. the cause of the positioning issue was not determined since product was not returned and insufficient clinical details. low pump flow: the pump was not returned. data log review showed flow trends down to 2.5l/min at p-7 where expected flows range from 2.9-3.3l/min. this occurred during numerous suction alarms. repositioning was reported to have occur numerous times with unknown resolve, however the exact timing of these repositions are not known. mc current spike occurs at 0650 and is followed by increase and elevated mc and pump flows, even though the pump flows still remain within the p-6 flow range of 2.5-2.9l/min. the repositioning of the pump, deteriorating patient condition / worsening heart function, and mc spike are all potential contributing factors, and without the returned product and further clinical details one root cause cannot be determined. the cause of the low pump flow was not determined since product was not returned and insufficient clinical details. hemolysis: data log review showed no placement signal trends. suction alarms found in data logs. mc current spike occurs at 0650 and is followed by increase and elevated mc and pump flows, even though the pump flows still remain within the p-6 flow range of 2.5-2.9l/min. the repositioning of the pump, deteriorating patient condition / worsening heart function, and mc spike are all potential contributing factors, and without the returned product and further clinical details one root cause cannot be determined. the cause of the hemolysis was not determined since product was not returned and insufficient clinical information. h6 health effect - clinical code 1942 and component code 3038 were erroneously entered on the initial manufacturer device report number 1220648-2025-28270. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9534 | 1220648-2025-28271 | 21995804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21995804&pc=OZD | 04/19/2025 | 12/11/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that while the patient on impella support was ambulating and just off wall power, the automated impella controller (aic) alarmed battery level low and percentage dropped to 50% within minutes of walking. after about 30 seconds the battery became full again without being plugged in. the aic was exchanged and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. data review: console logs were consistent with what was reported in the complaint. device analysis: console was tested and the reported battery level low alarm issue could not be reproduced during testing on an engineering lab bench. there were no evident battery issues observed while running the batttest utility on telnet. conclusion: the likely cause of the reported battery level low issue was a momentary communication glitch between the power battery manager (pbm) and the batteries. b3 date of event was erroneously entered on the initial manufacturer device report number. h8 was erroneously selected on the initial manufacturer device report number. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrected information provided in b1, b2, h1 and h6.
9535 | 1220648-2025-28272 | 21995983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21995983&pc=PYX | 04/18/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026614334 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a male patient was placed on bipella support. the patient labs were hemolyzing. a decision was made to remove the impella rp flex. upon removal, a clot was noted at the pigtail of the device. further information stated that care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9536 | 1220648-2025-28273 | 21996028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996028&pc=OZD | 04/11/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/11/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618566 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella 5.5 support, post-operative aortic and mitral valve replacement surgery five days prior was implanted with an impella rp flex device via the right femoral vein for bipella mechanical circulatory support. prior to bipella support, heparin was stopped due to nosebleed. thereafter, approximately two days after start of bipella support, the patient's urine was noted to be port wine color, haptoglobin low and labs were hemolyzing. the following day, an assessment was made for further patient bleeding (outcome of the assessment unnoted). hemolysis was still present, however, with low haptoglobin, elevated lactate dehydrogenase, and red urine. additionally, the impella rp flows were not within normal limits for given the performance level, and there was an abrupt change in motor current/flows noted at 06:00 morning. previous posterial nasal bleeding led to clots, respiratory failure, and reintubation, and the team was, therefore, hesitant to restart anticoagulation. the next day the patient vomited and aspirated overnight requiring reintubation. an orogastric tube was placed and set to suction with 500ml dark red output. three units of packed red blood cells and two units of fresh frozen plasma were given. the patient lost pulses in right lower extremity, systemic heparin drip was initiated. the impella rp flex flows increased from 1.8l/min to 2.4l/min. the patient had been started on continuous renal replacement therapy. on chest x-ray, the impella rp flex outlet appeared deep in left pulmonary artery. the physician was notified and bipella support continued. the following day, the patient was taken to the operating room for veno-arterial extracorporeal membrane oxygenation (va ecmo) cannulation and impella rp flex explant. the impella 5.5 remained implanted and in place as verified by echocardiogram. however, three days later the patient would expire. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted the patient experienced adverse events with the impella 5.5 and is serious injury. however, there is not enough evidence to exclude the impella rp flex as an associated factor to the patient's outcome (demise) with report of inadequate pump flow and suboptimal positioning in addition to adverse event. the impella 5.5 is one of two devices associated with the event. note the following: -refer to manufacturing report number 1220648-2025-28268 for the impella rp flex report. the impella 5.5 device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9537 | 1220648-2025-28274 | 21996179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996179&pc=OZD | 04/18/2025 | 05/27/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569901 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the patient was peri procedurally shocked twice for ventricular tachycardia (vt). | The complainant had placed a impella cp to support a patient admitted in with acute myocardial infarction and cardiogenic shock. it was reported that an echocardiogram showed the impella cp in poor position. it was reported that there was signs of hemolysis with the patients plasma free hemoglobin reported to be 210. decision was made for care to be withdrawn and the patient expired. |  |  |  |  | The investigation of positioning issues, hemolysis, and vt has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28274. b5 additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28274. h6 health effect - clinical code 2095 and health effect - impact code 4647 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28274. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9538 | 1220648-2025-28275 | 21996219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996219&pc=OZD | 04/19/2025 | 05/22/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594817 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support (mcs) and experienced ventricular fibrillation requiring intervention. the patient had coronary artery bypass graft surgery x 4 and was unable to come off the cardiopulmonary bypass (cpb) pump. an intra-aortic balloon pump (iabp) was placed without sufficient support. the decision was made to do an axillary cutdown for impella cp placement which was successfully completed. the iabp was left on standby. cpb was weaned and removed. the patient was transferred to the unit on support. overnight, the patient ventricular fibrillation arrested requiring multiple shocks and resuscitation. return of spontaneous circulation was achieved. however, the patient remained in guarded condition with continued bleeding from the chest tubes. blood products were continued to be given, and the patient was prepped for the operating room to manage the bleeding. the patient returned from operating room and no tamponade was noted. the source of the bleeding was found and resolved. t he chest was left open and the impella position was verified. the patient continued on impella mcs. four days later, the chest was closed and two days later, the patient was successfully completely weaned and the impella cp was explanted with a survived outcome. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9539 | 1220648-2025-28276 | 21996303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996303&pc=OZD | 04/19/2025 | 06/16/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614317 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella cp pump in a 68-year-old female patient was unable to get flows above .4 l/min. it was noted that active cpr was being performed due to complete cardiac standstill leading up to and throughout the event. the pump was replaced, but the patient later passed away. there is not enough evidence to disassociate the impella with the patient's outcome. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow issue was most likely patient condition related based on the data log investigation showing low flows from start of case and placement signal (ps) trending down, and clinical reports of patient's deteriorating condition requiring extracorporeal membrane oxygenation (ecmo).. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9540 | 1220648-2025-28277 | 21996436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996436&pc=OZD | 04/19/2025 | 06/06/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025596956 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 as elective placement for mechanical circulatory support and the following day the patient experienced arrhythmias, the patient remained intubated due to continued arrhythmias. the following day, the patient experienced ventricular tachycardia causing decreased flow and suction events. the patient was defibrillated two times. lidocaine drip was started. the next day, the patient underwent percutaneous coronary intervention. two stents were placed in the circumflex artery. four days later, the patient was extubated. the patient was awake and alert. suction events were noted, however, and the patient was still having runs of ventricular tachycardia correlating with coughing spells. the plan was to explant the impella cp device the following morning which was successfully completed with a survived outcome. the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9541 | 1220648-2025-28278 | 21996714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996714&pc=OZD | 04/13/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/13/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620637 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following implantation of an impella 5.5, high purge pressures appeared on the automated impella controller. the purge cassette was replaced, but the issue persisted. tissue plasminogen activator was subsequently run through the purge line and the issue resolved. hemodynamics were trending down as expected for the situation and care was withdrawn. abiomed's medical safety officer reviewed the event and determined this to be a product malfunction. there was no pump exchange, no change in pump flow or therapy was provided to the patient during impella 5.5 support. the patient was very ill and in advanced heart failure stage and was eligible for advance therapies to treat. | [5 empty] | The impella device was not returned, and the investigation has been completed. the clinical information stated that shortly after starting the pump there was high purge pressure alarms with no resolution by changing the purge cassette. tissue plasminogen activator use was instructed but it was unknown if it was added. there was no product or data logs returned. due to the lack of sufficient evidence, the root cause of the high purge pressure could not be determined. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9542 | 1220648-2025-28279 | 21996805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996805&pc=OZD | 04/19/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639641 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 73 year old male patient was on impella cp support and during support, the patient developed a left ventricle thrombus. no further information was provided. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9543 | 1220648-2025-28280 | 21996924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21996924&pc=OZD | 04/19/2025 | 06/12/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629213 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 60 year old female patient was on impella cp support and during support, the patient had oozing at the access site. intervention is unknown. support continued. additionally, the pump was sitting shallow and there was an apical thrombus. staff did not want to reposition due to the thrombus. furthermore, the patient had hemolysis. staff decided to explant the pump. upon explanting, the perclose sutures were tightened due to bleeding. angioseal was inserted to help with bleeding. the patient survived the procedure. | [5 empty] | The investigation of hemolysis, positioning issues, access site bleeding, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9544 | 1220648-2025-28281 | 21997195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997195&pc=OZD | 04/19/2025 | 06/16/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573441 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following impella cp implant, the pump encountered immediate suction. volume was given, but the flow remained low and the suction alarms returned. the staff believed that further care was futile and the patient passed away. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. data logs were investigated and low pump flow and suction alarm was confirmed at case start. the placement signal was observed to be marginally drifting down. therefore, as per the clinical information and data log review, the root cause of the low pump flow issue was most likely spontaneous ventricular rupture patient condition related since low flows were observed since pump start even after giving volume. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9545 | 1220648-2025-28282 | 21997261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997261&pc=OZD | 04/19/2025 | 06/09/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619077 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 as elective placement for mechanical circulatory support. the patient was planned for coronary artery bypass graft surgery x 4, atrial valve replacement and tricuspid valve replacement which were completed. post surgery, the patient was critically ill, on multiple high dose vasopressors and, however, "impressively" clearing lactate. approximately two days later, urine was less than 30 and amber red. the nurse, however, noted meth blue was given so coloration could likely be impacted by this. the plan was to start the patient back on systemic heparin and also wean pressors, bridge to recovery. the next day, the patient was noted to be doing better although still critical, started back up on systemic heparin. urine was still concentrated and pink but clearing. no hemolysis concern per the care team. mcs was continued. three days thereafter, the patient was noted to be heparin induced thrombocytopenia positive, and the following day there was increasing signs of hemolysis which led to explant of the impella 5.5 in which thrombus was noted adhered to the cannula. the patient was noted to have survived the explant. ejection fraction recovered to 20% from preoperative 5%. | [5 empty] | The pump was returned for investigation. data logs were not provided for this case run. a tiny amount of biomaterial was noted on the outflow cage. no other evidence of physical damage was noted on the returned pump. the pump was run according to specifications in a bucket of water. the pump was sent for hemolysis testing, and it passed the test with an mih value of 2.22. the cause of the hemolysis issue was not determined since the issue did not reproduce on the returned product, and the field pump run activity could not be confirmed without data log review. the cause of the thrombocytopenia issue was not determined due to insufficient clinical details. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9546 | 1220648-2025-28283 | 21997307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997307&pc=OZD | 04/19/2025 | 05/09/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635385 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported 65-year-old female patient on a 5.5 pump for support due to cardiomyopathy and cardiogenic shock. the pump was supporting when after 9 days the patient went to ct for imaging of an access site hematoma. the next day all heparin was held systemically. the purge was running with sodium bicarb. the pump remains on for support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output - use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9547 | 1220648-2025-28284 | 21997346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997346&pc=OZD | 04/20/2025 | 06/12/2025 |  | 05/09/2025 |  |  | 04/20/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168791 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a controller failure message appeared on the automated impella controller (aic) during impella support. an aic to aic transfer was completed and planned for support to continue. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: on 4/14, the case before the complaint, cp pump (b)(6) was connected to (b)(6) during case wizard. during initial priming, there were optical communication issues followed by the optical bench power cycling. this was followed by failed optical calibration. before the pump is started and throughout the case, there are frequent power battery manger (pbm)2 communication glitches. the real time (rt) logs for this case shows the raw optical sensor toggling between 50 mmhg and -1638.4mmhg but all other optical signals including calculated placement signal were as expected. the complaint case began on 4/17 with 5.5 pump (b)(6) with the completion of case start with no log abnormalities. starting at 14:54 on 4/20, there are 1045 errors for invalid placement signal. these twice resulted in controller error alarms for invalid optical crcs. the rt logs show that starting around 21:50 on 4/18, the raw optical sensor toggled between 70mmhg and -1638.4mmhg. at the time the 1045 errors began on 4/20 in the logs, the optical signals changed to all optical signals oscillating to -1638.4 and calculated placement signal triggered to 3000 mmhg. device analysis: elements of the clinical performance were reproduced during testing. while running a test pump, the rt logs show consistent raw optical spikes down to -1638.4mmhg similar to the previous cases 5s and ccd arrays were tested but performed as expected. all of the optical signals' triggering to -1638.4, which happened on 4/20, was not reproduced during testing. root cause: there was no controller failure found during the case. the optical signal behavior was most likely caused by either the optical bench's power (pbm) or the optical bench. aic - controller failure (red alarm): there was no controller failure found during the case. optical signal issue: the optical signal behavior was most likely caused by either the optical bench's power (pbm) or the optical bench. aic - controller error (yellow alarm): the controller error was caused by invalid optical signals which were most likely caused by either the optical bench's power (pbm) or the optical bench. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9548 | 1220648-2025-28285 | 21997426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997426&pc=OZD | 04/19/2025 | 06/16/2025 |  | 05/09/2025 |  |  | 04/19/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615352 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 84 year old patient developed a hematoma near their impella access site during support. fem-stop was placed and heparin was put on hold to manage the condition. the following day, the pump began to experience intermittent suction issues. the pump was explanted the following day. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding - major: the cause of the access site bleeding - major issue most likely was due to anticoagulation management. low pump flow: the cause of low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9549 | 1220648-2025-28286 | 21997508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997508&pc=OZD | 04/20/2025 | 06/11/2025 |  | 05/09/2025 |  |  | 04/20/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129492 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) which was successfully completed. after arriving at the intensive care unit, the purge system block alarm triggered. the purge cassette was changed. however, the automated impella controller was unable to complete the steps of the cassette change. the automated impella controller was therefore exchanged and resolved the issue. the patient continued with impella mcs. there was no harm to the patient as a result of this event. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on the data and device analysis the purge issue was caused by inaccurate purge pressure sensing. the drastic jumps in purge pressure could not have been naturally caused (jumping from 2000 mmhg to 0 mmhg in a second then back to 2000 mmhg). | The automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
9550 | 1220648-2025-28287 | 21997756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21997756&pc=OZD | 04/20/2025 | 06/13/2025 |  | 05/09/2025 |  |  | 04/20/2025 | 05/09/2025 |  |  |  |  | 05/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642055 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day, the intensive care unit nurse accidentally pulled impella cp pump into aorta. the physician was called in to reposition. patient was stable and slightly hypertensive, so decision was made to remove impella at bedside. there was successful removal and hemostasis after 30 mins of manual compression. | [5 empty] | The investigation of positioning issue has been completed. no product was returned for analysis. no data log analysis required. clinical mentions the nurse accidentally pulled impella into aorta. the root cause of the positioning issue was most likely due to a use issue as evidenced by the clinical detail of staff pulling impella into aorta. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9551 | 1220648-2025-28196 | 21984126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984126&pc=OZD | 04/16/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/16/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563624 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and during the support, heparin & integrillin were stopped due to an access site hematoma. the hematoma resolved with compression and light pressure dressing. the patient was reported to be in stable condition following the event. the patient would be eventually escalated to an impella 5.5 device after approximately 13 hours of impella cp support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9552 | 1220648-2025-28197 | 21984223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984223&pc=OZD | 04/16/2025 | 05/22/2025 |  | 05/08/2025 |  |  | 04/16/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579116 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 device in a patient with acute myocardial infarction/cardiogenic shock was unsuccessful. access was gained to the right axillary, and the impella 5.5 device was unable to pass through the axillary. an angiogram found the patient to have a large amount of calcium in the axillary artery. multiple balloon angioplasties were performed, and multiple attempts were made to cross the lesion without success. consequently, the case was aborted, and the patient was sent back to the unit in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to patient was having large amount of calcium in axillary artery.
9553 | 1220648-2025-28198 | 21984321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984321&pc=OZD | 04/17/2025 | 07/02/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570676 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that according to a cardiologist, there was no purge flow at some point before insertion of the impella cp on a patient. however, the purge pressure was ok. they could not see the purge solution exiting near the motor. however, they decided to implant the impella cp anyway after trying to reposition the purge cassette. after the implantation everything seemed to look fine with normal pressure and flow. no alarm appeared and the patient was transferred to the intensive care unit (icu). in the icu, the patient suffered from hemolysis (hemolyzed blood in the blood samples and red/brown urine). however, they were no alarms on the pump. no echocardiogram or repositioning of the pump was performed due to the pump was planned to be removed in the following hours and no additional intervention was needed. the pump was successfully weaned and explanted. | [5 empty] | The pump and purge cassette were returned for investigation. no evidence of physical damage was found on the returned product. the pump was tested in a bucket of water without any issues, and priming was successful on the returned product, showing no abnormalities. the pump was further sent for hemolysis testing and passed the test with an mih value of 11.24. the data log shows no abnormalities reported related to hemolysis. additionally, purge parameters were observed to be normal during the initial priming steps. multiple case start complete steps were recorded, followed by de-airing and cassette change steps during the case start. imc log were no provided for the review. according to the clinical notes, the doctor was unable to see the purge fluid and proceeded with the case after normal troubleshooting. purge parameters were normal during case start. there was no priming issue found during the case. the device functioned/operated to specification. the root cause of the hemolysis issue was not determined due to lack of clinical information, returned product passed hemolysis testing, and data log investigation was inconclusive. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9554 | 1220648-2025-28199 | 21984438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984438&pc=OZD | 04/17/2025 | 09/23/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645679 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient with an impella cp pump placed via single access was sent from the community to tertiary center for care escalation. the impella cp was explanted after a day and escalated to a impella 5.5 pump due to hemolysis (urine production reduced) and alarm for "in aorta". there was no harm or injury from the removal and escalation. the patient survived the procedure. | [5 empty] | The pump was returned for investigation. data log was not provided and could not be retrieved from the remote server. no physical abnormalities were reported on the returned pump. the pump passed the laser continuity test for the optical sensor, and all the 5s optical parameters were within specification during optical bench testing. the pump was run as per the specifications in a bucket of water. the pump also passed the water bucket test and the pressure test on the optical sensor. the pump was sent for hemolysis testing, and the results passed with a value of 3.61 mih. the cause of the hemolysis issue was not determined without data log review. the cause of the optical signal issue was not determined without data log review. the cause of the positioning issue was not determined due to insufficient clinical information. h6 type of investigation codes 4112 and 4110 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28199-1. h6 investigation findings 3221 was erroneously entered on the manufacturer device report number 1220648-2025-28199-1. h1 inadvertently did not selected device evaluation on manufacturer device report number 1220648-2025-28199-1. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure." | H6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28199 d10 concomitant medical products and therapy dates updated. | The pump was returned for investigation. data log was not provided and could not be retrieved from the remote server. no physical abnormalities were reported on the returned pump. the pump passed the laser continuity test for the optical sensor, and all the 5s optical parameters were within specification during optical bench testing. the pump was run as per the specifications in a bucket of water. the pump also passed the water bucket test and the pressure test on the optical sensor. the pump was sent for hemolysis testing, and the results passed with a value of 3.61 mih. the cause of the hemolysis issue was not determined without data log review. the cause of the optical signal issue was not determined without data log review. the cause of the positioning issue was not determined due to insufficient clinical information.
9555 | 1220648-2025-28200 | 21984474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984474&pc=OZD | 04/16/2025 | 06/23/2025 |  | 05/08/2025 |  |  | 04/16/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570026 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | Unspecified Vascular Problem | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that that the patient while on impella support had to be defibrillated once for ventricular fibrillation (vf) and was shocked once for atrial fibrillation (af). four days later the patient was defibrillated for atrial tachycardia. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that it was difficult placing the impella 5.5. impella would not make the transfer from the subclavian artery into the aorta. this took about 20 minutes, in the end the impella was pulled back into the graft and advanced again under a different angle to make the pass. during the or for cabg, a full collapse of the left ventricle (lv) occurred. this could only be managed by changing to p0 to allow the lv to fill and then slowly ramping up the impella again. the position of impella was toward the septum but free in the lv. back in icu when the impella was ramped up the collapse happened again with another episode of severe hypotension and increasing cvd pressures. when it was ramped down the patient had severe hypotension and lactates continued to increase with a completely akinetic lv. the decision was made to escalate to ecpella. ecmo was placed on near to maximum flow. impella was still needed because the av was not opening. impella was successfully repositioned towards the apex. regulatory report required for this event. |  |  |  |  | The investigation of vascular damage - peripheral vessel, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. vascular damage - great vessel: clinical details reported that once the patient was back in the icu after a cabg and impella insertion, 1500cc of blood was lost through the pericardial drains. they were brought into the or for a thoracotomy, and diffuse bleeding was noted, but no apparent source was found. product was not returned for investigation. clinical details did not report the source of the bleed, and the patient had several procedures that have potential to be related. additionally product was not returned for investigation. for these reasons, the root cause of the vascular damage could not be determined. positioning issues: data logs were reviewed, and there was one position in aorta alarm during the case. however, it resolved after just 16 seconds, and it's unclear if it triggered around the time of repositioning or not. since insufficient clinical details were provided regarding the cause/timing of the complication and product wasn't returned, the root cause of the positioning issue could not be determined. vt/vf: clinical details reported that the patient had to be defibrillated on (b)(6) 2025 for vf and then later shocked for af. additionally, they had to be defibrillated on (b)(6) 2025 for at. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. b5 additional information inadvertently omitted on the initial manufacturer device report number 1220648-2025-28200. h6 health effect - clinical codes 1729, 2130, 2095 and health effect - impact code 4647 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28200. h6 health effect - clinical code 4441 was erroneously entered on the initial manufacturer device report number 1220648-2025-28200. new code entered 1888. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9556 | 1220648-2025-28201 | 21984574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984574&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603533 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a hematoma at the access site. a washout was performed and heparin was withheld. staff monitored. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9557 | 1220648-2025-28202 | 21984591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984591&pc=OZD | 04/07/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609979 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Long-term outcome and quality indicator (loqi) impella registry database was received and reported a right posterior cerebral artery territory stroke with hemorrhagic conversion. the pump was explanted six days prior after successful wean. relatedness is noted as unknown (undetermined) per loqi. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9558 | 1220648-2025-28203 | 21984672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984672&pc=OZD | 04/18/2025 | 06/05/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646599 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in the intensive care unit had their impella cp repositioned by pulling back one centimeter and hemolysis was resolved. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position since hemolysis was resolved after repositioning of the pump.
9559 | 1220648-2025-28204 | 21984740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984740&pc=OZD | 04/17/2025 | 05/19/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618605 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Gastrointestinal Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information noted the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a run of ventricular tachycardia and required shock a few times. it was further reported that the patient developed a gastrointestinal bleed. bival was stopped. the patient was scoped by gi md. an ulcer was found and was cauterized/clipped. |  |  |  |  | Corrected information provided in b2, b5, h1 and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9560 | 1220648-2025-28205 | 21984874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984874&pc=OZD | 04/19/2025 | 06/20/2025 |  | 05/08/2025 |  |  | 04/19/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1729674 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that after the change of the purge cassette and bag, the trend screen only would show on the automated impella controller (aic). attempts were made to hit other soft buttons and use toggle. however, nothing would select. the pump continued to run. the issue was resolved after changing the aic. there was no reported patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Automated imeplla controller (aic) - kbd/rotary knob issues: the cause of the unresponsive soft touch buttons was most likely a spill during the bag/cassette change that shorted the capacitive sensing on the soft touch buttons and made them unresponsive. h10 investigation results statement for automated imeplla controller (aic) - kbd/rotary knob issues was inadvertently omitted from supplemental report 1220648-2025-28205-1. h10 aic - display issue: there was no display issue found during the case. the device functioned/operated to specification statement was duplicated in narrative results statement on supplemental report 1220648-2025-28205-1. | The investigation of automated impella controller (aic) - display issue has been completed. aic - display issue: there was no display issue found during the case. the device functioned/operated to specification. the console was returned for testing and initially performed as expected with no issues with soft touch buttons. a solution was sprayed on the console's soft touch buttons and they became non-responsive, artificially reproducing the failure mode. communication between the kbd and epc was as expected initially and ribbon cable was seated properly. aic - display issue: there was no display issue found during the case. the device functioned/operated to specification. root cause: the cause of the unresponsive soft touch buttons was most likely a spill during the bag/cassette change that shorted the capacitive sensing on the soft touch buttons and made them unresponsive.
9561 | 1220648-2025-28206 | 21984997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21984997&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636853 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Hemorrhagic Stroke | [10 empty] | Death | Life Threatening | Required Intervention | Disability | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced ventricular tachycardia (vt), cerebral hemorrhage and acute kidney injury with a demise outcome. the right axillary cut down was performed. graft anastomosed, and introducer placed, secured with graft locks. dr sharp accessed lv accessed with .035 guidewire and pigtail catheter. the guidewire was exchanged for an .018 guidewire. the pigtail catheter was removed and the impella was primed, backloaded over wire and advanced into the left ventricular without difficulty. the guidewire was removed out and the impella was slowly ramped up. the incision was closed and impella secured with 3-point fixation. throughout the procedure, the patient was in and out of ventricular tachycardia. the patient was sent to the unit on support, and after arriving on the unit the patient was in ventricular tachycardia. milrinone was stopped, lidocaine was added, and an electrophysiologist was consulted. the patient continued on support, and approximately five days later, the patient went for a computed tomography scan for a left ventricular assist device workup, and a "large" brain bleed was noticed. anticoagulation was currently being held. neurological status was monitored and support continued. the next day, it was noted the patient remained critically ill, worsening oxygen and acidosis. two days thereafter, the patient had an acute kidney injury (aki) and was put on dialysis. it is unknown if the patient had a kidney injury prior to impella support and/or if the impella contributed to the aki. the patient was consulted for palliative care. the patient's condition and neurological status remained unchanged, was transitioned to comfort care and expired two days later. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
9562 | 1220648-2025-28207 | 21985143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985143&pc=OZD | 04/19/2025 | 05/30/2025 |  | 05/08/2025 |  |  | 04/19/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642096 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Ischemia | Septic Shock | Thrombosis/Thrombus | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was cathed and mitral valve coronary artery disease was found. it was determined that the patient needed cardiac surgery but the impella was not placed due to peripheral vascular disease. the patient was taken for possible impella 5.5 placement for cardiac procedure. however, the patient started decompensating so the patient was sent to lab for impella cp placement with vascular assistance. the impella cp was placed in usual fashion. limb ischemia was found. the physician knew there would be ischemia so he had vascular do the case with him from the start and the leg was bypassed. the cardiac procedure was not done due to the patient was stabilized with the impella cp and wanted kidney function to improve. during support, it was noted that the bypass sheath clotted off and the retrograde sheath was found to be kinked. a new repercussion sheath was placed and flow was restored. the patient ended up having the intervention of two main vessels with no issue. however, the patient seemed to be in florid septic shock and nonresponsive to vasopressors. the patient did grow bacteria on the bronchealveolar lavage (bal). antibiotics were being given. despite aggressive intervention, the patient could not maintain blood pressure. the patient coded and a decision was made to stop therapy. the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of limb ischemia, thrombus, and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9563 | 1220648-2025-28208 | 21985214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985214&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025612315 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 had a transthoracic echocardiogram which indicated that the pump was turned toward the mitral valve. the physician planned to flip it at bedside under echo. additionally, heparin was discontinued as a hematoma was developing around the axillary site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
9564 | 1220648-2025-28209 | 21985230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985230&pc=OZD | 04/20/2025 | 06/17/2025 |  | 05/08/2025 |  |  | 04/20/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601496 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Incomplete or Inadequate Priming | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported during preparation, the impella cp would not prime. no purge solution was exiting, and there were no issues with the purge system. the physician decided to use a new impella cp with no report or indication of harm to the patient. | [5 empty] | The investigation of priming issue has been completed. no product was returned for evaluation. data logs were reviewed and real time logs from case shows purge pressures and purge flows were normalizing over 1 minute log file. no purge or priming alarms noted in returned logs. clinical details noted that pump would not prime and no purge was exiting from purge gap. no problem was found with the pump during the case as no purge alarms were noted on returned logs and purge values were within normal range. | H6 investigation findings and investigation findings were erroneously entered on the initial manufacturer device report number 1220648-2025-28209-1. | Patient-specific information is unavailable at the time of this mdr writing. therefore, respective fields were left blank. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9565 | 1220648-2025-28210 | 21985368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985368&pc=OZD | 04/20/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/20/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594113 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient suffered a hemorrhagic stroke during impella cp support. the patient's eyes were found to be dilated and a computed tomography scan revealed a massive bleed and infarction. the decision was made by the family to withdraw care. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9566 | 1220648-2025-28211 | 21985418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985418&pc=OZD | 04/16/2025 | 06/17/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094061 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male | 126 KG | Non Hispanic | Asian | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient expired while on support. per the medical safety officer there is not enough evidence to exclude the automated impella controller (aic) as an associated factor in the patient's outcome given the shut down and loss of patient support. | It was reported that during use the automated impella console turned off for 35 seconds and rebooted. the procedure outcome was successful with another device. |  |  |  |  | Device returned 29-apr-2025, investigation still pending. based on the medical safety officer's (mso) review of this case the patient's demise is associated with the automated impella controller (aic) associated with manufacturer device report number 1220648-2025-28211 has been revised to death. the following fields have been revised. b1 revised to reflect both adverse event and product problem. b2 death and date of death added on recommendation of mso. b5 revised with mso statement. h1 revised to death. h6 health effect - clinical code and health effect - impact code updated based on recommendation of mso. the impella pump was revised to serious injury and updated accordingly under manufacturer device report number 11220648-2025-28159. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. based on the data and device analysis the unexpected reboot was caused by the calculator software subprocess crashing. the cause of the calculator crashing could not be determined since the failure mode could not be reproduced.
9567 | 1220648-2025-28212 | 21985501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985501&pc=OZD | 04/21/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/21/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A physician was participating in a human factors study and informed the company team during feedback about a patient requiring distal limb amputation following treatment with an impella cp. no further information was provided. | [5 empty] | Patient information, initial reporter, facility name and product-specific information along with initial reporter details are unavailable at the time of this mdr writing. therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9568 | 1220648-2025-28213 | 21985541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985541&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629018 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Malposition of Device | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after impella 5.5 was implanted a transthoracic echocardiogram done at bedside showed the impella was deep at 7cm and so it was pulled back to 5.7 cm. the next day the patient exhibited sinus arrhythmia. the patient remained tachycardic for most of the time on 5.5 support. on day eight of impella support the purge side arm was leaking small amount of purge fluid. no change was seen in the purge flows to pressure. the team monitored the purge flows and pressure until the patient had a transplant on day 11 of impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 impella 5.5with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9569 | 1220648-2025-28214 | 21985572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985572&pc=OZD | 04/21/2025 | 06/06/2025 |  | 05/08/2025 |  |  | 04/21/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately three days after start of the support, oozing around the impella access site was noted. froward tension sutures were applied, and angle matching was performed with improvement noted. the status of the patient after the event was unknown. however, the latest update noted approximately six days later the patient was successfully escalated to an impella 5.5 which is currently supporting. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, forward tension sutures and angle match dressing were done to achieve hemostasis.
9570 | 1220648-2025-28215 | 21985725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985725&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  | 03/31/2025 |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397878 |  | 0550-0004 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient had an impella 5.5 pump implanted as a bridge to transplant exhibited arrhythmic storm. the storm put the transplant on hold. the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9571 | 1220648-2025-28216 | 21985772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985772&pc=OZD | 04/22/2025 | 05/20/2025 |  | 05/08/2025 |  |  | 04/22/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633515 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with both the impella cp and impella rp flex devices for (bipella) mechanical circulatory support (mcs) and approximately four days after start of impella mcs, oozing at the groin site of the impella cp was noted and resolved by redressing the access site and suspending heparin. no blood products were needed. the patient was noted to be in stable condition following the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely anticoagulation management since oozing was resolved by discontinuing heparin. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9572 | 1220648-2025-28217 | 21985904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985904&pc=OZD | 04/16/2025 | 06/10/2025 |  | 05/08/2025 |  |  | 04/16/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603930 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 as an elective placement to a patient for support of coronary artery bypass graft surgery was unsuccessful due to vascular anatomy. consequently, the impella 5.5 implant was aborted. there was no report of harm, and the patient was reported to be in stable condition following the event. | [5 empty] | The root cause of delivery issue is determined to be patient condition because the catheters were unable to pass through despite multiple attempts with multiple wires and catheters. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9573 | 1220648-2025-28218 | 21985978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21985978&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026632519 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during the impella 5.5 implantation the pump was backloaded over the wire and advanced through the introducer. the device was successfully guided into the left ventrical at which point the patient went into ventricular fibrillation. the patient was defibrillated multiple times, the wire was removed, and support initiated and slowly increased to p9. on the second day of support the impella was repositioned as it was a little shallow. on the sixth day of support the patient had three episodes of ventricular tachycardia and was shocked three times. the patient was given lidocaine and survived to surgery. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9574 | 1220648-2025-28219 | 21986151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986151&pc=OZD | 04/16/2025 | 06/06/2025 |  | 05/08/2025 |  |  | 04/16/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1292764 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller had battery alarms. no adverse impact to patient management was reported. the pump was removed from service. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data review: console logs from were consistent with what was reported in the complaint. multiple instances of battery failure (transport connection blown/missing) alarms were observed in the imc logs. device analysis: console was tested on an engineering lab bench after arriving in field service. the reported battery failure alarm issue was able to be reproduced. results of running the batttest utility revealed that cell stack 1 in battery 1 had a voltage that was significantly less than the voltages of the rest of the cell stacks in the battery. isolative testing was performed by swapping battery 1 with a known good one which resolved the alarm. conclusion: the root cause of the battery failure was determined to be a cell imbalance in battery 1. repair: before sending this console back to the customer, field service was instructed to replace battery kit [0042-3016], test aic power system via sections 18-20 of pm procedure [0042-7309], perform a full functional check on the console and optical system [0042-7316].
9575 | 1220648-2025-28220 | 21986322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986322&pc=OZD | 04/22/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/22/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a loaner automated impella controller (aic) failed when attempting to switch controllers to put onto a patient. the aic alarmed "purge sensor not detected" and the team crew changed the purge set attempting to correct the issue. however, once reloaded, the device alarmed again. the team opted to transport the patient on the sending hospital's aic instead of the loaner aic. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing. therefore, respective fields reflect "unknown" or left blank accordingly. the investigation has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined after arriving. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
9576 | 1220648-2025-28221 | 21986391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986391&pc=PYX | 04/18/2025 | 09/19/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025515374 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 58 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additionally, the complainant reported that there were placement signal not reliable alarms with absent pulmonary artery pressure and impella flow calculations. support continued. | The complainant reported that a patient implanted with an impella rp had hemolysis concerns. the patient survived and the pump was successfully weaned. |  |  |  |  | The investigation of hemolysis been completed. biomaterial was observed on the top of the impeller. the data logs do not show any motor current spikes. the clinical details does not provide supporting information for hemolysis concerns. no date, time, indicators, and possible resolutions were provided. the pump was hemolysis tested and it passed with an mih of 16.13. the root cause of the hemolysis was not determined as no motor current spikes were observed in the data logs and insufficient information related to failure (date, time, resolution) was provided. b2 required intervention was erroneously selected on the initial manufacturer device report number 1220648-2025-28221. h6 health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2025-28221. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the data logs confirm a placement signal not reliable (psnr) alarm and show placement signal drifting down while pump flow was drifting up, starting six hours after implant. placement signal and calculated flow drift before were both going out of range. central venous pressure was also out of range while motor current remained stable. the identified pattern is characteristic of a dp sensor drift. the pump was flow loop tested, and sensor drifting was reproduced. the root cause of the placement signal issue was determined to be due to sensor drift as evidenced by data log pattern and reproduction on returned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28221. b1 product problem. b7 other relevant history, including preexisting medical conditions. d10 concomitant medical products and therapy dates. g1 reporting contact email updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9577 | 1220648-2025-28222 | 21986447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986447&pc=PYX | 04/16/2025 | 06/06/2025 |  | 05/08/2025 |  |  | 04/16/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025528703 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support (mcs). two days prior, the patient underwent coronary artery bypass graft and 3-valve replacement surgeries, and shortly after arriving to the cardiovascular intensive care unit (cvicu), the patient went into ventricular fibrillation and was coded briefly. an emergent sternotomy was performed at bedside and chest was left open for two days. the patient was taken to the operating room for surgical chest closure, and after closing the chest, the patient's right ventricle exhibited dysfunction on echocardiogram with central venous pressure of 20 and hypotension requiring inotropic support. at this time the decision was made to implant an impella rp flex device which was successfully completed. approximately two days after start of the impella mcs, it was noted flows slowly dropped over course of therapy from 3.0 liters per minute to 2.4 liters per minute. during this time motor current trended downward from 380¿s to 240¿s. pulmonary artery placement signal gradually increased into 70¿s mean. at this time, the physician decided to initiate tissue plasminogen activator (tpa) protocol for purge solution to prevent pump failure. 2 milligrams of tpa in 50 milliliters was given through purge system at which point flow restoration occurred with motor current returning to upper 300¿s and the pulmonary artery placement signal decreasing to mean of 40¿s. thereafter, it was noted the impella rp flex ran without issue since said resolution. the patient's condition continued to worsen despite aggressive therapy. a discussion was ultimately made with the family regarding prognosis, and the decision was made to withdrawal care. the impella was weaned. the patient was extubated and all drips were turned off and the patient passed away shortly thereafter. | [5 empty] | Although the decision was made to discontinue impellasupport, there is not enough evidence to exclude the impella device (increase in purge pressure with subsequent decreased pump flow) as an associated factor to the overall outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | High purge pressure: clinical details report that while the pump was experiencing low pump flows on (b)(6) 2025, the team decided to run tpa through the purge, and it resolved the complications. data logs were reviewed, and an increase in purge pressure was noted during the low pump flows. the first step up in purge pressure was small, but correlated with the sudden drop in mc. soon after, there was a larger step up in the purge pressure to ~600 mmhg, not high enough to trigger an alarm. there were several bag changes during this period, but none longer than 2 minutes. data logs also show sodium bicarb being taken out of the purge, presumably when tpa was added, and then purge pressure trending back down to ~400 mmhg. though tpa being added to the purge correlates with the resolution of high purge pressure, the behavior of the motor current during the complication does not correlate with biomaterial ingestion or build up. additionally, the nature of the rise in purge pressure does not align with the typical increase due to biomaterial. product was not returned for investigation. since data logs did not reveal sufficient evidence for a cause of the rise in purge pressure and product was not returned for investigation, the root cause of the high purge pressure could not be determined. low pump flow: clinical details report that on (b)(6) 2025 pump flows had slowly trended down from 3.0 to 2.4 l/min. during this time, motor current had also trended down, and pap trended up. it is reported that after tpa was added to the purge, all the previously mentioned pump metrics returned to normal. data logs were reviewed and the reported trends were confirmed late on (b)(6) 2025. there was a sudden drop in mc and step up in the dp ps while running at p-7. the mc continued to trend down, pump flows went below specification, and shortly after there was a large step up in purge pressure to 600 mmhg. all pump metrics did return to normal during the time that the purge data shows sodium bicarbonate was removed from the purge, presumably when tpa was added. though the issue resolved while tpa was being run in the purge, the drop in mc at the same time is not indicative of biomaterial. it was noted that roughly 17 hours prior to this complication, there was a motor current spike to 1600 ma that was indicative of a biomaterial ingestion event. however, mc and purge pressure resolved prior to the reported complication on (b)(6) 2025, so the two events cannot be directly correlated. finally, clinical details reported that the patient was on crrt with a lot of volume being pulled during the time of the complication, which has the potential to play a role in pump flows. product was not returned for investigation. since product was not returned for investigation and data logs did not have any clear patterns indicative of a cause, the root cause of the low pump flows could not be determined.
9578 | 1220648-2025-28223 | 21986617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986617&pc=OZD | 04/17/2025 | 06/06/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1607654 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) was not connecting to impellaconnect.com, was displaying five blue blinks for the blink code. configuration was checked and a factory reset was performed with no change. log files were downloaded and noticed that the dates for the log entries were either april 28th or jan 4th even though the actual dates the machine was turned on were aug 10th and 14th. the log files show that the tls certificate is being rejected when it tries to connect to the rlr endpoint. we are assuming that this is due to the discrepancy in dates. the issue was discovered during console reconfiguration and no patient was involved. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | After evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally.
9579 | 1220648-2025-28224 | 21986753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986753&pc=OZD | 04/11/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient presented with chest pain and dyspnea on exertion. the patient had diagnostic catheterization which showed microvascular decompression. the patient elected to go with percutaneous coronary intervention (pci) route. the patient was taken to the catheterization lab for impella assisted pci. it was noted that the patient developed a pontine stroke and had left side weakness. the physician is asking for stroke risk in patients that get an impella cp. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9580 | 1220648-2025-28225 | 21986879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986879&pc=OZD | 04/18/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026616528 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 52 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 pump was implanted support a patient who was awaiting transplant. during support the access site developed a hematoma which was washed out and sutured. the access site bleeding also prompted the team to stop the systemic heparin. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9581 | 1220648-2025-28226 | 21986907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21986907&pc=OZD | 04/17/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621661 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella cp support developed bleeding from the right groin site. reperfusion sheath, impella cp, and venous extracorporeal membrane oxygenation cannula was in the right groin. six units of blood products were administered. the angle was adjusted and pressure dressing applied. it was further reported that care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not returned, and the investigation has been completed. the root cause of access site bleeding was most likely anticoagulation management since the activated clotting time was out of range which is not recommended per instructions for use. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9582 | 1220648-2025-28227 | 21987213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987213&pc=OZD | 04/17/2025 | 06/06/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643547 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an attempted impella cp implant and during the implant, there was difficulty placing the impella across the aortic valve. the impella catheter kinked and buckled in the aortic root. the doctor elected to remove and replace the catheter with another impella cp. the patient survived the procedure. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | Reported having difficulty placing impella across av and the catheter kinked in the aortic root. pump was replaced with another cp. patient has tavr valve. pump was returned and no catheter or cannula kinks were found. delivery issue - the root cause of the delivery issue was most likely patient condition related as the pump had difficulty crossing the av as there was a tavr valve in place. catheter kink - the root cause of the product damage was most likely a catheter kink that occurred during an attempt to deliver the pump across the av with the tavr valve in place.
9583 | 1220648-2025-28228 | 21987266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987266&pc=OZD | 04/18/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025582374 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Other |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient had an impella 5.5 and ecmo support ongoing. after fourteen days of support the pump became malpositioned and was adjusted. the patient had care withdrawn and expired. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
9584 | 1220648-2025-28229 | 21987329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987329&pc=OZD | 04/17/2025 | 05/22/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced limb ischemia requiring intervention. the patient came to the hospital as an outpatient and was found to have a very severe left anterior descending (lad) lesion that required stenting. the physician was stenting the vessel, and a "very large" dissection was noted into the lad. the patient coded. the decision was made to place an impella cp device for mcs. a 6 french sheath was upsized to 14 french x 13 centimeter impella sheath and 6 french pigtail was placed across into the left ventricle (lv). a 0.018 guidewire was placed and the impella cp was placed. flows of 1.8 were noted with minimal pulse pressure. cardiopulmonary resuscitation (cpr) continued, and the patient was shocked approximately 14-16 times. the physicians attempted to repair the dissection and ultimately a stent was placed from the apex of the lad back to the left main artery. the patient became semi stable on support and started on levophed and epinephrine. the patient was also intubated prior to insertion of the impella cp device. during the cpr, the lad was wired and stented. the patient was semi stable on support. the impella sheath was removed, and the repositioning sheath was placed. pulses were absent via doppler and a 6 french braided sheath was placed for external femorofemoral bypass. the patient was transported to the cardiovascular intensive care unit in very critical but stable condition with the plan to rest the heart and hope for heart recovery. outlook was noted as "very grim" but there was hope for heart recovery. two days later, the patient was extubated and breathing on his own. the patient was awake and alert to his surroundings hemodynamics improved. support was continued and the next morning, the patient was slowing being weaned and down to p-6. th patient was tolerating everything very well and weaning continued over the next 24 hours. the patient was successfully weaned, and the device was explanted without incident and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9585 | 1220648-2025-28230 | 21987363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987363&pc=OZD | 04/08/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/08/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient had elective mitral valve repair six days prior and stayed in the hospital looking well until overnight. the started to have decreased urine output and was transferred to cardiovascular intensive care unit. the patient required increasing chemical support and had brief a code in the unit. the decision made to place the impella cp device, which was successfully completed. while in the catheterization lab the right ventricle was assessed, and a decision was made to place biventricular support (impella rp flex). soon after the implant the physician had concerns regarding perfusion to the right leg (impella cp leg). limb ischemia was indicated and the impella cp device was explanted and replaced with an impella 5.5 device for continued bipella mcs. now on impella 5.5 and impella rp flex, the following day the staff was still unable to get a doppler pulse on the lower extremities. the patient was continued on bipella support and continuous renal replacement therapy (crrt) was started. the next day an amiodarone drip was started due to tachycardia. a neurological assessment was completed the day after and there was no movement noted in the extremities. two days later, the patient expired. care was withdrawn. | [5 empty] | Product-specific information for the impella cp device is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. medical safety reviewed the event and noted it appears arterial flow was being occluded associated with limb ischemia on the impella cp., which appeared to be an unresolved issue (no information of resolution achieved) and is the leading issue associated with the patient's demise outcome. neuro assessment noted no movement in extremities (no stroke diagnosis), and it appears the issue is most likely due to the limb ischemia. arrhythmic event involving the impella rp flex and impella 5.5 occurred and appeared to have been resolved with amiodarone (this is a serious injury for the impella 5.5 and impella rp flex). the impella cp is one of three devices associated with the event. note the following: refer to initial medical device report for the impella rp flex-serial number: (b)(6) with date of event 10-apr-2025 and patient identifier ending in (b)(6). refer to initial medical device report for the impella 5.5-serial number: (b)(6) with date of event 10-apr-2025 and patient identifier ending in (b)(6). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9586 | 1220648-2025-28233 | 21987700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987700&pc=OZD | 04/18/2025 | 10/17/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025552407 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that the upon initial prime of the impella cp and calibration, placement signal was inaccurate. impella was placed and placement signal kept being an issue throughout the procedure. the physician was able to complete the procedure successfully. | The complainant reported that the upon initial prime of the impella cp, the placement signal was inaccurate. in response to suction alarms a 500ml bolus was given. the patient survived. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: the clinical stated that upon initial calibration the placement signal read -20/-56 and that the rep told the physician to not touch the metal cage upon insertion. the pump was returned and connected to the console and it read -36/-37 while flowing in a bucket of water. the 5s were taken and they were all within specification. no placement signal alarms were produced. the pump was uson tested and it reacted to the pressure but the was about 20mmhg higher then what the set pressure was. the pressure with the uson was started at high values instead of 0 which could have impacted the way the sensor was responding. though there are not data logs, the imc logs from the connection to the lab console show that the calibrated g0=16250 while the factor g0=16312, showing a difference of 70. based on the returned product and g0 values, it can be determined that the root cause of the optical signal issue is most likely due to an offset error however it is unknown how the offset became incorrect due to lack of field data logs provided, and the clinical stating that the doctor might have placed pressure on the sensor. low pump flow the pump was returned and there were no physical abnormalities found but was a slight film in the purge gap and some biomaterial around the impeller. the pump was soaked and it was ran for 26 hours without purging and the flows were slightly saturated, most likely due to the biomaterial present. low pump flows were not able to be reproduced. due to lack of data logs returned, the root cause of the low pump flow cannot be determined. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | B7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. there were no data logs returned for this case. thrombus was added since it was found on the returned product. the pump was returned and there were no physical abnormalities found but was a slight film in the purge gap and some biomaterial around the impeller. the pump was soaked and it was ran for 26 hours without purging and the flows were slightly saturated, most likely due to the biomaterial present. low pump flows were not able to be reproduced. due to lack of data logs returned, the root cause of the low pump flow cannot be determined. as the necessary information was not provided, the root cause of the thrombus was not determined h3 device not returned was erroneously selected yes on the initial manufacturer device report number 1220648-2025-28233. h6 type of investigation code 4114 was erroneously entered on the initial manufacturer device report number 1220648-2025-28233. h10 narrative erroneously stated that the device was not received on the initial manufacturer device report number 1220648-2025-28233. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9587 | 1220648-2025-28231 | 21987789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987789&pc=OZD | 04/17/2025 | 06/23/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630933 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and the next day after start of the impella mcs, slight oozing at the impella cp insertion site was noted. the patient required lysis, and a pressure bandage was placed at the site to manage the condition. the patient's outcome about resolution was indicated as unknown. however, latest update noted two days later, the patient was awake and well oriented. the patient was in the weaning process with an explant planned. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9588 | 1220648-2025-28234 | 21987800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987800&pc=OZD | 04/17/2025 | 05/30/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627551 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Unspecified Infection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and had a low grade fever. tylenol and antibiotics remained on board. additionally, the pigtail was abnormally directed towards the mitral valve and was deep 4.5 centimeters below the aortic valve. the apex showed an echogenic mass in the left ventricle apex indicating a thrombus. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of positioning issues, thrombus, and infection/sepsis has been completed. pump was not returned for investigation. therefore, the root cause of the positioning issue was not determined since product was not returned and it is unclear how the pump pulled out of position based on the clinical information provided. as the necessary information was not provided, the root cause of the thrombus and infection/sepsis was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9589 | 1220648-2025-28232 | 21987811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21987811&pc=OZD | 04/16/2025 | 05/21/2025 |  | 05/08/2025 |  |  | 04/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624681 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and approximately fifteen days into support, heparin was paused due to the patient's partial thromboplastin time. the mcs continued for another two days until planned explant following a successful bridge to transplant. shortly, thereafter, the patient had a brief period of left sided weakness in their left upper extremity. a stroke code was called, and a computed tomography scan revealed a subacute infarction to the right middle cerebral artery. the patient was stabilized within 30 minutes of the event. | [5 empty] | The impella device was not returned, and the investigation has been completed. as this clinical information was not provided, the true root cause of the stroke/cerebrovascular accident could not be determined. the root cause of thrombocytopenia could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9590 | 1220648-2025-28235 | 21988044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21988044&pc=PYX | 04/10/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/10/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625225 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient had elective mitral valve repair six days prior and stayed in the hospital looking well until overnight. the started to have decreased urine output and was transferred to cardiovascular intensive care unit. the patient required increasing chemical support and had brief a code in the unit. the decision made to place the impella cp device, which was successfully completed. while in the catheterization lab the right ventricle was assessed, and a decision was made to place biventricular support (impella rp flex). soon after the implant the physician had concerns regarding perfusion to the right leg (impella cp leg). limb ischemia was indicated and the impella cp device was explanted and replaced with an impella 5.5 device for continued bipella mcs. now on impella 5.5 and impella rp flex, the following day the staff was still unable to get a doppler pulse on the lower extremities. the patient was continued on bipella support and continuous renal replacement therapy (crrt) was started. the next day an amiodarone drip was started due to tachycardia. a neurological assessment was completed the day after and there was no movement noted in the extremities. two days later, the patient expired. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted it appears arterial flow was being occluded associated with limb ischemia on the impella cp., which appeared to be an unresolved issue (no information of resolution achieved) and is the leading issue associated with the patient's demise outcome. neuro assessment noted no movement in extremities (no stroke diagnosis), and it appears the issue is most likely due to the limb ischemia. arrhythmic event involving the impella rp flex and impella 5.5 occurred and appeared to have been resolved with amiodarone (this is a serious injury for the impella 5.5 and impella rp flex). the impella rp flex is one of three devices associated with the event. note the following: -refer to manufacturing report number 1220648-2025-28230 for the report on the impella cp. -refer to initial medical device report for the impella 5.5-serial number (b)(6) with date of event 10-apr-2025 and patient identifier ending in (b)(6). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smart assist system. section: potential adverse events: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
9591 | 1220648-2025-28236 | 21988171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21988171&pc=OZD | 03/14/2025 | 10/03/2025 |  | 05/08/2025 |  |  | 03/17/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601386 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Atrial Fibrillation | Hemolysis | Unspecified Infection | [10 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a concern regarding a potential optical sensor issue due to the placement signal waveform being incongruent with the patient¿s arterial line pressures. | The complainant reported a patient on impella 5.5 needing mechanical circulatory support. the pump had suction alarms occur during support. furthermore, the patient experienced atrial fibrillation and was cardioverted. the patient also had hemolysis during support, the pump was repositioned away from the mitral valve. moreover, the patient had an infection on support with a low-grade fever. the swan was discontinued and the relationship between the infection and the impella is unknown. later we received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening cardiogenic shock with a "definite" relationship to the impella device used: ¿(b)(6) 2025- ldh/pfh uptrending, starting integriin following c-shock call (b)(6) 2025-intermittent hypotension/suction alarms noted. ldh 1187-slowly trending down, remains on heparin/cangrelor (b)(6) 2025-r thoracentesis, 900 ml dark, red fluid removed, integrelin on hold (b)(6) 2025-presented to mrb, declined for oht/vad (b)(6) 2025-pt dnr (b)(6) 2025-pt made cmo, admitted to hospice. impella 5.5 turned off @ 1011.¿ the patient expired. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue, infection/sepsis, hemolysis, low pump flow, positioning issues, and atrial fibrillation. pump position was confirmed to be adequate. data logs were reviewed and the reported issues were confirmed. logs show these abnormalities occurring when the measured pressure by the optical sensor increased above a certain threshold. the pump was not returned for investigation. based on product not being returned and data log review not showing definitive cause, the root cause of the optical signal issue could not be determined. because this evidence has not been provided, the root cause of the infection cannot be determined. since product was not returned for investigation, data log review was not able to confirm if the pump metrics that had previously trended down recovered when pfhgb levels improved, and limited clinical details were provided, the root cause of hemolysis could not be determined. since product was not returned, lack of clinical details, and data log review did not establish a definitive cause, the root cause of the low pump flows could not be determined. since is it unclear from clinical details how the pump got into an improper position near the mitral valve, the root cause of the positioning issue could not be determined. the cause of the atrial fibrillation was unable to be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28236. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9592 | 1220648-2025-28237 | 21988247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21988247&pc=OZD | 04/10/2025 | 06/05/2025 |  | 05/08/2025 |  |  | 04/10/2025 | 05/08/2025 |  |  | 02/28/2026 |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474021 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient had elective mitral valve repair six days prior and stayed in the hospital looking well until overnight. the started to have decreased urine output and was transferred to cardiovascular intensive care unit. the patient required increasing chemical support and had brief a code in the unit. the decision made to place the impella cp device, which was successfully completed. while in the catheterization lab the right ventricle was assessed, and a decision was made to place biventricular support (impella rp flex). soon after the implant the physician had concerns regarding perfusion to the right leg (impella cp leg). limb ischemia was indicated and the impella cp device was explanted and replaced with an impella 5.5 device for continued bipella mcs. now on impella 5.5 and impella rp flex, the following day the staff was still unable to get a doppler pulse on the lower extremities. the patient was continued on bipella support and continuous renal replacement therapy (crrt) was started. the next day an amiodarone drip was started due to tachycardia. a neurological assessment was completed the day after and there was no movement noted in the extremities. two days later, the patient expired. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted it appears arterial flow was being occluded associated with limb ischemia on the impella cp., which appeared to be an unresolved issue (no information of resolution achieved) and is the leading issue associated with the patient's demise outcome. neuro assessment noted no movement in extremities (no stroke diagnosis), and it appears the issue is most likely due to the limb ischemia. arrhythmic event involving the impella rp flex and impella 5.5 occurred and appeared to have been resolved with amiodarone (this is a serious injury for the impella 5.5 and impella rp flex). the impella 5.5 is one of three devices associated with the event. note the following: -refer to manufacturing report number 1220648-2025-28230 for the report on the impella cp. -refer to manufacturing report number 1220648-2025-28235 for the report on the impella rp flex. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vt/vf has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 investigation findings and investigation conclusions codes were erroneously entered on the initial manufacturer device report number 1220648-2025-28237. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28237.
9593 | 1220648-2025-28238 | 21988635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21988635&pc=OZD | 04/18/2025 | 10/30/2025 |  | 05/08/2025 |  |  | 04/18/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | OZD |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026625180 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 83 year old male patient was on impella rp support and during support, placement signal pressures slowly increasing throughout the day, placement signal went above 200's and flows on impella showed 0.0. left sided device appears to be getting adequate preload from right side which would indicate device is still functioning properly and there is a differential sensor issue. motor current values are within normal range and holding steady, and cvp/papi are still being displayed indicating normal fiber optic sensor function. additionally, the patient had atrial fibrillation (af). the patient was cardioverted. the patient was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. a data log was not provided for this case run. clinical information revealed placement signal spikes during the case. motor current values were displayed as 0, and cvp was displayed during the time of the placement signal issue. this issue was confirmed to be related to the dp sensor of the device, which is addressed under the placement signal issue failure mode. the cause of the placement signal issue was not determined since product was not returned and sufficient clinical information was not provided including data log. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. a4 updated. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
9594 | 1220648-2025-28239 | 21989049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21989049&pc=OZD | 04/14/2025 | 05/08/2025 |  | 05/08/2025 |  |  | 04/14/2025 | 05/08/2025 |  |  |  |  | 05/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026630354 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male | 86 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Other |  |  |  |  | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced cardiac arrhythmia that was successfully handled. ultimately, the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9595 | 1220648-2025-28141 | 21975376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975376&pc=OZD | 04/15/2025 | 09/24/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026638658 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient on was impella 5.5 support. it was reported that the p-level was lowered due to the patient experiencing arrhythmias. the next day, volume was given, and the p-level was increased. the patient continued on support. | Long-term outcome and quality indicator (loqi) impella registry (loqi) also reported that after impella was removed: the patient endured a stroke - multiple territories with a "probable" relationship to the impella device used: ¿overnight (b)(6) 25, code gold and anesthesia stat called due to significant lethargy with acutely increasing o2 requirements brain imaging negative for acute stroke, repeat head ct with concerns for acute right cerebellar infarct. brain mri pending ct head without evidence of acute cranial pathology, cta head/neck with no high-grade stenosis or lvo. patient not a candidate for iv thrombolytics due to recent major procedure and evt not warranted due to no lvo on imaging. echo with ef 10-15%. severe global hypokinesis, no evidence of thrombus. mri brain with large acute right pica territory infarction involving the cerebellum and right side of the medulla with petechial hemorrhage, along with several other acute scattered strokes. eeg without evidence of seizure activity. tee without source of cardioembolism. currently on dapt per cardiology for 12 months in context of acs followed by single ap- asa 81 mg / plavix 75 mg. continue lipitor 40 mg. hvi¿. the patient outcome is not recovered/not resolved. | The complainant also reported that the impella had placement signal issue but there was no change in flows or motor current. |  |  |  | The investigation of vf/vt/af/at has been completed. the pump was returned and inspected. no physical abnormalities were observed and 5s parameters were nominal. the complete data logs were not recovered. as the necessary information was not provided, the root cause of the vt/vf was not determined. d6b: explant date added. d9: device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The cause of the stroke was not determined due to insufficient clinical details. b5 additional information inadvertently omitted on the initial manufacturer device report number 1220648-2025-28141-1. h6 additional coding was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28141-1. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. after one day on support, report of sudden change in placement signal and left ventricle (lv) waveforms. both reading much higher than they have been with stable motor current. the pump was returned and inspected. no physical abnormalities were observed and 5s parameters were nominal. uson testing showed the pump change with varying externally applied pressures, although with an offset of about 50 mmhg. although some data logs were recovered, no logs were recovered on the complaint occurrence day of (b)(6) 2025. the g0 calibration in the field did not change from the factory g0 of 15948. the cause of the optical signal issue was not determined since no issues were found with the returned pump and the complete data logs from the field were not recovered. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issues. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28141 d10 concomitant medical products and therapy dates updated.
9596 | 1220648-2025-28142 | 21975424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975424&pc=OZD | 04/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639040 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with an anterior myocardial infarction and ventricular fibrillation. an impella cp was implanted and percutaneous coronary intervention was performed for the left anterior descending artery. a suction event occurred which may have been caused by pump migration. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9597 | 1220648-2025-28143 | 21975491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975491&pc=OZD | 04/15/2025 | 06/12/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp to treat high-risk percutaneous coronary intervention. upon cp insertion pump unable to pass through the sheath despite multiple attempts. a crack noted in 25 cm sheath. short sheath attempted and was exchanged for dryseal, pump unable to pass. the patient had vagal response thought to be from excessive blood loss through sheath exchanges. one-unit of emergent blood given and emergent percutaneous coronary intervention started as the patient started to have active st-segment elevation myocardial infarction. the patient was stabilized and a decision made to abort impella insertion and groin was closed. additionally, there is a possible bend on inlet cage noted. the patient survived. | [5 empty] | The investigation of access site bleeding, delivery issues, difficulty inserting/removing pump through introducer, introducer damage ¿ mechanical, and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding is most likely due to use as bleeding was thought to have been caused from the amount of sheath exchanges. delivery issues: due to the lack of clinical details provided, the root cause of the delivery issues cannot be fully confirmed. difficulty inserting/removing pump through introducer: due to lack of clinical details, the true root cause cannot be determined. introducer damage - mechanical: due to lack of clinical details and no product returned, the root cause cannot be confirmed for how this occurred. product damage (cannula kink): due to lack of clinical details and no product returned, the root cause cannot be confirmed for how this occurred. h6 health effect - impact code 4604 and component code 4742 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28143. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9598 | 1220648-2025-28144 | 21975536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975536&pc=OZD | 04/16/2025 | 06/04/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648751 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp. during support, the pump was dislodged into the aorta during compressions. the pump was explanted and the patient was placed on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Reported impella cp dislodged into aorta during compressions. patient was placed on venoarterial extracorporeal membrane oxygenation (va ecmo) and impella was removed. the root cause of the positioning issue was most likely use related since the pump was dislodged into aorta during compressions. d6a and d6b were erroneously entered on the initial manufacturer device report number 1220648-2025-28144.
9599 | 1220648-2025-28145 | 21975544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975544&pc=OZD | 04/15/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025588063 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella rp flex was removed in cardiovascular catheterization laboratory after hemodynamics stabilized due to excessive access site bleeding. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
9600 | 1220648-2025-28146 | 21975566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975566&pc=OZD | 04/15/2025 | 06/10/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635367 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that multiple wires were used in attempts to implant the impella 5.5. they were unable to obtain left ventricle access. aortic valve not opening under trans-esophageal echo due to patient¿s anatomy and prior surgical procedures. there was not enough aorta length to implant new graft. the 5.5 insertion was aborted. there was no patient harm. | [5 empty] | The root cause of delivery issue is determined to be patient condition because the catheters were unable to pass through despite multiple attempts with multiple wires and catheters. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
9601 | 1220648-2025-28147 | 21975592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975592&pc=PYX | 04/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025567899 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella rp flex was implanted to support a patient admitted in with st-elevation myocardial infarction. an access site bleed occurred which was treated with quickclot. the oozing improved and the support was continued. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿
9602 | 1220648-2025-28148 | 21975648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975648&pc=OZD | 04/16/2025 | 06/20/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1707157 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported after support to a patient the automated impella controller was unplugged and with five minutes an alert displayed and triggered that the battery was low despite being fully charged when plugged into the outlet. there was no indication or report of harm to the patient. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. unable to identify any events that would correspond to complaint description, and data analysis and inspection of ic10325 indicated that this console was unlikely to have been associated with the complaint. the device functioned/operated to specification. h5 yes; was erroneously entered on the initial manufacturer device report number 1220648-2025-28148.h8 initial; was erroneously entered on the initial manufacturer device report number 1220648-2025-28148. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
9603 | 1220648-2025-28149 | 21975659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975659&pc=OZD | 04/16/2025 | 06/04/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640419 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp a patient experienced access site bleeding. the impella cp was restitched and the angle matched which resolved the bleeding. additionally, a chest x-ray showed a deeply positioned impella cp with a small kink in the cannula, and the pump was pulled back. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, positioning issues, and product damage (cannula kink) has been completed. access site bleeding: based on the clinical details the cause of the access site bleeding - major most likely was use related since epi drip, re-sutured and angled match improved the site oozing. product damage (cannula kink): the cause of the product damage (cannula kink) most likely was cannula kink due to unknown reason. positioning issues: based on the clinical details the cause of the positioning issues most likely was use related due to improper technique since chest x-ray showed a deeply positioned impella.
9604 | 1220648-2025-28150 | 21975752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975752&pc=OZD | 04/16/2025 | 06/13/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465870 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller was unplugged and an alert displayed and triggered that the battery. there was no indication or report of harm to the patient. | [5 empty] | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. the console logs from ic7938 do not show any events consistent with problem description. analysis of all data logs related to pump reveled that ic7938 was the second console used with the pump, and there was a battery-related issue on prior console. there were no issues observed with ic7938 during the case. console ic7938 was tested and no battery-related or charging-related anomalies were found, thus confirming that console ic7938 was erroneously identified as related to battery alarm. there was no issue found during the case. the device functioned/operated to specification. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed.
9605 | 1220648-2025-28151 | 21975864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975864&pc=OZD | 04/16/2025 | 05/22/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649446 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with an ejection fraction of 20% and multi-organ system failure. the patient coded multiple times and was cannulated for extracorporeal membrane oxygenation (ecmo). the decision was made to place an impella cp for left ventricle venting. during support, it was reported that the patient was converted to normal sinus rhythm from atrial tachycardia. it was noted of the plan to continue the course with ecpella (ecmo and impella cp) until lactate clear and kidney function recover. all labs were trending positively. the patient was escalated to an impella 5.5. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9606 | 1220648-2025-28152 | 21975866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975866&pc=PYX | 01/12/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 01/12/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561689 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, heart transplant rejection, was implanted with an impella rp flex device for mechanical circulatory support (mcs) and experienced low pump flow with some hemodynamic instability requiring intervention. the patient was initially on an impella cp mechanical circulatory support prior for heart and kidney transplant evaluation. that impella was successfully weaned. however, the patient had since decompensated requiring increased inotropes/pressors. the patient had a pulseless electrical activity arrest at the outside hospital requiring venoarterial extracorporeal membrane oxygenation (va ecmo) and impella cp placement (ecpella). the patient was also started on continuous renal replacement therapy. on day 8 of ecpella support, the patient was transitioned to bipella. the impella rp flex was implanted to support with the impella cp, and the va ecmo was removed. subsequently, approximately two days later, it was noted the patient was perfusing, although borderline with current bipella support (however it was noted to be at least better than impella cp mcs alone). the patient had impella rp flex suction alarm and support level had to be decreased. the physician was unable to increase the impella rp flex beyond p-7 without getting suction. also, per the registry notes, the impella cp was also decreased to p-7 for hemolysis. the patient was transfused with packed red blood cells due to increasing lactate and escalating norepinephrine. there was no bleeding noted but monitored for signs. per the registry, the patient/event (hemolysis) did not resolve. the patient would subsequently expire days later. it was noted the patient was not improving in heart and renal failure, and the patient asked for withdrawal of life-sustaining treatment. | [5 empty] | Medical safety reviewed the event and noted the impella rp flex was implanted (same day as impella cp-related hemolysis/against the septal wall adverse event) for bipella support. a subsequent issue occurred involving the rp flex (~2 days after) of inadequate flows and inability to increase the flow above p-7 without suction and is considered serious injury. the patient's demise is associated with the impella cp device. note: refer to manufacturing report number 1220648-2025-26875 for the report on the impella cp. the impella rp flex device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9607 | 1220648-2025-28153 | 21975874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975874&pc=OZD | 04/15/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885817 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that automated impella controller screen was missing two trends (cardiac output and native cardiac output) during a procedure. however, they were able to successfully use the console with no harm to the patient. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9608 | 1220648-2025-28154 | 21975962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975962&pc=OZD | 04/16/2025 | 07/14/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562206 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 10 YR |  | Female | 56 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | Malposition of Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that an impella 5.5 gave suction alarms while implanted in a patient. additionally, purge fluid was leaking onto component/pressure reservoir. staff was unsure of where the leak was coming from. the purge cassette was replaced. the following day, some dampness around the reservoir/chamber for the purge was noted. purge pressures and flows remained stable, and support continued. | [5 empty] | Cut pump was returned for investigation. the data log shows the placement signal trend and suction alarm during the start of the case without observing any motor current spikes or positioning-related alarms. the placement signal trend and flow were seen as normal in the one hour following the start of the case. the doctor reported that 1 unit of blood was administered during the suction event. additionally, some purge fluid leak was reported around the pressure reservoir without affecting the purge parameters. purge leak - pump: the returned pump was visually inspected, and no abnormalities were observed other than a cut catheter near the motor. the pump was purged after blocking the purge line's open end, and a reservoir leak was reproduced. the origin of the leak was not identified, so a higher concentration of colored fluid was used to pressurize the device overnight. the leak did not reappear until the next day when the reservoir was manually pressurized to approximately 1200-1400 mmhg, at which point the leak was reproduced across the radial seal of the purge reservoir. the cause of the purge leak issue was determined to be a sidearm reservoir issue, as the leak was reproduced from the reservoir radial seal on the returned product. low pump flow: the product was cut open, so no investigation could be performed on the low pump flow. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. d9 device returned to manufacturer date added. | The investigation of purge leak ¿ pump positioning issues, and low pump flow has been completed. the cause of the purge leak issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. d6b explant date added. d6a implant date was erroneously entered on the initial manufacturer device report number 1220648-2025-28154. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28154. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9609 | 1220648-2025-28155 | 21975986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21975986&pc=OZD | 04/16/2025 | 05/29/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642032 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a patient with st-elevated myocardial infarction and a placed stent. during support, a hematoma developed at the impella access site. pressure was held for 30 mins followed by femo stop placement. aggrastat was given and heparin was withheld. in addition, the patient had deteriorated and was requiring increased oxygen needs and presser support. continuous renal replacement therapy was started in response to anuria. extracorporeal membrane oxygenation cannulation was provided due to the patient's oxygenation status. however, the arterial cannula was not secured and somehow fell out resulting in very significant blood loss. the artery was re-cannulated and the vessel was repaired around the cannula, and multiple blood products were administered. ultimately, care was withdrawn and the patient expired. | [5 empty] | The investigation access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9610 | 1220648-2025-28156 | 21976106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976106&pc=OZD | 04/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634811 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 60 KG |  | White | Death | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had hematuria and hemolyzed labs, and the impella was repositioned. the patient also had sustained ventricular tachycardia (svt) and was successfully cardioverted. furthermore, there were critical in-aorta alarms and impella disconnected alarms. cpr was then in progress for pea arrest. the pump was then weaned and explanted. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
9611 | 1220648-2025-28157 | 21976178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976178&pc=OZD | 04/15/2025 | 06/09/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616819 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during the implant, the cannula folded up on itself in left ventricle and was unable to be straightened out until pulled back into descending aorta. attempts to recross unsuccessful. impella removed maintaining wire access and 14x25 sheath placed. impella re-inserted w/ flows of 2.9l/min. support continued. additionally, the patient had torsades requiring defibrillation and two episodes of ventricular fibrillation requiring defibrillation x5 during support. the patient was eventually weaned and explanted. the procedural outcome was the patient survived surgery. | [5 empty] | No product returned. positioning issues: clinical states that while exchanging for repo sheath the impella cannula folded up on itself in lv and was unable to be straightened out. pump was then pulled back into descending aorta which straightened the cannula. attempts to recross unsuccessful. impella re-inserted with no issues. the cause of the positioning issues was not determined due to lack of clinical information regarding why the pump was unable to be repositioned. product damage cannula kink: while exchanging for repo sheath impella cannula folded up on itself in lv and was unable to be straightened out until pulled back into descending aorta. the cause of the product damage cannula kink was most likely use related to the repo sheath exchange as it led to the pump being folded up on itself in the lv requiring it to be pulled out in order to straighten the cannula. vf/vt/af/at: patient had episodes of torsades and ventricular fibrillation. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9612 | 1220648-2025-28158 | 21976191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976191&pc=OZD | 04/17/2025 | 06/10/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601562 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed an impella 5.5 pump to support a non-st-segment elevation myocardial infarction patient who underwent coronary artery bypass graft. the pump was supporting when there was a user error in the purge concentration being infused. the team observed high purge pressure and the team tried to troubleshoot with purge cassette replacement, but instead the pump itself needed to be removed. there was no patient harm. | [5 empty] | After 2 days on support, high purge pressure alarm alarms occurred and troubleshooting could not resolve. night prior to alarms, pm nursing staff mistakenly changed purge solution from d5w w/bicarb to d5w w/potassium chloride. md decided to wean pt from impella support and remove 5.5. field logs were not recovered. the pump was returned and inspected. no visual abnormalities were found on the pump which was able to run for 10 minutes in the lab at nominal operating parameters. hpp was unable to be reproduced. the cause of the high purge pressure was not determined since data logs from the field were not returned and high purge pressure could not be reproduced on the returned pump. h!: device evaluated h3: device was returned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | After 2 days on support, high purge pressure alarm alarms occurred and troubleshooting could not resolve. night prior to alarms, pm nursing staff mistakenly changed purge solution from d5w w/bicarb to d5w w/potassium chloride. md decided to wean pt from impella support and remove 5.5. field logs were not recovered. the pump was returned and inspected. no visual abnormalities were found on the pump which was able to run for 10 minutes in the lab at nominal operating parameters. hpp was unable to be reproduced. the cause of the high purge pressure was not determined since data logs from the field were not returned and high purge pressure could not be reproduced on the returned pump.
9613 | 1220648-2025-28159 | 21976225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976225&pc=OZD | 04/16/2025 | 06/13/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645871 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male | 126 KG | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Per the medical safety officer review the impella pump related event is a serious injury. the patient's demise is associated with the automated impella controller (aic) accordingly under manufacturer device report number 1220648-2025-28211. | It was reported that an impella cp was implanted in a patient with a history of renal failure related to hypertension and dialysis. the patient presented with nstemi and heart failure symptoms. repositioning of the pump caused some ventricular tachycardia, so the physician elected to pull back the pump and leave it shallow. there was a notification of ongoing suction, in addition to positioning alarms. the ecmo cannula was inadvertently removed when transferring patient from bed to gurney. it was also noted that transesophageal echocardiography was performed at bedside and the pump was repositioned. the patient expired while on support. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. based on the medical safety officer's (mso) review of this case the patient's demise is associated with the automated impella controller (aic) accordingly under manufacturer device report number 1220648-2025-28211. the following fields have been revised. b2 death and date of death omitted based on recommendation of mso. b5 revised with mso statement. h1 revised to serious injury. h6 health effect - impact code 1802 death omitted based on recommendation of mso. | The investigation of positioning issues, vf/vt/af/at, and low pump flow has been completed. the impella device was not returned for evaluation. positioning issues: data logs were reviewed and "impella in ventricle" and "impella in aorta" alarms were triggered persistently twice during support. data logs at time of positioning issue on (b)(6) 2025 were not returned for evaluation. clinical details noted that during first instance of positioning alarms, patient was receiving bedside ecmo. the following day, the pump was repositioned and the day after, positioning alarms occurred again when patient was coding and ecmo cannula had been inadvertently removed when moving patient. the root cause of the positioning issues was most likely due to patient movement as first instance of positioning alarms occurred when patient was being cannulated for ecmo after patient was transferred to unit, and second instance of positioning alarms occurred when patient was being transferred to bed and ecmo cannula was also inadvertently removed. low pump flows: data logs were reviewed and suction alarms were present on first day of support. at time of suction alarms, both "impella in aorta" then "impella in ventricle" alarms were being triggered. suction alarms and low flows seem to suddenly resolve after about an hour. clinical details noted that ongoing suction alarms were occurring and patient was being placed on bedside ecmo. the root cause of the low pump flow was most likely due to improper positioning as the low flows occurred with positioning alarms. vf/vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28159. d6a and d6b were erroneously entered on the initial manufacturer device report number 1220648-2025-28159. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9614 | 1220648-2025-28160 | 21976268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976268&pc=OZD | 04/15/2025 | 06/06/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597112 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella 5.5 support had runs of ventricular tachycardia requiring two shocks. on day eight of 5.5 support, the patient went bradycardic requiring atropine administration. impella remained on p-4 with no alarms. additionally, impella kept having suction events which required lowering the p-level. the patient was made a "do not resuscitate" and expired on impella support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of vf/vt/af/ and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: clinical states that multiple times during the case, when the pump was attempted to go up in p-level, suction alarms were observed and the pump had to be reduced to the lower p-level. clinical also states that the patient had pulmonary edema and patient was going through dialysis. pump was not returned for investigation. data logs were returned and suction alarms were present and observed corresponding to the p-level change attempt. placement signal was seen trending downwards. no motor current spikes or purge issues were observed. therefore, based on the data log review and clinical information provided, the root cause of the low flow issue was most likely patient condition related. vt/ vf: as the necessary information was not provided, the root cause of the vt/vf was not determined b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28160. h6 health effect - impact code 4647 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28160.
9615 | 1220648-2025-28161 | 21976272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976272&pc=OZD | 04/17/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623421 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was implanted, and the site developed oozing around repo sheath with notable hematoma due to this and the patient being hypertensive. the impella cp was removed. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9616 | 1220648-2025-28162 | 21976338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976338&pc=OZD | 04/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634286 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 71 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted and an impella 5.5 pump was scheduled for delivery. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9617 | 1220648-2025-28163 | 21976462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976462&pc=OZD | 04/16/2025 | 05/19/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646554 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. the pump was unable to be advanced past the aorta iliac bifurcation; angiogram revealed calcification and narrowing of iliacs bilaterally. the pump was removed and place in sterile water, in surgical mode. vascular consulted and decision made to move forward with utilizing bilateral stenting. the pump was successfully able to advance through the newly stented iliacs and into the left ventricle. the pump was started and slowly ramped to maximum flows without complication. percutaneous coronary intervention was performed. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump unable to be advanced past the aorta iliac bifurcation due to calcification and narrowing of iliac. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9618 | 1220648-2025-28164 | 21976528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976528&pc=OZD | 04/23/2025 | 05/30/2025 |  | 05/07/2025 |  |  | 04/23/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621979 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Sense | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | Stroke/CVA | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that bilateral extremity doppler pulses could not be located on a patient implanted with an impella cp. during support, it was reported that suction alarms were present. flow and suction alarms resolved with volume. an echocardiogram revealed microvascular decompression and papillary muscle rupture, so an urgent mitraclip procedure was performed. blood was noted in the endotracheal tube, along with a drop in hemoglobin and low platelets, and one unit of packed red blood cells and two units of platelets were transfused. a head ct was performed and a left subarachnoid hemorrhage was identified. echo dictation found a thrombus on the impella which was treated with heparin. amio was restarted due to atrial fibrillation with loss of pulsitility. they also noted that the pump placement signal was low. a penumbra was then used multiple times with successful aspiration of the thrombus. an angiogram showed flow restored to the distal extremity. the access site was closed and doppler pulses were obtained. the patient survived. | [5 empty] | The investigation of stroke/cva, vf/vt/af/at, limb ischemia, thrombus, thrombocytopenia, and heart valve damage has been completed. the pump was received for investigation however, on the 23rd there was stated signs of clot ingestion. data logs reveal no change in purge pressure, purge flow or motor current on the date of the reported event. there were no other changes to those parameters throughout the case. the pump had biomaterial in the purge gap but could not be run as the pump was cut. as the necessary information was not provided, the root cause of thrombus, heart valve damage, stroke/cva, thrombocytopenia, and limb ischemia the issue was not determined due to insufficient clinical details. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-28164. h6 medical device problem code 1559 was erroneously entered on the initial manufacturer device report number 1220648-2025-28164. | The impella device was received from the customer, upon completion of the investigation a supplemental mdr will be filed. a.4 is unknown. impella cp with smart assist system: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9619 | 1220648-2025-28165 | 21976536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976536&pc=OZD | 04/17/2025 | 05/22/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623604 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon placing the impella 5.5 into the graft the doctor performed an arteriotomy to help place the impella across the aortic valve. once across the valve and what was visualized on transesophageal echocardiogram seem to be deep and the physician was instructed by the anesthesiologist to rotate the impella 5.5 and not to advance the impella 5.5 but to rotate the impella 5.5 to orient the inlet to the apex of the left ventricle (lv). the physician elected to advance while rotating the device through the arteriotomy site and in the chest through the graft resulting in the perforation of the lv. the physician then retracted the impella 5.5 and fixed the perforation. upon reattempting to place the impella 5.5 across for the second time the impella 5.5 again appeared to be in ideal depth and the doctor was instructed to rotate the impella 5.5 and the physician again advanced and rotated using the same exact technique as before and perforated the lv for a second time. the doctor again retracted the impella and fixed the perforation. on the third attempt, the physician seated the impella 5.5 shallow in the lv, and elected not to rotate the impella 5.5. the patient was then started on impella 5.5 support. bypass was weaned, graft was trimmed and peel-away sheath was removed. the impella 5.5 was secured to the upper chest. however, the patient's chest remained open. care was ultimately withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. the pump was not returned for investigation. according to the clinical notes, the left ventricle was found perforated two times while using the technique to advance and rotate the pump. on the third attempt, the doctor opted to leave the pump in a shallow position and proceeded with the case. the pump was utilized for 9 days following the left ventricle (lv) repair and third placement. the cause of the ventricle perforation issue was most likely use related, based on given clinical details.
9620 | 1220648-2025-28166 | 21976566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976566&pc=OZD | 04/16/2025 | 10/24/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623631 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the waveforms of an implanted impella 5.5 pump became irregular during patient support. an impella in aorta alarm appeared despite the device being properly positioned. it was additionally reported that the patient's labs had begun to hemolyze. the pump was replaced in response to the noted issues. the procedural outcome was the patient survived surgery. | [5 empty] | The fm thrombus was added based on data logs and returned product investigation. the fm positioning issues was added based on the investigation. hemolysis: the returned product had biomaterial on the impeller. the biomaterial was removed and the pump passed hemolysis testing with an mih of (b)(6). hemolysis was first reported from a csc call on (b)(6) 2025 and a confirmation from the representative on an internal call on (b)(6) 2025 confirmed that there was concern for hemolysis started on (b)(6) 2025 around 4:17 pm. the csc call also stated that there was md acknowledgement of an echo showing the pump pointing towards the mitral valve. they required more formal imaging after that. however, the representatives are unsure of the pumps positioning during this time. the data logs show that on (b)(6) around 8am there was a decrease in motor current and flows and impella position unknown alarms. the representatives on the internal call on (b)(6) 2025 were unsure of anything occurring at this time. around 4pm on (b)(6) there were positioning alarms with uncoupling of the ao and lv signal, but no motor current spikes or change in flows. this uncoupling shows that they are experiencing slight diastolic suction. on (b)(6) around 6:30pm there was ingestion event occurred, where there was a spike in motor current and flows became slightly saturated. the pump was explanted shortly after. due to hemolysis being reported on (b)(6), lack of confirmation from the representatives on positioning of the device, and the pump passing hemolysis testing, the root cause of the hemolysis cannot definitely be determined. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. optical signal issue: the clinical details state "waveforms became irregular. impella in aorta alarm despite proper position. hemolyzed labs and red urine. new pump improved everything." further information shows: the representatives stated on (b)(6) there was an "decoupling" of the ao and lv signal. the csc call on (b)(6) @4:17pm stated that an echo showed pointing towards the mitral valve. however, representatives could not confirm improper positioning. the data logs show stable 5s parameters, but a decoupling of the lv and ao ps as there were slight changes in motor current but the placement signal remained stable. there was a csc call on (b)(6) @5:21am where there was an ongoing impella position in aorta alarm where the ps became erratic but all other numbers did not change. the md acknowledged that the pump was near the papillary muscle. however, the representatives, do not confirm improper positioning. the data logs show that during this timeframe there were impella position in aorta alarms and suction and appear to be accurately triggered if the pump was near the papillary muscle. additionally, there is a cyclic pattern in the ps and raw optical sensor, where your snr and contrast decrease while your ps and raw optical increase. this cyclic pattern shortly resolves. there were no further patient related information provided and improper positioning was not confirmed by the representatives. additionally, no further devices said to be used during this time. the returned product showed no damages or abnormalities with the optical system. the 5s parameters were taken from the returned product, and all were within specification. the pump was ran in the lab and the optical sensor performed and held pressure. due to lack of confirmation regarding positioning or other patient related factors, no hard failures seen of the optical sensor, and the optical sensor adequately functioning on the returned product, the root cause of the optical signal issue cannot be determined. positioning issues: the pump was noted from the csc calls to be in not an optical position twice throughout the case. no other information was provided regarding it. due to lack of sufficient clinical details, the root cause of the positioning issues cannot be determined. | G3 ¿date received by manufacturer¿ corrected. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | D6b "if explanted, give date" corrected.
9621 | 1220648-2025-28167 | 21976771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976771&pc=OZD | 04/17/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611117 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that before an impella cp was implanted, a 12 french sheath was first placed to perform valvuloplasty. during this time, the patient became hemodynamically critical. the valvuloplasty was performed successfully, followed by the 14 french peel-away sheath, and the impella cp was implanted. after the impella cp was placed, it was discovered that the patient was bleeding from the femoral artery. a covered stent was inserted into the femoral artery and the impella cp was removed. the patient expired in the procedural area. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9622 | 1220648-2025-28168 | 21976815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976815&pc=OZD | 04/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026653483 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 53 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient implanted with an impella 5.5 experienced a significant amount of ventricular tachycardia one day into support. the patient was placed on amiodarone and lidocaine, and their condition improved. eight days later, anticoagulation was put on hold to help manage a diffuse alveolar hemorrhage. support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: & impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9623 | 1220648-2025-28169 | 21976852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976852&pc=OZD | 04/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629143 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella repo sheath was found to be occlusive in right femoral artery. the interventional cardiologist used distal reperfusion, using 7fr braided sheath in left femoral artery connected to 5fr braided sheath in right femoral artery distal to impella. leg color improved after placement. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9624 | 1220648-2025-28170 | 21976911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21976911&pc=OZD | 04/17/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640501 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after percutaneous coronary intervention with impella cp support, the impella cp was going to be left overnight to optimize hemodynamic. however, due to uncontrollable bleeding around the repositioning sheath, the decision was made to slowly wean th impella cp. the impella cp was explanted and two units of blood were administered. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9625 | 1220648-2025-28171 | 21977086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977086&pc=OZD | 04/17/2025 | 06/23/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626089 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA GUIDEWIRE 0.018 | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that at the start of the impella cp placement, the guidewire was spread, and the guidewire became unsterile. another guidewire was used, and the case was completed successfully. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use-related issue as evidenced by clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.4 is unknown.
9626 | 1220648-2025-28172 | 21977116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977116&pc=OZD | 04/17/2025 | 06/12/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  | 06/30/2025 |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357214 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, there were suction events and ventricular tachycardia during balloon inflations which resolved with deflation and dropping p level. additionally, the patient had access site bleeding. two units of packed red blood cells were given to the patient. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of low pump flow, vf/vt/af/at, and access site bleeding has been completed. the impella device was not returned for evaluation. the data logs were reviewed. no motor current (mc) spike outside of p-level changes observed. no positioning issue alarmed. fluctuating placement signal issue observed. suction alarms occurred on the first 3 hours of support. drop in flow observed at p-5 and 6 and resolved. low pump flow: some suction events and vt noted during intra-aortic balloon pump (iabp) balloon inflations, resolved with deflation and dropping p-level. the cause of the low pump flow was most likely patient condition related due to suction occurring during iabp inflations. as the necessary information was not provided, the root cause of the vt/vf and access site bleeding was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9627 | 1220648-2025-28173 | 21977148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977148&pc=OZD | 04/15/2025 | 10/24/2025 |  | 05/07/2025 |  |  | 04/15/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614085 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hematoma | Hemolysis | Ischemia | Ectopic Heartbeat | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that a patient impella 5.5 for hemodynamic support. during support the patient had some numbness in his right fingers however no pain. trans-thoracic echo was ordered and heparin was being held for 24 hr. hemolysis labs were also ordered. it was noted of some suction alarms, they reduced to p4 and had no more alarm. additionally the patient was having intermittent ectopy that resolved after p-level change. it was again noted of ectopy with movement however the team didn't want to reposition unless it becomes more of a problem hemodynamically/ sustained. giving fluids to help. it was also reported of lactate dehydronase hemolyzed. the impella's map readings were 10 points lower than patient's arterial line. additionally, there was a venous hematoma at the impella site, manual pressure was held and heparin was on hold. the patient continued on support. | Us complainant reported that a patient was placed on impella 5.5 for hemodynamic support. it was reported that the device exhibited an optimal signal issue. |  |  |  |  | D6b "if explanted, give date" corrected. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of vf/vt/af/at, low pump flow, limb ischemia, access site bleeding, and hemolysis has been completed. the impella device was not returned for evaluation. the cause of the low pump flow issue was most likely patient condition related, based on given clinical details and data log review. without additional clinical details, the root cause of vf/vt/af/at, limb ischemia, access site bleeding, and hemolysis could not be determined. d6b explant should have been left blank on manufacturer device report 1220648-2025-28173 as the pump has not been explanted h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28173. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. pump was not returned for investigation. optical signal issue: the placement signal was lower than the blood pressure, and the map is 10 points lower than the arterial line during the case run. the data log shows the placement signal was ~40-80 mmhg throughout the case. snr and contrast were dropped within specification range during mid case without affecting the placement signal reading. the root cause of the optical signal issue was not determined since product was not returned and no abnormalities were seen with the optical parameters in the data logs. low pump flow: a suction alarm was reported on 16-apr-2025, and the meantime patient also experienced intermittent ectopy, which resolved after the p-level change. the data log confirms the suction alarms and a sudden drop in pump flow without affecting the placement signal, 5s optical parameter, or the motor current. there were no suction alarms reported after reducing the p-level. the cause of the low pump flow issue was most likely patient condition related, based on given clinical details and data log review. hemolysis: ldh was hemolyzed multiple times during the case, 04-17 and 05-05. the data log did not report any motor current spikes or positioning issues during the case run. the cause of the hemolysis issue was not determined since without returned product investigation. access site bleeding - hematoma was reported on the 10th day of pump use, and the doctor stopped the heparin. no further information was given related to the hematoma. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. limb ischemia - reversible: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. vf/vt/af/at: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue for "2917" and additional information regarding the investigation of the optical signal.
9628 | 1220648-2025-28174 | 21977368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977368&pc=PYX | 04/18/2025 | 06/12/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025585355 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella rp flex support and while the doctor was closing the chest, the pa pressures went off the screen on the ic. the motor current went flat and the flows went to 0. the doctor re-opened the chest hoping that would help, and it did not. the doctor then removed the impella rp flex and found a foreign material in the inlet of the pump. it was decided not to place another pump. the patient survived the procedure. | [5 empty] | The investigation of pump stop and thrombus has been completed. the impella device was not returned for evaluation. upon reviewing the data logs, there was no pump stop issue found during the case. the device functioned/operated to specification. the cause of the low pump flow issue was most likely biomaterial based on data log review. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
9629 | 1220648-2025-28175 | 21977458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977458&pc=OZD | 04/17/2025 | 06/12/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631186 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the 78-year-old male patient admitted in with known ventricular tachycardia (vt) arrest and in need of an electrophysiology procedure with the support of a 5.5 pump. the delivery of the 5.5 was complicated by the damage to the native vasculature. the placement was aborted. the vessel was repaired and had blood loss that was treated by blood product infusion. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding, delivery issues, vascular damage - great vessel, and product delivery issues has been completed. the impella device was not returned for evaluation. delivery issue: as the 5.5 cannula was passing the anastomosis the artery was compromised and device was removed before making it to the subclavian. the root cause of the delivery issue was most likely use related, as the 5.5 cannula was passing the anastomosis the artery was compromised and device was removed before making it to the subclavian. access site bleeding: the graft anastomosis separated and compromised the axillary artery ,bleeding. the root cause of the access site bleeding issue was most likely use related, the 5.5 cannula was going through the beveled graft and the artery was compromised. vascular damage -great vessel: artery cross clamped and vessel repaired. the root cause of the vascular damage -great vessel issue was not determined since product was not returned and insufficient clinical details provided. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h10 instruction for use were incomplete on the initial manufacturer device report number 1220648-2025-28175. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9630 | 1220648-2025-28176 | 21977696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977696&pc=OZD | 04/18/2025 | 06/20/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600250 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump was placed to support a patient admitted in with cardiogenic shock and organ failure. the impella cp accessed leg was ischemic. the ischemia was not treated by femorofemoral bypass or premature explant as the patient could not survive off the pump. further details stated that the patient expired on support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9631 | 1220648-2025-28177 | 21977910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977910&pc=OZD | 04/17/2025 | 06/18/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814768 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella aic started to boot up but would not progress beyond the boot up process. lines across screen while console was attempting to boot hard drive. also, the console was not able to turn off via pressing the side button. it was turned off via the recessed power buttons on bottom side. there was no patient involvement. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. there are no logs available for the complaint device analysis: this console was tested after arriving at the engineering lab. the boot-up issue and the line across the screen could not be reproduced. the console was able to run the optical test pump for 1 hour at p-level p9, as intended. the input voltage at the 4-in-1 connector x25 is measuring 23.7v under ac power, as intended. root cause: the root cause of the boot-up issue and the screen flicker could not be determined due to the inability to reproduce.
9632 | 1220648-2025-28178 | 21977959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21977959&pc=OZD | 04/18/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643602 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was placed emergent post arrest during percutaneous coronary intervention in the catheterization lab. however, post sheath exchange, a hematoma was noted at the left femoral artery. vascular was consulted. manual pressure and blood products were administered with initial improvement. the patient continued to decompensate. massive transfusion protocol was initiated and the patient was taken for vascular cut down of the right and left femoral artery and the impella cp removal. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9633 | 1220648-2025-28179 | 21978032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978032&pc=OZD | 04/18/2025 | 05/20/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620067 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 support and developed bleeding from the anastomoses site and through the graft. bleeding was also from multiple sites despite activated clotting time of 135 and discontinuation of heparin. the patient lost three units of blood during transport from the operating room to the intensive care unit. several units of blood products have been administered. bleeding remained aggressive. a hemashield platinum woven graft was used, but blood was then leaking through the graft itself. the decision was made to reopen the axillary pocket and tie off the graft to prevent further blood flow through it. the jp drain was left in place. bleeding reduced significantly but minimal oozing was still noticed at the graft anastomosed site. the pocket was closed and dressed. it was further reported that the patient went into tachycardia episode. esmolol was started and cardiovert total sinus rhythm. the next day, the patient entered supraventricular tachycardia again overnight and was cardioverted back to sinus rhythm. the patient remained on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding, vascular damage - peripheral vessel, and vf/vt/af/at has been completed. the root cause of access site bleeding was most likely patient condition since patient had bleeding from multiple sites. the root cause of vascular damage peripheral vessel and vt/vf cannot be determined since the product was not returned and insufficient clinical details were provided.
9634 | 1220648-2025-28180 | 21978274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978274&pc=OZD | 04/18/2025 | 06/06/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591248 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, a clot was ingested. the impella cp dropped flows from 3.5 to 1.2 and then to 0.8. a clot was ingested by the impella cp pump. the impella cp was removed and replaced with a new pump. the patient expired on support of the replacement pump. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. clinical stated reported clot during case. pump was returned and no damages or abnormalities were found. as the necessary information was not provided, the root cause of the thrombus was not determined.
9635 | 1220648-2025-28181 | 21978299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978299&pc=OZD | 04/12/2025 | 06/27/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The medical safety officer reviewed and determined the following: controller failure¿ appeared for "1 second" then alarm condition resolved without any intervention; there was no aic exchange and no report or indication of interruption/change to in flow. support continued until patient was terminally weaned. patient expired from events unrelated to the aic. the demise outcome is associated with the impella pump under manufacturer device report number 1220648-2025-28062. | The user facility reported that during routine purge cassette change at bedside an icu controller failure briefly appeared but resolved without any intervention. when alarm history was reviewed at explant, this alarm was discovered. rn who had the patient that day stated that there were no changes to pump performance or flow noted. there was no patient harm reported. |  |  |  |  | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the root cause of the controller failure alarm was not determined due to the inability to reproduce the issue. it could be attributed to an open purge line/leakage. the medical safety officer reviewed and determined the following: controller failure¿ appeared for "1 second" then alarm condition resolved without any intervention; there was no aic exchange and no report or indication of interruption/change to in flow. support continued until patient was terminally weaned. patient expired from events unrelated to the aic. the demise outcome is associated with the impella pump under manufacturer device report number 1220648-2025-28062 b5 and h10 was inadvertently omitted medical safety officer statement on the initial manufacturer device report number 1220648-2025-28181. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
9636 | 1220648-2025-28182 | 21978456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978456&pc=OZD | 04/18/2025 | 10/01/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642649 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was implanted with an impella cp pump for support during a high-risk procedure and cardiogenic shock. the patient was suffering from pulseless electrical activity (pea) and had cardiac intervention. during the cardiac procedure and while on support, there was low flows and ps issues but, the pump was not replaced. after all measures and under four hours of support, care was withdrawn and the patient expired. the patient was not an extracorporeal membrane oxygenation or impella rp candidate. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product or data logs were returned for evaluation. the root cause of the optical signal issue could not be definitively determined as no product or logs were returned for evaluation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28182. h4 device manufacture date corrected | The investigation of low pump flow has been completed. the root cause of the low pump flow could not be definitively determined as no product or logs were returned for evaluation.
9637 | 1220648-2025-28183 | 21978571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978571&pc=OZD | 04/18/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615146 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant escalated an impella cp to an impella 5.5 pump to support a patient admitted in for coronary artery bypass graft and mitral valve replacement. the medical doctor was adjusting the pump position when the pump came out of position in the left and could not be re-crossed. the pump was replaced and there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.